## Paper Tables
System | Error Source | Nominal/Mean Value/Range | Error Limits (±) | Probability Distribution | Uncertainty Type (A/B) | Relative Standard Uncertainty | Sensitivity Coefficient | Degrees of Freedom, v | Subsystem Combined Uncertainty | Effective Degrees of Freedom, veff | System | Subsystem Combined Uncertainty | System Combined Uncertainty (RSS) | System Expanded Uncertainty (k=2) | Subsystem Expanded Uncertainty (k=2) | System Combined Uncertainty (RSS) | System Expanded Uncertainty (k=2) | System | Subsystem Combined Uncertainty | System Combined Uncertainty (RSS)
FLIR SC7600 Camera | tAtm | 293.05 - 295.05K | 0.1 | Normal | A | 0.5926 | 0.2017 | 999 | 0.1267 | 1020 | FLIR SC7600 Camera | 0.1267 | 0.1816 | 0.3632 | 0.2534 | 0.3632 | 0.7264 | FLIR SC7600 Camera | 0.1203 | 0.1816
| Humidity | 59-62% | 0.1 | Normal | A | 0.0105 | 0.0173 | 999 |  |  | Flash Box | 0.0014 |  |  | 0.0028 |  |  | Flash Box | 0.0014
| Distance | 0.26m | 0.01 | Uniform | A | 0.0058 | 2.2216 | 999 |  |  | Sample | 0.1295 |  |  | 0.259 |  |  | Sample | 0.1295
| TSR reading | 5.13 | 0.001 | Normal | A | 0.0005 | 0.0112 | 999 |  |  | Computer | 0.0141 |  |  | 0.0282 |  |  | Computer | 0.0141
| Positioning | 0 | 0.001 | Uniform | B | 0.0005 | 1.7006 | ∞
Flash Box | Flash intensity | 75% | 0.001 | Uniform | B | 0.0005 | 1.7399 | ∞ | 0.0014 | ∞ | [8 empty] | SQRT(SUMSQ(W3:W6)) | 0.17732
| Synchronisation | 0s | 0.001 | Uniform | B | 0.0005 | 1.7166 | ∞
Sample | Emissivity | 85% | 0.02 | Uniform | A | 0.0114 | 1.7591 | 999 | 0.1295 | 1147
| Temperature | 293.05 - 295.05K | 0.1 | Normal | A | 0.5926 | 0.2017 | 999
| Mount Positioning | 0 | 0.001 | Uniform | B | 0.0005 | 1.732 | ∞
| Replacement moving | 0 | 0.02 | Uniform | B | 0.0116 | 1.7218 | ∞
Computer | Operator bias | 0 | 0.01 | Uniform | B | 0.0058 | 1.7276 | ∞ | 0.0141 | ∞
| Computation error | 0 | 0.01 | Uniform | B | 0.0058 | 1.7273 | ∞
Influence of error sources on resulting uncertainties
System | Error Source | Uncertainty Type (A/B) | Influence on Resulting Uncertainty
IR Camera | Atmospheric temp. | A | Greater variations in temp. increase uncertainty in readings
| Humidity | A | Greater variations in humidity increase uncertainty in readings
| Distance | A | Changes in distance will move ROI and alter camera focus - keep constant
| Measured reading | A | Small, constant temperature fluctuations in air result in variations in values per frame
| Positioning | B | Changes in position and angle will move ROI and alter camera focus - keep constant
Flash Box | Flash intensity | B | Variations in flash intensity will cause changes in sample surface temperature per recording
| Synchronisation | B | Flash must occur for the first recorded frame to give constant accurate recordings
Sample | Emissivity | A | Sample emissivity is constant - difficult to accurately obtain values for composite materials
| Temperature | A | Influenced by flash - must be allowed to cool to room temperature before starting next recording
| Mount positioning | B | Changes in position and angle will move ROI and alter camera focus - keep constant
| Replacement moving | B | Held in place by mount - changes in position and angle may still occur
Computer | Operator bias | B | Operator may input inaccurate readings or make incorrect measurements
| Computation error | B | Operator may make incorrect calculations

## Conditions
Run | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Mean | StdDev
Atmo Temp (°C) | 21.4 | 21.8 | 21.9 | 19.9 | 20.6 | 21 | 20.9 | 21.4 | 21.4 | 21.6 | 21.19 | 0.57524
Atmo Temp (K) | 294.55 | 294.95 | 295.05 | 293.05 | 293.75 | 294.15 | 294.05 | 294.55 | 294.55 | 294.75 | 294.34 | 0.57524
Humidity (%) | 0.59 | 0.61 | 0.62 | 0.6 | 0.62 | 0.62 | 0.61 | 0.6 | 0.6 | 0.6 | 0.607 | 0.01005

## Readings
|  | Run1 |  |  | Run2 |  |  | Run3 |  |  | Run4 |  |  | Run5 |  |  | Run6 |  |  | Run7 |  |  | Run8 |  |  | Run9 |  |  | Run10
Frame | 1 | 2 | 3 | 1 | 2 | 3 | 1 | 2 | 3 | 1 | 2 | 3 | 1 | 2 | 3 | 1 | 2 | 3 | 1 | 2 | 3 | 1 | 2 | 3 | 1 | 2 | 3 | 1 | 2 | 3 | Mean | StdDev | NormStd
1 | 7.7669 | 7.7703 | 7.7319 | 7.7396 | 7.7553 | 7.7604 | 7.7554 | 7.7655 | 7.77 | 7.7149 | 7.7294 | 7.7331 | 7.7155 | 7.717 | 7.7211 | 7.7384 | 7.7445 | 7.7354 | 7.7292 | 7.7399 | 7.7354 | 7.7415 | 7.7445 | 7.7393 | 7.7384 | 7.743 | 7.7513 | 7.751 | 7.7559 | 7.759 | 7.7431 | 0.015507 | 0.0020027
2 | 7.4368 | 7.4347 | 7.399 | 7.4042 | 7.4217 | 7.4274 | 7.4239 | 7.4352 | 7.4387 | 7.3865 | 7.3942 | 7.4013 | 7.3853 | 7.3883 | 7.3907 | 7.403 | 7.4107 | 7.4043 | 7.3976 | 7.4044 | 7.4027 | 7.4124 | 7.4113 | 7.4101 | 7.4065 | 7.4117 | 7.4152 | 7.4193 | 7.4219 | 7.4258 | 7.4108 | 0.014985 | 0.0020221
3 | 7.2347 | 7.2405 | 7.2057 | 7.2091 | 7.2279 | 7.2332 | 7.2305 | 7.2421 | 7.2446 | 7.1895 | 7.2001 | 7.207 | 7.1939 | 7.1972 | 7.1993 | 7.208 | 7.2145 | 7.211 | 7.2026 | 7.2094 | 7.2092 | 7.2205 | 7.2183 | 7.2169 | 7.2121 | 7.2178 | 7.22 | 7.2258 | 7.2293 | 7.2314 | 7.2167 | 0.014523 | 0.0020124
4 | 7.0923 | 7.1008 | 7.0664 | 7.0689 | 7.0883 | 7.0933 | 7.0911 | 7.1028 | 7.1044 | 7.0483 | 7.0604 | 7.0668 | 7.0558 | 7.059 | 7.0612 | 7.0679 | 7.0731 | 7.0714 | 7.062 | 7.0692 | 7.0698 | 7.0818 | 7.0792 | 7.0775 | 7.0719 | 7.0781 | 7.0797 | 7.0863 | 7.0905 | 7.0915 | 7.077 | 0.014403 | 0.0020352
5 | 6.9828 | 6.991 | 6.9569 | 6.9587 | 6.9785 | 6.9833 | 6.9814 | 6.993 | 6.994 | 6.9381 | 6.9505 | 6.9565 | 6.9469 | 6.95 | 6.9523 | 6.9579 | 6.9622 | 6.9615 | 6.9516 | 6.959 | 6.9601 | 6.9724 | 6.9697 | 6.9678 | 6.9618 | 6.9684 | 6.9696 | 6.9765 | 6.981 | 6.9816 | 6.9672 | 0.014372 | 0.0020628
6 | 6.8938 | 6.9003 | 6.8664 | 6.8679 | 6.8879 | 6.8925 | 6.8907 | 6.9023 | 6.9029 | 6.8477 | 6.8597 | 6.8654 | 6.8567 | 6.8598 | 6.8622 | 6.8672 | 6.8707 | 6.8707 | 6.8605 | 6.8681 | 6.8695 | 6.8819 | 6.8793 | 6.8772 | 6.871 | 6.8779 | 6.8788 | 6.8858 | 6.8905 | 6.8909 | 6.8765 | 0.014369 | 0.0020896
7 | 6.8185 | 6.823 | 6.7893 | 6.7905 | 6.8106 | 6.8151 | 6.8134 | 6.825 | 6.8253 | 6.771 | 6.7822 | 6.7878 | 6.7797 | 6.7828 | 6.7853 | 6.79 | 6.793 | 6.7933 | 6.783 | 6.7908 | 6.7923 | 6.8047 | 6.8021 | 6.8 | 6.7936 | 6.8008 | 6.8015 | 6.8085 | 6.8132 | 6.8136 | 6.7993 | 0.014382 | 0.0021152
8 | 6.7531 | 6.7557 | 6.7221 | 6.7232 | 6.7432 | 6.7477 | 6.7461 | 6.7576 | 6.7576 | 6.7043 | 6.7147 | 6.7202 | 6.7126 | 6.7156 | 6.7181 | 6.7228 | 6.7253 | 6.7259 | 6.7156 | 6.7234 | 6.725 | 6.7373 | 6.7348 | 6.7328 | 6.7263 | 6.7337 | 6.7342 | 6.7411 | 6.7458 | 6.7463 | 6.7321 | 0.014402 | 0.0021393
9 | 6.6951 | 6.6961 | 6.6626 | 6.6635 | 6.6835 | 6.688 | 6.6864 | 6.6979 | 6.6977 | 6.6452 | 6.6548 | 6.6603 | 6.653 | 6.656 | 6.6586 | 6.6633 | 6.6655 | 6.6661 | 6.6559 | 6.6637 | 6.6654 | 6.6775 | 6.6752 | 6.6732 | 6.6667 | 6.6742 | 6.6746 | 6.6814 | 6.686 | 6.6867 | 6.6725 | 0.014425 | 0.0021619
10 | 6.643 | 6.6426 | 6.6091 | 6.6099 | 6.6299 | 6.6344 | 6.6328 | 6.6442 | 6.644 | 6.5923 | 6.601 | 6.6066 | 6.5994 | 6.6025 | 6.6052 | 6.6099 | 6.6118 | 6.6125 | 6.6024 | 6.6102 | 6.6119 | 6.6238 | 6.6218 | 6.6198 | 6.6132 | 6.6209 | 6.6211 | 6.6278 | 6.6323 | 6.6331 | 6.619 | 0.014446 | 0.0021825
11 | 6.5956 | 6.594 | 6.5606 | 6.5614 | 6.5813 | 6.5858 | 6.5842 | 6.5956 | 6.5953 | 6.5442 | 6.5523 | 6.5579 | 6.5509 | 6.554 | 6.5566 | 6.5615 | 6.5632 | 6.5639 | 6.5539 | 6.5617 | 6.5633 | 6.5751 | 6.5732 | 6.5714 | 6.5647 | 6.5725 | 6.5727 | 6.5792 | 6.5835 | 6.5846 | 6.5705 | 0.014463 | 0.0022011
12 | 6.552 | 6.5497 | 6.5163 | 6.5169 | 6.5369 | 6.5413 | 6.5397 | 6.5512 | 6.5508 | 6.5003 | 6.5076 | 6.5134 | 6.5064 | 6.5096 | 6.5122 | 6.5172 | 6.5188 | 6.5194 | 6.5096 | 6.5173 | 6.519 | 6.5306 | 6.5289 | 6.5271 | 6.5203 | 6.5283 | 6.5284 | 6.5347 | 6.5389 | 6.5402 | 6.5261 | 0.014473 | 0.0022176
13 | 6.5116 | 6.5088 | 6.4754 | 6.476 | 6.496 | 6.5003 | 6.4988 | 6.5102 | 6.5098 | 6.4597 | 6.4665 | 6.4725 | 6.4655 | 6.4687 | 6.4713 | 6.4764 | 6.4779 | 6.4785 | 6.4688 | 6.4765 | 6.4781 | 6.4895 | 6.488 | 6.4863 | 6.4795 | 6.4876 | 6.4876 | 6.4938 | 6.4978 | 6.4992 | 6.4852 | 0.014476 | 0.0022321
14 | 6.474 | 6.4709 | 6.4376 | 6.4381 | 6.458 | 6.4624 | 6.4608 | 6.4723 | 6.4718 | 6.4221 | 6.4284 | 6.4345 | 6.4275 | 6.4308 | 6.4334 | 6.4386 | 6.44 | 6.4405 | 6.431 | 6.4386 | 6.4402 | 6.4515 | 6.4501 | 6.4485 | 6.4417 | 6.4499 | 6.4498 | 6.4558 | 6.4597 | 6.4613 | 6.4473 | 0.014472 | 0.0022447
15 | 6.4388 | 6.4357 | 6.4023 | 6.4027 | 6.4227 | 6.427 | 6.4255 | 6.4369 | 6.4365 | 6.3869 | 6.3929 | 6.3991 | 6.3921 | 6.3954 | 6.3981 | 6.4034 | 6.4048 | 6.4052 | 6.3958 | 6.4033 | 6.4049 | 6.416 | 6.4148 | 6.4133 | 6.4064 | 6.4147 | 6.4146 | 6.4204 | 6.4242 | 6.4259 | 6.412 | 0.014463 | 0.0022557
16 | 6.4057 | 6.4027 | 6.3693 | 6.3696 | 6.3896 | 6.3939 | 6.3924 | 6.4038 | 6.4034 | 6.354 | 6.3596 | 6.366 | 6.3589 | 6.3624 | 6.365 | 6.3704 | 6.3718 | 6.3721 | 6.3628 | 6.3703 | 6.3718 | 6.3829 | 6.3818 | 6.3804 | 6.3734 | 6.3818 | 6.3817 | 6.3873 | 6.391 | 6.3928 | 6.379 | 0.01445 | 0.0022652
17 | 6.3744 | 6.3716 | 6.3382 | 6.3385 | 6.3585 | 6.3628 | 6.3613 | 6.3728 | 6.3723 | 6.323 | 6.3283 | 6.3349 | 6.3278 | 6.3313 | 6.3339 | 6.3394 | 6.3408 | 6.341 | 6.3319 | 6.3393 | 6.3408 | 6.3517 | 6.3508 | 6.3494 | 6.3424 | 6.3509 | 6.3508 | 6.3562 | 6.3598 | 6.3617 | 6.3479 | 0.014432 | 0.0022735
18 | 6.3448 | 6.3424 | 6.309 | 6.3092 | 6.3292 | 6.3335 | 6.3319 | 6.3435 | 6.343 | 6.2937 | 6.2989 | 6.3056 | 6.2984 | 6.302 | 6.3045 | 6.3102 | 6.3116 | 6.3117 | 6.3027 | 6.3101 | 6.3115 | 6.3223 | 6.3215 | 6.3202 | 6.3132 | 6.3217 | 6.3216 | 6.3268 | 6.3304 | 6.3324 | 6.3186 | 0.014412 | 0.0022808
19 | 6.3167 | 6.3148 | 6.2813 | 6.2814 | 6.3014 | 6.3057 | 6.3042 | 6.3157 | 6.3153 | 6.266 | 6.271 | 6.2779 | 6.2707 | 6.2743 | 6.2768 | 6.2826 | 6.2839 | 6.284 | 6.2751 | 6.2824 | 6.2838 | 6.2945 | 6.2939 | 6.2926 | 6.2856 | 6.2941 | 6.294 | 6.2991 | 6.3025 | 6.3046 | 6.2909 | 0.01439 | 0.0022874
20 | 6.2899 | 6.2885 | 6.255 | 6.2551 | 6.2751 | 6.2794 | 6.2779 | 6.2895 | 6.289 | 6.2396 | 6.2445 | 6.2516 | 6.2443 | 6.248 | 6.2505 | 6.2563 | 6.2577 | 6.2577 | 6.2489 | 6.2562 | 6.2575 | 6.2681 | 6.2676 | 6.2664 | 6.2593 | 6.268 | 6.2678 | 6.2727 | 6.2761 | 6.2783 | 6.2646 | 0.014367 | 0.0022933
21 | 6.2643 | 6.2636 | 6.23 | 6.23 | 6.2501 | 6.2543 | 6.2529 | 6.2645 | 6.2641 | 6.2145 | 6.2193 | 6.2265 | 6.2192 | 6.223 | 6.2254 | 6.2314 | 6.2328 | 6.2327 | 6.224 | 6.2312 | 6.2325 | 6.2431 | 6.2427 | 6.2415 | 6.2344 | 6.2431 | 6.2429 | 6.2477 | 6.251 | 6.2532 | 6.2395 | 0.014344 | 0.0022989
22 | 6.2399 | 6.2398 | 6.2062 | 6.2061 | 6.2263 | 6.2304 | 6.229 | 6.2406 | 6.2403 | 6.1905 | 6.1953 | 6.2027 | 6.1953 | 6.1991 | 6.2016 | 6.2076 | 6.209 | 6.2088 | 6.2003 | 6.2075 | 6.2087 | 6.2192 | 6.2189 | 6.2178 | 6.2106 | 6.2193 | 6.2192 | 6.2238 | 6.227 | 6.2293 | 6.2157 | 0.014322 | 0.0023042
23 | 6.2165 | 6.2171 | 6.1834 | 6.1833 | 6.2035 | 6.2076 | 6.2062 | 6.2179 | 6.2175 | 6.1676 | 6.1724 | 6.1799 | 6.1725 | 6.1764 | 6.1788 | 6.1849 | 6.1863 | 6.1861 | 6.1776 | 6.1847 | 6.1859 | 6.1963 | 6.1962 | 6.1951 | 6.1879 | 6.1966 | 6.1965 | 6.201 | 6.2041 | 6.2065 | 6.1929 | 0.014301 | 0.0023093
24 | 6.194 | 6.1954 | 6.1616 | 6.1615 | 6.1817 | 6.1858 | 6.1844 | 6.1961 | 6.1958 | 6.1456 | 6.1505 | 6.1581 | 6.1506 | 6.1546 | 6.1569 | 6.1631 | 6.1646 | 6.1643 | 6.1559 | 6.163 | 6.1641 | 6.1745 | 6.1744 | 6.1734 | 6.1661 | 6.1749 | 6.1748 | 6.1791 | 6.1822 | 6.1846 | 6.1711 | 0.014282 | 0.0023143
25 | 6.1724 | 6.1745 | 6.1407 | 6.1405 | 6.1607 | 6.1648 | 6.1635 | 6.1752 | 6.1749 | 6.1245 | 6.1294 | 6.1372 | 6.1297 | 6.1337 | 6.136 | 6.1422 | 6.1438 | 6.1434 | 6.1351 | 6.1421 | 6.1432 | 6.1536 | 6.1536 | 6.1526 | 6.1453 | 6.1541 | 6.154 | 6.1582 | 6.1612 | 6.1636 | 6.1501 | 0.014265 | 0.0023194
26 | 6.1517 | 6.1545 | 6.1206 | 6.1204 | 6.1407 | 6.1447 | 6.1434 | 6.1552 | 6.1549 | 6.1042 | 6.1092 | 6.1171 | 6.1095 | 6.1136 | 6.1159 | 6.1222 | 6.1238 | 6.1233 | 6.1151 | 6.1221 | 6.1232 | 6.1335 | 6.1335 | 6.1326 | 6.1253 | 6.1341 | 6.134 | 6.138 | 6.1411 | 6.1435 | 6.13 | 0.01425 | 0.0023246
27 | 6.1316 | 6.1353 | 6.1013 | 6.101 | 6.1213 | 6.1254 | 6.1241 | 6.1359 | 6.1356 | 6.0847 | 6.0898 | 6.0978 | 6.0902 | 6.0943 | 6.0966 | 6.1029 | 6.1045 | 6.104 | 6.0959 | 6.1028 | 6.1038 | 6.1141 | 6.1143 | 6.1134 | 6.106 | 6.1149 | 6.1148 | 6.1187 | 6.1217 | 6.1241 | 6.1107 | 0.014237 | 0.0023299
28 | 6.1124 | 6.1168 | 6.0827 | 6.0823 | 6.1027 | 6.1067 | 6.1054 | 6.1173 | 6.1171 | 6.0659 | 6.0711 | 6.0792 | 6.0715 | 6.0757 | 6.078 | 6.0843 | 6.086 | 6.0855 | 6.0774 | 6.0843 | 6.0852 | 6.0955 | 6.0957 | 6.0949 | 6.0874 | 6.0963 | 6.0963 | 6.1 | 6.103 | 6.1054 | 6.0921 | 0.014227 | 0.0023353
29 | 6.0937 | 6.0989 | 6.0647 | 6.0643 | 6.0848 | 6.0887 | 6.0875 | 6.0994 | 6.0992 | 6.0477 | 6.053 | 6.0613 | 6.0535 | 6.0578 | 6.06 | 6.0664 | 6.0681 | 6.0675 | 6.0595 | 6.0664 | 6.0673 | 6.0776 | 6.0778 | 6.077 | 6.0695 | 6.0785 | 6.0784 | 6.0821 | 6.085 | 6.0874 | 6.0741 | 0.014219 | 0.002341
30 | 6.0758 | 6.0817 | 6.0474 | 6.047 | 6.0675 | 6.0714 | 6.0702 | 6.0821 | 6.0819 | 6.0301 | 6.0356 | 6.0439 | 6.0362 | 6.0405 | 6.0426 | 6.0491 | 6.0508 | 6.0502 | 6.0423 | 6.0491 | 6.05 | 6.0602 | 6.0605 | 6.0598 | 6.0523 | 6.0612 | 6.0612 | 6.0647 | 6.0676 | 6.07 | 6.0568 | 0.014214 | 0.0023468
31 | 6.0584 | 6.065 | 6.0306 | 6.0302 | 6.0507 | 6.0546 | 6.0534 | 6.0654 | 6.0652 | 6.0131 | 6.0187 | 6.0272 | 6.0194 | 6.0237 | 6.0259 | 6.0323 | 6.0341 | 6.0335 | 6.0256 | 6.0324 | 6.0332 | 6.0435 | 6.0438 | 6.0431 | 6.0355 | 6.0445 | 6.0445 | 6.0479 | 6.0508 | 6.0532 | 6.04 | 0.014212 | 0.0023529
32 | 6.0415 | 6.0489 | 6.0144 | 6.0139 | 6.0345 | 6.0383 | 6.0372 | 6.0493 | 6.049 | 5.9967 | 6.0024 | 6.011 | 6.0031 | 6.0075 | 6.0097 | 6.0161 | 6.018 | 6.0173 | 6.0095 | 6.0162 | 6.017 | 6.0273 | 6.0277 | 6.0269 | 6.0194 | 6.0284 | 6.0283 | 6.0317 | 6.0345 | 6.0369 | 6.0238 | 0.014211 | 0.0023592
33 | 6.0252 | 6.0332 | 5.9987 | 5.9981 | 6.0188 | 6.0226 | 6.0215 | 6.0336 | 6.0334 | 5.9807 | 5.9867 | 5.9953 | 5.9874 | 5.9918 | 5.9939 | 6.0004 | 6.0023 | 6.0016 | 5.9939 | 6.0005 | 6.0013 | 6.0116 | 6.012 | 6.0113 | 6.0037 | 6.0127 | 6.0127 | 6.0159 | 6.0188 | 6.0212 | 6.008 | 0.014213 | 0.0023657
34 | 6.0094 | 6.0181 | 5.9834 | 5.9829 | 6.0036 | 6.0073 | 6.0063 | 6.0184 | 6.0182 | 5.9653 | 5.9713 | 5.9801 | 5.9721 | 5.9766 | 5.9787 | 5.9852 | 5.9871 | 5.9864 | 5.9787 | 5.9853 | 5.9861 | 5.9964 | 5.9969 | 5.9962 | 5.9885 | 5.9975 | 5.9976 | 6.0007 | 6.0035 | 6.0059 | 5.9928 | 0.014217 | 0.0023724
35 | 5.9941 | 6.0034 | 5.9686 | 5.968 | 5.9888 | 5.9925 | 5.9915 | 6.0037 | 6.0035 | 5.9502 | 5.9565 | 5.9653 | 5.9573 | 5.9619 | 5.9639 | 5.9704 | 5.9724 | 5.9717 | 5.964 | 5.9706 | 5.9713 | 5.9817 | 5.9821 | 5.9815 | 5.9738 | 5.9828 | 5.9828 | 5.9859 | 5.9887 | 5.991 | 5.978 | 0.014223 | 0.0023793
36 | 5.9792 | 5.9891 | 5.9543 | 5.9536 | 5.9744 | 5.9781 | 5.9772 | 5.9894 | 5.9892 | 5.9357 | 5.9421 | 5.9509 | 5.943 | 5.9475 | 5.9496 | 5.9561 | 5.9581 | 5.9574 | 5.9497 | 5.9563 | 5.957 | 5.9674 | 5.9678 | 5.9672 | 5.9595 | 5.9685 | 5.9686 | 5.9715 | 5.9743 | 5.9766 | 5.9636 | 0.014231 | 0.0023863
37 | 5.9648 | 5.9753 | 5.9403 | 5.9396 | 5.9605 | 5.9642 | 5.9632 | 5.9755 | 5.9753 | 5.9215 | 5.9281 | 5.937 | 5.929 | 5.9336 | 5.9356 | 5.9421 | 5.9442 | 5.9435 | 5.9359 | 5.9424 | 5.943 | 5.9535 | 5.9539 | 5.9533 | 5.9455 | 5.9546 | 5.9547 | 5.9575 | 5.9603 | 5.9626 | 5.9497 | 0.014241 | 0.0023935
38 | 5.9508 | 5.9618 | 5.9267 | 5.926 | 5.9469 | 5.9506 | 5.9497 | 5.962 | 5.9618 | 5.9077 | 5.9144 | 5.9234 | 5.9154 | 5.92 | 5.922 | 5.9286 | 5.9307 | 5.9299 | 5.9223 | 5.9288 | 5.9295 | 5.9399 | 5.9404 | 5.9398 | 5.932 | 5.9411 | 5.9412 | 5.9439 | 5.9467 | 5.949 | 5.9361 | 0.014252 | 0.0024009
39 | 5.9371 | 5.9487 | 5.9135 | 5.9128 | 5.9337 | 5.9373 | 5.9365 | 5.9489 | 5.9487 | 5.8943 | 5.9012 | 5.9102 | 5.9022 | 5.9069 | 5.9088 | 5.9154 | 5.9175 | 5.9167 | 5.9092 | 5.9156 | 5.9162 | 5.9268 | 5.9273 | 5.9267 | 5.9188 | 5.9279 | 5.928 | 5.9307 | 5.9334 | 5.9358 | 5.9229 | 0.014264 | 0.0024084
40 | 5.9239 | 5.9359 | 5.9006 | 5.8999 | 5.9209 | 5.9244 | 5.9236 | 5.936 | 5.9359 | 5.8812 | 5.8882 | 5.8974 | 5.8893 | 5.894 | 5.896 | 5.9025 | 5.9047 | 5.9039 | 5.8964 | 5.9028 | 5.9034 | 5.914 | 5.9144 | 5.9139 | 5.906 | 5.9151 | 5.9152 | 5.9178 | 5.9205 | 5.9228 | 5.91 | 0.014278 | 0.0024159
41 | 5.911 | 5.9234 | 5.888 | 5.8873 | 5.9083 | 5.9119 | 5.9111 | 5.9236 | 5.9234 | 5.8684 | 5.8757 | 5.8848 | 5.8767 | 5.8815 | 5.8834 | 5.89 | 5.8922 | 5.8914 | 5.8839 | 5.8903 | 5.8909 | 5.9015 | 5.902 | 5.9014 | 5.8935 | 5.9026 | 5.9027 | 5.9052 | 5.908 | 5.9103 | 5.8975 | 0.014293 | 0.0024236
42 | 5.8984 | 5.9113 | 5.8758 | 5.875 | 5.8961 | 5.8996 | 5.8989 | 5.9114 | 5.9112 | 5.856 | 5.8634 | 5.8726 | 5.8645 | 5.8693 | 5.8712 | 5.8778 | 5.88 | 5.8792 | 5.8718 | 5.8781 | 5.8786 | 5.8893 | 5.8898 | 5.8893 | 5.8813 | 5.8904 | 5.8906 | 5.893 | 5.8957 | 5.898 | 5.8853 | 0.014309 | 0.0024314
43 | 5.8862 | 5.8995 | 5.8638 | 5.8631 | 5.8842 | 5.8877 | 5.887 | 5.8996 | 5.8994 | 5.8439 | 5.8514 | 5.8607 | 5.8526 | 5.8574 | 5.8593 | 5.8658 | 5.8681 | 5.8673 | 5.8599 | 5.8662 | 5.8667 | 5.8774 | 5.8779 | 5.8774 | 5.8694 | 5.8785 | 5.8787 | 5.881 | 5.8838 | 5.886 | 5.8733 | 0.014326 | 0.0024392
44 | 5.8743 | 5.8879 | 5.8522 | 5.8514 | 5.8726 | 5.876 | 5.8754 | 5.888 | 5.8878 | 5.8321 | 5.8397 | 5.8491 | 5.8409 | 5.8457 | 5.8476 | 5.8542 | 5.8565 | 5.8557 | 5.8483 | 5.8546 | 5.8551 | 5.8659 | 5.8663 | 5.8659 | 5.8578 | 5.8669 | 5.8671 | 5.8694 | 5.8721 | 5.8743 | 5.8617 | 0.014344 | 0.002447
45 | 5.8627 | 5.8766 | 5.8408 | 5.84 | 5.8612 | 5.8646 | 5.864 | 5.8767 | 5.8765 | 5.8205 | 5.8283 | 5.8377 | 5.8295 | 5.8344 | 5.8363 | 5.8428 | 5.8452 | 5.8444 | 5.837 | 5.8433 | 5.8437 | 5.8546 | 5.855 | 5.8546 | 5.8465 | 5.8556 | 5.8558 | 5.858 | 5.8607 | 5.8629 | 5.8503 | 0.014362 | 0.0024549
46 | 5.8514 | 5.8656 | 5.8296 | 5.8288 | 5.8501 | 5.8535 | 5.8529 | 5.8656 | 5.8654 | 5.8093 | 5.8172 | 5.8266 | 5.8184 | 5.8233 | 5.8251 | 5.8317 | 5.8341 | 5.8333 | 5.826 | 5.8322 | 5.8326 | 5.8435 | 5.8439 | 5.8435 | 5.8354 | 5.8445 | 5.8447 | 5.8469 | 5.8496 | 5.8518 | 5.8392 | 0.014381 | 0.0024628
47 | 5.8403 | 5.8548 | 5.8188 | 5.818 | 5.8393 | 5.8426 | 5.8421 | 5.8548 | 5.8546 | 5.7982 | 5.8063 | 5.8157 | 5.8075 | 5.8124 | 5.8143 | 5.8208 | 5.8233 | 5.8225 | 5.8152 | 5.8214 | 5.8217 | 5.8327 | 5.8331 | 5.8327 | 5.8246 | 5.8336 | 5.8339 | 5.836 | 5.8387 | 5.8409 | 5.8284 | 0.0144 | 0.0024707
48 | 5.8296 | 5.8443 | 5.8081 | 5.8073 | 5.8287 | 5.8319 | 5.8315 | 5.8443 | 5.844 | 5.7875 | 5.7957 | 5.8051 | 5.7969 | 5.8018 | 5.8037 | 5.8102 | 5.8127 | 5.8119 | 5.8046 | 5.8108 | 5.8111 | 5.8222 | 5.8225 | 5.8222 | 5.814 | 5.8231 | 5.8233 | 5.8254 | 5.828 | 5.8302 | 5.8177 | 0.01442 | 0.0024787
49 | 5.819 | 5.834 | 5.7977 | 5.7969 | 5.8183 | 5.8215 | 5.8211 | 5.834 | 5.8337 | 5.7769 | 5.7852 | 5.7947 | 5.7865 | 5.7915 | 5.7933 | 5.7998 | 5.8023 | 5.8016 | 5.7943 | 5.8004 | 5.8007 | 5.8118 | 5.8122 | 5.8119 | 5.8036 | 5.8127 | 5.813 | 5.8149 | 5.8176 | 5.8198 | 5.8074 | 0.01444 | 0.0024866
50 | 5.8088 | 5.8239 | 5.7875 | 5.7867 | 5.8081 | 5.8113 | 5.811 | 5.8239 | 5.8236 | 5.7666 | 5.7751 | 5.7846 | 5.7763 | 5.7813 | 5.7831 | 5.7897 | 5.7922 | 5.7914 | 5.7841 | 5.7903 | 5.7906 | 5.8018 | 5.802 | 5.8018 | 5.7935 | 5.8026 | 5.8028 | 5.8048 | 5.8074 | 5.8096 | 5.7972 | 0.014461 | 0.0024944
51 | 5.7988 | 5.814 | 5.7775 | 5.7767 | 5.7982 | 5.8014 | 5.8011 | 5.814 | 5.8137 | 5.7565 | 5.7651 | 5.7746 | 5.7663 | 5.7713 | 5.7732 | 5.7797 | 5.7822 | 5.7815 | 5.7742 | 5.7803 | 5.7806 | 5.7919 | 5.7921 | 5.7919 | 5.7836 | 5.7926 | 5.7929 | 5.7948 | 5.7974 | 5.7996 | 5.7873 | 0.014481 | 0.0025023
52 | 5.789 | 5.8044 | 5.7677 | 5.7669 | 5.7884 | 5.7916 | 5.7913 | 5.8043 | 5.804 | 5.7467 | 5.7553 | 5.7649 | 5.7566 | 5.7616 | 5.7634 | 5.7699 | 5.7725 | 5.7718 | 5.7645 | 5.7706 | 5.7709 | 5.7822 | 5.7824 | 5.7822 | 5.7739 | 5.7829 | 5.7832 | 5.785 | 5.7877 | 5.7898 | 5.7775 | 0.014502 | 0.0025101
53 | 5.7794 | 5.7949 | 5.7581 | 5.7573 | 5.7789 | 5.782 | 5.7818 | 5.7948 | 5.7945 | 5.737 | 5.7457 | 5.7553 | 5.747 | 5.7521 | 5.7538 | 5.7604 | 5.763 | 5.7623 | 5.755 | 5.7611 | 5.7613 | 5.7727 | 5.7729 | 5.7727 | 5.7643 | 5.7734 | 5.7737 | 5.7755 | 5.7781 | 5.7802 | 5.768 | 0.014523 | 0.0025179
54 | 5.77 | 5.7856 | 5.7487 | 5.748 | 5.7695 | 5.7726 | 5.7724 | 5.7855 | 5.7852 | 5.7275 | 5.7364 | 5.746 | 5.7376 | 5.7427 | 5.7445 | 5.751 | 5.7536 | 5.7529 | 5.7457 | 5.7517 | 5.7519 | 5.7634 | 5.7636 | 5.7634 | 5.755 | 5.764 | 5.7644 | 5.7661 | 5.7687 | 5.7708 | 5.7586 | 0.014544 | 0.0025257
55 | 5.7609 | 5.7765 | 5.7395 | 5.7388 | 5.7604 | 5.7634 | 5.7633 | 5.7764 | 5.776 | 5.7183 | 5.7272 | 5.7368 | 5.7284 | 5.7335 | 5.7353 | 5.7418 | 5.7444 | 5.7438 | 5.7366 | 5.7426 | 5.7428 | 5.7543 | 5.7544 | 5.7543 | 5.7459 | 5.7549 | 5.7552 | 5.7569 | 5.7595 | 5.7616 | 5.7495 | 0.014566 | 0.0025334
56 | 5.752 | 5.7676 | 5.7305 | 5.7297 | 5.7514 | 5.7544 | 5.7543 | 5.7674 | 5.7671 | 5.7092 | 5.7182 | 5.7278 | 5.7194 | 5.7245 | 5.7263 | 5.7328 | 5.7355 | 5.7348 | 5.7276 | 5.7336 | 5.7338 | 5.7454 | 5.7455 | 5.7453 | 5.7369 | 5.7459 | 5.7463 | 5.7479 | 5.7505 | 5.7526 | 5.7405 | 0.014587 | 0.0025411
57 | 5.7432 | 5.7588 | 5.7216 | 5.7209 | 5.7426 | 5.7456 | 5.7455 | 5.7587 | 5.7583 | 5.7002 | 5.7093 | 5.719 | 5.7106 | 5.7157 | 5.7175 | 5.724 | 5.7266 | 5.726 | 5.7189 | 5.7248 | 5.7249 | 5.7366 | 5.7367 | 5.7366 | 5.7281 | 5.7371 | 5.7375 | 5.739 | 5.7416 | 5.7437 | 5.7316 | 0.014608 | 0.0025487
58 | 5.7346 | 5.7503 | 5.7129 | 5.7122 | 5.7339 | 5.7369 | 5.7369 | 5.7501 | 5.7497 | 5.6915 | 5.7006 | 5.7103 | 5.7019 | 5.707 | 5.7088 | 5.7153 | 5.718 | 5.7174 | 5.7102 | 5.7161 | 5.7163 | 5.7281 | 5.728 | 5.728 | 5.7194 | 5.7284 | 5.7288 | 5.7303 | 5.733 | 5.735 | 5.723 | 0.01463 | 0.0025563
59 | 5.7263 | 5.7418 | 5.7044 | 5.7037 | 5.7254 | 5.7283 | 5.7284 | 5.7416 | 5.7412 | 5.6829 | 5.6921 | 5.7018 | 5.6934 | 5.6985 | 5.7003 | 5.7068 | 5.7095 | 5.7089 | 5.7018 | 5.7076 | 5.7078 | 5.7196 | 5.7196 | 5.7195 | 5.711 | 5.7199 | 5.7203 | 5.7218 | 5.7244 | 5.7265 | 5.7145 | 0.014651 | 0.0025638
60 | 5.7181 | 5.7336 | 5.696 | 5.6953 | 5.7171 | 5.72 | 5.72 | 5.7334 | 5.7329 | 5.6745 | 5.6838 | 5.6934 | 5.685 | 5.6902 | 5.6919 | 5.6985 | 5.7012 | 5.7006 | 5.6935 | 5.6993 | 5.6994 | 5.7114 | 5.7113 | 5.7112 | 5.7026 | 5.7116 | 5.712 | 5.7135 | 5.7161 | 5.7181 | 5.7062 | 0.014672 | 0.0025713
61 | 5.71 | 5.7254 | 5.6877 | 5.6871 | 5.7089 | 5.7117 | 5.7119 | 5.7252 | 5.7248 | 5.6663 | 5.6756 | 5.6852 | 5.6768 | 5.682 | 5.6837 | 5.6902 | 5.693 | 5.6924 | 5.6853 | 5.6911 | 5.6912 | 5.7032 | 5.7031 | 5.7031 | 5.6945 | 5.7034 | 5.7038 | 5.7053 | 5.7078 | 5.7099 | 5.698 | 0.014694 | 0.0025788
62 | 5.7022 | 5.7175 | 5.6797 | 5.679 | 5.7008 | 5.7037 | 5.7038 | 5.7172 | 5.7167 | 5.6582 | 5.6675 | 5.6772 | 5.6688 | 5.6739 | 5.6757 | 5.6822 | 5.6849 | 5.6844 | 5.6773 | 5.6831 | 5.6831 | 5.6953 | 5.6951 | 5.6951 | 5.6864 | 5.6954 | 5.6958 | 5.6972 | 5.6998 | 5.7018 | 5.69 | 0.014715 | 0.0025862
63 | 5.6945 | 5.7096 | 5.6717 | 5.6711 | 5.6929 | 5.6957 | 5.6959 | 5.7094 | 5.7089 | 5.6502 | 5.6596 | 5.6693 | 5.6608 | 5.666 | 5.6677 | 5.6743 | 5.677 | 5.6765 | 5.6694 | 5.6752 | 5.6752 | 5.6874 | 5.6872 | 5.6872 | 5.6786 | 5.6875 | 5.6879 | 5.6893 | 5.6918 | 5.6939 | 5.6821 | 0.014737 | 0.0025936
64 | 5.6869 | 5.7019 | 5.6639 | 5.6633 | 5.6851 | 5.6879 | 5.6882 | 5.7017 | 5.7011 | 5.6424 | 5.6518 | 5.6615 | 5.6531 | 5.6582 | 5.66 | 5.6665 | 5.6693 | 5.6688 | 5.6617 | 5.6674 | 5.6675 | 5.6797 | 5.6794 | 5.6795 | 5.6708 | 5.6797 | 5.6802 | 5.6815 | 5.684 | 5.6861 | 5.6743 | 0.014759 | 0.0026009
65 | 5.6795 | 5.6944 | 5.6562 | 5.6556 | 5.6775 | 5.6803 | 5.6806 | 5.6941 | 5.6935 | 5.6347 | 5.6442 | 5.6538 | 5.6454 | 5.6506 | 5.6523 | 5.6588 | 5.6616 | 5.6611 | 5.6541 | 5.6598 | 5.6598 | 5.6722 | 5.6718 | 5.6719 | 5.6632 | 5.6721 | 5.6726 | 5.6738 | 5.6764 | 5.6784 | 5.6667 | 0.01478 | 0.0026083
66 | 5.6723 | 5.6869 | 5.6486 | 5.6481 | 5.67 | 5.6727 | 5.6731 | 5.6866 | 5.6861 | 5.6272 | 5.6367 | 5.6463 | 5.6379 | 5.6431 | 5.6448 | 5.6513 | 5.6541 | 5.6537 | 5.6466 | 5.6523 | 5.6523 | 5.6647 | 5.6643 | 5.6645 | 5.6557 | 5.6646 | 5.6651 | 5.6663 | 5.6688 | 5.6708 | 5.6592 | 0.014802 | 0.0026156
67 | 5.6651 | 5.6796 | 5.6412 | 5.6407 | 5.6626 | 5.6653 | 5.6657 | 5.6793 | 5.6787 | 5.6198 | 5.6293 | 5.6389 | 5.6305 | 5.6357 | 5.6374 | 5.6439 | 5.6467 | 5.6463 | 5.6393 | 5.6449 | 5.6449 | 5.6574 | 5.657 | 5.6571 | 5.6483 | 5.6572 | 5.6577 | 5.6589 | 5.6614 | 5.6634 | 5.6518 | 0.014824 | 0.0026228
68 | 5.6582 | 5.6724 | 5.6339 | 5.6334 | 5.6554 | 5.6581 | 5.6585 | 5.6721 | 5.6715 | 5.6125 | 5.622 | 5.6317 | 5.6232 | 5.6284 | 5.6301 | 5.6366 | 5.6395 | 5.639 | 5.632 | 5.6376 | 5.6376 | 5.6503 | 5.6497 | 5.6499 | 5.6411 | 5.6499 | 5.6505 | 5.6516 | 5.6541 | 5.6562 | 5.6446 | 0.014846 | 0.0026301
69 | 5.6513 | 5.6654 | 5.6267 | 5.6263 | 5.6482 | 5.6509 | 5.6513 | 5.665 | 5.6644 | 5.6054 | 5.6149 | 5.6245 | 5.616 | 5.6213 | 5.623 | 5.6295 | 5.6323 | 5.6319 | 5.6249 | 5.6305 | 5.6305 | 5.6432 | 5.6426 | 5.6428 | 5.634 | 5.6428 | 5.6433 | 5.6445 | 5.647 | 5.649 | 5.6374 | 0.014868 | 0.0026373
70 | 5.6446 | 5.6584 | 5.6196 | 5.6192 | 5.6412 | 5.6438 | 5.6443 | 5.658 | 5.6574 | 5.5983 | 5.6079 | 5.6175 | 5.609 | 5.6142 | 5.6159 | 5.6225 | 5.6253 | 5.6249 | 5.6179 | 5.6235 | 5.6234 | 5.6362 | 5.6356 | 5.6359 | 5.627 | 5.6358 | 5.6363 | 5.6374 | 5.6399 | 5.6419 | 5.6304 | 0.01489 | 0.0026445
71 | 5.638 | 5.6515 | 5.6127 | 5.6123 | 5.6343 | 5.6369 | 5.6374 | 5.6512 | 5.6505 | 5.5914 | 5.601 | 5.6106 | 5.6021 | 5.6073 | 5.609 | 5.6155 | 5.6184 | 5.618 | 5.611 | 5.6166 | 5.6165 | 5.6294 | 5.6287 | 5.629 | 5.6201 | 5.6289 | 5.6294 | 5.6305 | 5.633 | 5.635 | 5.6235 | 0.014912 | 0.0026517
72 | 5.6316 | 5.6448 | 5.6058 | 5.6055 | 5.6275 | 5.6301 | 5.6306 | 5.6444 | 5.6437 | 5.5846 | 5.5941 | 5.6037 | 5.5952 | 5.6005 | 5.6022 | 5.6087 | 5.6116 | 5.6112 | 5.6043 | 5.6098 | 5.6097 | 5.6227 | 5.6219 | 5.6222 | 5.6133 | 5.6221 | 5.6226 | 5.6237 | 5.6261 | 5.6282 | 5.6167 | 0.014935 | 0.002659
73 | 5.6252 | 5.6382 | 5.5991 | 5.5988 | 5.6208 | 5.6233 | 5.6239 | 5.6378 | 5.637 | 5.5779 | 5.5875 | 5.597 | 5.5885 | 5.5938 | 5.5955 | 5.602 | 5.6049 | 5.6046 | 5.5976 | 5.6031 | 5.603 | 5.616 | 5.6152 | 5.6156 | 5.6066 | 5.6154 | 5.6159 | 5.617 | 5.6194 | 5.6214 | 5.6101 | 0.014957 | 0.0026662
74 | 5.619 | 5.6316 | 5.5924 | 5.5922 | 5.6142 | 5.6167 | 5.6174 | 5.6312 | 5.6305 | 5.5714 | 5.5809 | 5.5904 | 5.5819 | 5.5871 | 5.5889 | 5.5954 | 5.5983 | 5.598 | 5.591 | 5.5965 | 5.5964 | 5.6095 | 5.6086 | 5.609 | 5.6 | 5.6088 | 5.6094 | 5.6103 | 5.6128 | 5.6148 | 5.6035 | 0.01498 | 0.0026734
75 | 5.6129 | 5.6252 | 5.5859 | 5.5857 | 5.6077 | 5.6102 | 5.6109 | 5.6248 | 5.624 | 5.5649 | 5.5744 | 5.5839 | 5.5754 | 5.5806 | 5.5824 | 5.5889 | 5.5918 | 5.5915 | 5.5846 | 5.59 | 5.5899 | 5.6031 | 5.6022 | 5.6026 | 5.5936 | 5.6023 | 5.6029 | 5.6038 | 5.6063 | 5.6083 | 5.597 | 0.015003 | 0.0026806
76 | 5.6069 | 5.6189 | 5.5794 | 5.5793 | 5.6013 | 5.6038 | 5.6045 | 5.6185 | 5.6176 | 5.5585 | 5.568 | 5.5775 | 5.569 | 5.5742 | 5.5759 | 5.5825 | 5.5853 | 5.5851 | 5.5782 | 5.5836 | 5.5835 | 5.5968 | 5.5958 | 5.5962 | 5.5872 | 5.5959 | 5.5965 | 5.5974 | 5.5998 | 5.6019 | 5.5906 | 0.015027 | 0.0026878
77 | 5.601 | 5.6127 | 5.5731 | 5.573 | 5.595 | 5.5975 | 5.5982 | 5.6122 | 5.6113 | 5.5523 | 5.5617 | 5.5712 | 5.5627 | 5.5679 | 5.5696 | 5.5762 | 5.579 | 5.5788 | 5.5719 | 5.5773 | 5.5772 | 5.5906 | 5.5895 | 5.59 | 5.5809 | 5.5896 | 5.5902 | 5.5911 | 5.5935 | 5.5956 | 5.5844 | 0.01505 | 0.002695
78 | 5.5952 | 5.6065 | 5.5668 | 5.5668 | 5.5888 | 5.5912 | 5.592 | 5.6061 | 5.6052 | 5.5461 | 5.5555 | 5.565 | 5.5564 | 5.5617 | 5.5634 | 5.5699 | 5.5728 | 5.5727 | 5.5658 | 5.5711 | 5.571 | 5.5844 | 5.5833 | 5.5838 | 5.5747 | 5.5834 | 5.584 | 5.5849 | 5.5873 | 5.5893 | 5.5782 | 0.015074 | 0.0027023
79 | 5.5895 | 5.6005 | 5.5607 | 5.5606 | 5.5827 | 5.5851 | 5.5859 | 5.6 | 5.5991 | 5.54 | 5.5494 | 5.5589 | 5.5503 | 5.5556 | 5.5573 | 5.5638 | 5.5667 | 5.5666 | 5.5597 | 5.565 | 5.5648 | 5.5784 | 5.5772 | 5.5777 | 5.5686 | 5.5773 | 5.5779 | 5.5788 | 5.5811 | 5.5832 | 5.5721 | 0.015098 | 0.0027096
80 | 5.5839 | 5.5945 | 5.5546 | 5.5546 | 5.5766 | 5.5791 | 5.5799 | 5.594 | 5.5931 | 5.5341 | 5.5434 | 5.5529 | 5.5443 | 5.5495 | 5.5512 | 5.5578 | 5.5607 | 5.5606 | 5.5537 | 5.5589 | 5.5588 | 5.5724 | 5.5712 | 5.5718 | 5.5626 | 5.5713 | 5.5719 | 5.5727 | 5.5751 | 5.5772 | 5.5661 | 0.015122 | 0.0027169
81 | 5.5784 | 5.5886 | 5.5486 | 5.5487 | 5.5707 | 5.5731 | 5.574 | 5.5881 | 5.5872 | 5.5282 | 5.5375 | 5.5469 | 5.5383 | 5.5436 | 5.5453 | 5.5518 | 5.5547 | 5.5546 | 5.5478 | 5.553 | 5.5529 | 5.5666 | 5.5653 | 5.5659 | 5.5567 | 5.5653 | 5.566 | 5.5668 | 5.5691 | 5.5712 | 5.5602 | 0.015147 | 0.0027242
82 | 5.573 | 5.5828 | 5.5427 | 5.5428 | 5.5649 | 5.5672 | 5.5681 | 5.5823 | 5.5813 | 5.5224 | 5.5316 | 5.541 | 5.5324 | 5.5377 | 5.5394 | 5.546 | 5.5489 | 5.5488 | 5.5419 | 5.5472 | 5.547 | 5.5608 | 5.5594 | 5.5601 | 5.5509 | 5.5595 | 5.5601 | 5.5609 | 5.5633 | 5.5653 | 5.5543 | 0.015172 | 0.0027316
83 | 5.5677 | 5.5771 | 5.5369 | 5.537 | 5.5591 | 5.5614 | 5.5624 | 5.5766 | 5.5756 | 5.5167 | 5.5259 | 5.5353 | 5.5266 | 5.5319 | 5.5336 | 5.5402 | 5.5431 | 5.543 | 5.5362 | 5.5414 | 5.5412 | 5.5551 | 5.5537 | 5.5543 | 5.5451 | 5.5537 | 5.5544 | 5.5551 | 5.5575 | 5.5595 | 5.5486 | 0.015197 | 0.002739
84 | 5.5625 | 5.5715 | 5.5311 | 5.5313 | 5.5534 | 5.5557 | 5.5567 | 5.571 | 5.5699 | 5.5111 | 5.5202 | 5.5296 | 5.5209 | 5.5262 | 5.5279 | 5.5345 | 5.5374 | 5.5374 | 5.5305 | 5.5357 | 5.5355 | 5.5495 | 5.548 | 5.5487 | 5.5394 | 5.548 | 5.5487 | 5.5494 | 5.5517 | 5.5538 | 5.5429 | 0.015223 | 0.0027464
85 | 5.5573 | 5.566 | 5.5255 | 5.5257 | 5.5478 | 5.5501 | 5.5511 | 5.5654 | 5.5643 | 5.5055 | 5.5146 | 5.524 | 5.5153 | 5.5206 | 5.5223 | 5.5289 | 5.5318 | 5.5318 | 5.525 | 5.5301 | 5.5299 | 5.5439 | 5.5424 | 5.5431 | 5.5339 | 5.5424 | 5.5431 | 5.5438 | 5.5461 | 5.5482 | 5.5373 | 0.015249 | 0.0027539
86 | 5.5523 | 5.5605 | 5.5199 | 5.5202 | 5.5423 | 5.5446 | 5.5456 | 5.5599 | 5.5588 | 5.5001 | 5.5091 | 5.5184 | 5.5098 | 5.515 | 5.5168 | 5.5233 | 5.5262 | 5.5263 | 5.5195 | 5.5246 | 5.5244 | 5.5385 | 5.5369 | 5.5376 | 5.5283 | 5.5369 | 5.5376 | 5.5383 | 5.5406 | 5.5427 | 5.5318 | 0.015276 | 0.0027614
87 | 5.5473 | 5.5551 | 5.5144 | 5.5148 | 5.5368 | 5.5391 | 5.5402 | 5.5545 | 5.5534 | 5.4947 | 5.5037 | 5.513 | 5.5043 | 5.5096 | 5.5113 | 5.5179 | 5.5208 | 5.5208 | 5.514 | 5.5191 | 5.5189 | 5.5331 | 5.5314 | 5.5322 | 5.5229 | 5.5314 | 5.5321 | 5.5328 | 5.5351 | 5.5372 | 5.5264 | 0.015302 | 0.0027689
88 | 5.5424 | 5.5498 | 5.509 | 5.5094 | 5.5314 | 5.5337 | 5.5348 | 5.5492 | 5.548 | 5.4894 | 5.4983 | 5.5076 | 5.4989 | 5.5042 | 5.5059 | 5.5125 | 5.5154 | 5.5155 | 5.5087 | 5.5137 | 5.5136 | 5.5278 | 5.5261 | 5.5269 | 5.5175 | 5.5261 | 5.5267 | 5.5274 | 5.5297 | 5.5318 | 5.521 | 0.015329 | 0.0027765
89 | 5.5376 | 5.5445 | 5.5036 | 5.5041 | 5.5261 | 5.5284 | 5.5295 | 5.5439 | 5.5427 | 5.4842 | 5.493 | 5.5023 | 5.4936 | 5.4988 | 5.5006 | 5.5072 | 5.5101 | 5.5102 | 5.5034 | 5.5084 | 5.5082 | 5.5226 | 5.5208 | 5.5216 | 5.5123 | 5.5208 | 5.5214 | 5.5221 | 5.5243 | 5.5265 | 5.5158 | 0.015357 | 0.0027842
90 | 5.5329 | 5.5394 | 5.4983 | 5.4988 | 5.5209 | 5.5231 | 5.5243 | 5.5388 | 5.5375 | 5.4791 | 5.4878 | 5.497 | 5.4883 | 5.4936 | 5.4953 | 5.5019 | 5.5048 | 5.505 | 5.4982 | 5.5032 | 5.503 | 5.5174 | 5.5156 | 5.5164 | 5.507 | 5.5155 | 5.5162 | 5.5169 | 5.5191 | 5.5212 | 5.5106 | 0.015385 | 0.0027919
91 | 5.5282 | 5.5343 | 5.4931 | 5.4937 | 5.5157 | 5.5179 | 5.5191 | 5.5336 | 5.5324 | 5.474 | 5.4827 | 5.4919 | 5.4831 | 5.4884 | 5.4901 | 5.4968 | 5.4997 | 5.4998 | 5.4931 | 5.498 | 5.4978 | 5.5123 | 5.5104 | 5.5113 | 5.5019 | 5.5104 | 5.5111 | 5.5117 | 5.5139 | 5.5161 | 5.5054 | 0.015413 | 0.0027996
92 | 5.5236 | 5.5292 | 5.488 | 5.4886 | 5.5106 | 5.5128 | 5.5141 | 5.5286 | 5.5273 | 5.469 | 5.4776 | 5.4868 | 5.478 | 5.4833 | 5.485 | 5.4917 | 5.4946 | 5.4947 | 5.488 | 5.4929 | 5.4927 | 5.5073 | 5.5053 | 5.5063 | 5.4968 | 5.5053 | 5.506 | 5.5066 | 5.5088 | 5.511 | 5.5003 | 0.015442 | 0.0028074
93 | 5.5191 | 5.5243 | 5.4829 | 5.4836 | 5.5056 | 5.5078 | 5.509 | 5.5236 | 5.5223 | 5.4641 | 5.4726 | 5.4817 | 5.473 | 5.4782 | 5.48 | 5.4866 | 5.4895 | 5.4897 | 5.483 | 5.4879 | 5.4877 | 5.5023 | 5.5003 | 5.5013 | 5.4918 | 5.5002 | 5.501 | 5.5016 | 5.5037 | 5.5059 | 5.4953 | 0.015471 | 0.0028153
94 | 5.5146 | 5.5193 | 5.4779 | 5.4786 | 5.5006 | 5.5028 | 5.5041 | 5.5187 | 5.5173 | 5.4592 | 5.4676 | 5.4768 | 5.468 | 5.4732 | 5.475 | 5.4817 | 5.4845 | 5.4848 | 5.4781 | 5.4829 | 5.4827 | 5.4974 | 5.4954 | 5.4964 | 5.4869 | 5.4953 | 5.496 | 5.4966 | 5.4987 | 5.501 | 5.4904 | 0.0155 | 0.0028232
95 | 5.5103 | 5.5145 | 5.4729 | 5.4737 | 5.4957 | 5.4979 | 5.4992 | 5.5138 | 5.5125 | 5.4544 | 5.4628 | 5.4719 | 5.4631 | 5.4683 | 5.4701 | 5.4768 | 5.4796 | 5.4799 | 5.4732 | 5.478 | 5.4778 | 5.4926 | 5.4905 | 5.4915 | 5.482 | 5.4904 | 5.4911 | 5.4917 | 5.4938 | 5.496 | 5.4855 | 0.01553 | 0.0028311
96 | 5.5059 | 5.5097 | 5.468 | 5.4689 | 5.4909 | 5.493 | 5.4944 | 5.5091 | 5.5076 | 5.4497 | 5.458 | 5.467 | 5.4582 | 5.4635 | 5.4652 | 5.4719 | 5.4748 | 5.4751 | 5.4684 | 5.4732 | 5.473 | 5.4878 | 5.4856 | 5.4867 | 5.4772 | 5.4855 | 5.4863 | 5.4868 | 5.4889 | 5.4912 | 5.4807 | 0.01556 | 0.0028391
97 | 5.5017 | 5.505 | 5.4632 | 5.4641 | 5.4861 | 5.4882 | 5.4896 | 5.5043 | 5.5029 | 5.445 | 5.4532 | 5.4623 | 5.4534 | 5.4587 | 5.4605 | 5.4671 | 5.47 | 5.4703 | 5.4636 | 5.4684 | 5.4682 | 5.4831 | 5.4809 | 5.482 | 5.4724 | 5.4808 | 5.4815 | 5.4821 | 5.4841 | 5.4864 | 5.476 | 0.015591 | 0.0028472
98 | 5.4975 | 5.5004 | 5.4584 | 5.4594 | 5.4814 | 5.4835 | 5.4849 | 5.4997 | 5.4982 | 5.4404 | 5.4485 | 5.4576 | 5.4487 | 5.454 | 5.4557 | 5.4624 | 5.4653 | 5.4656 | 5.459 | 5.4637 | 5.4635 | 5.4785 | 5.4762 | 5.4773 | 5.4677 | 5.4761 | 5.4768 | 5.4773 | 5.4794 | 5.4817 | 5.4713 | 0.015622 | 0.0028553
99 | 5.4934 | 5.4958 | 5.4537 | 5.4548 | 5.4768 | 5.4788 | 5.4803 | 5.4951 | 5.4936 | 5.4359 | 5.4439 | 5.4529 | 5.444 | 5.4493 | 5.4511 | 5.4578 | 5.4606 | 5.461 | 5.4543 | 5.4591 | 5.4588 | 5.4739 | 5.4715 | 5.4727 | 5.4631 | 5.4714 | 5.4722 | 5.4727 | 5.4747 | 5.477 | 5.4667 | 0.015653 | 0.0028634
100 | 5.4893 | 5.4912 | 5.4491 | 5.4502 | 5.4722 | 5.4742 | 5.4757 | 5.4905 | 5.489 | 5.4314 | 5.4393 | 5.4483 | 5.4394 | 5.4447 | 5.4464 | 5.4532 | 5.456 | 5.4564 | 5.4498 | 5.4545 | 5.4542 | 5.4694 | 5.467 | 5.4681 | 5.4585 | 5.4668 | 5.4676 | 5.4681 | 5.4701 | 5.4724 | 5.4621 | 0.015685 | 0.0028716
101 | 5.4853 | 5.4868 | 5.4445 | 5.4457 | 5.4676 | 5.4697 | 5.4712 | 5.486 | 5.4845 | 5.427 | 5.4348 | 5.4438 | 5.4349 | 5.4401 | 5.4419 | 5.4486 | 5.4515 | 5.4519 | 5.4453 | 5.4499 | 5.4497 | 5.4649 | 5.4624 | 5.4636 | 5.454 | 5.4623 | 5.4631 | 5.4636 | 5.4655 | 5.4679 | 5.4576 | 0.015717 | 0.0028799
102 | 5.4814 | 5.4823 | 5.44 | 5.4412 | 5.4632 | 5.4652 | 5.4667 | 5.4816 | 5.48 | 5.4227 | 5.4304 | 5.4393 | 5.4304 | 5.4356 | 5.4374 | 5.4442 | 5.447 | 5.4474 | 5.4408 | 5.4454 | 5.4452 | 5.4605 | 5.458 | 5.4592 | 5.4496 | 5.4578 | 5.4586 | 5.4591 | 5.461 | 5.4634 | 5.4531 | 0.01575 | 0.0028882
103 | 5.4775 | 5.478 | 5.4355 | 5.4368 | 5.4587 | 5.4607 | 5.4623 | 5.4772 | 5.4756 | 5.4184 | 5.426 | 5.4349 | 5.4259 | 5.4312 | 5.433 | 5.4397 | 5.4426 | 5.443 | 5.4364 | 5.441 | 5.4408 | 5.4561 | 5.4536 | 5.4548 | 5.4452 | 5.4534 | 5.4542 | 5.4546 | 5.4566 | 5.459 | 5.4488 | 0.015783 | 0.0028966
104 | 5.4737 | 5.4737 | 5.4311 | 5.4324 | 5.4544 | 5.4564 | 5.4579 | 5.4729 | 5.4712 | 5.4142 | 5.4216 | 5.4305 | 5.4216 | 5.4268 | 5.4286 | 5.4354 | 5.4382 | 5.4386 | 5.4321 | 5.4367 | 5.4364 | 5.4518 | 5.4492 | 5.4505 | 5.4408 | 5.449 | 5.4498 | 5.4503 | 5.4522 | 5.4546 | 5.4444 | 0.015816 | 0.002905
105 | 5.4699 | 5.4694 | 5.4267 | 5.4281 | 5.45 | 5.452 | 5.4536 | 5.4686 | 5.4669 | 5.41 | 5.4174 | 5.4262 | 5.4172 | 5.4224 | 5.4242 | 5.431 | 5.4338 | 5.4343 | 5.4278 | 5.4323 | 5.4321 | 5.4476 | 5.4449 | 5.4462 | 5.4365 | 5.4447 | 5.4455 | 5.4459 | 5.4478 | 5.4503 | 5.4401 | 0.01585 | 0.0029135
106 | 5.4661 | 5.4652 | 5.4224 | 5.4239 | 5.4458 | 5.4478 | 5.4494 | 5.4644 | 5.4627 | 5.4058 | 5.4131 | 5.4219 | 5.4129 | 5.4182 | 5.42 | 5.4268 | 5.4296 | 5.4301 | 5.4236 | 5.4281 | 5.4279 | 5.4434 | 5.4406 | 5.442 | 5.4323 | 5.4404 | 5.4412 | 5.4417 | 5.4436 | 5.446 | 5.4359 | 0.015884 | 0.002922
107 | 5.4625 | 5.461 | 5.4181 | 5.4197 | 5.4416 | 5.4435 | 5.4452 | 5.4602 | 5.4585 | 5.4018 | 5.4089 | 5.4177 | 5.4087 | 5.4139 | 5.4157 | 5.4226 | 5.4254 | 5.4259 | 5.4194 | 5.4239 | 5.4236 | 5.4392 | 5.4364 | 5.4378 | 5.4281 | 5.4362 | 5.437 | 5.4375 | 5.4393 | 5.4418 | 5.4317 | 0.015918 | 0.0029305
108 | 5.4588 | 5.4569 | 5.4139 | 5.4155 | 5.4374 | 5.4394 | 5.441 | 5.4561 | 5.4543 | 5.3977 | 5.4048 | 5.4136 | 5.4045 | 5.4097 | 5.4116 | 5.4184 | 5.4212 | 5.4217 | 5.4152 | 5.4197 | 5.4195 | 5.4351 | 5.4323 | 5.4337 | 5.424 | 5.4321 | 5.4329 | 5.4333 | 5.4351 | 5.4376 | 5.4276 | 0.015952 | 0.0029391
109 | 5.4553 | 5.4529 | 5.4098 | 5.4114 | 5.4333 | 5.4352 | 5.4369 | 5.452 | 5.4502 | 5.3938 | 5.4007 | 5.4095 | 5.4004 | 5.4056 | 5.4074 | 5.4143 | 5.4171 | 5.4176 | 5.4112 | 5.4156 | 5.4154 | 5.4311 | 5.4282 | 5.4296 | 5.4199 | 5.428 | 5.4288 | 5.4292 | 5.431 | 5.4335 | 5.4235 | 0.015987 | 0.0029478
110 | 5.4517 | 5.4489 | 5.4057 | 5.4074 | 5.4292 | 5.4312 | 5.4329 | 5.448 | 5.4462 | 5.3899 | 5.3967 | 5.4054 | 5.3963 | 5.4015 | 5.4033 | 5.4102 | 5.413 | 5.4136 | 5.4071 | 5.4115 | 5.4113 | 5.4271 | 5.4241 | 5.4256 | 5.4158 | 5.4239 | 5.4247 | 5.4251 | 5.4269 | 5.4294 | 5.4195 | 0.016022 | 0.0029565
111 | 5.4483 | 5.4449 | 5.4016 | 5.4034 | 5.4252 | 5.4271 | 5.4289 | 5.444 | 5.4422 | 5.386 | 5.3927 | 5.4014 | 5.3923 | 5.3975 | 5.3993 | 5.4062 | 5.409 | 5.4096 | 5.4031 | 5.4075 | 5.4073 | 5.4231 | 5.4201 | 5.4216 | 5.4118 | 5.4199 | 5.4207 | 5.4211 | 5.4229 | 5.4254 | 5.4155 | 0.016058 | 0.0029652
112 | 5.4448 | 5.441 | 5.3976 | 5.3994 | 5.4213 | 5.4231 | 5.4249 | 5.4401 | 5.4382 | 5.3822 | 5.3888 | 5.3974 | 5.3883 | 5.3935 | 5.3953 | 5.4022 | 5.405 | 5.4056 | 5.3992 | 5.4035 | 5.4033 | 5.4192 | 5.4162 | 5.4177 | 5.4079 | 5.4159 | 5.4168 | 5.4171 | 5.4189 | 5.4215 | 5.4115 | 0.016094 | 0.0029739
113 | 5.4414 | 5.4371 | 5.3936 | 5.3955 | 5.4173 | 5.4192 | 5.421 | 5.4362 | 5.4343 | 5.3784 | 5.3849 | 5.3935 | 5.3844 | 5.3896 | 5.3914 | 5.3983 | 5.4011 | 5.4017 | 5.3953 | 5.3996 | 5.3994 | 5.4154 | 5.4123 | 5.4138 | 5.404 | 5.412 | 5.4128 | 5.4132 | 5.4149 | 5.4175 | 5.4076 | 0.01613 | 0.0029827
114 | 5.4381 | 5.4333 | 5.3897 | 5.3917 | 5.4135 | 5.4153 | 5.4172 | 5.4324 | 5.4305 | 5.3747 | 5.3811 | 5.3896 | 5.3805 | 5.3857 | 5.3875 | 5.3944 | 5.3972 | 5.3979 | 5.3915 | 5.3957 | 5.3955 | 5.4116 | 5.4084 | 5.41 | 5.4001 | 5.4081 | 5.409 | 5.4093 | 5.411 | 5.4137 | 5.4038 | 0.016166 | 0.0029916
115 | 5.4348 | 5.4295 | 5.3858 | 5.3879 | 5.4096 | 5.4115 | 5.4134 | 5.4286 | 5.4267 | 5.371 | 5.3773 | 5.3858 | 5.3767 | 5.3818 | 5.3837 | 5.3906 | 5.3933 | 5.3941 | 5.3877 | 5.3919 | 5.3917 | 5.4078 | 5.4046 | 5.4062 | 5.3963 | 5.4043 | 5.4051 | 5.4055 | 5.4072 | 5.4098 | 5.4 | 0.016202 | 0.0030004
116 | 5.4315 | 5.4258 | 5.382 | 5.3841 | 5.4058 | 5.4077 | 5.4096 | 5.4249 | 5.4229 | 5.3674 | 5.3735 | 5.382 | 5.3729 | 5.378 | 5.3799 | 5.3868 | 5.3895 | 5.3903 | 5.3839 | 5.3881 | 5.3879 | 5.4041 | 5.4008 | 5.4025 | 5.3926 | 5.4005 | 5.4014 | 5.4017 | 5.4034 | 5.406 | 5.3962 | 0.016239 | 0.0030093
117 | 5.4283 | 5.4221 | 5.3782 | 5.3804 | 5.4021 | 5.4039 | 5.4059 | 5.4212 | 5.4192 | 5.3638 | 5.3698 | 5.3783 | 5.3691 | 5.3743 | 5.3761 | 5.3831 | 5.3858 | 5.3866 | 5.3802 | 5.3844 | 5.3842 | 5.4004 | 5.3971 | 5.3988 | 5.3889 | 5.3968 | 5.3976 | 5.398 | 5.3996 | 5.4023 | 5.3925 | 0.016276 | 0.0030182
118 | 5.4251 | 5.4185 | 5.3744 | 5.3767 | 5.3984 | 5.4002 | 5.4022 | 5.4175 | 5.4155 | 5.3602 | 5.3661 | 5.3746 | 5.3654 | 5.3705 | 5.3724 | 5.3794 | 5.3821 | 5.3829 | 5.3765 | 5.3807 | 5.3805 | 5.3968 | 5.3934 | 5.3951 | 5.3852 | 5.3931 | 5.3939 | 5.3943 | 5.3959 | 5.3986 | 5.3889 | 0.016313 | 0.0030272
119 | 5.422 | 5.4149 | 5.3707 | 5.3731 | 5.3947 | 5.3966 | 5.3985 | 5.4139 | 5.4118 | 5.3567 | 5.3625 | 5.371 | 5.3617 | 5.3669 | 5.3687 | 5.3757 | 5.3784 | 5.3792 | 5.3729 | 5.377 | 5.3768 | 5.3932 | 5.3897 | 5.3915 | 5.3815 | 5.3894 | 5.3903 | 5.3906 | 5.3922 | 5.3949 | 5.3852 | 0.01635 | 0.0030362
120 | 5.4189 | 5.4113 | 5.3671 | 5.3695 | 5.3911 | 5.3929 | 5.3949 | 5.4103 | 5.4082 | 5.3533 | 5.3589 | 5.3674 | 5.3581 | 5.3632 | 5.3651 | 5.3721 | 5.3748 | 5.3756 | 5.3693 | 5.3734 | 5.3732 | 5.3896 | 5.3861 | 5.3879 | 5.378 | 5.3858 | 5.3867 | 5.387 | 5.3886 | 5.3913 | 5.3817 | 0.016388 | 0.0030451
121 | 5.4159 | 5.4078 | 5.3635 | 5.3659 | 5.3876 | 5.3894 | 5.3914 | 5.4068 | 5.4047 | 5.3499 | 5.3554 | 5.3638 | 5.3545 | 5.3596 | 5.3615 | 5.3685 | 5.3712 | 5.3721 | 5.3658 | 5.3698 | 5.3696 | 5.3861 | 5.3826 | 5.3844 | 5.3744 | 5.3822 | 5.3831 | 5.3834 | 5.385 | 5.3877 | 5.3781 | 0.016426 | 0.0030541
122 | 5.4128 | 5.4043 | 5.3599 | 5.3624 | 5.384 | 5.3858 | 5.3878 | 5.4033 | 5.4011 | 5.3465 | 5.3519 | 5.3603 | 5.351 | 5.3561 | 5.3579 | 5.365 | 5.3677 | 5.3685 | 5.3623 | 5.3663 | 5.3661 | 5.3826 | 5.379 | 5.3809 | 5.3709 | 5.3787 | 5.3796 | 5.3799 | 5.3814 | 5.3842 | 5.3746 | 0.016463 | 0.0030632
123 | 5.4098 | 5.4009 | 5.3563 | 5.3589 | 5.3805 | 5.3823 | 5.3844 | 5.3998 | 5.3977 | 5.3432 | 5.3485 | 5.3568 | 5.3475 | 5.3526 | 5.3544 | 5.3615 | 5.3642 | 5.3651 | 5.3588 | 5.3628 | 5.3626 | 5.3792 | 5.3756 | 5.3774 | 5.3674 | 5.3752 | 5.3761 | 5.3764 | 5.3779 | 5.3807 | 5.3711 | 0.016501 | 0.0030722
124 | 5.4069 | 5.3975 | 5.3528 | 5.3555 | 5.3771 | 5.3788 | 5.3809 | 5.3964 | 5.3942 | 5.3399 | 5.345 | 5.3533 | 5.344 | 5.3491 | 5.351 | 5.3581 | 5.3607 | 5.3616 | 5.3554 | 5.3593 | 5.3592 | 5.3758 | 5.3721 | 5.374 | 5.364 | 5.3718 | 5.3727 | 5.3729 | 5.3744 | 5.3773 | 5.3677 | 0.016539 | 0.0030812
125 | 5.404 | 5.3941 | 5.3494 | 5.3521 | 5.3737 | 5.3754 | 5.3775 | 5.393 | 5.3908 | 5.3366 | 5.3417 | 5.3499 | 5.3406 | 5.3457 | 5.3475 | 5.3547 | 5.3573 | 5.3582 | 5.352 | 5.3559 | 5.3557 | 5.3724 | 5.3687 | 5.3706 | 5.3606 | 5.3683 | 5.3692 | 5.3695 | 5.371 | 5.3738 | 5.3643 | 0.016577 | 0.0030903
126 | 5.4011 | 5.3908 | 5.3459 | 5.3487 | 5.3703 | 5.372 | 5.3741 | 5.3897 | 5.3874 | 5.3334 | 5.3383 | 5.3466 | 5.3372 | 5.3423 | 5.3441 | 5.3513 | 5.3539 | 5.3548 | 5.3486 | 5.3525 | 5.3524 | 5.3691 | 5.3653 | 5.3673 | 5.3572 | 5.365 | 5.3659 | 5.3661 | 5.3676 | 5.3705 | 5.361 | 0.016616 | 0.0030993
127 | 5.3983 | 5.3875 | 5.3426 | 5.3454 | 5.3669 | 5.3687 | 5.3708 | 5.3864 | 5.3841 | 5.3302 | 5.335 | 5.3432 | 5.3338 | 5.3389 | 5.3408 | 5.3479 | 5.3505 | 5.3515 | 5.3453 | 5.3492 | 5.349 | 5.3658 | 5.362 | 5.364 | 5.3539 | 5.3616 | 5.3625 | 5.3628 | 5.3642 | 5.3671 | 5.3577 | 0.016654 | 0.0031084
128 | 5.3955 | 5.3842 | 5.3392 | 5.3421 | 5.3636 | 5.3654 | 5.3675 | 5.3831 | 5.3808 | 5.3271 | 5.3318 | 5.3399 | 5.3305 | 5.3356 | 5.3375 | 5.3446 | 5.3472 | 5.3482 | 5.342 | 5.3459 | 5.3457 | 5.3625 | 5.3587 | 5.3607 | 5.3506 | 5.3583 | 5.3592 | 5.3595 | 5.3609 | 5.3638 | 5.3544 | 0.016692 | 0.0031175
129 | 5.3927 | 5.381 | 5.3359 | 5.3389 | 5.3604 | 5.3621 | 5.3643 | 5.3799 | 5.3776 | 5.324 | 5.3285 | 5.3367 | 5.3272 | 5.3323 | 5.3342 | 5.3414 | 5.3439 | 5.345 | 5.3388 | 5.3426 | 5.3425 | 5.3593 | 5.3554 | 5.3575 | 5.3474 | 5.355 | 5.356 | 5.3562 | 5.3576 | 5.3605 | 5.3512 | 0.01673 | 0.0031265
130 | 5.3899 | 5.3778 | 5.3326 | 5.3357 | 5.3571 | 5.3588 | 5.361 | 5.3767 | 5.3743 | 5.3209 | 5.3253 | 5.3335 | 5.324 | 5.3291 | 5.3309 | 5.3381 | 5.3407 | 5.3417 | 5.3356 | 5.3394 | 5.3392 | 5.3562 | 5.3522 | 5.3543 | 5.3442 | 5.3518 | 5.3528 | 5.353 | 5.3543 | 5.3573 | 5.348 | 0.016769 | 0.0031356
131 | 5.3872 | 5.3747 | 5.3294 | 5.3325 | 5.3539 | 5.3556 | 5.3578 | 5.3735 | 5.3711 | 5.3179 | 5.3222 | 5.3303 | 5.3208 | 5.3258 | 5.3277 | 5.3349 | 5.3375 | 5.3386 | 5.3324 | 5.3362 | 5.336 | 5.353 | 5.349 | 5.3511 | 5.341 | 5.3486 | 5.3496 | 5.3498 | 5.3511 | 5.3541 | 5.3448 | 0.016807 | 0.0031446
132 | 5.3846 | 5.3716 | 5.3262 | 5.3293 | 5.3508 | 5.3525 | 5.3547 | 5.3704 | 5.368 | 5.3149 | 5.319 | 5.3271 | 5.3176 | 5.3227 | 5.3245 | 5.3318 | 5.3343 | 5.3354 | 5.3293 | 5.333 | 5.3329 | 5.3499 | 5.3459 | 5.348 | 5.3379 | 5.3455 | 5.3464 | 5.3466 | 5.3479 | 5.351 | 5.3416 | 0.016846 | 0.0031537
133 | 5.3819 | 5.3685 | 5.323 | 5.3262 | 5.3476 | 5.3493 | 5.3516 | 5.3673 | 5.3649 | 5.3119 | 5.316 | 5.324 | 5.3144 | 5.3195 | 5.3214 | 5.3287 | 5.3312 | 5.3323 | 5.3262 | 5.3299 | 5.3298 | 5.3468 | 5.3428 | 5.3449 | 5.3348 | 5.3423 | 5.3433 | 5.3435 | 5.3448 | 5.3478 | 5.3385 | 0.016884 | 0.0031627
134 | 5.3793 | 5.3654 | 5.3198 | 5.3232 | 5.3445 | 5.3462 | 5.3485 | 5.3642 | 5.3618 | 5.309 | 5.3129 | 5.3209 | 5.3113 | 5.3164 | 5.3183 | 5.3256 | 5.3281 | 5.3292 | 5.3231 | 5.3268 | 5.3267 | 5.3438 | 5.3397 | 5.3418 | 5.3317 | 5.3392 | 5.3402 | 5.3404 | 5.3417 | 5.3447 | 5.3355 | 0.016923 | 0.0031717
135 | 5.3767 | 5.3624 | 5.3167 | 5.3201 | 5.3415 | 5.3432 | 5.3454 | 5.3612 | 5.3587 | 5.3061 | 5.3099 | 5.3179 | 5.3083 | 5.3133 | 5.3152 | 5.3225 | 5.325 | 5.3262 | 5.3201 | 5.3237 | 5.3236 | 5.3408 | 5.3366 | 5.3388 | 5.3286 | 5.3362 | 5.3371 | 5.3374 | 5.3386 | 5.3417 | 5.3324 | 0.016961 | 0.0031807
136 | 5.3741 | 5.3595 | 5.3137 | 5.3171 | 5.3385 | 5.3401 | 5.3424 | 5.3582 | 5.3557 | 5.3032 | 5.3069 | 5.3148 | 5.3052 | 5.3103 | 5.3121 | 5.3195 | 5.3219 | 5.3231 | 5.3171 | 5.3207 | 5.3206 | 5.3378 | 5.3336 | 5.3358 | 5.3256 | 5.3331 | 5.3341 | 5.3343 | 5.3355 | 5.3386 | 5.3294 | 0.016999 | 0.0031897
137 | 5.3716 | 5.3565 | 5.3106 | 5.3141 | 5.3355 | 5.3371 | 5.3394 | 5.3552 | 5.3527 | 5.3004 | 5.3039 | 5.3119 | 5.3022 | 5.3073 | 5.3091 | 5.3165 | 5.3189 | 5.3202 | 5.3141 | 5.3177 | 5.3176 | 5.3348 | 5.3306 | 5.3329 | 5.3227 | 5.3302 | 5.3311 | 5.3313 | 5.3325 | 5.3356 | 5.3265 | 0.017037 | 0.0031986
138 | 5.3691 | 5.3536 | 5.3076 | 5.3112 | 5.3325 | 5.3341 | 5.3365 | 5.3523 | 5.3498 | 5.2976 | 5.301 | 5.3089 | 5.2992 | 5.3043 | 5.3062 | 5.3135 | 5.316 | 5.3172 | 5.3112 | 5.3147 | 5.3146 | 5.3319 | 5.3277 | 5.3299 | 5.3197 | 5.3272 | 5.3282 | 5.3284 | 5.3295 | 5.3327 | 5.3235 | 0.017076 | 0.0032075
139 | 5.3666 | 5.3507 | 5.3046 | 5.3083 | 5.3296 | 5.3312 | 5.3336 | 5.3494 | 5.3468 | 5.2948 | 5.2981 | 5.306 | 5.2963 | 5.3013 | 5.3032 | 5.3106 | 5.313 | 5.3143 | 5.3083 | 5.3118 | 5.3117 | 5.329 | 5.3247 | 5.327 | 5.3168 | 5.3243 | 5.3252 | 5.3254 | 5.3265 | 5.3297 | 5.3206 | 0.017114 | 0.0032165
140 | 5.3642 | 5.3478 | 5.3017 | 5.3054 | 5.3266 | 5.3283 | 5.3307 | 5.3465 | 5.344 | 5.2921 | 5.2953 | 5.3031 | 5.2934 | 5.2984 | 5.3003 | 5.3077 | 5.3101 | 5.3114 | 5.3054 | 5.3089 | 5.3088 | 5.3262 | 5.3218 | 5.3242 | 5.3139 | 5.3214 | 5.3223 | 5.3225 | 5.3236 | 5.3268 | 5.3178 | 0.017151 | 0.0032253
141 | 5.3617 | 5.345 | 5.2988 | 5.3026 | 5.3238 | 5.3254 | 5.3278 | 5.3437 | 5.3411 | 5.2894 | 5.2924 | 5.3002 | 5.2905 | 5.2955 | 5.2974 | 5.3048 | 5.3072 | 5.3085 | 5.3026 | 5.306 | 5.3059 | 5.3234 | 5.319 | 5.3213 | 5.3111 | 5.3185 | 5.3195 | 5.3197 | 5.3207 | 5.324 | 5.3149 | 0.017189 | 0.0032342
142 | 5.3593 | 5.3422 | 5.2959 | 5.2997 | 5.3209 | 5.3225 | 5.325 | 5.3409 | 5.3383 | 5.2867 | 5.2896 | 5.2974 | 5.2876 | 5.2926 | 5.2945 | 5.302 | 5.3044 | 5.3057 | 5.2997 | 5.3032 | 5.3031 | 5.3206 | 5.3161 | 5.3185 | 5.3082 | 5.3156 | 5.3166 | 5.3168 | 5.3179 | 5.3211 | 5.3121 | 0.017227 | 0.003243
143 | 5.357 | 5.3395 | 5.293 | 5.2969 | 5.3181 | 5.3197 | 5.3222 | 5.3381 | 5.3355 | 5.284 | 5.2869 | 5.2946 | 5.2848 | 5.2898 | 5.2917 | 5.2992 | 5.3015 | 5.3029 | 5.297 | 5.3003 | 5.3003 | 5.3178 | 5.3133 | 5.3157 | 5.3055 | 5.3128 | 5.3138 | 5.314 | 5.315 | 5.3183 | 5.3093 | 0.017265 | 0.0032518
144 | 5.3546 | 5.3367 | 5.2902 | 5.2942 | 5.3153 | 5.3169 | 5.3194 | 5.3353 | 5.3327 | 5.2814 | 5.2841 | 5.2918 | 5.282 | 5.287 | 5.2889 | 5.2964 | 5.2988 | 5.3001 | 5.2942 | 5.2976 | 5.2975 | 5.3151 | 5.3106 | 5.313 | 5.3027 | 5.31 | 5.3111 | 5.3112 | 5.3122 | 5.3155 | 5.3066 | 0.017302 | 0.0032605
145 | 5.3523 | 5.334 | 5.2874 | 5.2915 | 5.3126 | 5.3142 | 5.3166 | 5.3326 | 5.3299 | 5.2788 | 5.2814 | 5.2891 | 5.2792 | 5.2842 | 5.2861 | 5.2937 | 5.296 | 5.2974 | 5.2915 | 5.2948 | 5.2948 | 5.3124 | 5.3078 | 5.3103 | 5.3 | 5.3073 | 5.3083 | 5.3085 | 5.3094 | 5.3128 | 5.3038 | 0.017339 | 0.0032692
146 | 5.35 | 5.3313 | 5.2846 | 5.2888 | 5.3098 | 5.3114 | 5.3139 | 5.3299 | 5.3272 | 5.2763 | 5.2787 | 5.2864 | 5.2765 | 5.2815 | 5.2834 | 5.2909 | 5.2933 | 5.2946 | 5.2888 | 5.2921 | 5.292 | 5.3097 | 5.3051 | 5.3076 | 5.2973 | 5.3046 | 5.3056 | 5.3057 | 5.3067 | 5.31 | 5.3011 | 0.017376 | 0.0032778
147 | 5.3477 | 5.3287 | 5.2819 | 5.2861 | 5.3072 | 5.3087 | 5.3112 | 5.3273 | 5.3245 | 5.2737 | 5.2761 | 5.2837 | 5.2738 | 5.2788 | 5.2807 | 5.2882 | 5.2905 | 5.292 | 5.2861 | 5.2894 | 5.2893 | 5.307 | 5.3024 | 5.3049 | 5.2946 | 5.3019 | 5.3029 | 5.303 | 5.304 | 5.3073 | 5.2985 | 0.017413 | 0.0032864
148 | 5.3455 | 5.3261 | 5.2792 | 5.2834 | 5.3045 | 5.3061 | 5.3086 | 5.3246 | 5.3219 | 5.2712 | 5.2734 | 5.2811 | 5.2711 | 5.2761 | 5.278 | 5.2856 | 5.2879 | 5.2893 | 5.2835 | 5.2867 | 5.2867 | 5.3044 | 5.2998 | 5.3023 | 5.2919 | 5.2992 | 5.3002 | 5.3004 | 5.3013 | 5.3047 | 5.2958 | 0.01745 | 0.003295
149 | 5.3432 | 5.3235 | 5.2765 | 5.2808 | 5.3018 | 5.3034 | 5.3059 | 5.322 | 5.3192 | 5.2687 | 5.2708 | 5.2784 | 5.2685 | 5.2734 | 5.2753 | 5.2829 | 5.2852 | 5.2867 | 5.2808 | 5.2841 | 5.284 | 5.3018 | 5.2971 | 5.2997 | 5.2893 | 5.2966 | 5.2976 | 5.2977 | 5.2986 | 5.302 | 5.2932 | 0.017486 | 0.0033035
150 | 5.341 | 5.3209 | 5.2738 | 5.2782 | 5.2992 | 5.3008 | 5.3033 | 5.3194 | 5.3166 | 5.2663 | 5.2683 | 5.2758 | 5.2658 | 5.2708 | 5.2727 | 5.2803 | 5.2826 | 5.2841 | 5.2783 | 5.2814 | 5.2814 | 5.2992 | 5.2945 | 5.2971 | 5.2867 | 5.2939 | 5.295 | 5.2951 | 5.296 | 5.2994 | 5.2906 | 0.017522 | 0.003312
151 | 5.3388 | 5.3183 | 5.2712 | 5.2757 | 5.2966 | 5.2982 | 5.3007 | 5.3168 | 5.314 | 5.2639 | 5.2657 | 5.2733 | 5.2632 | 5.2682 | 5.2701 | 5.2777 | 5.28 | 5.2815 | 5.2757 | 5.2788 | 5.2788 | 5.2967 | 5.2919 | 5.2945 | 5.2841 | 5.2913 | 5.2924 | 5.2925 | 5.2934 | 5.2968 | 5.288 | 0.017558 | 0.0033204
152 | 5.3367 | 5.3158 | 5.2686 | 5.2731 | 5.2941 | 5.2956 | 5.2982 | 5.3143 | 5.3115 | 5.2615 | 5.2632 | 5.2707 | 5.2607 | 5.2656 | 5.2675 | 5.2752 | 5.2774 | 5.2789 | 5.2732 | 5.2763 | 5.2763 | 5.2941 | 5.2894 | 5.292 | 5.2816 | 5.2888 | 5.2898 | 5.29 | 5.2908 | 5.2943 | 5.2855 | 0.017594 | 0.0033288
153 | 5.3345 | 5.3133 | 5.266 | 5.2706 | 5.2915 | 5.2931 | 5.2957 | 5.3118 | 5.3089 | 5.2591 | 5.2607 | 5.2682 | 5.2581 | 5.2631 | 5.265 | 5.2727 | 5.2749 | 5.2764 | 5.2707 | 5.2737 | 5.2737 | 5.2917 | 5.2869 | 5.2895 | 5.2791 | 5.2862 | 5.2873 | 5.2874 | 5.2882 | 5.2917 | 5.283 | 0.01763 | 0.0033371
154 | 5.3324 | 5.3109 | 5.2634 | 5.2681 | 5.289 | 5.2905 | 5.2931 | 5.3093 | 5.3064 | 5.2567 | 5.2582 | 5.2657 | 5.2556 | 5.2605 | 5.2624 | 5.2702 | 5.2723 | 5.2739 | 5.2682 | 5.2712 | 5.2712 | 5.2892 | 5.2844 | 5.287 | 5.2766 | 5.2837 | 5.2848 | 5.2849 | 5.2857 | 5.2892 | 5.2805 | 0.017665 | 0.0033453
155 | 5.3303 | 5.3084 | 5.2609 | 5.2657 | 5.2865 | 5.2881 | 5.2907 | 5.3069 | 5.304 | 5.2544 | 5.2558 | 5.2632 | 5.2531 | 5.258 | 5.2599 | 5.2677 | 5.2698 | 5.2714 | 5.2657 | 5.2687 | 5.2688 | 5.2867 | 5.2819 | 5.2846 | 5.2741 | 5.2812 | 5.2823 | 5.2824 | 5.2832 | 5.2867 | 5.278 | 0.0177 | 0.0033535
156 | 5.3282 | 5.306 | 5.2584 | 5.2632 | 5.2841 | 5.2856 | 5.2882 | 5.3044 | 5.3015 | 5.2521 | 5.2534 | 5.2608 | 5.2506 | 5.2556 | 5.2575 | 5.2652 | 5.2674 | 5.269 | 5.2633 | 5.2663 | 5.2663 | 5.2843 | 5.2794 | 5.2821 | 5.2717 | 5.2788 | 5.2799 | 5.28 | 5.2807 | 5.2843 | 5.2756 | 0.017735 | 0.0033616
157 | 5.3261 | 5.3036 | 5.2559 | 5.2608 | 5.2816 | 5.2831 | 5.2858 | 5.302 | 5.2991 | 5.2498 | 5.251 | 5.2584 | 5.2482 | 5.2531 | 5.255 | 5.2628 | 5.2649 | 5.2665 | 5.2609 | 5.2638 | 5.2639 | 5.2819 | 5.277 | 5.2797 | 5.2693 | 5.2763 | 5.2774 | 5.2775 | 5.2782 | 5.2818 | 5.2732 | 0.017769 | 0.0033697
158 | 5.3241 | 5.3012 | 5.2535 | 5.2584 | 5.2792 | 5.2807 | 5.2834 | 5.2996 | 5.2967 | 5.2476 | 5.2486 | 5.256 | 5.2458 | 5.2507 | 5.2526 | 5.2604 | 5.2625 | 5.2641 | 5.2585 | 5.2614 | 5.2615 | 5.2795 | 5.2746 | 5.2773 | 5.2669 | 5.2739 | 5.275 | 5.2751 | 5.2758 | 5.2794 | 5.2708 | 0.017803 | 0.0033777
159 | 5.3221 | 5.2989 | 5.251 | 5.256 | 5.2768 | 5.2783 | 5.281 | 5.2973 | 5.2943 | 5.2453 | 5.2463 | 5.2536 | 5.2434 | 5.2483 | 5.2502 | 5.258 | 5.2601 | 5.2617 | 5.2561 | 5.259 | 5.2591 | 5.2772 | 5.2722 | 5.275 | 5.2645 | 5.2715 | 5.2726 | 5.2727 | 5.2734 | 5.277 | 5.2684 | 0.017837 | 0.0033856
160 | 5.3201 | 5.2966 | 5.2486 | 5.2537 | 5.2744 | 5.276 | 5.2786 | 5.2949 | 5.2919 | 5.2431 | 5.2439 | 5.2513 | 5.241 | 5.2459 | 5.2478 | 5.2556 | 5.2577 | 5.2594 | 5.2538 | 5.2566 | 5.2567 | 5.2749 | 5.2699 | 5.2726 | 5.2622 | 5.2692 | 5.2703 | 5.2704 | 5.271 | 5.2747 | 5.2661 | 0.01787 | 0.0033935
161 | 5.3181 | 5.2943 | 5.2462 | 5.2514 | 5.2721 | 5.2736 | 5.2763 | 5.2926 | 5.2896 | 5.2409 | 5.2416 | 5.2489 | 5.2386 | 5.2435 | 5.2454 | 5.2533 | 5.2554 | 5.2571 | 5.2515 | 5.2543 | 5.2544 | 5.2726 | 5.2675 | 5.2703 | 5.2598 | 5.2668 | 5.2679 | 5.268 | 5.2686 | 5.2723 | 5.2638 | 0.017904 | 0.0034013
162 | 5.3161 | 5.292 | 5.2439 | 5.2491 | 5.2698 | 5.2713 | 5.274 | 5.2903 | 5.2873 | 5.2388 | 5.2394 | 5.2466 | 5.2363 | 5.2412 | 5.2431 | 5.251 | 5.253 | 5.2547 | 5.2492 | 5.252 | 5.2521 | 5.2703 | 5.2652 | 5.268 | 5.2575 | 5.2645 | 5.2656 | 5.2657 | 5.2663 | 5.27 | 5.2615 | 0.017936 | 0.003409
163 | 5.3142 | 5.2898 | 5.2415 | 5.2468 | 5.2675 | 5.269 | 5.2717 | 5.288 | 5.285 | 5.2366 | 5.2371 | 5.2444 | 5.234 | 5.2389 | 5.2408 | 5.2487 | 5.2507 | 5.2525 | 5.2469 | 5.2497 | 5.2498 | 5.268 | 5.2629 | 5.2658 | 5.2553 | 5.2622 | 5.2633 | 5.2634 | 5.264 | 5.2677 | 5.2592 | 0.017969 | 0.0034167
164 | 5.3122 | 5.2875 | 5.2392 | 5.2446 | 5.2652 | 5.2667 | 5.2694 | 5.2858 | 5.2827 | 5.2345 | 5.2349 | 5.2421 | 5.2317 | 5.2366 | 5.2385 | 5.2464 | 5.2484 | 5.2502 | 5.2447 | 5.2474 | 5.2475 | 5.2658 | 5.2607 | 5.2635 | 5.253 | 5.26 | 5.2611 | 5.2612 | 5.2617 | 5.2654 | 5.257 | 0.018001 | 0.0034243
165 | 5.3103 | 5.2853 | 5.2369 | 5.2423 | 5.263 | 5.2644 | 5.2672 | 5.2836 | 5.2805 | 5.2324 | 5.2327 | 5.2399 | 5.2295 | 5.2343 | 5.2362 | 5.2442 | 5.2462 | 5.2479 | 5.2424 | 5.2451 | 5.2453 | 5.2636 | 5.2584 | 5.2613 | 5.2508 | 5.2577 | 5.2588 | 5.2589 | 5.2594 | 5.2632 | 5.2547 | 0.018033 | 0.0034318
166 | 5.3084 | 5.2831 | 5.2347 | 5.2401 | 5.2607 | 5.2622 | 5.265 | 5.2814 | 5.2782 | 5.2304 | 5.2305 | 5.2377 | 5.2272 | 5.2321 | 5.234 | 5.242 | 5.2439 | 5.2457 | 5.2402 | 5.2429 | 5.243 | 5.2614 | 5.2562 | 5.2591 | 5.2486 | 5.2555 | 5.2566 | 5.2567 | 5.2572 | 5.261 | 5.2525 | 0.018064 | 0.0034392
167 | 5.3065 | 5.281 | 5.2324 | 5.238 | 5.2585 | 5.26 | 5.2628 | 5.2792 | 5.276 | 5.2283 | 5.2283 | 5.2355 | 5.225 | 5.2299 | 5.2318 | 5.2398 | 5.2417 | 5.2435 | 5.238 | 5.2407 | 5.2408 | 5.2592 | 5.254 | 5.2569 | 5.2464 | 5.2533 | 5.2544 | 5.2545 | 5.2549 | 5.2588 | 5.2503 | 0.018096 | 0.0034465
168 | 5.3047 | 5.2788 | 5.2302 | 5.2358 | 5.2563 | 5.2578 | 5.2606 | 5.277 | 5.2739 | 5.2263 | 5.2262 | 5.2333 | 5.2228 | 5.2277 | 5.2296 | 5.2376 | 5.2395 | 5.2413 | 5.2359 | 5.2385 | 5.2387 | 5.257 | 5.2518 | 5.2548 | 5.2442 | 5.2511 | 5.2522 | 5.2523 | 5.2527 | 5.2566 | 5.2482 | 0.018126 | 0.0034538
169 | 5.3028 | 5.2767 | 5.228 | 5.2337 | 5.2542 | 5.2556 | 5.2584 | 5.2749 | 5.2717 | 5.2243 | 5.2241 | 5.2312 | 5.2206 | 5.2255 | 5.2274 | 5.2354 | 5.2374 | 5.2392 | 5.2337 | 5.2363 | 5.2365 | 5.2549 | 5.2497 | 5.2527 | 5.2421 | 5.2489 | 5.2501 | 5.2501 | 5.2505 | 5.2544 | 5.246 | 0.018157 | 0.003461
170 | 5.301 | 5.2746 | 5.2258 | 5.2315 | 5.252 | 5.2535 | 5.2563 | 5.2727 | 5.2696 | 5.2223 | 5.222 | 5.229 | 5.2185 | 5.2233 | 5.2252 | 5.2333 | 5.2352 | 5.2371 | 5.2316 | 5.2342 | 5.2344 | 5.2528 | 5.2475 | 5.2505 | 5.2399 | 5.2468 | 5.2479 | 5.248 | 5.2484 | 5.2523 | 5.2439 | 0.018187 | 0.0034681
171 | 5.2992 | 5.2725 | 5.2236 | 5.2294 | 5.2499 | 5.2513 | 5.2542 | 5.2706 | 5.2675 | 5.2203 | 5.2199 | 5.2269 | 5.2164 | 5.2212 | 5.2231 | 5.2312 | 5.2331 | 5.2349 | 5.2295 | 5.2321 | 5.2322 | 5.2507 | 5.2454 | 5.2484 | 5.2378 | 5.2446 | 5.2458 | 5.2459 | 5.2462 | 5.2501 | 5.2418 | 0.018216 | 0.0034752
172 | 5.2974 | 5.2704 | 5.2215 | 5.2274 | 5.2478 | 5.2492 | 5.2521 | 5.2686 | 5.2654 | 5.2183 | 5.2178 | 5.2249 | 5.2142 | 5.2191 | 5.221 | 5.2291 | 5.231 | 5.2328 | 5.2275 | 5.23 | 5.2302 | 5.2486 | 5.2433 | 5.2464 | 5.2358 | 5.2425 | 5.2437 | 5.2438 | 5.2441 | 5.248 | 5.2397 | 0.018245 | 0.0034821
173 | 5.2956 | 5.2684 | 5.2194 | 5.2253 | 5.2457 | 5.2471 | 5.25 | 5.2665 | 5.2633 | 5.2164 | 5.2158 | 5.2228 | 5.2122 | 5.217 | 5.2189 | 5.227 | 5.2289 | 5.2308 | 5.2254 | 5.2279 | 5.2281 | 5.2466 | 5.2413 | 5.2443 | 5.2337 | 5.2405 | 5.2416 | 5.2417 | 5.242 | 5.246 | 5.2377 | 0.018274 | 0.003489
174 | 5.2938 | 5.2664 | 5.2173 | 5.2233 | 5.2436 | 5.2451 | 5.2479 | 5.2645 | 5.2612 | 5.2145 | 5.2138 | 5.2208 | 5.2101 | 5.2149 | 5.2168 | 5.2249 | 5.2268 | 5.2287 | 5.2234 | 5.2258 | 5.226 | 5.2446 | 5.2392 | 5.2423 | 5.2317 | 5.2384 | 5.2396 | 5.2396 | 5.2399 | 5.2439 | 5.2356 | 0.018303 | 0.0034958
175 | 5.292 | 5.2644 | 5.2152 | 5.2212 | 5.2416 | 5.243 | 5.2459 | 5.2624 | 5.2592 | 5.2126 | 5.2118 | 5.2187 | 5.208 | 5.2129 | 5.2148 | 5.2229 | 5.2247 | 5.2267 | 5.2213 | 5.2238 | 5.224 | 5.2426 | 5.2372 | 5.2403 | 5.2296 | 5.2364 | 5.2375 | 5.2376 | 5.2379 | 5.2418 | 5.2336 | 0.018331 | 0.0035025
176 | 5.2903 | 5.2624 | 5.2131 | 5.2192 | 5.2396 | 5.241 | 5.2439 | 5.2604 | 5.2572 | 5.2107 | 5.2098 | 5.2167 | 5.206 | 5.2108 | 5.2127 | 5.2209 | 5.2227 | 5.2247 | 5.2193 | 5.2217 | 5.222 | 5.2406 | 5.2352 | 5.2383 | 5.2276 | 5.2343 | 5.2355 | 5.2356 | 5.2358 | 5.2398 | 5.2316 | 0.018358 | 0.0035091
177 | 5.2886 | 5.2604 | 5.2111 | 5.2173 | 5.2376 | 5.239 | 5.2419 | 5.2584 | 5.2552 | 5.2089 | 5.2078 | 5.2148 | 5.204 | 5.2088 | 5.2107 | 5.2189 | 5.2207 | 5.2227 | 5.2174 | 5.2197 | 5.22 | 5.2386 | 5.2332 | 5.2363 | 5.2256 | 5.2323 | 5.2335 | 5.2336 | 5.2338 | 5.2378 | 5.2296 | 0.018385 | 0.0035156
178 | 5.2869 | 5.2585 | 5.2091 | 5.2153 | 5.2356 | 5.237 | 5.2399 | 5.2565 | 5.2532 | 5.207 | 5.2059 | 5.2128 | 5.202 | 5.2068 | 5.2087 | 5.2169 | 5.2187 | 5.2207 | 5.2154 | 5.2177 | 5.218 | 5.2366 | 5.2312 | 5.2343 | 5.2237 | 5.2303 | 5.2315 | 5.2316 | 5.2318 | 5.2358 | 5.2276 | 0.018412 | 0.0035221
179 | 5.2852 | 5.2566 | 5.207 | 5.2134 | 5.2336 | 5.235 | 5.2379 | 5.2545 | 5.2512 | 5.2052 | 5.204 | 5.2109 | 5.2 | 5.2048 | 5.2067 | 5.2149 | 5.2167 | 5.2187 | 5.2135 | 5.2158 | 5.216 | 5.2347 | 5.2292 | 5.2324 | 5.2217 | 5.2284 | 5.2296 | 5.2296 | 5.2298 | 5.2339 | 5.2257 | 0.018438 | 0.0035284
180 | 5.2835 | 5.2546 | 5.2051 | 5.2114 | 5.2317 | 5.2331 | 5.236 | 5.2526 | 5.2493 | 5.2034 | 5.2021 | 5.2089 | 5.1981 | 5.2029 | 5.2048 | 5.213 | 5.2148 | 5.2168 | 5.2115 | 5.2138 | 5.2141 | 5.2328 | 5.2273 | 5.2305 | 5.2198 | 5.2264 | 5.2276 | 5.2277 | 5.2278 | 5.2319 | 5.2238 | 0.018464 | 0.0035347
181 | 5.2818 | 5.2528 | 5.2031 | 5.2095 | 5.2297 | 5.2311 | 5.2341 | 5.2507 | 5.2474 | 5.2016 | 5.2002 | 5.207 | 5.1961 | 5.2009 | 5.2028 | 5.2111 | 5.2128 | 5.2149 | 5.2096 | 5.2119 | 5.2122 | 5.2309 | 5.2254 | 5.2286 | 5.2179 | 5.2245 | 5.2257 | 5.2257 | 5.2259 | 5.23 | 5.2219 | 0.01849 | 0.0035408
182 | 5.2801 | 5.2509 | 5.2011 | 5.2076 | 5.2278 | 5.2292 | 5.2322 | 5.2488 | 5.2454 | 5.1998 | 5.1983 | 5.2051 | 5.1942 | 5.199 | 5.2009 | 5.2092 | 5.2109 | 5.213 | 5.2078 | 5.21 | 5.2103 | 5.229 | 5.2235 | 5.2267 | 5.216 | 5.2226 | 5.2238 | 5.2238 | 5.2239 | 5.2281 | 5.22 | 0.018515 | 0.0035469
183 | 5.2785 | 5.249 | 5.1992 | 5.2058 | 5.2259 | 5.2273 | 5.2303 | 5.2469 | 5.2436 | 5.1981 | 5.1965 | 5.2033 | 5.1923 | 5.1971 | 5.199 | 5.2073 | 5.209 | 5.2111 | 5.2059 | 5.2081 | 5.2084 | 5.2271 | 5.2216 | 5.2248 | 5.2141 | 5.2207 | 5.2219 | 5.2219 | 5.222 | 5.2262 | 5.2181 | 0.018539 | 0.0035529
184 | 5.2769 | 5.2472 | 5.1973 | 5.2039 | 5.224 | 5.2254 | 5.2284 | 5.2451 | 5.2417 | 5.1963 | 5.1946 | 5.2014 | 5.1904 | 5.1952 | 5.1971 | 5.2054 | 5.2071 | 5.2092 | 5.204 | 5.2062 | 5.2065 | 5.2253 | 5.2197 | 5.223 | 5.2123 | 5.2188 | 5.22 | 5.2201 | 5.2201 | 5.2243 | 5.2162 | 0.018564 | 0.0035588
185 | 5.2752 | 5.2454 | 5.1954 | 5.2021 | 5.2222 | 5.2236 | 5.2266 | 5.2432 | 5.2398 | 5.1946 | 5.1928 | 5.1996 | 5.1886 | 5.1933 | 5.1952 | 5.2036 | 5.2052 | 5.2073 | 5.2022 | 5.2043 | 5.2047 | 5.2235 | 5.2179 | 5.2212 | 5.2104 | 5.217 | 5.2182 | 5.2182 | 5.2183 | 5.2224 | 5.2144 | 0.018587 | 0.0035646
186 | 5.2736 | 5.2436 | 5.1935 | 5.2003 | 5.2203 | 5.2217 | 5.2247 | 5.2414 | 5.238 | 5.1929 | 5.191 | 5.1978 | 5.1867 | 5.1915 | 5.1934 | 5.2017 | 5.2034 | 5.2055 | 5.2004 | 5.2025 | 5.2028 | 5.2216 | 5.2161 | 5.2193 | 5.2086 | 5.2151 | 5.2163 | 5.2164 | 5.2164 | 5.2206 | 5.2126 | 0.018611 | 0.0035703
187 | 5.272 | 5.2418 | 5.1917 | 5.1985 | 5.2185 | 5.2199 | 5.2229 | 5.2396 | 5.2362 | 5.1912 | 5.1893 | 5.196 | 5.1849 | 5.1897 | 5.1915 | 5.1999 | 5.2015 | 5.2037 | 5.1986 | 5.2007 | 5.201 | 5.2198 | 5.2143 | 5.2175 | 5.2068 | 5.2133 | 5.2145 | 5.2145 | 5.2146 | 5.2188 | 5.2108 | 0.018633 | 0.0035759
188 | 5.2705 | 5.24 | 5.1898 | 5.1967 | 5.2167 | 5.2181 | 5.2211 | 5.2378 | 5.2344 | 5.1896 | 5.1875 | 5.1942 | 5.1831 | 5.1878 | 5.1897 | 5.1981 | 5.1997 | 5.2019 | 5.1968 | 5.1989 | 5.1992 | 5.2181 | 5.2125 | 5.2158 | 5.205 | 5.2115 | 5.2127 | 5.2127 | 5.2127 | 5.217 | 5.209 | 0.018656 | 0.0035815
189 | 5.2689 | 5.2383 | 5.188 | 5.1949 | 5.2149 | 5.2163 | 5.2193 | 5.236 | 5.2326 | 5.1879 | 5.1858 | 5.1924 | 5.1813 | 5.186 | 5.1879 | 5.1963 | 5.1979 | 5.2001 | 5.195 | 5.1971 | 5.1974 | 5.2163 | 5.2107 | 5.214 | 5.2032 | 5.2097 | 5.2109 | 5.2109 | 5.2109 | 5.2152 | 5.2072 | 0.018678 | 0.0035869
190 | 5.2673 | 5.2365 | 5.1862 | 5.1932 | 5.2131 | 5.2145 | 5.2175 | 5.2343 | 5.2308 | 5.1863 | 5.184 | 5.1907 | 5.1795 | 5.1843 | 5.1861 | 5.1946 | 5.1961 | 5.1983 | 5.1933 | 5.1953 | 5.1957 | 5.2145 | 5.2089 | 5.2123 | 5.2015 | 5.2079 | 5.2092 | 5.2092 | 5.2091 | 5.2134 | 5.2055 | 0.018699 | 0.0035922
191 | 5.2658 | 5.2348 | 5.1844 | 5.1914 | 5.2114 | 5.2127 | 5.2158 | 5.2325 | 5.2291 | 5.1846 | 5.1823 | 5.1889 | 5.1778 | 5.1825 | 5.1844 | 5.1928 | 5.1944 | 5.1966 | 5.1916 | 5.1935 | 5.1939 | 5.2128 | 5.2072 | 5.2105 | 5.1997 | 5.2062 | 5.2074 | 5.2074 | 5.2073 | 5.2117 | 5.2037 | 0.01872 | 0.0035975
192 | 5.2643 | 5.2331 | 5.1826 | 5.1897 | 5.2096 | 5.211 | 5.2141 | 5.2308 | 5.2273 | 5.183 | 5.1806 | 5.1872 | 5.176 | 5.1807 | 5.1826 | 5.1911 | 5.1926 | 5.1949 | 5.1899 | 5.1918 | 5.1922 | 5.2111 | 5.2054 | 5.2088 | 5.198 | 5.2044 | 5.2057 | 5.2057 | 5.2056 | 5.2099 | 5.202 | 0.018741 | 0.0036026
193 | 5.2627 | 5.2314 | 5.1808 | 5.188 | 5.2079 | 5.2093 | 5.2123 | 5.2291 | 5.2256 | 5.1814 | 5.1789 | 5.1855 | 5.1743 | 5.179 | 5.1809 | 5.1894 | 5.1909 | 5.1931 | 5.1882 | 5.19 | 5.1904 | 5.2094 | 5.2037 | 5.2071 | 5.1963 | 5.2027 | 5.204 | 5.2039 | 5.2038 | 5.2082 | 5.2003 | 0.018761 | 0.0036077
194 | 5.2612 | 5.2297 | 5.1791 | 5.1863 | 5.2062 | 5.2076 | 5.2106 | 5.2274 | 5.2239 | 5.1798 | 5.1773 | 5.1838 | 5.1726 | 5.1773 | 5.1792 | 5.1877 | 5.1892 | 5.1914 | 5.1865 | 5.1883 | 5.1887 | 5.2077 | 5.202 | 5.2054 | 5.1946 | 5.201 | 5.2023 | 5.2022 | 5.2021 | 5.2065 | 5.1986 | 0.018781 | 0.0036127
195 | 5.2597 | 5.2281 | 5.1774 | 5.1847 | 5.2045 | 5.2059 | 5.2089 | 5.2257 | 5.2222 | 5.1783 | 5.1756 | 5.1822 | 5.1709 | 5.1756 | 5.1775 | 5.186 | 5.1875 | 5.1897 | 5.1848 | 5.1866 | 5.1871 | 5.206 | 5.2003 | 5.2038 | 5.1929 | 5.1993 | 5.2006 | 5.2005 | 5.2004 | 5.2048 | 5.1969 | 0.0188 | 0.0036175
196 | 5.2582 | 5.2264 | 5.1756 | 5.183 | 5.2028 | 5.2042 | 5.2073 | 5.2241 | 5.2205 | 5.1767 | 5.174 | 5.1805 | 5.1692 | 5.1739 | 5.1758 | 5.1843 | 5.1858 | 5.1881 | 5.1831 | 5.185 | 5.1854 | 5.2044 | 5.1987 | 5.2021 | 5.1913 | 5.1976 | 5.1989 | 5.1989 | 5.1987 | 5.2031 | 5.1953 | 0.018819 | 0.0036223
197 | 5.2567 | 5.2248 | 5.1739 | 5.1814 | 5.2011 | 5.2025 | 5.2056 | 5.2224 | 5.2189 | 5.1752 | 5.1724 | 5.1789 | 5.1675 | 5.1722 | 5.1741 | 5.1827 | 5.1841 | 5.1864 | 5.1815 | 5.1833 | 5.1837 | 5.2027 | 5.197 | 5.2005 | 5.1896 | 5.1959 | 5.1972 | 5.1972 | 5.197 | 5.2014 | 5.1936 | 0.018837 | 0.003627
198 | 5.2553 | 5.2232 | 5.1723 | 5.1797 | 5.1995 | 5.2009 | 5.204 | 5.2208 | 5.2172 | 5.1737 | 5.1708 | 5.1773 | 5.1659 | 5.1706 | 5.1724 | 5.181 | 5.1824 | 5.1848 | 5.1799 | 5.1816 | 5.1821 | 5.2011 | 5.1954 | 5.1988 | 5.188 | 5.1943 | 5.1956 | 5.1956 | 5.1954 | 5.1998 | 5.192 | 0.018855 | 0.0036315
199 | 5.2538 | 5.2216 | 5.1706 | 5.1781 | 5.1979 | 5.1992 | 5.2023 | 5.2192 | 5.2156 | 5.1722 | 5.1692 | 5.1757 | 5.1642 | 5.1689 | 5.1708 | 5.1794 | 5.1808 | 5.1831 | 5.1783 | 5.18 | 5.1805 | 5.1995 | 5.1938 | 5.1972 | 5.1864 | 5.1927 | 5.194 | 5.1939 | 5.1937 | 5.1982 | 5.1904 | 0.018872 | 0.003636
200 | 5.2524 | 5.22 | 5.1689 | 5.1765 | 5.1962 | 5.1976 | 5.2007 | 5.2176 | 5.214 | 5.1707 | 5.1676 | 5.1741 | 5.1626 | 5.1673 | 5.1692 | 5.1778 | 5.1792 | 5.1815 | 5.1767 | 5.1784 | 5.1789 | 5.1979 | 5.1922 | 5.1956 | 5.1848 | 5.191 | 5.1923 | 5.1923 | 5.1921 | 5.1965 | 5.1888 | 0.018889 | 0.0036404
201 | 5.2509 | 5.2185 | 5.1673 | 5.175 | 5.1946 | 5.196 | 5.1991 | 5.216 | 5.2124 | 5.1692 | 5.1661 | 5.1725 | 5.161 | 5.1657 | 5.1675 | 5.1762 | 5.1776 | 5.1799 | 5.1751 | 5.1768 | 5.1773 | 5.1963 | 5.1906 | 5.1941 | 5.1832 | 5.1894 | 5.1907 | 5.1907 | 5.1905 | 5.1949 | 5.1872 | 0.018906 | 0.0036447
202 | 5.2495 | 5.2169 | 5.1657 | 5.1734 | 5.193 | 5.1944 | 5.1975 | 5.2144 | 5.2108 | 5.1677 | 5.1645 | 5.1709 | 5.1594 | 5.1641 | 5.166 | 5.1746 | 5.176 | 5.1784 | 5.1735 | 5.1752 | 5.1757 | 5.1948 | 5.189 | 5.1925 | 5.1816 | 5.1878 | 5.1892 | 5.1891 | 5.1888 | 5.1933 | 5.1856 | 0.018922 | 0.0036489
203 | 5.2481 | 5.2154 | 5.1641 | 5.1718 | 5.1915 | 5.1928 | 5.196 | 5.2129 | 5.2092 | 5.1663 | 5.163 | 5.1694 | 5.1578 | 5.1625 | 5.1644 | 5.173 | 5.1744 | 5.1768 | 5.172 | 5.1736 | 5.1741 | 5.1932 | 5.1874 | 5.191 | 5.1801 | 5.1863 | 5.1876 | 5.1876 | 5.1873 | 5.1918 | 5.184 | 0.018937 | 0.003653
204 | 5.2467 | 5.2138 | 5.1625 | 5.1703 | 5.1899 | 5.1912 | 5.1944 | 5.2113 | 5.2077 | 5.1648 | 5.1615 | 5.1679 | 5.1563 | 5.1609 | 5.1628 | 5.1715 | 5.1728 | 5.1752 | 5.1705 | 5.172 | 5.1726 | 5.1917 | 5.1859 | 5.1894 | 5.1785 | 5.1847 | 5.186 | 5.186 | 5.1857 | 5.1902 | 5.1825 | 0.018952 | 0.003657
205 | 5.2453 | 5.2123 | 5.1609 | 5.1688 | 5.1884 | 5.1897 | 5.1929 | 5.2098 | 5.2061 | 5.1634 | 5.16 | 5.1663 | 5.1547 | 5.1594 | 5.1612 | 5.1699 | 5.1712 | 5.1737 | 5.1689 | 5.1705 | 5.171 | 5.1901 | 5.1843 | 5.1879 | 5.177 | 5.1832 | 5.1845 | 5.1844 | 5.1841 | 5.1886 | 5.181 | 0.018967 | 0.0036609
206 | 5.2439 | 5.2108 | 5.1593 | 5.1673 | 5.1868 | 5.1882 | 5.1913 | 5.2083 | 5.2046 | 5.162 | 5.1585 | 5.1648 | 5.1532 | 5.1578 | 5.1597 | 5.1684 | 5.1697 | 5.1722 | 5.1674 | 5.169 | 5.1695 | 5.1886 | 5.1828 | 5.1864 | 5.1755 | 5.1816 | 5.183 | 5.1829 | 5.1826 | 5.1871 | 5.1794 | 0.018981 | 0.0036647
207 | 5.2425 | 5.2093 | 5.1578 | 5.1658 | 5.1853 | 5.1866 | 5.1898 | 5.2068 | 5.2031 | 5.1606 | 5.157 | 5.1633 | 5.1517 | 5.1563 | 5.1582 | 5.1669 | 5.1682 | 5.1706 | 5.1659 | 5.1674 | 5.168 | 5.1871 | 5.1813 | 5.1849 | 5.174 | 5.1801 | 5.1814 | 5.1814 | 5.181 | 5.1856 | 5.1779 | 0.018995 | 0.0036684
208 | 5.2411 | 5.2079 | 5.1562 | 5.1643 | 5.1838 | 5.1851 | 5.1883 | 5.2053 | 5.2016 | 5.1592 | 5.1556 | 5.1619 | 5.1502 | 5.1548 | 5.1567 | 5.1654 | 5.1667 | 5.1691 | 5.1645 | 5.1659 | 5.1665 | 5.1856 | 5.1798 | 5.1834 | 5.1725 | 5.1786 | 5.1799 | 5.1799 | 5.1795 | 5.1841 | 5.1764 | 0.019008 | 0.003672
209 | 5.2398 | 5.2064 | 5.1547 | 5.1628 | 5.1823 | 5.1836 | 5.1868 | 5.2038 | 5.2001 | 5.1578 | 5.1541 | 5.1604 | 5.1487 | 5.1533 | 5.1552 | 5.1639 | 5.1652 | 5.1677 | 5.163 | 5.1644 | 5.165 | 5.1842 | 5.1783 | 5.1819 | 5.171 | 5.1771 | 5.1785 | 5.1784 | 5.178 | 5.1826 | 5.175 | 0.019021 | 0.0036756
210 | 5.2384 | 5.205 | 5.1532 | 5.1614 | 5.1808 | 5.1821 | 5.1853 | 5.2023 | 5.1986 | 5.1564 | 5.1527 | 5.159 | 5.1472 | 5.1518 | 5.1537 | 5.1625 | 5.1637 | 5.1662 | 5.1615 | 5.1629 | 5.1635 | 5.1827 | 5.1769 | 5.1805 | 5.1695 | 5.1756 | 5.177 | 5.1769 | 5.1765 | 5.1811 | 5.1735 | 0.019033 | 0.003679
211 | 5.2371 | 5.2035 | 5.1517 | 5.1599 | 5.1793 | 5.1807 | 5.1839 | 5.2009 | 5.1972 | 5.1551 | 5.1513 | 5.1575 | 5.1457 | 5.1503 | 5.1522 | 5.161 | 5.1622 | 5.1647 | 5.1601 | 5.1615 | 5.1621 | 5.1813 | 5.1754 | 5.1791 | 5.1681 | 5.1742 | 5.1755 | 5.1755 | 5.175 | 5.1796 | 5.172 | 0.019045 | 0.0036823
212 | 5.2357 | 5.2021 | 5.1502 | 5.1585 | 5.1779 | 5.1792 | 5.1824 | 5.1994 | 5.1957 | 5.1537 | 5.1499 | 5.1561 | 5.1443 | 5.1489 | 5.1507 | 5.1595 | 5.1607 | 5.1633 | 5.1587 | 5.16 | 5.1606 | 5.1798 | 5.174 | 5.1776 | 5.1667 | 5.1727 | 5.1741 | 5.174 | 5.1735 | 5.1782 | 5.1706 | 0.019057 | 0.0036856
213 | 5.2344 | 5.2007 | 5.1487 | 5.1571 | 5.1765 | 5.1778 | 5.181 | 5.198 | 5.1943 | 5.1524 | 5.1485 | 5.1547 | 5.1428 | 5.1474 | 5.1493 | 5.1581 | 5.1593 | 5.1618 | 5.1572 | 5.1586 | 5.1592 | 5.1784 | 5.1726 | 5.1762 | 5.1652 | 5.1713 | 5.1726 | 5.1726 | 5.1721 | 5.1767 | 5.1692 | 0.019068 | 0.0036887
214 | 5.2331 | 5.1993 | 5.1472 | 5.1557 | 5.175 | 5.1763 | 5.1796 | 5.1966 | 5.1929 | 5.1511 | 5.1471 | 5.1533 | 5.1414 | 5.146 | 5.1478 | 5.1567 | 5.1578 | 5.1604 | 5.1558 | 5.1571 | 5.1578 | 5.177 | 5.1711 | 5.1748 | 5.1638 | 5.1698 | 5.1712 | 5.1711 | 5.1706 | 5.1753 | 5.1678 | 0.019078 | 0.0036917
215 | 5.2318 | 5.1979 | 5.1458 | 5.1543 | 5.1736 | 5.1749 | 5.1782 | 5.1952 | 5.1914 | 5.1498 | 5.1457 | 5.1519 | 5.14 | 5.1446 | 5.1464 | 5.1553 | 5.1564 | 5.159 | 5.1545 | 5.1557 | 5.1564 | 5.1756 | 5.1697 | 5.1734 | 5.1624 | 5.1684 | 5.1698 | 5.1697 | 5.1692 | 5.1739 | 5.1664 | 0.019088 | 0.0036947
216 | 5.2305 | 5.1966 | 5.1444 | 5.1529 | 5.1722 | 5.1735 | 5.1768 | 5.1938 | 5.19 | 5.1485 | 5.1444 | 5.1505 | 5.1386 | 5.1432 | 5.145 | 5.1539 | 5.155 | 5.1576 | 5.1531 | 5.1543 | 5.155 | 5.1742 | 5.1684 | 5.172 | 5.161 | 5.167 | 5.1684 | 5.1683 | 5.1678 | 5.1725 | 5.165 | 0.019098 | 0.0036976
217 | 5.2292 | 5.1952 | 5.1429 | 5.1515 | 5.1708 | 5.1721 | 5.1754 | 5.1924 | 5.1887 | 5.1472 | 5.143 | 5.1492 | 5.1372 | 5.1418 | 5.1436 | 5.1525 | 5.1536 | 5.1562 | 5.1517 | 5.1529 | 5.1536 | 5.1728 | 5.167 | 5.1707 | 5.1597 | 5.1656 | 5.167 | 5.1669 | 5.1664 | 5.1711 | 5.1636 | 0.019107 | 0.0037003
218 | 5.2279 | 5.1939 | 5.1415 | 5.1502 | 5.1694 | 5.1707 | 5.174 | 5.1911 | 5.1873 | 5.1459 | 5.1417 | 5.1478 | 5.1358 | 5.1404 | 5.1422 | 5.1511 | 5.1522 | 5.1549 | 5.1504 | 5.1515 | 5.1522 | 5.1715 | 5.1656 | 5.1693 | 5.1583 | 5.1643 | 5.1656 | 5.1656 | 5.165 | 5.1697 | 5.1622 | 0.019116 | 0.003703
219 | 5.2266 | 5.1925 | 5.1401 | 5.1488 | 5.1681 | 5.1694 | 5.1726 | 5.1897 | 5.1859 | 5.1447 | 5.1404 | 5.1465 | 5.1345 | 5.139 | 5.1409 | 5.1498 | 5.1508 | 5.1535 | 5.149 | 5.1502 | 5.1509 | 5.1701 | 5.1643 | 5.168 | 5.1569 | 5.1629 | 5.1643 | 5.1642 | 5.1636 | 5.1684 | 5.1609 | 0.019124 | 0.0037056
220 | 5.2253 | 5.1912 | 5.1387 | 5.1475 | 5.1667 | 5.168 | 5.1713 | 5.1884 | 5.1846 | 5.1434 | 5.1391 | 5.1452 | 5.1331 | 5.1377 | 5.1395 | 5.1484 | 5.1495 | 5.1521 | 5.1477 | 5.1488 | 5.1495 | 5.1688 | 5.1629 | 5.1666 | 5.1556 | 5.1615 | 5.1629 | 5.1628 | 5.1622 | 5.167 | 5.1595 | 0.019132 | 0.003708
221 | 5.2241 | 5.1899 | 5.1374 | 5.1462 | 5.1654 | 5.1667 | 5.1699 | 5.187 | 5.1832 | 5.1422 | 5.1378 | 5.1439 | 5.1318 | 5.1363 | 5.1381 | 5.1471 | 5.1481 | 5.1508 | 5.1464 | 5.1475 | 5.1482 | 5.1675 | 5.1616 | 5.1653 | 5.1543 | 5.1602 | 5.1616 | 5.1615 | 5.1609 | 5.1657 | 5.1582 | 0.019139 | 0.0037104
222 | 5.2228 | 5.1886 | 5.136 | 5.1449 | 5.164 | 5.1653 | 5.1686 | 5.1857 | 5.1819 | 5.141 | 5.1365 | 5.1426 | 5.1304 | 5.135 | 5.1368 | 5.1458 | 5.1468 | 5.1495 | 5.1451 | 5.1461 | 5.1469 | 5.1662 | 5.1603 | 5.164 | 5.153 | 5.1589 | 5.1603 | 5.1602 | 5.1595 | 5.1643 | 5.1569 | 0.019146 | 0.0037127
223 | 5.2216 | 5.1873 | 5.1347 | 5.1436 | 5.1627 | 5.164 | 5.1673 | 5.1844 | 5.1806 | 5.1397 | 5.1352 | 5.1413 | 5.1291 | 5.1337 | 5.1355 | 5.1445 | 5.1455 | 5.1482 | 5.1438 | 5.1448 | 5.1455 | 5.1649 | 5.159 | 5.1627 | 5.1517 | 5.1575 | 5.159 | 5.1589 | 5.1582 | 5.163 | 5.1556 | 0.019153 | 0.0037149
224 | 5.2203 | 5.1861 | 5.1333 | 5.1423 | 5.1614 | 5.1627 | 5.166 | 5.1831 | 5.1793 | 5.1385 | 5.134 | 5.14 | 5.1278 | 5.1323 | 5.1342 | 5.1432 | 5.1441 | 5.1469 | 5.1425 | 5.1435 | 5.1442 | 5.1636 | 5.1577 | 5.1614 | 5.1504 | 5.1562 | 5.1577 | 5.1576 | 5.1569 | 5.1617 | 5.1543 | 0.019159 | 0.0037171
225 | 5.2191 | 5.1848 | 5.132 | 5.141 | 5.1601 | 5.1614 | 5.1647 | 5.1818 | 5.178 | 5.1373 | 5.1327 | 5.1387 | 5.1265 | 5.131 | 5.1329 | 5.1419 | 5.1428 | 5.1456 | 5.1412 | 5.1422 | 5.143 | 5.1623 | 5.1564 | 5.1602 | 5.1491 | 5.1549 | 5.1564 | 5.1563 | 5.1556 | 5.1604 | 5.153 | 0.019164 | 0.0037191
226 | 5.2179 | 5.1835 | 5.1307 | 5.1397 | 5.1588 | 5.1601 | 5.1634 | 5.1806 | 5.1767 | 5.1361 | 5.1315 | 5.1375 | 5.1252 | 5.1298 | 5.1316 | 5.1406 | 5.1415 | 5.1443 | 5.1399 | 5.1409 | 5.1417 | 5.161 | 5.1551 | 5.1589 | 5.1478 | 5.1537 | 5.1551 | 5.155 | 5.1543 | 5.1591 | 5.1517 | 0.01917 | 0.003721
227 | 5.2167 | 5.1823 | 5.1294 | 5.1385 | 5.1575 | 5.1588 | 5.1621 | 5.1793 | 5.1754 | 5.135 | 5.1302 | 5.1362 | 5.124 | 5.1285 | 5.1303 | 5.1393 | 5.1403 | 5.143 | 5.1387 | 5.1396 | 5.1404 | 5.1597 | 5.1539 | 5.1576 | 5.1466 | 5.1524 | 5.1538 | 5.1537 | 5.153 | 5.1578 | 5.1505 | 0.019175 | 0.0037229
228 | 5.2154 | 5.1811 | 5.1281 | 5.1372 | 5.1563 | 5.1576 | 5.1609 | 5.178 | 5.1742 | 5.1338 | 5.129 | 5.135 | 5.1227 | 5.1272 | 5.129 | 5.1381 | 5.139 | 5.1418 | 5.1375 | 5.1384 | 5.1392 | 5.1585 | 5.1526 | 5.1564 | 5.1453 | 5.1511 | 5.1526 | 5.1524 | 5.1517 | 5.1566 | 5.1492 | 0.019179 | 0.0037246
229 | 5.2142 | 5.1798 | 5.1268 | 5.136 | 5.155 | 5.1563 | 5.1596 | 5.1768 | 5.1729 | 5.1326 | 5.1278 | 5.1338 | 5.1215 | 5.126 | 5.1278 | 5.1368 | 5.1377 | 5.1405 | 5.1362 | 5.1371 | 5.1379 | 5.1573 | 5.1514 | 5.1552 | 5.1441 | 5.1499 | 5.1513 | 5.1512 | 5.1504 | 5.1553 | 5.148 | 0.019183 | 0.0037263
230 | 5.213 | 5.1786 | 5.1255 | 5.1348 | 5.1538 | 5.1551 | 5.1584 | 5.1756 | 5.1717 | 5.1315 | 5.1266 | 5.1326 | 5.1202 | 5.1247 | 5.1265 | 5.1356 | 5.1365 | 5.1393 | 5.135 | 5.1359 | 5.1367 | 5.156 | 5.1501 | 5.154 | 5.1428 | 5.1486 | 5.1501 | 5.15 | 5.1492 | 5.1541 | 5.1467 | 0.019186 | 0.0037279
231 | 5.2119 | 5.1774 | 5.1242 | 5.1336 | 5.1525 | 5.1538 | 5.1572 | 5.1743 | 5.1705 | 5.1303 | 5.1255 | 5.1314 | 5.119 | 5.1235 | 5.1253 | 5.1344 | 5.1352 | 5.1381 | 5.1338 | 5.1346 | 5.1355 | 5.1548 | 5.1489 | 5.1527 | 5.1416 | 5.1474 | 5.1488 | 5.1487 | 5.1479 | 5.1528 | 5.1455 | 0.019189 | 0.0037294
232 | 5.2107 | 5.1763 | 5.123 | 5.1324 | 5.1513 | 5.1526 | 5.156 | 5.1731 | 5.1693 | 5.1292 | 5.1243 | 5.1302 | 5.1178 | 5.1223 | 5.124 | 5.1332 | 5.134 | 5.1368 | 5.1326 | 5.1334 | 5.1342 | 5.1536 | 5.1477 | 5.1515 | 5.1404 | 5.1462 | 5.1476 | 5.1475 | 5.1467 | 5.1516 | 5.1443 | 0.019192 | 0.0037308
233 | 5.2095 | 5.1751 | 5.1218 | 5.1312 | 5.1501 | 5.1514 | 5.1547 | 5.1719 | 5.1681 | 5.1281 | 5.1231 | 5.129 | 5.1166 | 5.121 | 5.1228 | 5.132 | 5.1328 | 5.1356 | 5.1314 | 5.1322 | 5.133 | 5.1524 | 5.1465 | 5.1504 | 5.1392 | 5.145 | 5.1464 | 5.1463 | 5.1455 | 5.1504 | 5.1431 | 0.019194 | 0.0037321
234 | 5.2083 | 5.1739 | 5.1205 | 5.13 | 5.1489 | 5.1502 | 5.1536 | 5.1708 | 5.1669 | 5.127 | 5.122 | 5.1278 | 5.1154 | 5.1198 | 5.1216 | 5.1308 | 5.1316 | 5.1344 | 5.1303 | 5.131 | 5.1318 | 5.1512 | 5.1453 | 5.1492 | 5.138 | 5.1438 | 5.1452 | 5.1451 | 5.1443 | 5.1492 | 5.1419 | 0.019196 | 0.0037333
235 | 5.2072 | 5.1728 | 5.1193 | 5.1288 | 5.1477 | 5.149 | 5.1524 | 5.1696 | 5.1657 | 5.1259 | 5.1208 | 5.1267 | 5.1142 | 5.1186 | 5.1204 | 5.1296 | 5.1304 | 5.1333 | 5.1291 | 5.1298 | 5.1307 | 5.15 | 5.1442 | 5.148 | 5.1369 | 5.1426 | 5.144 | 5.1439 | 5.1431 | 5.148 | 5.1407 | 0.019198 | 0.0037344
236 | 5.206 | 5.1716 | 5.1181 | 5.1277 | 5.1465 | 5.1478 | 5.1512 | 5.1684 | 5.1645 | 5.1248 | 5.1197 | 5.1255 | 5.113 | 5.1175 | 5.1192 | 5.1284 | 5.1292 | 5.1321 | 5.1279 | 5.1286 | 5.1295 | 5.1489 | 5.143 | 5.1469 | 5.1357 | 5.1414 | 5.1429 | 5.1427 | 5.1419 | 5.1468 | 5.1396 | 0.019199 | 0.0037355
237 | 5.2049 | 5.1705 | 5.1169 | 5.1265 | 5.1454 | 5.1466 | 5.15 | 5.1672 | 5.1633 | 5.1237 | 5.1186 | 5.1244 | 5.1118 | 5.1163 | 5.1181 | 5.1273 | 5.128 | 5.1309 | 5.1268 | 5.1275 | 5.1283 | 5.1477 | 5.1418 | 5.1457 | 5.1346 | 5.1402 | 5.1417 | 5.1416 | 5.1407 | 5.1456 | 5.1384 | 0.019199 | 0.0037365
238 | 5.2037 | 5.1693 | 5.1157 | 5.1254 | 5.1442 | 5.1455 | 5.1489 | 5.1661 | 5.1622 | 5.1226 | 5.1175 | 5.1233 | 5.1107 | 5.1151 | 5.1169 | 5.1261 | 5.1268 | 5.1298 | 5.1256 | 5.1263 | 5.1272 | 5.1466 | 5.1407 | 5.1446 | 5.1334 | 5.1391 | 5.1405 | 5.1404 | 5.1395 | 5.1445 | 5.1373 | 0.0192 | 0.0037373
239 | 5.2026 | 5.1682 | 5.1145 | 5.1243 | 5.1431 | 5.1443 | 5.1477 | 5.165 | 5.161 | 5.1216 | 5.1164 | 5.1222 | 5.1095 | 5.114 | 5.1158 | 5.125 | 5.1257 | 5.1286 | 5.1245 | 5.1252 | 5.126 | 5.1454 | 5.1396 | 5.1434 | 5.1323 | 5.1379 | 5.1394 | 5.1393 | 5.1383 | 5.1433 | 5.1361 | 0.0192 | 0.0037381
240 | 5.2014 | 5.1671 | 5.1134 | 5.1231 | 5.1419 | 5.1432 | 5.1466 | 5.1638 | 5.1599 | 5.1205 | 5.1153 | 5.1211 | 5.1084 | 5.1128 | 5.1146 | 5.1238 | 5.1245 | 5.1275 | 5.1234 | 5.124 | 5.1249 | 5.1443 | 5.1384 | 5.1423 | 5.1311 | 5.1368 | 5.1383 | 5.1381 | 5.1372 | 5.1422 | 5.135 | 0.019199 | 0.0037389
241 | 5.2003 | 5.166 | 5.1122 | 5.122 | 5.1408 | 5.142 | 5.1455 | 5.1627 | 5.1588 | 5.1195 | 5.1142 | 5.12 | 5.1073 | 5.1117 | 5.1135 | 5.1227 | 5.1234 | 5.1264 | 5.1223 | 5.1229 | 5.1238 | 5.1432 | 5.1373 | 5.1412 | 5.13 | 5.1356 | 5.1371 | 5.137 | 5.136 | 5.1411 | 5.1339 | 0.019198 | 0.0037395
242 | 5.1992 | 5.1649 | 5.111 | 5.1209 | 5.1397 | 5.1409 | 5.1443 | 5.1616 | 5.1577 | 5.1184 | 5.1131 | 5.1189 | 5.1062 | 5.1106 | 5.1123 | 5.1216 | 5.1223 | 5.1252 | 5.1212 | 5.1218 | 5.1227 | 5.1421 | 5.1362 | 5.1401 | 5.1289 | 5.1345 | 5.136 | 5.1359 | 5.1349 | 5.1399 | 5.1328 | 0.019197 | 0.00374
243 | 5.1981 | 5.1638 | 5.1099 | 5.1198 | 5.1386 | 5.1398 | 5.1432 | 5.1605 | 5.1565 | 5.1174 | 5.112 | 5.1178 | 5.1051 | 5.1095 | 5.1112 | 5.1205 | 5.1211 | 5.1241 | 5.1201 | 5.1206 | 5.1216 | 5.141 | 5.1351 | 5.139 | 5.1278 | 5.1334 | 5.1349 | 5.1348 | 5.1338 | 5.1388 | 5.1317 | 0.019195 | 0.0037405
244 | 5.197 | 5.1628 | 5.1088 | 5.1188 | 5.1374 | 5.1387 | 5.1421 | 5.1594 | 5.1554 | 5.1164 | 5.111 | 5.1167 | 5.104 | 5.1084 | 5.1101 | 5.1194 | 5.12 | 5.123 | 5.119 | 5.1195 | 5.1205 | 5.1399 | 5.134 | 5.1379 | 5.1267 | 5.1323 | 5.1338 | 5.1337 | 5.1327 | 5.1377 | 5.1306 | 0.019193 | 0.0037409
245 | 5.1959 | 5.1617 | 5.1076 | 5.1177 | 5.1364 | 5.1376 | 5.141 | 5.1583 | 5.1544 | 5.1153 | 5.1099 | 5.1157 | 5.1029 | 5.1073 | 5.109 | 5.1183 | 5.1189 | 5.1219 | 5.118 | 5.1184 | 5.1194 | 5.1388 | 5.1329 | 5.1369 | 5.1257 | 5.1312 | 5.1327 | 5.1326 | 5.1316 | 5.1366 | 5.1295 | 0.019191 | 0.0037412
246 | 5.1948 | 5.1607 | 5.1065 | 5.1166 | 5.1353 | 5.1365 | 5.14 | 5.1572 | 5.1533 | 5.1143 | 5.1089 | 5.1146 | 5.1018 | 5.1062 | 5.1079 | 5.1173 | 5.1178 | 5.1209 | 5.1169 | 5.1174 | 5.1183 | 5.1377 | 5.1319 | 5.1358 | 5.1246 | 5.1301 | 5.1316 | 5.1315 | 5.1305 | 5.1355 | 5.1284 | 0.019188 | 0.0037414
247 | 5.1937 | 5.1596 | 5.1054 | 5.1156 | 5.1342 | 5.1354 | 5.1389 | 5.1562 | 5.1522 | 5.1133 | 5.1079 | 5.1136 | 5.1007 | 5.1051 | 5.1069 | 5.1162 | 5.1167 | 5.1198 | 5.1159 | 5.1163 | 5.1173 | 5.1367 | 5.1308 | 5.1348 | 5.1235 | 5.1291 | 5.1306 | 5.1304 | 5.1294 | 5.1344 | 5.1273 | 0.019184 | 0.0037416
248 | 5.1926 | 5.1586 | 5.1043 | 5.1145 | 5.1331 | 5.1344 | 5.1378 | 5.1551 | 5.1511 | 5.1123 | 5.1069 | 5.1126 | 5.0997 | 5.104 | 5.1058 | 5.1151 | 5.1157 | 5.1187 | 5.1148 | 5.1152 | 5.1162 | 5.1356 | 5.1298 | 5.1337 | 5.1225 | 5.128 | 5.1295 | 5.1293 | 5.1283 | 5.1334 | 5.1263 | 0.019181 | 0.0037417
249 | 5.1915 | 5.1576 | 5.1032 | 5.1135 | 5.1321 | 5.1333 | 5.1368 | 5.1541 | 5.1501 | 5.1114 | 5.1059 | 5.1115 | 5.0986 | 5.103 | 5.1047 | 5.1141 | 5.1146 | 5.1177 | 5.1138 | 5.1142 | 5.1151 | 5.1346 | 5.1287 | 5.1327 | 5.1214 | 5.1269 | 5.1285 | 5.1283 | 5.1272 | 5.1323 | 5.1252 | 0.019177 | 0.0037416
250 | 5.1905 | 5.1565 | 5.1022 | 5.1125 | 5.131 | 5.1323 | 5.1357 | 5.153 | 5.1491 | 5.1104 | 5.1049 | 5.1105 | 5.0976 | 5.1019 | 5.1037 | 5.113 | 5.1136 | 5.1166 | 5.1128 | 5.1131 | 5.1141 | 5.1335 | 5.1277 | 5.1316 | 5.1204 | 5.1259 | 5.1274 | 5.1272 | 5.1262 | 5.1313 | 5.1242 | 0.019173 | 0.0037416
251 | 5.1894 | 5.1555 | 5.1011 | 5.1114 | 5.13 | 5.1312 | 5.1347 | 5.152 | 5.148 | 5.1094 | 5.1039 | 5.1095 | 5.0965 | 5.1009 | 5.1026 | 5.112 | 5.1125 | 5.1156 | 5.1117 | 5.1121 | 5.1131 | 5.1325 | 5.1267 | 5.1306 | 5.1194 | 5.1248 | 5.1264 | 5.1262 | 5.1251 | 5.1302 | 5.1232 | 0.019168 | 0.0037414
252 | 5.1883 | 5.1545 | 5.1 | 5.1104 | 5.129 | 5.1302 | 5.1337 | 5.151 | 5.147 | 5.1085 | 5.1029 | 5.1085 | 5.0955 | 5.0999 | 5.1016 | 5.111 | 5.1115 | 5.1146 | 5.1107 | 5.111 | 5.112 | 5.1315 | 5.1256 | 5.1296 | 5.1184 | 5.1238 | 5.1254 | 5.1252 | 5.1241 | 5.1292 | 5.1221 | 0.019163 | 0.0037412
253 | 5.1873 | 5.1535 | 5.099 | 5.1094 | 5.1279 | 5.1292 | 5.1326 | 5.15 | 5.146 | 5.1075 | 5.1019 | 5.1075 | 5.0945 | 5.0988 | 5.1006 | 5.11 | 5.1104 | 5.1136 | 5.1097 | 5.11 | 5.111 | 5.1305 | 5.1246 | 5.1286 | 5.1174 | 5.1228 | 5.1243 | 5.1242 | 5.123 | 5.1282 | 5.1211 | 0.019157 | 0.0037408
254 | 5.1862 | 5.1525 | 5.098 | 5.1084 | 5.1269 | 5.1281 | 5.1316 | 5.149 | 5.145 | 5.1066 | 5.1009 | 5.1066 | 5.0935 | 5.0978 | 5.0996 | 5.109 | 5.1094 | 5.1126 | 5.1087 | 5.109 | 5.11 | 5.1295 | 5.1236 | 5.1276 | 5.1164 | 5.1218 | 5.1233 | 5.1231 | 5.122 | 5.1271 | 5.1201 | 0.019152 | 0.0037405
255 | 5.1852 | 5.1516 | 5.0969 | 5.1074 | 5.1259 | 5.1271 | 5.1306 | 5.148 | 5.144 | 5.1056 | 5.1 | 5.1056 | 5.0925 | 5.0968 | 5.0986 | 5.108 | 5.1084 | 5.1116 | 5.1078 | 5.108 | 5.109 | 5.1285 | 5.1226 | 5.1266 | 5.1154 | 5.1208 | 5.1223 | 5.1221 | 5.121 | 5.1261 | 5.1191 | 0.019146 | 0.00374
256 | 5.1842 | 5.1506 | 5.0959 | 5.1065 | 5.1249 | 5.1261 | 5.1296 | 5.147 | 5.143 | 5.1047 | 5.099 | 5.1046 | 5.0915 | 5.0958 | 5.0976 | 5.107 | 5.1074 | 5.1106 | 5.1068 | 5.107 | 5.108 | 5.1275 | 5.1217 | 5.1257 | 5.1144 | 5.1198 | 5.1213 | 5.1211 | 5.12 | 5.1251 | 5.1181 | 0.019139 | 0.0037395
257 | 5.1831 | 5.1496 | 5.0949 | 5.1055 | 5.1239 | 5.1251 | 5.1286 | 5.146 | 5.142 | 5.1038 | 5.0981 | 5.1037 | 5.0905 | 5.0948 | 5.0966 | 5.106 | 5.1064 | 5.1096 | 5.1058 | 5.106 | 5.1071 | 5.1265 | 5.1207 | 5.1247 | 5.1134 | 5.1188 | 5.1204 | 5.1201 | 5.119 | 5.1241 | 5.1172 | 0.019132 | 0.0037388
258 | 5.1821 | 5.1487 | 5.0939 | 5.1045 | 5.1229 | 5.1242 | 5.1277 | 5.145 | 5.141 | 5.1028 | 5.0972 | 5.1027 | 5.0895 | 5.0939 | 5.0956 | 5.105 | 5.1054 | 5.1086 | 5.1049 | 5.105 | 5.1061 | 5.1255 | 5.1197 | 5.1237 | 5.1124 | 5.1178 | 5.1194 | 5.1192 | 5.118 | 5.1232 | 5.1162 | 0.019125 | 0.0037382
259 | 5.1811 | 5.1478 | 5.0929 | 5.1036 | 5.122 | 5.1232 | 5.1267 | 5.144 | 5.14 | 5.1019 | 5.0962 | 5.1018 | 5.0886 | 5.0929 | 5.0946 | 5.1041 | 5.1044 | 5.1076 | 5.1039 | 5.104 | 5.1051 | 5.1245 | 5.1188 | 5.1228 | 5.1115 | 5.1168 | 5.1184 | 5.1182 | 5.117 | 5.1222 | 5.1152 | 0.019118 | 0.0037374
260 | 5.18 | 5.1468 | 5.0919 | 5.1026 | 5.121 | 5.1222 | 5.1257 | 5.1431 | 5.1391 | 5.101 | 5.0953 | 5.1009 | 5.0876 | 5.0919 | 5.0936 | 5.1031 | 5.1034 | 5.1067 | 5.103 | 5.1031 | 5.1042 | 5.1236 | 5.1178 | 5.1218 | 5.1105 | 5.1159 | 5.1174 | 5.1172 | 5.116 | 5.1212 | 5.1143 | 0.01911 | 0.0037366
261 | 5.179 | 5.1459 | 5.0909 | 5.1017 | 5.12 | 5.1213 | 5.1248 | 5.1421 | 5.1381 | 5.1001 | 5.0944 | 5.0999 | 5.0867 | 5.091 | 5.0927 | 5.1022 | 5.1025 | 5.1057 | 5.102 | 5.1021 | 5.1032 | 5.1226 | 5.1169 | 5.1209 | 5.1096 | 5.1149 | 5.1165 | 5.1163 | 5.1151 | 5.1203 | 5.1133 | 0.019102 | 0.0037357
262 | 5.178 | 5.145 | 5.0899 | 5.1008 | 5.1191 | 5.1203 | 5.1238 | 5.1412 | 5.1372 | 5.0993 | 5.0935 | 5.099 | 5.0857 | 5.09 | 5.0917 | 5.1012 | 5.1015 | 5.1048 | 5.1011 | 5.1012 | 5.1023 | 5.1217 | 5.1159 | 5.12 | 5.1086 | 5.114 | 5.1155 | 5.1153 | 5.1141 | 5.1193 | 5.1124 | 0.019093 | 0.0037347
263 | 5.177 | 5.1441 | 5.0889 | 5.0998 | 5.1181 | 5.1194 | 5.1229 | 5.1403 | 5.1362 | 5.0984 | 5.0926 | 5.0981 | 5.0848 | 5.0891 | 5.0908 | 5.1003 | 5.1006 | 5.1038 | 5.1002 | 5.1002 | 5.1013 | 5.1208 | 5.115 | 5.119 | 5.1077 | 5.113 | 5.1146 | 5.1144 | 5.1131 | 5.1184 | 5.1114 | 0.019085 | 0.0037337
264 | 5.176 | 5.1432 | 5.088 | 5.0989 | 5.1172 | 5.1184 | 5.122 | 5.1393 | 5.1353 | 5.0975 | 5.0917 | 5.0972 | 5.0839 | 5.0882 | 5.0898 | 5.0994 | 5.0996 | 5.1029 | 5.0993 | 5.0993 | 5.1004 | 5.1198 | 5.1141 | 5.1181 | 5.1068 | 5.1121 | 5.1137 | 5.1135 | 5.1122 | 5.1174 | 5.1105 | 0.019075 | 0.0037326
265 | 5.175 | 5.1423 | 5.087 | 5.098 | 5.1163 | 5.1175 | 5.121 | 5.1384 | 5.1344 | 5.0966 | 5.0909 | 5.0963 | 5.0829 | 5.0872 | 5.0889 | 5.0984 | 5.0987 | 5.102 | 5.0984 | 5.0984 | 5.0995 | 5.1189 | 5.1132 | 5.1172 | 5.1059 | 5.1112 | 5.1128 | 5.1125 | 5.1113 | 5.1165 | 5.1096 | 0.019066 | 0.0037314
266 | 5.174 | 5.1414 | 5.0861 | 5.0971 | 5.1154 | 5.1166 | 5.1201 | 5.1375 | 5.1335 | 5.0958 | 5.09 | 5.0954 | 5.082 | 5.0863 | 5.088 | 5.0975 | 5.0978 | 5.1011 | 5.0975 | 5.0974 | 5.0986 | 5.118 | 5.1123 | 5.1163 | 5.105 | 5.1102 | 5.1118 | 5.1116 | 5.1103 | 5.1156 | 5.1087 | 0.019056 | 0.0037302
267 | 5.173 | 5.1405 | 5.0851 | 5.0962 | 5.1145 | 5.1157 | 5.1192 | 5.1366 | 5.1326 | 5.0949 | 5.0891 | 5.0946 | 5.0811 | 5.0854 | 5.0871 | 5.0966 | 5.0968 | 5.1002 | 5.0966 | 5.0965 | 5.0977 | 5.1171 | 5.1114 | 5.1154 | 5.1041 | 5.1093 | 5.1109 | 5.1107 | 5.1094 | 5.1146 | 5.1078 | 0.019046 | 0.0037289
268 | 5.1721 | 5.1396 | 5.0842 | 5.0953 | 5.1136 | 5.1147 | 5.1183 | 5.1357 | 5.1317 | 5.0941 | 5.0883 | 5.0937 | 5.0802 | 5.0845 | 5.0862 | 5.0957 | 5.0959 | 5.0993 | 5.0957 | 5.0956 | 5.0968 | 5.1162 | 5.1105 | 5.1145 | 5.1032 | 5.1084 | 5.11 | 5.1098 | 5.1085 | 5.1137 | 5.1069 | 0.019036 | 0.0037275
269 | 5.1711 | 5.1387 | 5.0833 | 5.0945 | 5.1127 | 5.1138 | 5.1174 | 5.1348 | 5.1308 | 5.0932 | 5.0874 | 5.0928 | 5.0793 | 5.0836 | 5.0853 | 5.0948 | 5.095 | 5.0984 | 5.0948 | 5.0947 | 5.0959 | 5.1153 | 5.1096 | 5.1137 | 5.1023 | 5.1075 | 5.1091 | 5.1089 | 5.1076 | 5.1128 | 5.106 | 0.019025 | 0.0037261
270 | 5.1701 | 5.1379 | 5.0824 | 5.0936 | 5.1118 | 5.113 | 5.1165 | 5.1339 | 5.1299 | 5.0924 | 5.0866 | 5.092 | 5.0785 | 5.0827 | 5.0844 | 5.094 | 5.0941 | 5.0975 | 5.094 | 5.0938 | 5.095 | 5.1144 | 5.1087 | 5.1128 | 5.1014 | 5.1066 | 5.1083 | 5.108 | 5.1067 | 5.1119 | 5.1051 | 0.019014 | 0.0037246
271 | 5.1691 | 5.137 | 5.0815 | 5.0927 | 5.1109 | 5.1121 | 5.1156 | 5.1331 | 5.129 | 5.0916 | 5.0857 | 5.0911 | 5.0776 | 5.0818 | 5.0835 | 5.0931 | 5.0932 | 5.0966 | 5.0931 | 5.0929 | 5.0941 | 5.1136 | 5.1079 | 5.1119 | 5.1006 | 5.1058 | 5.1074 | 5.1071 | 5.1058 | 5.1111 | 5.1042 | 0.019003 | 0.003723
272 | 5.1682 | 5.1362 | 5.0806 | 5.0919 | 5.11 | 5.1112 | 5.1148 | 5.1322 | 5.1281 | 5.0907 | 5.0849 | 5.0903 | 5.0767 | 5.081 | 5.0826 | 5.0922 | 5.0924 | 5.0957 | 5.0923 | 5.0921 | 5.0933 | 5.1127 | 5.107 | 5.1111 | 5.0997 | 5.1049 | 5.1065 | 5.1063 | 5.1049 | 5.1102 | 5.1034 | 0.018992 | 0.0037214
273 | 5.1672 | 5.1353 | 5.0797 | 5.091 | 5.1091 | 5.1103 | 5.1139 | 5.1313 | 5.1273 | 5.0899 | 5.0841 | 5.0895 | 5.0759 | 5.0801 | 5.0818 | 5.0914 | 5.0915 | 5.0949 | 5.0914 | 5.0912 | 5.0924 | 5.1118 | 5.1062 | 5.1102 | 5.0988 | 5.104 | 5.1056 | 5.1054 | 5.104 | 5.1093 | 5.1025 | 0.01898 | 0.0037197
274 | 5.1663 | 5.1345 | 5.0788 | 5.0902 | 5.1083 | 5.1095 | 5.1131 | 5.1305 | 5.1264 | 5.0891 | 5.0833 | 5.0887 | 5.075 | 5.0792 | 5.0809 | 5.0905 | 5.0906 | 5.094 | 5.0906 | 5.0903 | 5.0916 | 5.111 | 5.1053 | 5.1094 | 5.098 | 5.1032 | 5.1048 | 5.1045 | 5.1032 | 5.1084 | 5.1016 | 0.018968 | 0.003718
275 | 5.1653 | 5.1337 | 5.0779 | 5.0893 | 5.1074 | 5.1086 | 5.1122 | 5.1296 | 5.1256 | 5.0883 | 5.0825 | 5.0878 | 5.0742 | 5.0784 | 5.08 | 5.0897 | 5.0898 | 5.0932 | 5.0898 | 5.0895 | 5.0907 | 5.1101 | 5.1045 | 5.1086 | 5.0972 | 5.1023 | 5.1039 | 5.1037 | 5.1023 | 5.1076 | 5.1008 | 0.018955 | 0.0037162
276 | 5.1644 | 5.1329 | 5.077 | 5.0885 | 5.1066 | 5.1078 | 5.1114 | 5.1288 | 5.1247 | 5.0875 | 5.0817 | 5.087 | 5.0733 | 5.0775 | 5.0792 | 5.0888 | 5.0889 | 5.0923 | 5.0889 | 5.0886 | 5.0899 | 5.1093 | 5.1036 | 5.1077 | 5.0963 | 5.1015 | 5.1031 | 5.1028 | 5.1014 | 5.1067 | 5.0999 | 0.018943 | 0.0037143
277 | 5.1634 | 5.132 | 5.0762 | 5.0877 | 5.1057 | 5.1069 | 5.1105 | 5.1279 | 5.1239 | 5.0867 | 5.0809 | 5.0862 | 5.0725 | 5.0767 | 5.0784 | 5.088 | 5.088 | 5.0915 | 5.0881 | 5.0878 | 5.089 | 5.1084 | 5.1028 | 5.1069 | 5.0955 | 5.1006 | 5.1022 | 5.102 | 5.1006 | 5.1059 | 5.0991 | 0.01893 | 0.0037124
278 | 5.1625 | 5.1312 | 5.0753 | 5.0869 | 5.1049 | 5.1061 | 5.1097 | 5.1271 | 5.1231 | 5.0859 | 5.0801 | 5.0854 | 5.0717 | 5.0759 | 5.0775 | 5.0872 | 5.0872 | 5.0907 | 5.0873 | 5.087 | 5.0882 | 5.1076 | 5.102 | 5.1061 | 5.0947 | 5.0998 | 5.1014 | 5.1011 | 5.0997 | 5.105 | 5.0983 | 0.018917 | 0.0037104
279 | 5.1616 | 5.1304 | 5.0745 | 5.086 | 5.1041 | 5.1052 | 5.1089 | 5.1263 | 5.1222 | 5.0851 | 5.0793 | 5.0846 | 5.0709 | 5.0751 | 5.0767 | 5.0864 | 5.0864 | 5.0898 | 5.0865 | 5.0861 | 5.0874 | 5.1068 | 5.1012 | 5.1053 | 5.0939 | 5.099 | 5.1006 | 5.1003 | 5.0989 | 5.1042 | 5.0975 | 0.018903 | 0.0037084
280 | 5.1606 | 5.1297 | 5.0736 | 5.0852 | 5.1033 | 5.1044 | 5.108 | 5.1255 | 5.1214 | 5.0844 | 5.0786 | 5.0839 | 5.0701 | 5.0742 | 5.0759 | 5.0855 | 5.0855 | 5.089 | 5.0857 | 5.0853 | 5.0866 | 5.106 | 5.1004 | 5.1045 | 5.093 | 5.0981 | 5.0998 | 5.0995 | 5.0981 | 5.1034 | 5.0966 | 0.018889 | 0.0037063
281 | 5.1597 | 5.1289 | 5.0728 | 5.0844 | 5.1024 | 5.1036 | 5.1072 | 5.1247 | 5.1206 | 5.0836 | 5.0778 | 5.0831 | 5.0692 | 5.0734 | 5.0751 | 5.0847 | 5.0847 | 5.0882 | 5.0849 | 5.0845 | 5.0858 | 5.1052 | 5.0996 | 5.1037 | 5.0922 | 5.0973 | 5.099 | 5.0987 | 5.0972 | 5.1026 | 5.0958 | 0.018875 | 0.0037041
282 | 5.1588 | 5.1281 | 5.0719 | 5.0836 | 5.1016 | 5.1028 | 5.1064 | 5.1239 | 5.1198 | 5.0828 | 5.077 | 5.0823 | 5.0684 | 5.0726 | 5.0743 | 5.0839 | 5.0839 | 5.0874 | 5.0841 | 5.0837 | 5.085 | 5.1044 | 5.0988 | 5.1029 | 5.0914 | 5.0965 | 5.0982 | 5.0979 | 5.0964 | 5.1018 | 5.095 | 0.018861 | 0.0037019
283 | 5.1579 | 5.1273 | 5.0711 | 5.0829 | 5.1008 | 5.102 | 5.1056 | 5.1231 | 5.119 | 5.0821 | 5.0763 | 5.0815 | 5.0677 | 5.0718 | 5.0735 | 5.0831 | 5.0831 | 5.0866 | 5.0833 | 5.0829 | 5.0842 | 5.1036 | 5.098 | 5.1021 | 5.0907 | 5.0957 | 5.0974 | 5.0971 | 5.0956 | 5.101 | 5.0942 | 0.018847 | 0.0036996
284 | 5.157 | 5.1265 | 5.0703 | 5.0821 | 5.1 | 5.1012 | 5.1048 | 5.1223 | 5.1182 | 5.0813 | 5.0755 | 5.0808 | 5.0669 | 5.071 | 5.0727 | 5.0824 | 5.0823 | 5.0858 | 5.0826 | 5.0821 | 5.0834 | 5.1028 | 5.0972 | 5.1013 | 5.0899 | 5.0949 | 5.0966 | 5.0963 | 5.0948 | 5.1002 | 5.0934 | 0.018832 | 0.0036973
285 | 5.156 | 5.1258 | 5.0695 | 5.0813 | 5.0992 | 5.1004 | 5.104 | 5.1215 | 5.1174 | 5.0806 | 5.0748 | 5.08 | 5.0661 | 5.0702 | 5.0719 | 5.0816 | 5.0815 | 5.085 | 5.0818 | 5.0813 | 5.0826 | 5.102 | 5.0965 | 5.1006 | 5.0891 | 5.0941 | 5.0958 | 5.0955 | 5.094 | 5.0994 | 5.0927 | 0.018817 | 0.0036949
286 | 5.1551 | 5.125 | 5.0687 | 5.0805 | 5.0985 | 5.0996 | 5.1033 | 5.1207 | 5.1167 | 5.0798 | 5.0741 | 5.0793 | 5.0653 | 5.0695 | 5.0711 | 5.0808 | 5.0807 | 5.0842 | 5.081 | 5.0805 | 5.0818 | 5.1012 | 5.0957 | 5.0998 | 5.0883 | 5.0933 | 5.095 | 5.0947 | 5.0932 | 5.0986 | 5.0919 | 0.018802 | 0.0036925
287 | 5.1542 | 5.1243 | 5.0679 | 5.0798 | 5.0977 | 5.0988 | 5.1025 | 5.12 | 5.1159 | 5.0791 | 5.0733 | 5.0786 | 5.0646 | 5.0687 | 5.0703 | 5.08 | 5.0799 | 5.0835 | 5.0803 | 5.0797 | 5.0811 | 5.1004 | 5.0949 | 5.099 | 5.0876 | 5.0926 | 5.0942 | 5.0939 | 5.0924 | 5.0978 | 5.0911 | 0.018786 | 0.00369
288 | 5.1533 | 5.1235 | 5.0671 | 5.079 | 5.0969 | 5.0981 | 5.1017 | 5.1192 | 5.1151 | 5.0783 | 5.0726 | 5.0778 | 5.0638 | 5.0679 | 5.0695 | 5.0793 | 5.0791 | 5.0827 | 5.0795 | 5.079 | 5.0803 | 5.0997 | 5.0942 | 5.0983 | 5.0868 | 5.0918 | 5.0935 | 5.0932 | 5.0917 | 5.097 | 5.0903 | 0.018771 | 0.0036875
289 | 5.1524 | 5.1228 | 5.0663 | 5.0783 | 5.0961 | 5.0973 | 5.101 | 5.1184 | 5.1144 | 5.0776 | 5.0719 | 5.0771 | 5.063 | 5.0672 | 5.0688 | 5.0785 | 5.0784 | 5.0819 | 5.0788 | 5.0782 | 5.0795 | 5.0989 | 5.0934 | 5.0975 | 5.0861 | 5.091 | 5.0927 | 5.0924 | 5.0909 | 5.0962 | 5.0896 | 0.018755 | 0.0036849
290 | 5.1516 | 5.1221 | 5.0655 | 5.0775 | 5.0954 | 5.0965 | 5.1002 | 5.1177 | 5.1136 | 5.0769 | 5.0712 | 5.0764 | 5.0623 | 5.0664 | 5.068 | 5.0778 | 5.0776 | 5.0812 | 5.0781 | 5.0774 | 5.0788 | 5.0981 | 5.0927 | 5.0968 | 5.0853 | 5.0903 | 5.092 | 5.0916 | 5.0901 | 5.0955 | 5.0888 | 0.018739 | 0.0036823
291 | 5.1507 | 5.1213 | 5.0647 | 5.0768 | 5.0946 | 5.0958 | 5.0995 | 5.1169 | 5.1129 | 5.0762 | 5.0705 | 5.0757 | 5.0616 | 5.0657 | 5.0673 | 5.077 | 5.0768 | 5.0804 | 5.0773 | 5.0767 | 5.078 | 5.0974 | 5.0919 | 5.0961 | 5.0846 | 5.0895 | 5.0912 | 5.0909 | 5.0894 | 5.0947 | 5.0881 | 0.018722 | 0.0036796
292 | 5.1498 | 5.1206 | 5.064 | 5.0761 | 5.0939 | 5.095 | 5.0987 | 5.1162 | 5.1121 | 5.0755 | 5.0698 | 5.0749 | 5.0608 | 5.0649 | 5.0665 | 5.0763 | 5.0761 | 5.0797 | 5.0766 | 5.0759 | 5.0773 | 5.0966 | 5.0912 | 5.0953 | 5.0838 | 5.0888 | 5.0905 | 5.0901 | 5.0886 | 5.094 | 5.0873 | 0.018706 | 0.0036769
293 | 5.1489 | 5.1199 | 5.0632 | 5.0753 | 5.0932 | 5.0943 | 5.098 | 5.1155 | 5.1114 | 5.0748 | 5.0691 | 5.0742 | 5.0601 | 5.0642 | 5.0658 | 5.0755 | 5.0753 | 5.079 | 5.0759 | 5.0752 | 5.0766 | 5.0959 | 5.0905 | 5.0946 | 5.0831 | 5.088 | 5.0897 | 5.0894 | 5.0878 | 5.0932 | 5.0866 | 0.018689 | 0.0036742
294 | 5.148 | 5.1192 | 5.0624 | 5.0746 | 5.0924 | 5.0936 | 5.0972 | 5.1147 | 5.1107 | 5.0741 | 5.0684 | 5.0735 | 5.0594 | 5.0635 | 5.0651 | 5.0748 | 5.0746 | 5.0782 | 5.0752 | 5.0745 | 5.0758 | 5.0952 | 5.0898 | 5.0939 | 5.0824 | 5.0873 | 5.089 | 5.0887 | 5.0871 | 5.0925 | 5.0859 | 0.018672 | 0.0036713
295 | 5.1472 | 5.1185 | 5.0617 | 5.0739 | 5.0917 | 5.0928 | 5.0965 | 5.114 | 5.1099 | 5.0734 | 5.0677 | 5.0729 | 5.0587 | 5.0627 | 5.0643 | 5.0741 | 5.0739 | 5.0775 | 5.0745 | 5.0737 | 5.0751 | 5.0944 | 5.0891 | 5.0932 | 5.0817 | 5.0866 | 5.0883 | 5.0879 | 5.0864 | 5.0918 | 5.0851 | 0.018655 | 0.0036685
296 | 5.1463 | 5.1178 | 5.0609 | 5.0732 | 5.091 | 5.0921 | 5.0958 | 5.1133 | 5.1092 | 5.0727 | 5.067 | 5.0722 | 5.0579 | 5.062 | 5.0636 | 5.0734 | 5.0731 | 5.0768 | 5.0738 | 5.073 | 5.0744 | 5.0937 | 5.0883 | 5.0925 | 5.0809 | 5.0858 | 5.0875 | 5.0872 | 5.0856 | 5.091 | 5.0844 | 0.018637 | 0.0036656
297 | 5.1454 | 5.1171 | 5.0602 | 5.0725 | 5.0902 | 5.0914 | 5.0951 | 5.1126 | 5.1085 | 5.072 | 5.0664 | 5.0715 | 5.0572 | 5.0613 | 5.0629 | 5.0727 | 5.0724 | 5.0761 | 5.0731 | 5.0723 | 5.0737 | 5.093 | 5.0876 | 5.0918 | 5.0802 | 5.0851 | 5.0868 | 5.0865 | 5.0849 | 5.0903 | 5.0837 | 0.01862 | 0.0036626
298 | 5.1446 | 5.1164 | 5.0595 | 5.0718 | 5.0895 | 5.0907 | 5.0944 | 5.1119 | 5.1078 | 5.0713 | 5.0657 | 5.0708 | 5.0565 | 5.0606 | 5.0622 | 5.072 | 5.0717 | 5.0754 | 5.0724 | 5.0716 | 5.073 | 5.0923 | 5.0869 | 5.0911 | 5.0795 | 5.0844 | 5.0861 | 5.0858 | 5.0842 | 5.0896 | 5.083 | 0.018602 | 0.0036596
299 | 5.1437 | 5.1157 | 5.0587 | 5.0711 | 5.0888 | 5.09 | 5.0937 | 5.1112 | 5.1071 | 5.0706 | 5.065 | 5.0702 | 5.0558 | 5.0599 | 5.0615 | 5.0713 | 5.071 | 5.0747 | 5.0717 | 5.0709 | 5.0723 | 5.0916 | 5.0863 | 5.0904 | 5.0788 | 5.0837 | 5.0854 | 5.0851 | 5.0835 | 5.0889 | 5.0823 | 0.018584 | 0.0036566
300 | 5.1429 | 5.1151 | 5.058 | 5.0704 | 5.0881 | 5.0893 | 5.093 | 5.1105 | 5.1064 | 5.0699 | 5.0644 | 5.0695 | 5.0551 | 5.0592 | 5.0608 | 5.0706 | 5.0703 | 5.074 | 5.071 | 5.0702 | 5.0716 | 5.0909 | 5.0856 | 5.0897 | 5.0782 | 5.083 | 5.0847 | 5.0844 | 5.0828 | 5.0882 | 5.0816 | 0.018565 | 0.0036535
301 | 5.142 | 5.1144 | 5.0573 | 5.0697 | 5.0874 | 5.0886 | 5.0923 | 5.1098 | 5.1057 | 5.0693 | 5.0637 | 5.0688 | 5.0545 | 5.0585 | 5.0601 | 5.0699 | 5.0696 | 5.0733 | 5.0703 | 5.0695 | 5.0709 | 5.0902 | 5.0849 | 5.089 | 5.0775 | 5.0823 | 5.084 | 5.0837 | 5.082 | 5.0875 | 5.0809 | 0.018547 | 0.0036504
302 | 5.1412 | 5.1137 | 5.0566 | 5.0691 | 5.0867 | 5.0879 | 5.0916 | 5.1091 | 5.105 | 5.0686 | 5.0631 | 5.0682 | 5.0538 | 5.0578 | 5.0594 | 5.0692 | 5.0689 | 5.0726 | 5.0697 | 5.0688 | 5.0702 | 5.0895 | 5.0842 | 5.0884 | 5.0768 | 5.0816 | 5.0833 | 5.083 | 5.0813 | 5.0868 | 5.0802 | 0.018528 | 0.0036472
303 | 5.1403 | 5.1131 | 5.0559 | 5.0684 | 5.0861 | 5.0872 | 5.0909 | 5.1084 | 5.1043 | 5.068 | 5.0625 | 5.0675 | 5.0531 | 5.0571 | 5.0587 | 5.0685 | 5.0682 | 5.0719 | 5.069 | 5.0681 | 5.0695 | 5.0888 | 5.0835 | 5.0877 | 5.0761 | 5.0809 | 5.0827 | 5.0823 | 5.0807 | 5.0861 | 5.0795 | 0.01851 | 0.003644
304 | 5.1395 | 5.1124 | 5.0552 | 5.0677 | 5.0854 | 5.0865 | 5.0902 | 5.1077 | 5.1037 | 5.0673 | 5.0618 | 5.0669 | 5.0524 | 5.0565 | 5.058 | 5.0679 | 5.0675 | 5.0712 | 5.0684 | 5.0674 | 5.0689 | 5.0881 | 5.0829 | 5.087 | 5.0754 | 5.0803 | 5.082 | 5.0816 | 5.08 | 5.0854 | 5.0788 | 0.018491 | 0.0036407
305 | 5.1386 | 5.1118 | 5.0545 | 5.0671 | 5.0847 | 5.0858 | 5.0896 | 5.1071 | 5.103 | 5.0667 | 5.0612 | 5.0663 | 5.0518 | 5.0558 | 5.0574 | 5.0672 | 5.0668 | 5.0705 | 5.0677 | 5.0667 | 5.0682 | 5.0875 | 5.0822 | 5.0864 | 5.0748 | 5.0796 | 5.0813 | 5.0809 | 5.0793 | 5.0847 | 5.0782 | 0.018472 | 0.0036375
306 | 5.1378 | 5.1111 | 5.0538 | 5.0664 | 5.084 | 5.0852 | 5.0889 | 5.1064 | 5.1023 | 5.066 | 5.0606 | 5.0656 | 5.0511 | 5.0551 | 5.0567 | 5.0665 | 5.0661 | 5.0699 | 5.0671 | 5.0661 | 5.0675 | 5.0868 | 5.0816 | 5.0857 | 5.0741 | 5.0789 | 5.0806 | 5.0803 | 5.0786 | 5.084 | 5.0775 | 0.018452 | 0.0036341
307 | 5.137 | 5.1105 | 5.0531 | 5.0658 | 5.0834 | 5.0845 | 5.0882 | 5.1057 | 5.1017 | 5.0654 | 5.06 | 5.065 | 5.0505 | 5.0545 | 5.056 | 5.0659 | 5.0654 | 5.0692 | 5.0664 | 5.0654 | 5.0669 | 5.0861 | 5.0809 | 5.0851 | 5.0735 | 5.0783 | 5.08 | 5.0796 | 5.0779 | 5.0834 | 5.0768 | 0.018433 | 0.0036308
308 | 5.1361 | 5.1099 | 5.0524 | 5.0651 | 5.0827 | 5.0838 | 5.0876 | 5.1051 | 5.101 | 5.0647 | 5.0594 | 5.0644 | 5.0498 | 5.0538 | 5.0554 | 5.0652 | 5.0648 | 5.0686 | 5.0658 | 5.0647 | 5.0662 | 5.0855 | 5.0803 | 5.0844 | 5.0728 | 5.0776 | 5.0793 | 5.0789 | 5.0773 | 5.0827 | 5.0762 | 0.018413 | 0.0036274
309 | 5.1353 | 5.1092 | 5.0517 | 5.0645 | 5.0821 | 5.0832 | 5.0869 | 5.1044 | 5.1004 | 5.0641 | 5.0588 | 5.0638 | 5.0492 | 5.0532 | 5.0547 | 5.0646 | 5.0641 | 5.0679 | 5.0651 | 5.0641 | 5.0656 | 5.0848 | 5.0796 | 5.0838 | 5.0722 | 5.0769 | 5.0787 | 5.0783 | 5.0766 | 5.082 | 5.0755 | 0.018393 | 0.0036239
310 | 5.1345 | 5.1086 | 5.0511 | 5.0638 | 5.0814 | 5.0825 | 5.0863 | 5.1038 | 5.0997 | 5.0635 | 5.0582 | 5.0631 | 5.0486 | 5.0525 | 5.0541 | 5.0639 | 5.0635 | 5.0673 | 5.0645 | 5.0634 | 5.0649 | 5.0842 | 5.079 | 5.0831 | 5.0715 | 5.0763 | 5.078 | 5.0776 | 5.0759 | 5.0814 | 5.0749 | 0.018373 | 0.0036205
311 | 5.1337 | 5.108 | 5.0504 | 5.0632 | 5.0808 | 5.0819 | 5.0856 | 5.1031 | 5.0991 | 5.0628 | 5.0576 | 5.0625 | 5.0479 | 5.0519 | 5.0534 | 5.0633 | 5.0628 | 5.0666 | 5.0639 | 5.0628 | 5.0643 | 5.0835 | 5.0784 | 5.0825 | 5.0709 | 5.0756 | 5.0774 | 5.077 | 5.0753 | 5.0807 | 5.0742 | 0.018353 | 0.003617
312 | 5.1329 | 5.1074 | 5.0497 | 5.0626 | 5.0801 | 5.0812 | 5.085 | 5.1025 | 5.0984 | 5.0622 | 5.057 | 5.0619 | 5.0473 | 5.0513 | 5.0528 | 5.0627 | 5.0622 | 5.066 | 5.0633 | 5.0621 | 5.0637 | 5.0829 | 5.0778 | 5.0819 | 5.0703 | 5.075 | 5.0767 | 5.0763 | 5.0746 | 5.0801 | 5.0736 | 0.018333 | 0.0036134
313 | 5.132 | 5.1068 | 5.0491 | 5.0619 | 5.0795 | 5.0806 | 5.0844 | 5.1019 | 5.0978 | 5.0616 | 5.0564 | 5.0613 | 5.0467 | 5.0506 | 5.0522 | 5.062 | 5.0615 | 5.0654 | 5.0627 | 5.0615 | 5.063 | 5.0823 | 5.0771 | 5.0813 | 5.0697 | 5.0744 | 5.0761 | 5.0757 | 5.074 | 5.0794 | 5.073 | 0.018313 | 0.0036099
314 | 5.1312 | 5.1062 | 5.0484 | 5.0613 | 5.0789 | 5.08 | 5.0837 | 5.1012 | 5.0972 | 5.061 | 5.0558 | 5.0608 | 5.0461 | 5.05 | 5.0515 | 5.0614 | 5.0609 | 5.0647 | 5.062 | 5.0609 | 5.0624 | 5.0816 | 5.0765 | 5.0807 | 5.069 | 5.0737 | 5.0755 | 5.0751 | 5.0734 | 5.0788 | 5.0723 | 0.018292 | 0.0036063
315 | 5.1304 | 5.1056 | 5.0478 | 5.0607 | 5.0782 | 5.0793 | 5.0831 | 5.1006 | 5.0966 | 5.0604 | 5.0552 | 5.0602 | 5.0454 | 5.0494 | 5.0509 | 5.0608 | 5.0602 | 5.0641 | 5.0614 | 5.0602 | 5.0618 | 5.081 | 5.0759 | 5.0801 | 5.0684 | 5.0731 | 5.0749 | 5.0745 | 5.0727 | 5.0782 | 5.0717 | 0.018271 | 0.0036026
316 | 5.1296 | 5.105 | 5.0471 | 5.0601 | 5.0776 | 5.0787 | 5.0825 | 5.1 | 5.096 | 5.0598 | 5.0546 | 5.0596 | 5.0448 | 5.0488 | 5.0503 | 5.0602 | 5.0596 | 5.0635 | 5.0608 | 5.0596 | 5.0612 | 5.0804 | 5.0753 | 5.0794 | 5.0678 | 5.0725 | 5.0742 | 5.0738 | 5.0721 | 5.0776 | 5.0711 | 0.018251 | 0.003599
317 | 5.1288 | 5.1044 | 5.0465 | 5.0595 | 5.077 | 5.0781 | 5.0819 | 5.0994 | 5.0953 | 5.0592 | 5.0541 | 5.059 | 5.0442 | 5.0482 | 5.0497 | 5.0596 | 5.059 | 5.0629 | 5.0603 | 5.059 | 5.0606 | 5.0798 | 5.0747 | 5.0788 | 5.0672 | 5.0719 | 5.0736 | 5.0732 | 5.0715 | 5.0769 | 5.0705 | 0.01823 | 0.0035953
318 | 5.128 | 5.1038 | 5.0459 | 5.0589 | 5.0764 | 5.0775 | 5.0813 | 5.0988 | 5.0947 | 5.0586 | 5.0535 | 5.0584 | 5.0436 | 5.0476 | 5.0491 | 5.059 | 5.0584 | 5.0623 | 5.0597 | 5.0584 | 5.06 | 5.0792 | 5.0741 | 5.0783 | 5.0666 | 5.0713 | 5.073 | 5.0726 | 5.0709 | 5.0763 | 5.0699 | 0.018209 | 0.0035915
319 | 5.1272 | 5.1032 | 5.0452 | 5.0583 | 5.0758 | 5.0769 | 5.0807 | 5.0982 | 5.0941 | 5.058 | 5.053 | 5.0579 | 5.043 | 5.047 | 5.0485 | 5.0584 | 5.0578 | 5.0617 | 5.0591 | 5.0578 | 5.0594 | 5.0785 | 5.0735 | 5.0777 | 5.066 | 5.0707 | 5.0724 | 5.072 | 5.0702 | 5.0757 | 5.0693 | 0.018187 | 0.0035878
320 | 5.1264 | 5.1026 | 5.0446 | 5.0577 | 5.0752 | 5.0763 | 5.0801 | 5.0976 | 5.0935 | 5.0574 | 5.0524 | 5.0573 | 5.0425 | 5.0464 | 5.0479 | 5.0578 | 5.0571 | 5.0611 | 5.0585 | 5.0572 | 5.0588 | 5.0779 | 5.0729 | 5.0771 | 5.0654 | 5.0701 | 5.0718 | 5.0714 | 5.0696 | 5.0751 | 5.0687 | 0.018166 | 0.003584
321 | 5.1257 | 5.102 | 5.044 | 5.0571 | 5.0746 | 5.0757 | 5.0795 | 5.097 | 5.0929 | 5.0568 | 5.0519 | 5.0567 | 5.0419 | 5.0458 | 5.0473 | 5.0572 | 5.0565 | 5.0605 | 5.0579 | 5.0566 | 5.0582 | 5.0773 | 5.0724 | 5.0765 | 5.0648 | 5.0695 | 5.0712 | 5.0708 | 5.069 | 5.0745 | 5.0681 | 0.018145 | 0.0035802
322 | 5.1249 | 5.1015 | 5.0434 | 5.0566 | 5.074 | 5.0751 | 5.0789 | 5.0964 | 5.0924 | 5.0562 | 5.0513 | 5.0562 | 5.0413 | 5.0452 | 5.0467 | 5.0566 | 5.0559 | 5.0599 | 5.0573 | 5.056 | 5.0576 | 5.0768 | 5.0718 | 5.0759 | 5.0642 | 5.0689 | 5.0706 | 5.0702 | 5.0684 | 5.0739 | 5.0675 | 0.018123 | 0.0035763
323 | 5.1241 | 5.1009 | 5.0428 | 5.056 | 5.0734 | 5.0745 | 5.0783 | 5.0958 | 5.0918 | 5.0557 | 5.0508 | 5.0556 | 5.0407 | 5.0446 | 5.0461 | 5.056 | 5.0553 | 5.0593 | 5.0568 | 5.0554 | 5.057 | 5.0762 | 5.0712 | 5.0753 | 5.0637 | 5.0683 | 5.07 | 5.0696 | 5.0678 | 5.0733 | 5.0669 | 0.018101 | 0.0035725
324 | 5.1233 | 5.1004 | 5.0422 | 5.0554 | 5.0728 | 5.0739 | 5.0777 | 5.0952 | 5.0912 | 5.0551 | 5.0502 | 5.0551 | 5.0402 | 5.044 | 5.0455 | 5.0554 | 5.0547 | 5.0587 | 5.0562 | 5.0548 | 5.0564 | 5.0756 | 5.0706 | 5.0748 | 5.0631 | 5.0677 | 5.0695 | 5.069 | 5.0672 | 5.0727 | 5.0663 | 0.018079 | 0.0035686
325 | 5.1225 | 5.0998 | 5.0416 | 5.0548 | 5.0722 | 5.0733 | 5.0771 | 5.0946 | 5.0906 | 5.0545 | 5.0497 | 5.0546 | 5.0396 | 5.0435 | 5.045 | 5.0549 | 5.0542 | 5.0581 | 5.0557 | 5.0542 | 5.0558 | 5.075 | 5.0701 | 5.0742 | 5.0625 | 5.0671 | 5.0689 | 5.0684 | 5.0666 | 5.0721 | 5.0657 | 0.018057 | 0.0035647
326 | 5.1218 | 5.0992 | 5.041 | 5.0543 | 5.0717 | 5.0727 | 5.0766 | 5.0941 | 5.09 | 5.054 | 5.0492 | 5.054 | 5.039 | 5.0429 | 5.0444 | 5.0543 | 5.0536 | 5.0576 | 5.0551 | 5.0537 | 5.0553 | 5.0744 | 5.0695 | 5.0737 | 5.062 | 5.0665 | 5.0683 | 5.0679 | 5.0661 | 5.0715 | 5.0651 | 0.018035 | 0.0035607
327 | 5.121 | 5.0987 | 5.0404 | 5.0537 | 5.0711 | 5.0722 | 5.076 | 5.0935 | 5.0895 | 5.0534 | 5.0486 | 5.0535 | 5.0385 | 5.0423 | 5.0438 | 5.0537 | 5.053 | 5.057 | 5.0546 | 5.0531 | 5.0547 | 5.0738 | 5.069 | 5.0731 | 5.0614 | 5.066 | 5.0677 | 5.0673 | 5.0655 | 5.071 | 5.0646 | 0.018013 | 0.0035567
328 | 5.1202 | 5.0982 | 5.0398 | 5.0532 | 5.0705 | 5.0716 | 5.0754 | 5.0929 | 5.0889 | 5.0528 | 5.0481 | 5.053 | 5.0379 | 5.0418 | 5.0433 | 5.0532 | 5.0524 | 5.0564 | 5.054 | 5.0525 | 5.0542 | 5.0733 | 5.0684 | 5.0725 | 5.0608 | 5.0654 | 5.0672 | 5.0667 | 5.0649 | 5.0704 | 5.064 | 0.017991 | 0.0035528
329 | 5.1195 | 5.0976 | 5.0392 | 5.0526 | 5.07 | 5.071 | 5.0749 | 5.0924 | 5.0884 | 5.0523 | 5.0476 | 5.0524 | 5.0374 | 5.0412 | 5.0427 | 5.0526 | 5.0519 | 5.0559 | 5.0535 | 5.052 | 5.0536 | 5.0727 | 5.0679 | 5.072 | 5.0603 | 5.0648 | 5.0666 | 5.0662 | 5.0643 | 5.0698 | 5.0634 | 0.017969 | 0.0035487
330 | 5.1187 | 5.0971 | 5.0386 | 5.0521 | 5.0694 | 5.0705 | 5.0743 | 5.0918 | 5.0878 | 5.0517 | 5.0471 | 5.0519 | 5.0368 | 5.0407 | 5.0422 | 5.0521 | 5.0513 | 5.0553 | 5.0529 | 5.0514 | 5.053 | 5.0722 | 5.0673 | 5.0715 | 5.0597 | 5.0643 | 5.066 | 5.0656 | 5.0638 | 5.0693 | 5.0629 | 0.017946 | 0.0035447
331 | 5.1179 | 5.0966 | 5.038 | 5.0515 | 5.0688 | 5.0699 | 5.0738 | 5.0913 | 5.0873 | 5.0512 | 5.0466 | 5.0514 | 5.0363 | 5.0401 | 5.0416 | 5.0515 | 5.0507 | 5.0548 | 5.0524 | 5.0509 | 5.0525 | 5.0716 | 5.0668 | 5.0709 | 5.0592 | 5.0637 | 5.0655 | 5.065 | 5.0632 | 5.0687 | 5.0623 | 0.017924 | 0.0035406
332 | 5.1172 | 5.096 | 5.0375 | 5.051 | 5.0683 | 5.0694 | 5.0732 | 5.0907 | 5.0867 | 5.0506 | 5.0461 | 5.0509 | 5.0357 | 5.0396 | 5.0411 | 5.051 | 5.0502 | 5.0542 | 5.0519 | 5.0503 | 5.0519 | 5.071 | 5.0662 | 5.0704 | 5.0586 | 5.0632 | 5.0649 | 5.0645 | 5.0626 | 5.0681 | 5.0618 | 0.017901 | 0.0035366
333 | 5.1164 | 5.0955 | 5.0369 | 5.0504 | 5.0678 | 5.0688 | 5.0727 | 5.0902 | 5.0862 | 5.0501 | 5.0456 | 5.0504 | 5.0352 | 5.039 | 5.0405 | 5.0504 | 5.0496 | 5.0537 | 5.0513 | 5.0498 | 5.0514 | 5.0705 | 5.0657 | 5.0698 | 5.0581 | 5.0626 | 5.0644 | 5.0639 | 5.0621 | 5.0676 | 5.0612 | 0.017879 | 0.0035325
334 | 5.1157 | 5.095 | 5.0363 | 5.0499 | 5.0672 | 5.0683 | 5.0721 | 5.0896 | 5.0856 | 5.0496 | 5.0451 | 5.0499 | 5.0347 | 5.0385 | 5.04 | 5.0499 | 5.0491 | 5.0531 | 5.0508 | 5.0492 | 5.0509 | 5.0699 | 5.0652 | 5.0693 | 5.0576 | 5.0621 | 5.0639 | 5.0634 | 5.0615 | 5.067 | 5.0607 | 0.017856 | 0.0035284
335 | 5.1149 | 5.0945 | 5.0358 | 5.0494 | 5.0667 | 5.0677 | 5.0716 | 5.0891 | 5.0851 | 5.049 | 5.0446 | 5.0494 | 5.0342 | 5.038 | 5.0394 | 5.0494 | 5.0485 | 5.0526 | 5.0503 | 5.0487 | 5.0503 | 5.0694 | 5.0647 | 5.0688 | 5.057 | 5.0616 | 5.0633 | 5.0628 | 5.061 | 5.0665 | 5.0601 | 0.017833 | 0.0035242
336 | 5.1142 | 5.094 | 5.0352 | 5.0489 | 5.0661 | 5.0672 | 5.0711 | 5.0886 | 5.0846 | 5.0485 | 5.0441 | 5.0489 | 5.0336 | 5.0374 | 5.0389 | 5.0489 | 5.048 | 5.0521 | 5.0498 | 5.0481 | 5.0498 | 5.0689 | 5.0641 | 5.0683 | 5.0565 | 5.061 | 5.0628 | 5.0623 | 5.0604 | 5.0659 | 5.0596 | 0.01781 | 0.0035201
337 | 5.1134 | 5.0934 | 5.0347 | 5.0483 | 5.0656 | 5.0667 | 5.0705 | 5.088 | 5.084 | 5.048 | 5.0436 | 5.0484 | 5.0331 | 5.0369 | 5.0384 | 5.0483 | 5.0474 | 5.0515 | 5.0493 | 5.0476 | 5.0493 | 5.0683 | 5.0636 | 5.0678 | 5.056 | 5.0605 | 5.0623 | 5.0618 | 5.0599 | 5.0654 | 5.0591 | 0.017787 | 0.0035159
338 | 5.1127 | 5.0929 | 5.0341 | 5.0478 | 5.0651 | 5.0661 | 5.07 | 5.0875 | 5.0835 | 5.0474 | 5.0432 | 5.0479 | 5.0326 | 5.0364 | 5.0379 | 5.0478 | 5.0469 | 5.051 | 5.0488 | 5.0471 | 5.0488 | 5.0678 | 5.0631 | 5.0672 | 5.0555 | 5.06 | 5.0617 | 5.0612 | 5.0594 | 5.0649 | 5.0585 | 0.017764 | 0.0035117
339 | 5.1119 | 5.0924 | 5.0336 | 5.0473 | 5.0646 | 5.0656 | 5.0695 | 5.087 | 5.083 | 5.0469 | 5.0427 | 5.0474 | 5.0321 | 5.0359 | 5.0374 | 5.0473 | 5.0464 | 5.0505 | 5.0483 | 5.0466 | 5.0482 | 5.0673 | 5.0626 | 5.0667 | 5.055 | 5.0594 | 5.0612 | 5.0607 | 5.0588 | 5.0643 | 5.058 | 0.017741 | 0.0035075
340 | 5.1112 | 5.092 | 5.033 | 5.0468 | 5.064 | 5.0651 | 5.069 | 5.0865 | 5.0825 | 5.0464 | 5.0422 | 5.0469 | 5.0316 | 5.0354 | 5.0368 | 5.0468 | 5.0458 | 5.05 | 5.0478 | 5.046 | 5.0477 | 5.0668 | 5.0621 | 5.0662 | 5.0545 | 5.0589 | 5.0607 | 5.0602 | 5.0583 | 5.0638 | 5.0575 | 0.017718 | 0.0035033
341 | 5.1105 | 5.0915 | 5.0325 | 5.0463 | 5.0635 | 5.0646 | 5.0685 | 5.086 | 5.082 | 5.0459 | 5.0417 | 5.0465 | 5.0311 | 5.0349 | 5.0363 | 5.0463 | 5.0453 | 5.0495 | 5.0473 | 5.0455 | 5.0472 | 5.0662 | 5.0616 | 5.0657 | 5.054 | 5.0584 | 5.0602 | 5.0597 | 5.0578 | 5.0633 | 5.057 | 0.017695 | 0.0034991
342 | 5.1097 | 5.091 | 5.032 | 5.0458 | 5.063 | 5.0641 | 5.0679 | 5.0855 | 5.0815 | 5.0454 | 5.0413 | 5.046 | 5.0306 | 5.0344 | 5.0358 | 5.0458 | 5.0448 | 5.0489 | 5.0468 | 5.045 | 5.0467 | 5.0657 | 5.0611 | 5.0652 | 5.0535 | 5.0579 | 5.0597 | 5.0592 | 5.0573 | 5.0628 | 5.0565 | 0.017671 | 0.0034948
343 | 5.109 | 5.0905 | 5.0314 | 5.0453 | 5.0625 | 5.0635 | 5.0674 | 5.085 | 5.081 | 5.0449 | 5.0408 | 5.0455 | 5.0301 | 5.0339 | 5.0353 | 5.0453 | 5.0443 | 5.0484 | 5.0463 | 5.0445 | 5.0462 | 5.0652 | 5.0606 | 5.0647 | 5.053 | 5.0574 | 5.0592 | 5.0587 | 5.0568 | 5.0623 | 5.056 | 0.017648 | 0.0034905
344 | 5.1083 | 5.09 | 5.0309 | 5.0448 | 5.062 | 5.063 | 5.0669 | 5.0844 | 5.0805 | 5.0444 | 5.0404 | 5.0451 | 5.0296 | 5.0334 | 5.0348 | 5.0448 | 5.0438 | 5.0479 | 5.0458 | 5.044 | 5.0457 | 5.0647 | 5.0601 | 5.0643 | 5.0525 | 5.0569 | 5.0587 | 5.0582 | 5.0562 | 5.0618 | 5.0555 | 0.017625 | 0.0034863
345 | 5.1076 | 5.0895 | 5.0304 | 5.0443 | 5.0615 | 5.0625 | 5.0664 | 5.0839 | 5.08 | 5.0438 | 5.0399 | 5.0446 | 5.0292 | 5.0329 | 5.0343 | 5.0443 | 5.0433 | 5.0474 | 5.0453 | 5.0435 | 5.0452 | 5.0642 | 5.0596 | 5.0638 | 5.052 | 5.0564 | 5.0582 | 5.0576 | 5.0557 | 5.0612 | 5.055 | 0.017601 | 0.003482
346 | 5.1068 | 5.0891 | 5.0299 | 5.0438 | 5.061 | 5.062 | 5.0659 | 5.0835 | 5.0795 | 5.0433 | 5.0395 | 5.0441 | 5.0287 | 5.0324 | 5.0338 | 5.0438 | 5.0428 | 5.0469 | 5.0449 | 5.043 | 5.0447 | 5.0637 | 5.0592 | 5.0633 | 5.0515 | 5.0559 | 5.0577 | 5.0571 | 5.0552 | 5.0607 | 5.0545 | 0.017578 | 0.0034777
347 | 5.1061 | 5.0886 | 5.0294 | 5.0433 | 5.0605 | 5.0615 | 5.0655 | 5.083 | 5.079 | 5.0429 | 5.039 | 5.0437 | 5.0282 | 5.0319 | 5.0334 | 5.0433 | 5.0423 | 5.0465 | 5.0444 | 5.0425 | 5.0442 | 5.0632 | 5.0587 | 5.0628 | 5.051 | 5.0554 | 5.0572 | 5.0567 | 5.0547 | 5.0602 | 5.054 | 0.017554 | 0.0034734
348 | 5.1054 | 5.0881 | 5.0289 | 5.0429 | 5.06 | 5.0611 | 5.065 | 5.0825 | 5.0785 | 5.0424 | 5.0386 | 5.0432 | 5.0277 | 5.0314 | 5.0329 | 5.0428 | 5.0418 | 5.046 | 5.0439 | 5.042 | 5.0437 | 5.0627 | 5.0582 | 5.0623 | 5.0505 | 5.0549 | 5.0567 | 5.0562 | 5.0542 | 5.0597 | 5.0535 | 0.017531 | 0.003469
349 | 5.1047 | 5.0876 | 5.0284 | 5.0424 | 5.0595 | 5.0606 | 5.0645 | 5.082 | 5.078 | 5.0419 | 5.0381 | 5.0428 | 5.0272 | 5.0309 | 5.0324 | 5.0423 | 5.0413 | 5.0455 | 5.0434 | 5.0415 | 5.0433 | 5.0622 | 5.0577 | 5.0619 | 5.05 | 5.0544 | 5.0562 | 5.0557 | 5.0537 | 5.0592 | 5.053 | 0.017507 | 0.0034647
350 | 5.104 | 5.0872 | 5.0279 | 5.0419 | 5.0591 | 5.0601 | 5.064 | 5.0815 | 5.0775 | 5.0414 | 5.0377 | 5.0423 | 5.0268 | 5.0305 | 5.0319 | 5.0419 | 5.0408 | 5.045 | 5.043 | 5.041 | 5.0428 | 5.0617 | 5.0573 | 5.0614 | 5.0496 | 5.0539 | 5.0557 | 5.0552 | 5.0533 | 5.0588 | 5.0525 | 0.017484 | 0.0034604
351 | 5.1033 | 5.0867 | 5.0274 | 5.0414 | 5.0586 | 5.0596 | 5.0635 | 5.081 | 5.0771 | 5.0409 | 5.0373 | 5.0419 | 5.0263 | 5.03 | 5.0314 | 5.0414 | 5.0403 | 5.0445 | 5.0425 | 5.0406 | 5.0423 | 5.0613 | 5.0568 | 5.0609 | 5.0491 | 5.0535 | 5.0552 | 5.0547 | 5.0528 | 5.0583 | 5.052 | 0.01746 | 0.003456
352 | 5.1026 | 5.0863 | 5.0269 | 5.041 | 5.0581 | 5.0591 | 5.0631 | 5.0806 | 5.0766 | 5.0404 | 5.0368 | 5.0415 | 5.0259 | 5.0295 | 5.031 | 5.0409 | 5.0398 | 5.044 | 5.0421 | 5.0401 | 5.0418 | 5.0608 | 5.0563 | 5.0605 | 5.0486 | 5.053 | 5.0548 | 5.0542 | 5.0523 | 5.0578 | 5.0515 | 0.017436 | 0.0034517
353 | 5.1019 | 5.0858 | 5.0264 | 5.0405 | 5.0576 | 5.0587 | 5.0626 | 5.0801 | 5.0761 | 5.0399 | 5.0364 | 5.041 | 5.0254 | 5.029 | 5.0305 | 5.0405 | 5.0393 | 5.0436 | 5.0416 | 5.0396 | 5.0414 | 5.0603 | 5.0559 | 5.06 | 5.0482 | 5.0525 | 5.0543 | 5.0537 | 5.0518 | 5.0573 | 5.0511 | 0.017413 | 0.0034473
354 | 5.1012 | 5.0854 | 5.0259 | 5.04 | 5.0572 | 5.0582 | 5.0621 | 5.0796 | 5.0757 | 5.0395 | 5.036 | 5.0406 | 5.025 | 5.0286 | 5.03 | 5.04 | 5.0388 | 5.0431 | 5.0412 | 5.0392 | 5.0409 | 5.0598 | 5.0554 | 5.0595 | 5.0477 | 5.052 | 5.0538 | 5.0533 | 5.0513 | 5.0568 | 5.0506 | 0.017389 | 0.0034429
355 | 5.1005 | 5.0849 | 5.0254 | 5.0396 | 5.0567 | 5.0577 | 5.0617 | 5.0792 | 5.0752 | 5.039 | 5.0356 | 5.0402 | 5.0245 | 5.0281 | 5.0296 | 5.0395 | 5.0384 | 5.0426 | 5.0407 | 5.0387 | 5.0404 | 5.0594 | 5.055 | 5.0591 | 5.0472 | 5.0516 | 5.0534 | 5.0528 | 5.0509 | 5.0564 | 5.0501 | 0.017365 | 0.0034386
356 | 5.0998 | 5.0845 | 5.0249 | 5.0391 | 5.0562 | 5.0572 | 5.0612 | 5.0787 | 5.0747 | 5.0385 | 5.0351 | 5.0397 | 5.0241 | 5.0277 | 5.0291 | 5.0391 | 5.0379 | 5.0422 | 5.0403 | 5.0382 | 5.04 | 5.0589 | 5.0545 | 5.0586 | 5.0468 | 5.0511 | 5.0529 | 5.0523 | 5.0504 | 5.0559 | 5.0497 | 0.017341 | 0.0034342
357 | 5.0991 | 5.0841 | 5.0244 | 5.0387 | 5.0558 | 5.0568 | 5.0607 | 5.0782 | 5.0743 | 5.038 | 5.0347 | 5.0393 | 5.0236 | 5.0272 | 5.0287 | 5.0386 | 5.0374 | 5.0417 | 5.0398 | 5.0378 | 5.0395 | 5.0584 | 5.0541 | 5.0582 | 5.0463 | 5.0507 | 5.0524 | 5.0519 | 5.0499 | 5.0554 | 5.0492 | 0.017318 | 0.0034298
358 | 5.0984 | 5.0836 | 5.024 | 5.0382 | 5.0553 | 5.0563 | 5.0603 | 5.0778 | 5.0738 | 5.0376 | 5.0343 | 5.0389 | 5.0232 | 5.0268 | 5.0282 | 5.0382 | 5.037 | 5.0413 | 5.0394 | 5.0373 | 5.0391 | 5.058 | 5.0536 | 5.0578 | 5.0459 | 5.0502 | 5.052 | 5.0514 | 5.0495 | 5.055 | 5.0487 | 0.017294 | 0.0034254
359 | 5.0977 | 5.0832 | 5.0235 | 5.0378 | 5.0549 | 5.0559 | 5.0598 | 5.0773 | 5.0734 | 5.0371 | 5.0339 | 5.0385 | 5.0227 | 5.0263 | 5.0278 | 5.0377 | 5.0365 | 5.0408 | 5.039 | 5.0369 | 5.0386 | 5.0575 | 5.0532 | 5.0573 | 5.0454 | 5.0498 | 5.0515 | 5.051 | 5.049 | 5.0545 | 5.0483 | 0.01727 | 0.003421
360 | 5.097 | 5.0828 | 5.023 | 5.0373 | 5.0544 | 5.0554 | 5.0594 | 5.0769 | 5.0729 | 5.0367 | 5.0335 | 5.0381 | 5.0223 | 5.0259 | 5.0273 | 5.0373 | 5.0361 | 5.0404 | 5.0385 | 5.0364 | 5.0382 | 5.0571 | 5.0528 | 5.0569 | 5.045 | 5.0493 | 5.0511 | 5.0505 | 5.0485 | 5.054 | 5.0478 | 0.017246 | 0.0034166
361 | 5.0963 | 5.0823 | 5.0226 | 5.0369 | 5.054 | 5.055 | 5.0589 | 5.0764 | 5.0725 | 5.0362 | 5.0331 | 5.0377 | 5.0219 | 5.0255 | 5.0269 | 5.0368 | 5.0356 | 5.0399 | 5.0381 | 5.036 | 5.0377 | 5.0566 | 5.0523 | 5.0564 | 5.0446 | 5.0489 | 5.0506 | 5.0501 | 5.0481 | 5.0536 | 5.0474 | 0.017223 | 0.0034122
362 | 5.0956 | 5.0819 | 5.0221 | 5.0365 | 5.0535 | 5.0545 | 5.0585 | 5.076 | 5.0721 | 5.0357 | 5.0327 | 5.0373 | 5.0214 | 5.025 | 5.0265 | 5.0364 | 5.0351 | 5.0395 | 5.0377 | 5.0355 | 5.0373 | 5.0562 | 5.0519 | 5.056 | 5.0441 | 5.0484 | 5.0502 | 5.0496 | 5.0476 | 5.0531 | 5.0469 | 0.017199 | 0.0034078
363 | 5.0949 | 5.0815 | 5.0216 | 5.036 | 5.0531 | 5.0541 | 5.0581 | 5.0755 | 5.0716 | 5.0353 | 5.0323 | 5.0369 | 5.021 | 5.0246 | 5.026 | 5.036 | 5.0347 | 5.039 | 5.0373 | 5.0351 | 5.0368 | 5.0557 | 5.0515 | 5.0556 | 5.0437 | 5.048 | 5.0497 | 5.0492 | 5.0472 | 5.0527 | 5.0465 | 0.017175 | 0.0034033
364 | 5.0943 | 5.0811 | 5.0212 | 5.0356 | 5.0527 | 5.0536 | 5.0576 | 5.0751 | 5.0712 | 5.0348 | 5.0319 | 5.0365 | 5.0206 | 5.0242 | 5.0256 | 5.0355 | 5.0343 | 5.0386 | 5.0369 | 5.0346 | 5.0364 | 5.0553 | 5.0511 | 5.0552 | 5.0433 | 5.0475 | 5.0493 | 5.0487 | 5.0467 | 5.0522 | 5.0461 | 0.017151 | 0.0033989
365 | 5.0936 | 5.0807 | 5.0207 | 5.0352 | 5.0522 | 5.0532 | 5.0572 | 5.0747 | 5.0708 | 5.0344 | 5.0316 | 5.0361 | 5.0202 | 5.0237 | 5.0252 | 5.0351 | 5.0338 | 5.0382 | 5.0364 | 5.0342 | 5.036 | 5.0548 | 5.0507 | 5.0547 | 5.0428 | 5.0471 | 5.0489 | 5.0483 | 5.0463 | 5.0518 | 5.0456 | 0.017127 | 0.0033945
366 | 5.0929 | 5.0803 | 5.0203 | 5.0348 | 5.0518 | 5.0528 | 5.0568 | 5.0742 | 5.0703 | 5.034 | 5.0312 | 5.0357 | 5.0198 | 5.0233 | 5.0247 | 5.0347 | 5.0334 | 5.0377 | 5.036 | 5.0338 | 5.0356 | 5.0544 | 5.0502 | 5.0543 | 5.0424 | 5.0467 | 5.0484 | 5.0479 | 5.0459 | 5.0514 | 5.0452 | 0.017104 | 0.0033901
367 | 5.0922 | 5.0799 | 5.0198 | 5.0343 | 5.0514 | 5.0524 | 5.0563 | 5.0738 | 5.0699 | 5.0335 | 5.0308 | 5.0353 | 5.0194 | 5.0229 | 5.0243 | 5.0343 | 5.0329 | 5.0373 | 5.0356 | 5.0333 | 5.0351 | 5.054 | 5.0498 | 5.0539 | 5.042 | 5.0462 | 5.048 | 5.0474 | 5.0454 | 5.0509 | 5.0448 | 0.01708 | 0.0033857
368 | 5.0916 | 5.0795 | 5.0194 | 5.0339 | 5.0509 | 5.0519 | 5.0559 | 5.0734 | 5.0695 | 5.0331 | 5.0304 | 5.0349 | 5.019 | 5.0225 | 5.0239 | 5.0339 | 5.0325 | 5.0369 | 5.0352 | 5.0329 | 5.0347 | 5.0535 | 5.0494 | 5.0535 | 5.0416 | 5.0458 | 5.0476 | 5.047 | 5.045 | 5.0505 | 5.0443 | 0.017056 | 0.0033813
369 | 5.0909 | 5.0791 | 5.0189 | 5.0335 | 5.0505 | 5.0515 | 5.0555 | 5.073 | 5.0691 | 5.0326 | 5.03 | 5.0345 | 5.0186 | 5.0221 | 5.0235 | 5.0334 | 5.0321 | 5.0365 | 5.0348 | 5.0325 | 5.0343 | 5.0531 | 5.049 | 5.0531 | 5.0412 | 5.0454 | 5.0472 | 5.0466 | 5.0446 | 5.0501 | 5.0439 | 0.017033 | 0.0033769
370 | 5.0902 | 5.0787 | 5.0185 | 5.0331 | 5.0501 | 5.0511 | 5.0551 | 5.0726 | 5.0687 | 5.0322 | 5.0297 | 5.0341 | 5.0182 | 5.0217 | 5.0231 | 5.033 | 5.0316 | 5.0361 | 5.0344 | 5.0321 | 5.0339 | 5.0527 | 5.0486 | 5.0527 | 5.0408 | 5.045 | 5.0467 | 5.0461 | 5.0441 | 5.0496 | 5.0435 | 0.017009 | 0.0033725
371 | 5.0896 | 5.0783 | 5.0181 | 5.0327 | 5.0497 | 5.0507 | 5.0547 | 5.0721 | 5.0683 | 5.0318 | 5.0293 | 5.0338 | 5.0178 | 5.0212 | 5.0227 | 5.0326 | 5.0312 | 5.0356 | 5.034 | 5.0317 | 5.0335 | 5.0523 | 5.0482 | 5.0523 | 5.0403 | 5.0446 | 5.0463 | 5.0457 | 5.0437 | 5.0492 | 5.0431 | 0.016985 | 0.0033681
372 | 5.0889 | 5.0779 | 5.0176 | 5.0323 | 5.0493 | 5.0502 | 5.0543 | 5.0717 | 5.0678 | 5.0313 | 5.0289 | 5.0334 | 5.0174 | 5.0208 | 5.0223 | 5.0322 | 5.0308 | 5.0352 | 5.0336 | 5.0312 | 5.033 | 5.0518 | 5.0478 | 5.0519 | 5.0399 | 5.0442 | 5.0459 | 5.0453 | 5.0433 | 5.0488 | 5.0426 | 0.016962 | 0.0033636
373 | 5.0882 | 5.0775 | 5.0172 | 5.0319 | 5.0489 | 5.0498 | 5.0539 | 5.0713 | 5.0674 | 5.0309 | 5.0286 | 5.033 | 5.017 | 5.0204 | 5.0219 | 5.0318 | 5.0304 | 5.0348 | 5.0332 | 5.0308 | 5.0326 | 5.0514 | 5.0474 | 5.0515 | 5.0395 | 5.0437 | 5.0455 | 5.0449 | 5.0429 | 5.0484 | 5.0422 | 0.016938 | 0.0033593
374 | 5.0876 | 5.0771 | 5.0168 | 5.0315 | 5.0485 | 5.0494 | 5.0534 | 5.0709 | 5.067 | 5.0305 | 5.0282 | 5.0327 | 5.0166 | 5.02 | 5.0215 | 5.0314 | 5.03 | 5.0344 | 5.0328 | 5.0304 | 5.0322 | 5.051 | 5.047 | 5.0511 | 5.0391 | 5.0433 | 5.0451 | 5.0445 | 5.0425 | 5.048 | 5.0418 | 0.016915 | 0.0033549
375 | 5.0869 | 5.0767 | 5.0164 | 5.0311 | 5.048 | 5.049 | 5.053 | 5.0705 | 5.0666 | 5.0301 | 5.0278 | 5.0323 | 5.0162 | 5.0196 | 5.0211 | 5.031 | 5.0296 | 5.034 | 5.0325 | 5.03 | 5.0318 | 5.0506 | 5.0466 | 5.0507 | 5.0387 | 5.0429 | 5.0447 | 5.0441 | 5.042 | 5.0476 | 5.0414 | 0.016891 | 0.0033505
376 | 5.0863 | 5.0763 | 5.0159 | 5.0307 | 5.0476 | 5.0486 | 5.0526 | 5.0701 | 5.0662 | 5.0296 | 5.0275 | 5.0319 | 5.0158 | 5.0192 | 5.0207 | 5.0306 | 5.0291 | 5.0336 | 5.0321 | 5.0296 | 5.0314 | 5.0502 | 5.0462 | 5.0503 | 5.0383 | 5.0425 | 5.0443 | 5.0437 | 5.0416 | 5.0471 | 5.041 | 0.016868 | 0.0033461
377 | 5.0856 | 5.0759 | 5.0155 | 5.0303 | 5.0472 | 5.0482 | 5.0523 | 5.0697 | 5.0658 | 5.0292 | 5.0271 | 5.0316 | 5.0154 | 5.0189 | 5.0203 | 5.0302 | 5.0287 | 5.0332 | 5.0317 | 5.0292 | 5.031 | 5.0498 | 5.0458 | 5.0499 | 5.038 | 5.0421 | 5.0439 | 5.0433 | 5.0412 | 5.0467 | 5.0406 | 0.016844 | 0.0033417
378 | 5.085 | 5.0756 | 5.0151 | 5.0299 | 5.0469 | 5.0478 | 5.0519 | 5.0693 | 5.0654 | 5.0288 | 5.0268 | 5.0312 | 5.015 | 5.0185 | 5.0199 | 5.0298 | 5.0283 | 5.0328 | 5.0313 | 5.0288 | 5.0306 | 5.0494 | 5.0455 | 5.0495 | 5.0376 | 5.0417 | 5.0435 | 5.0429 | 5.0408 | 5.0463 | 5.0402 | 0.016821 | 0.0033373
379 | 5.0843 | 5.0752 | 5.0147 | 5.0295 | 5.0465 | 5.0474 | 5.0515 | 5.0689 | 5.0651 | 5.0284 | 5.0264 | 5.0308 | 5.0147 | 5.0181 | 5.0195 | 5.0294 | 5.0279 | 5.0324 | 5.0309 | 5.0284 | 5.0303 | 5.049 | 5.0451 | 5.0492 | 5.0372 | 5.0413 | 5.0431 | 5.0425 | 5.0404 | 5.0459 | 5.0398 | 0.016797 | 0.0033329
380 | 5.0837 | 5.0748 | 5.0143 | 5.0291 | 5.0461 | 5.047 | 5.0511 | 5.0685 | 5.0647 | 5.028 | 5.0261 | 5.0305 | 5.0143 | 5.0177 | 5.0191 | 5.029 | 5.0275 | 5.032 | 5.0306 | 5.028 | 5.0299 | 5.0486 | 5.0447 | 5.0488 | 5.0368 | 5.0409 | 5.0427 | 5.0421 | 5.04 | 5.0455 | 5.0394 | 0.016774 | 0.0033286
381 | 5.083 | 5.0745 | 5.0139 | 5.0287 | 5.0457 | 5.0466 | 5.0507 | 5.0681 | 5.0643 | 5.0276 | 5.0257 | 5.0301 | 5.0139 | 5.0173 | 5.0187 | 5.0287 | 5.0271 | 5.0316 | 5.0302 | 5.0277 | 5.0295 | 5.0482 | 5.0443 | 5.0484 | 5.0364 | 5.0406 | 5.0423 | 5.0417 | 5.0396 | 5.0451 | 5.039 | 0.016751 | 0.0033242
382 | 5.0824 | 5.0741 | 5.0135 | 5.0284 | 5.0453 | 5.0462 | 5.0503 | 5.0677 | 5.0639 | 5.0272 | 5.0254 | 5.0298 | 5.0136 | 5.0169 | 5.0184 | 5.0283 | 5.0267 | 5.0313 | 5.0298 | 5.0273 | 5.0291 | 5.0478 | 5.044 | 5.048 | 5.036 | 5.0402 | 5.0419 | 5.0413 | 5.0392 | 5.0447 | 5.0386 | 0.016728 | 0.0033199
383 | 5.0817 | 5.0737 | 5.0131 | 5.028 | 5.0449 | 5.0459 | 5.0499 | 5.0674 | 5.0635 | 5.0268 | 5.0251 | 5.0294 | 5.0132 | 5.0166 | 5.018 | 5.0279 | 5.0264 | 5.0309 | 5.0295 | 5.0269 | 5.0287 | 5.0474 | 5.0436 | 5.0477 | 5.0357 | 5.0398 | 5.0415 | 5.0409 | 5.0388 | 5.0443 | 5.0382 | 0.016704 | 0.0033155
384 | 5.0811 | 5.0734 | 5.0127 | 5.0276 | 5.0445 | 5.0455 | 5.0495 | 5.067 | 5.0631 | 5.0264 | 5.0247 | 5.0291 | 5.0128 | 5.0162 | 5.0176 | 5.0275 | 5.026 | 5.0305 | 5.0291 | 5.0265 | 5.0283 | 5.047 | 5.0432 | 5.0473 | 5.0353 | 5.0394 | 5.0412 | 5.0405 | 5.0385 | 5.044 | 5.0379 | 0.016681 | 0.0033112
385 | 5.0804 | 5.073 | 5.0123 | 5.0272 | 5.0442 | 5.0451 | 5.0492 | 5.0666 | 5.0628 | 5.026 | 5.0244 | 5.0288 | 5.0125 | 5.0158 | 5.0172 | 5.0272 | 5.0256 | 5.0301 | 5.0287 | 5.0261 | 5.028 | 5.0466 | 5.0429 | 5.0469 | 5.0349 | 5.039 | 5.0408 | 5.0401 | 5.0381 | 5.0436 | 5.0375 | 0.016658 | 0.0033069
386 | 5.0798 | 5.0727 | 5.0119 | 5.0269 | 5.0438 | 5.0447 | 5.0488 | 5.0662 | 5.0624 | 5.0256 | 5.0241 | 5.0284 | 5.0121 | 5.0155 | 5.0169 | 5.0268 | 5.0252 | 5.0297 | 5.0284 | 5.0257 | 5.0276 | 5.0463 | 5.0425 | 5.0466 | 5.0345 | 5.0387 | 5.0404 | 5.0397 | 5.0377 | 5.0432 | 5.0371 | 0.016635 | 0.0033026
387 | 5.0792 | 5.0723 | 5.0115 | 5.0265 | 5.0434 | 5.0443 | 5.0484 | 5.0659 | 5.062 | 5.0252 | 5.0237 | 5.0281 | 5.0118 | 5.0151 | 5.0165 | 5.0264 | 5.0248 | 5.0294 | 5.028 | 5.0254 | 5.0272 | 5.0459 | 5.0421 | 5.0462 | 5.0342 | 5.0383 | 5.04 | 5.0394 | 5.0373 | 5.0428 | 5.0367 | 0.016612 | 0.0032983
388 | 5.0786 | 5.072 | 5.0111 | 5.0261 | 5.043 | 5.044 | 5.0481 | 5.0655 | 5.0617 | 5.0248 | 5.0234 | 5.0278 | 5.0114 | 5.0147 | 5.0161 | 5.0261 | 5.0244 | 5.029 | 5.0277 | 5.025 | 5.0268 | 5.0455 | 5.0418 | 5.0458 | 5.0338 | 5.0379 | 5.0397 | 5.039 | 5.0369 | 5.0424 | 5.0363 | 0.01659 | 0.003294
389 | 5.0779 | 5.0716 | 5.0107 | 5.0258 | 5.0427 | 5.0436 | 5.0477 | 5.0651 | 5.0613 | 5.0244 | 5.0231 | 5.0274 | 5.011 | 5.0144 | 5.0158 | 5.0257 | 5.0241 | 5.0286 | 5.0273 | 5.0246 | 5.0265 | 5.0451 | 5.0414 | 5.0455 | 5.0335 | 5.0375 | 5.0393 | 5.0386 | 5.0366 | 5.0421 | 5.036 | 0.016567 | 0.0032897
390 | 5.0773 | 5.0713 | 5.0103 | 5.0254 | 5.0423 | 5.0432 | 5.0473 | 5.0648 | 5.0609 | 5.024 | 5.0228 | 5.0271 | 5.0107 | 5.014 | 5.0154 | 5.0253 | 5.0237 | 5.0283 | 5.027 | 5.0243 | 5.0261 | 5.0448 | 5.0411 | 5.0451 | 5.0331 | 5.0372 | 5.0389 | 5.0382 | 5.0362 | 5.0417 | 5.0356 | 0.016544 | 0.0032854
391 | 5.0767 | 5.0709 | 5.0099 | 5.0251 | 5.0419 | 5.0429 | 5.047 | 5.0644 | 5.0606 | 5.0236 | 5.0225 | 5.0268 | 5.0104 | 5.0137 | 5.0151 | 5.025 | 5.0233 | 5.0279 | 5.0266 | 5.0239 | 5.0257 | 5.0444 | 5.0407 | 5.0448 | 5.0327 | 5.0368 | 5.0386 | 5.0379 | 5.0358 | 5.0413 | 5.0352 | 0.016521 | 0.0032812
392 | 5.076 | 5.0706 | 5.0096 | 5.0247 | 5.0416 | 5.0425 | 5.0466 | 5.064 | 5.0602 | 5.0232 | 5.0221 | 5.0265 | 5.01 | 5.0133 | 5.0147 | 5.0246 | 5.0229 | 5.0275 | 5.0263 | 5.0235 | 5.0254 | 5.044 | 5.0404 | 5.0444 | 5.0324 | 5.0365 | 5.0382 | 5.0375 | 5.0354 | 5.0409 | 5.0349 | 0.016499 | 0.0032769
393 | 5.0754 | 5.0702 | 5.0092 | 5.0243 | 5.0412 | 5.0421 | 5.0463 | 5.0637 | 5.0599 | 5.0228 | 5.0218 | 5.0261 | 5.0097 | 5.013 | 5.0144 | 5.0243 | 5.0226 | 5.0272 | 5.026 | 5.0232 | 5.025 | 5.0437 | 5.04 | 5.0441 | 5.032 | 5.0361 | 5.0378 | 5.0371 | 5.0351 | 5.0406 | 5.0345 | 0.016476 | 0.0032727
394 | 5.0748 | 5.0699 | 5.0088 | 5.024 | 5.0409 | 5.0418 | 5.0459 | 5.0633 | 5.0595 | 5.0225 | 5.0215 | 5.0258 | 5.0093 | 5.0126 | 5.014 | 5.0239 | 5.0222 | 5.0268 | 5.0256 | 5.0228 | 5.0247 | 5.0433 | 5.0397 | 5.0437 | 5.0317 | 5.0357 | 5.0375 | 5.0368 | 5.0347 | 5.0402 | 5.0341 | 0.016454 | 0.0032685
395 | 5.0742 | 5.0696 | 5.0084 | 5.0236 | 5.0405 | 5.0414 | 5.0456 | 5.063 | 5.0592 | 5.0221 | 5.0212 | 5.0255 | 5.009 | 5.0123 | 5.0137 | 5.0236 | 5.0219 | 5.0265 | 5.0253 | 5.0225 | 5.0243 | 5.0429 | 5.0394 | 5.0434 | 5.0313 | 5.0354 | 5.0371 | 5.0364 | 5.0344 | 5.0398 | 5.0338 | 0.016432 | 0.0032643
396 | 5.0736 | 5.0692 | 5.0081 | 5.0233 | 5.0402 | 5.0411 | 5.0452 | 5.0626 | 5.0588 | 5.0217 | 5.0209 | 5.0252 | 5.0087 | 5.0119 | 5.0133 | 5.0232 | 5.0215 | 5.0261 | 5.025 | 5.0221 | 5.024 | 5.0426 | 5.039 | 5.043 | 5.031 | 5.035 | 5.0368 | 5.0361 | 5.034 | 5.0395 | 5.0334 | 0.016409 | 0.0032601
397 | 5.073 | 5.0689 | 5.0077 | 5.023 | 5.0398 | 5.0407 | 5.0449 | 5.0623 | 5.0585 | 5.0213 | 5.0206 | 5.0249 | 5.0083 | 5.0116 | 5.013 | 5.0229 | 5.0211 | 5.0258 | 5.0246 | 5.0218 | 5.0236 | 5.0422 | 5.0387 | 5.0427 | 5.0306 | 5.0347 | 5.0364 | 5.0357 | 5.0336 | 5.0391 | 5.0331 | 0.016387 | 0.0032559
398 | 5.0723 | 5.0686 | 5.0073 | 5.0226 | 5.0395 | 5.0404 | 5.0445 | 5.0619 | 5.0581 | 5.021 | 5.0203 | 5.0246 | 5.008 | 5.0113 | 5.0127 | 5.0225 | 5.0208 | 5.0254 | 5.0243 | 5.0214 | 5.0233 | 5.0419 | 5.0384 | 5.0424 | 5.0303 | 5.0343 | 5.0361 | 5.0354 | 5.0333 | 5.0388 | 5.0327 | 0.016365 | 0.0032518
399 | 5.0717 | 5.0682 | 5.007 | 5.0223 | 5.0391 | 5.04 | 5.0442 | 5.0616 | 5.0578 | 5.0206 | 5.02 | 5.0243 | 5.0077 | 5.0109 | 5.0123 | 5.0222 | 5.0204 | 5.0251 | 5.024 | 5.0211 | 5.0229 | 5.0415 | 5.038 | 5.042 | 5.03 | 5.034 | 5.0357 | 5.035 | 5.0329 | 5.0384 | 5.0324 | 0.016343 | 0.0032476
400 | 5.0711 | 5.0679 | 5.0066 | 5.0219 | 5.0388 | 5.0397 | 5.0438 | 5.0612 | 5.0575 | 5.0202 | 5.0197 | 5.024 | 5.0074 | 5.0106 | 5.012 | 5.0219 | 5.0201 | 5.0247 | 5.0236 | 5.0207 | 5.0226 | 5.0412 | 5.0377 | 5.0417 | 5.0296 | 5.0337 | 5.0354 | 5.0347 | 5.0326 | 5.0381 | 5.032 | 0.016321 | 0.0032435
401 | 5.0705 | 5.0676 | 5.0063 | 5.0216 | 5.0384 | 5.0393 | 5.0435 | 5.0609 | 5.0571 | 5.0198 | 5.0194 | 5.0237 | 5.007 | 5.0103 | 5.0117 | 5.0215 | 5.0197 | 5.0244 | 5.0233 | 5.0204 | 5.0223 | 5.0408 | 5.0374 | 5.0414 | 5.0293 | 5.0333 | 5.035 | 5.0343 | 5.0322 | 5.0377 | 5.0317 | 0.0163 | 0.0032394
402 | 5.0699 | 5.0673 | 5.0059 | 5.0213 | 5.0381 | 5.039 | 5.0432 | 5.0605 | 5.0568 | 5.0195 | 5.0191 | 5.0234 | 5.0067 | 5.0099 | 5.0113 | 5.0212 | 5.0194 | 5.0241 | 5.023 | 5.0201 | 5.0219 | 5.0405 | 5.037 | 5.0411 | 5.029 | 5.033 | 5.0347 | 5.034 | 5.0319 | 5.0374 | 5.0313 | 0.016278 | 0.0032353
403 | 5.0693 | 5.067 | 5.0056 | 5.0209 | 5.0378 | 5.0387 | 5.0428 | 5.0602 | 5.0565 | 5.0191 | 5.0188 | 5.0231 | 5.0064 | 5.0096 | 5.011 | 5.0209 | 5.019 | 5.0237 | 5.0227 | 5.0197 | 5.0216 | 5.0401 | 5.0367 | 5.0407 | 5.0286 | 5.0326 | 5.0344 | 5.0336 | 5.0315 | 5.037 | 5.031 | 0.016256 | 0.0032312
404 | 5.0687 | 5.0667 | 5.0052 | 5.0206 | 5.0374 | 5.0383 | 5.0425 | 5.0599 | 5.0561 | 5.0188 | 5.0185 | 5.0228 | 5.0061 | 5.0093 | 5.0107 | 5.0205 | 5.0187 | 5.0234 | 5.0224 | 5.0194 | 5.0213 | 5.0398 | 5.0364 | 5.0404 | 5.0283 | 5.0323 | 5.034 | 5.0333 | 5.0312 | 5.0367 | 5.0307 | 0.016235 | 0.0032272
405 | 5.0681 | 5.0663 | 5.0049 | 5.0203 | 5.0371 | 5.038 | 5.0422 | 5.0595 | 5.0558 | 5.0184 | 5.0183 | 5.0225 | 5.0058 | 5.009 | 5.0104 | 5.0202 | 5.0184 | 5.0231 | 5.0221 | 5.0191 | 5.0209 | 5.0395 | 5.0361 | 5.0401 | 5.028 | 5.032 | 5.0337 | 5.033 | 5.0309 | 5.0364 | 5.0303 | 0.016213 | 0.0032231
406 | 5.0675 | 5.066 | 5.0045 | 5.02 | 5.0368 | 5.0377 | 5.0418 | 5.0592 | 5.0555 | 5.018 | 5.018 | 5.0222 | 5.0055 | 5.0086 | 5.01 | 5.0199 | 5.018 | 5.0227 | 5.0218 | 5.0187 | 5.0206 | 5.0391 | 5.0358 | 5.0398 | 5.0277 | 5.0317 | 5.0334 | 5.0326 | 5.0305 | 5.036 | 5.03 | 0.016192 | 0.0032191
407 | 5.0669 | 5.0657 | 5.0042 | 5.0196 | 5.0365 | 5.0373 | 5.0415 | 5.0589 | 5.0552 | 5.0177 | 5.0177 | 5.0219 | 5.0052 | 5.0083 | 5.0097 | 5.0196 | 5.0177 | 5.0224 | 5.0214 | 5.0184 | 5.0203 | 5.0388 | 5.0355 | 5.0395 | 5.0273 | 5.0313 | 5.033 | 5.0323 | 5.0302 | 5.0357 | 5.0297 | 0.016171 | 0.0032151
408 | 5.0663 | 5.0654 | 5.0038 | 5.0193 | 5.0361 | 5.037 | 5.0412 | 5.0586 | 5.0548 | 5.0173 | 5.0174 | 5.0216 | 5.0049 | 5.008 | 5.0094 | 5.0192 | 5.0174 | 5.0221 | 5.0211 | 5.0181 | 5.0199 | 5.0385 | 5.0351 | 5.0391 | 5.027 | 5.031 | 5.0327 | 5.032 | 5.0299 | 5.0353 | 5.0293 | 0.01615 | 0.0032111
409 | 5.0657 | 5.0651 | 5.0035 | 5.019 | 5.0358 | 5.0367 | 5.0409 | 5.0582 | 5.0545 | 5.017 | 5.0171 | 5.0213 | 5.0045 | 5.0077 | 5.0091 | 5.0189 | 5.017 | 5.0218 | 5.0208 | 5.0178 | 5.0196 | 5.0381 | 5.0348 | 5.0388 | 5.0267 | 5.0307 | 5.0324 | 5.0316 | 5.0295 | 5.035 | 5.029 | 0.016129 | 0.0032071
410 | 5.0651 | 5.0648 | 5.0031 | 5.0187 | 5.0355 | 5.0364 | 5.0406 | 5.0579 | 5.0542 | 5.0166 | 5.0169 | 5.021 | 5.0042 | 5.0074 | 5.0088 | 5.0186 | 5.0167 | 5.0214 | 5.0205 | 5.0174 | 5.0193 | 5.0378 | 5.0345 | 5.0385 | 5.0264 | 5.0304 | 5.0321 | 5.0313 | 5.0292 | 5.0347 | 5.0287 | 0.016108 | 0.0032032
411 | 5.0645 | 5.0645 | 5.0028 | 5.0184 | 5.0352 | 5.036 | 5.0402 | 5.0576 | 5.0539 | 5.0163 | 5.0166 | 5.0207 | 5.0039 | 5.0071 | 5.0085 | 5.0183 | 5.0164 | 5.0211 | 5.0202 | 5.0171 | 5.019 | 5.0375 | 5.0342 | 5.0382 | 5.0261 | 5.03 | 5.0317 | 5.031 | 5.0289 | 5.0344 | 5.0283 | 0.016087 | 0.0031993
412 | 5.0639 | 5.0642 | 5.0025 | 5.018 | 5.0349 | 5.0357 | 5.0399 | 5.0573 | 5.0536 | 5.0159 | 5.0163 | 5.0205 | 5.0037 | 5.0068 | 5.0082 | 5.018 | 5.016 | 5.0208 | 5.0199 | 5.0168 | 5.0187 | 5.0371 | 5.0339 | 5.0379 | 5.0258 | 5.0297 | 5.0314 | 5.0307 | 5.0286 | 5.034 | 5.028 | 0.016066 | 0.0031954
413 | 5.0634 | 5.0639 | 5.0021 | 5.0177 | 5.0345 | 5.0354 | 5.0396 | 5.057 | 5.0533 | 5.0156 | 5.016 | 5.0202 | 5.0034 | 5.0064 | 5.0078 | 5.0177 | 5.0157 | 5.0205 | 5.0196 | 5.0165 | 5.0184 | 5.0368 | 5.0336 | 5.0376 | 5.0255 | 5.0294 | 5.0311 | 5.0303 | 5.0282 | 5.0337 | 5.0277 | 0.016046 | 0.0031915
414 | 5.0628 | 5.0636 | 5.0018 | 5.0174 | 5.0342 | 5.0351 | 5.0393 | 5.0566 | 5.053 | 5.0152 | 5.0158 | 5.0199 | 5.0031 | 5.0061 | 5.0075 | 5.0174 | 5.0154 | 5.0202 | 5.0193 | 5.0162 | 5.018 | 5.0365 | 5.0333 | 5.0373 | 5.0252 | 5.0291 | 5.0308 | 5.03 | 5.0279 | 5.0334 | 5.0274 | 0.016025 | 0.0031876
415 | 5.0622 | 5.0633 | 5.0015 | 5.0171 | 5.0339 | 5.0348 | 5.039 | 5.0563 | 5.0526 | 5.0149 | 5.0155 | 5.0196 | 5.0028 | 5.0058 | 5.0072 | 5.0171 | 5.0151 | 5.0199 | 5.019 | 5.0159 | 5.0177 | 5.0362 | 5.033 | 5.037 | 5.0249 | 5.0288 | 5.0305 | 5.0297 | 5.0276 | 5.0331 | 5.0271 | 0.016005 | 0.0031838
416 | 5.0616 | 5.063 | 5.0012 | 5.0168 | 5.0336 | 5.0345 | 5.0387 | 5.056 | 5.0523 | 5.0145 | 5.0153 | 5.0194 | 5.0025 | 5.0055 | 5.0069 | 5.0168 | 5.0148 | 5.0196 | 5.0188 | 5.0156 | 5.0174 | 5.0359 | 5.0327 | 5.0367 | 5.0246 | 5.0285 | 5.0302 | 5.0294 | 5.0273 | 5.0328 | 5.0268 | 0.015985 | 0.00318
417 | 5.061 | 5.0628 | 5.0008 | 5.0165 | 5.0333 | 5.0342 | 5.0384 | 5.0557 | 5.052 | 5.0142 | 5.015 | 5.0191 | 5.0022 | 5.0052 | 5.0066 | 5.0165 | 5.0145 | 5.0193 | 5.0185 | 5.0153 | 5.0171 | 5.0356 | 5.0324 | 5.0364 | 5.0242 | 5.0282 | 5.0299 | 5.0291 | 5.027 | 5.0324 | 5.0264 | 0.015965 | 0.0031762
418 | 5.0605 | 5.0625 | 5.0005 | 5.0162 | 5.033 | 5.0339 | 5.0381 | 5.0554 | 5.0517 | 5.0139 | 5.0147 | 5.0188 | 5.0019 | 5.0049 | 5.0064 | 5.0162 | 5.0141 | 5.019 | 5.0182 | 5.0149 | 5.0168 | 5.0352 | 5.0321 | 5.0361 | 5.024 | 5.0279 | 5.0296 | 5.0288 | 5.0267 | 5.0321 | 5.0261 | 0.015945 | 0.0031724
419 | 5.0599 | 5.0622 | 5.0002 | 5.0159 | 5.0327 | 5.0336 | 5.0378 | 5.0551 | 5.0514 | 5.0135 | 5.0145 | 5.0186 | 5.0016 | 5.0047 | 5.0061 | 5.0159 | 5.0138 | 5.0187 | 5.0179 | 5.0146 | 5.0165 | 5.0349 | 5.0319 | 5.0358 | 5.0237 | 5.0276 | 5.0292 | 5.0285 | 5.0263 | 5.0318 | 5.0258 | 0.015925 | 0.0031687
420 | 5.0593 | 5.0619 | 4.9999 | 5.0156 | 5.0324 | 5.0332 | 5.0375 | 5.0548 | 5.0511 | 5.0132 | 5.0142 | 5.0183 | 5.0013 | 5.0044 | 5.0058 | 5.0156 | 5.0135 | 5.0184 | 5.0176 | 5.0143 | 5.0162 | 5.0346 | 5.0316 | 5.0355 | 5.0234 | 5.0273 | 5.0289 | 5.0281 | 5.026 | 5.0315 | 5.0255 | 0.015905 | 0.0031649
421 | 5.0587 | 5.0616 | 4.9996 | 5.0153 | 5.0321 | 5.0329 | 5.0372 | 5.0545 | 5.0508 | 5.0129 | 5.014 | 5.018 | 5.0011 | 5.0041 | 5.0055 | 5.0153 | 5.0132 | 5.0181 | 5.0173 | 5.014 | 5.0159 | 5.0343 | 5.0313 | 5.0352 | 5.0231 | 5.027 | 5.0286 | 5.0278 | 5.0257 | 5.0312 | 5.0252 | 0.015886 | 0.0031612
422 | 5.0582 | 5.0613 | 4.9993 | 5.015 | 5.0318 | 5.0326 | 5.0369 | 5.0542 | 5.0506 | 5.0125 | 5.0137 | 5.0178 | 5.0008 | 5.0038 | 5.0052 | 5.015 | 5.0129 | 5.0178 | 5.0171 | 5.0137 | 5.0156 | 5.034 | 5.031 | 5.035 | 5.0228 | 5.0267 | 5.0283 | 5.0275 | 5.0254 | 5.0309 | 5.0249 | 0.015866 | 0.0031575
423 | 5.0576 | 5.0611 | 4.9989 | 5.0147 | 5.0315 | 5.0324 | 5.0366 | 5.0539 | 5.0503 | 5.0122 | 5.0135 | 5.0175 | 5.0005 | 5.0035 | 5.0049 | 5.0147 | 5.0126 | 5.0175 | 5.0168 | 5.0135 | 5.0153 | 5.0337 | 5.0307 | 5.0347 | 5.0225 | 5.0264 | 5.028 | 5.0272 | 5.0251 | 5.0306 | 5.0246 | 0.015847 | 0.0031539
424 | 5.057 | 5.0608 | 4.9986 | 5.0144 | 5.0312 | 5.0321 | 5.0363 | 5.0536 | 5.05 | 5.0119 | 5.0132 | 5.0173 | 5.0002 | 5.0032 | 5.0046 | 5.0144 | 5.0123 | 5.0172 | 5.0165 | 5.0132 | 5.015 | 5.0334 | 5.0304 | 5.0344 | 5.0222 | 5.0261 | 5.0277 | 5.0269 | 5.0248 | 5.0303 | 5.0243 | 0.015828 | 0.0031503
425 | 5.0565 | 5.0605 | 4.9983 | 5.0141 | 5.0309 | 5.0318 | 5.036 | 5.0533 | 5.0497 | 5.0115 | 5.013 | 5.017 | 5 | 5.0029 | 5.0043 | 5.0141 | 5.012 | 5.0169 | 5.0162 | 5.0129 | 5.0147 | 5.0331 | 5.0302 | 5.0341 | 5.0219 | 5.0258 | 5.0275 | 5.0266 | 5.0245 | 5.03 | 5.024 | 0.015809 | 0.0031467
426 | 5.0559 | 5.0602 | 4.998 | 5.0139 | 5.0306 | 5.0315 | 5.0357 | 5.053 | 5.0494 | 5.0112 | 5.0127 | 5.0168 | 4.9997 | 5.0027 | 5.0041 | 5.0138 | 5.0117 | 5.0166 | 5.016 | 5.0126 | 5.0144 | 5.0328 | 5.0299 | 5.0338 | 5.0216 | 5.0255 | 5.0272 | 5.0263 | 5.0242 | 5.0297 | 5.0237 | 0.01579 | 0.0031431
427 | 5.0553 | 5.06 | 4.9977 | 5.0136 | 5.0303 | 5.0312 | 5.0355 | 5.0527 | 5.0491 | 5.0109 | 5.0125 | 5.0165 | 4.9994 | 5.0024 | 5.0038 | 5.0135 | 5.0114 | 5.0163 | 5.0157 | 5.0123 | 5.0141 | 5.0325 | 5.0296 | 5.0336 | 5.0213 | 5.0252 | 5.0269 | 5.026 | 5.0239 | 5.0294 | 5.0234 | 0.015771 | 0.0031395
428 | 5.0548 | 5.0597 | 4.9974 | 5.0133 | 5.0301 | 5.0309 | 5.0352 | 5.0525 | 5.0488 | 5.0106 | 5.0122 | 5.0163 | 4.9992 | 5.0021 | 5.0035 | 5.0133 | 5.0111 | 5.016 | 5.0154 | 5.012 | 5.0139 | 5.0322 | 5.0293 | 5.0333 | 5.0211 | 5.0249 | 5.0266 | 5.0258 | 5.0236 | 5.0291 | 5.0231 | 0.015753 | 0.003136
429 | 5.0542 | 5.0594 | 4.9971 | 5.013 | 5.0298 | 5.0306 | 5.0349 | 5.0522 | 5.0485 | 5.0102 | 5.012 | 5.016 | 4.9989 | 5.0018 | 5.0032 | 5.013 | 5.0108 | 5.0157 | 5.0151 | 5.0117 | 5.0136 | 5.0319 | 5.0291 | 5.033 | 5.0208 | 5.0246 | 5.0263 | 5.0255 | 5.0233 | 5.0288 | 5.0228 | 0.015734 | 0.0031325
430 | 5.0536 | 5.0592 | 4.9968 | 5.0127 | 5.0295 | 5.0303 | 5.0346 | 5.0519 | 5.0483 | 5.0099 | 5.0118 | 5.0158 | 4.9986 | 5.0015 | 5.003 | 5.0127 | 5.0105 | 5.0154 | 5.0149 | 5.0114 | 5.0133 | 5.0316 | 5.0288 | 5.0327 | 5.0205 | 5.0244 | 5.026 | 5.0252 | 5.023 | 5.0285 | 5.0226 | 0.015716 | 0.003129
431 | 5.0531 | 5.0589 | 4.9965 | 5.0124 | 5.0292 | 5.03 | 5.0343 | 5.0516 | 5.048 | 5.0096 | 5.0115 | 5.0155 | 4.9984 | 5.0013 | 5.0027 | 5.0124 | 5.0102 | 5.0152 | 5.0146 | 5.0112 | 5.013 | 5.0314 | 5.0285 | 5.0325 | 5.0202 | 5.0241 | 5.0257 | 5.0249 | 5.0228 | 5.0282 | 5.0223 | 0.015698 | 0.0031256
432 | 5.0525 | 5.0586 | 4.9962 | 5.0122 | 5.0289 | 5.0298 | 5.0341 | 5.0513 | 5.0477 | 5.0093 | 5.0113 | 5.0153 | 4.9981 | 5.001 | 5.0024 | 5.0121 | 5.01 | 5.0149 | 5.0144 | 5.0109 | 5.0127 | 5.0311 | 5.0283 | 5.0322 | 5.02 | 5.0238 | 5.0254 | 5.0246 | 5.0225 | 5.0279 | 5.022 | 0.01568 | 0.0031222
433 | 5.052 | 5.0584 | 4.9959 | 5.0119 | 5.0287 | 5.0295 | 5.0338 | 5.051 | 5.0474 | 5.009 | 5.0111 | 5.015 | 4.9979 | 5.0007 | 5.0021 | 5.0119 | 5.0097 | 5.0146 | 5.0141 | 5.0106 | 5.0124 | 5.0308 | 5.028 | 5.0319 | 5.0197 | 5.0235 | 5.0252 | 5.0243 | 5.0222 | 5.0277 | 5.0217 | 0.015662 | 0.0031188
434 | 5.0514 | 5.0581 | 4.9956 | 5.0116 | 5.0284 | 5.0292 | 5.0335 | 5.0508 | 5.0472 | 5.0087 | 5.0108 | 5.0148 | 4.9976 | 5.0005 | 5.0019 | 5.0116 | 5.0094 | 5.0143 | 5.0138 | 5.0103 | 5.0122 | 5.0305 | 5.0277 | 5.0317 | 5.0194 | 5.0233 | 5.0249 | 5.024 | 5.0219 | 5.0274 | 5.0214 | 0.015644 | 0.0031154
435 | 5.0509 | 5.0579 | 4.9953 | 5.0113 | 5.0281 | 5.0289 | 5.0332 | 5.0505 | 5.0469 | 5.0083 | 5.0106 | 5.0146 | 4.9973 | 5.0002 | 5.0016 | 5.0113 | 5.0091 | 5.014 | 5.0136 | 5.01 | 5.0119 | 5.0302 | 5.0275 | 5.0314 | 5.0191 | 5.023 | 5.0246 | 5.0237 | 5.0216 | 5.0271 | 5.0211 | 0.015626 | 0.0031121
436 | 5.0503 | 5.0576 | 4.995 | 5.0111 | 5.0278 | 5.0286 | 5.033 | 5.0502 | 5.0466 | 5.008 | 5.0104 | 5.0143 | 4.9971 | 4.9999 | 5.0014 | 5.0111 | 5.0088 | 5.0138 | 5.0133 | 5.0098 | 5.0116 | 5.0299 | 5.0272 | 5.0311 | 5.0189 | 5.0227 | 5.0243 | 5.0235 | 5.0213 | 5.0268 | 5.0209 | 0.015609 | 0.0031088
437 | 5.0498 | 5.0574 | 4.9948 | 5.0108 | 5.0276 | 5.0284 | 5.0327 | 5.0499 | 5.0464 | 5.0077 | 5.0101 | 5.0141 | 4.9968 | 4.9997 | 5.0011 | 5.0108 | 5.0085 | 5.0135 | 5.0131 | 5.0095 | 5.0113 | 5.0297 | 5.0269 | 5.0309 | 5.0186 | 5.0224 | 5.0241 | 5.0232 | 5.0211 | 5.0265 | 5.0206 | 0.015591 | 0.0031055
438 | 5.0492 | 5.0571 | 4.9945 | 5.0105 | 5.0273 | 5.0281 | 5.0324 | 5.0497 | 5.0461 | 5.0074 | 5.0099 | 5.0139 | 4.9966 | 4.9994 | 5.0008 | 5.0105 | 5.0083 | 5.0132 | 5.0128 | 5.0092 | 5.0111 | 5.0294 | 5.0267 | 5.0306 | 5.0183 | 5.0222 | 5.0238 | 5.0229 | 5.0208 | 5.0262 | 5.0203 | 0.015574 | 0.0031023
439 | 5.0487 | 5.0569 | 4.9942 | 5.0103 | 5.027 | 5.0278 | 5.0322 | 5.0494 | 5.0458 | 5.0071 | 5.0097 | 5.0136 | 4.9963 | 4.9992 | 5.0006 | 5.0103 | 5.008 | 5.013 | 5.0126 | 5.009 | 5.0108 | 5.0291 | 5.0264 | 5.0304 | 5.0181 | 5.0219 | 5.0235 | 5.0226 | 5.0205 | 5.026 | 5.02 | 0.015557 | 0.003099
440 | 5.0481 | 5.0566 | 4.9939 | 5.01 | 5.0268 | 5.0276 | 5.0319 | 5.0491 | 5.0456 | 5.0068 | 5.0095 | 5.0134 | 4.9961 | 4.9989 | 5.0003 | 5.01 | 5.0077 | 5.0127 | 5.0123 | 5.0087 | 5.0105 | 5.0288 | 5.0262 | 5.0301 | 5.0178 | 5.0216 | 5.0232 | 5.0224 | 5.0202 | 5.0257 | 5.0198 | 0.01554 | 0.0030959
441 | 5.0476 | 5.0564 | 4.9936 | 5.0097 | 5.0265 | 5.0273 | 5.0316 | 5.0489 | 5.0453 | 5.0065 | 5.0092 | 5.0132 | 4.9958 | 4.9986 | 5.0001 | 5.0097 | 5.0074 | 5.0124 | 5.0121 | 5.0084 | 5.0103 | 5.0286 | 5.0259 | 5.0298 | 5.0176 | 5.0214 | 5.023 | 5.0221 | 5.02 | 5.0254 | 5.0195 | 0.015524 | 0.0030927
442 | 5.0471 | 5.0561 | 4.9933 | 5.0095 | 5.0262 | 5.027 | 5.0314 | 5.0486 | 5.0451 | 5.0062 | 5.009 | 5.0129 | 4.9956 | 4.9984 | 4.9998 | 5.0095 | 5.0072 | 5.0122 | 5.0118 | 5.0082 | 5.01 | 5.0283 | 5.0257 | 5.0296 | 5.0173 | 5.0211 | 5.0227 | 5.0218 | 5.0197 | 5.0251 | 5.0192 | 0.015507 | 0.0030896
443 | 5.0465 | 5.0559 | 4.9931 | 5.0092 | 5.026 | 5.0268 | 5.0311 | 5.0483 | 5.0448 | 5.0059 | 5.0088 | 5.0127 | 4.9954 | 4.9981 | 4.9996 | 5.0092 | 5.0069 | 5.0119 | 5.0116 | 5.0079 | 5.0097 | 5.028 | 5.0254 | 5.0293 | 5.0171 | 5.0208 | 5.0224 | 5.0216 | 5.0194 | 5.0249 | 5.0189 | 0.015491 | 0.0030865
444 | 5.046 | 5.0556 | 4.9928 | 5.009 | 5.0257 | 5.0265 | 5.0309 | 5.0481 | 5.0445 | 5.0056 | 5.0086 | 5.0125 | 4.9951 | 4.9979 | 4.9993 | 5.009 | 5.0066 | 5.0116 | 5.0113 | 5.0077 | 5.0095 | 5.0277 | 5.0252 | 5.0291 | 5.0168 | 5.0206 | 5.0222 | 5.0213 | 5.0192 | 5.0246 | 5.0187 | 0.015475 | 0.0030834
445 | 5.0454 | 5.0554 | 4.9925 | 5.0087 | 5.0255 | 5.0262 | 5.0306 | 5.0478 | 5.0443 | 5.0053 | 5.0084 | 5.0123 | 4.9949 | 4.9976 | 4.9991 | 5.0087 | 5.0064 | 5.0114 | 5.0111 | 5.0074 | 5.0092 | 5.0275 | 5.0249 | 5.0288 | 5.0165 | 5.0203 | 5.0219 | 5.021 | 5.0189 | 5.0243 | 5.0184 | 0.015458 | 0.0030803
446 | 5.0449 | 5.0551 | 4.9922 | 5.0085 | 5.0252 | 5.026 | 5.0304 | 5.0476 | 5.044 | 5.005 | 5.0082 | 5.012 | 4.9946 | 4.9974 | 4.9988 | 5.0085 | 5.0061 | 5.0111 | 5.0109 | 5.0071 | 5.009 | 5.0272 | 5.0247 | 5.0286 | 5.0163 | 5.0201 | 5.0217 | 5.0208 | 5.0186 | 5.0241 | 5.0182 | 0.015443 | 0.0030773
447 | 5.0444 | 5.0549 | 4.992 | 5.0082 | 5.025 | 5.0257 | 5.0301 | 5.0473 | 5.0438 | 5.0047 | 5.008 | 5.0118 | 4.9944 | 4.9971 | 4.9986 | 5.0082 | 5.0058 | 5.0109 | 5.0106 | 5.0069 | 5.0087 | 5.0269 | 5.0244 | 5.0283 | 5.016 | 5.0198 | 5.0214 | 5.0205 | 5.0184 | 5.0238 | 5.0179 | 0.015427 | 0.0030744
448 | 5.0438 | 5.0547 | 4.9917 | 5.0079 | 5.0247 | 5.0255 | 5.0299 | 5.047 | 5.0435 | 5.0044 | 5.0078 | 5.0116 | 4.9942 | 4.9969 | 4.9983 | 5.008 | 5.0056 | 5.0106 | 5.0104 | 5.0066 | 5.0085 | 5.0267 | 5.0242 | 5.0281 | 5.0158 | 5.0196 | 5.0211 | 5.0202 | 5.0181 | 5.0236 | 5.0176 | 0.015411 | 0.0030714
449 | 5.0433 | 5.0544 | 4.9914 | 5.0077 | 5.0244 | 5.0252 | 5.0296 | 5.0468 | 5.0433 | 5.0041 | 5.0075 | 5.0114 | 4.9939 | 4.9967 | 4.9981 | 5.0077 | 5.0053 | 5.0104 | 5.0102 | 5.0064 | 5.0082 | 5.0264 | 5.024 | 5.0279 | 5.0155 | 5.0193 | 5.0209 | 5.02 | 5.0178 | 5.0233 | 5.0174 | 0.015396 | 0.0030685
450 | 5.0428 | 5.0542 | 4.9912 | 5.0074 | 5.0242 | 5.025 | 5.0294 | 5.0465 | 5.043 | 5.0038 | 5.0073 | 5.0112 | 4.9937 | 4.9964 | 4.9978 | 5.0075 | 5.0051 | 5.0101 | 5.0099 | 5.0061 | 5.0079 | 5.0262 | 5.0237 | 5.0276 | 5.0153 | 5.0191 | 5.0206 | 5.0197 | 5.0176 | 5.023 | 5.0171 | 0.015381 | 0.0030656
451 | 5.0423 | 5.054 | 4.9909 | 5.0072 | 5.0239 | 5.0247 | 5.0291 | 5.0463 | 5.0428 | 5.0035 | 5.0071 | 5.011 | 4.9935 | 4.9962 | 4.9976 | 5.0072 | 5.0048 | 5.0099 | 5.0097 | 5.0059 | 5.0077 | 5.0259 | 5.0235 | 5.0274 | 5.0151 | 5.0188 | 5.0204 | 5.0195 | 5.0173 | 5.0228 | 5.0169 | 0.015365 | 0.0030628
452 | 5.0417 | 5.0537 | 4.9906 | 5.007 | 5.0237 | 5.0245 | 5.0289 | 5.046 | 5.0425 | 5.0032 | 5.0069 | 5.0108 | 4.9933 | 4.9959 | 4.9974 | 5.007 | 5.0045 | 5.0096 | 5.0095 | 5.0056 | 5.0074 | 5.0257 | 5.0232 | 5.0271 | 5.0148 | 5.0186 | 5.0201 | 5.0192 | 5.0171 | 5.0225 | 5.0166 | 0.015351 | 0.0030599
453 | 5.0412 | 5.0535 | 4.9904 | 5.0067 | 5.0235 | 5.0242 | 5.0286 | 5.0458 | 5.0423 | 5.0029 | 5.0067 | 5.0105 | 4.993 | 4.9957 | 4.9971 | 5.0067 | 5.0043 | 5.0094 | 5.0092 | 5.0054 | 5.0072 | 5.0254 | 5.023 | 5.0269 | 5.0146 | 5.0183 | 5.0199 | 5.0189 | 5.0168 | 5.0223 | 5.0164 | 0.015336 | 0.0030572
454 | 5.0407 | 5.0533 | 4.9901 | 5.0065 | 5.0232 | 5.024 | 5.0284 | 5.0455 | 5.0421 | 5.0027 | 5.0065 | 5.0103 | 4.9928 | 4.9955 | 4.9969 | 5.0065 | 5.004 | 5.0091 | 5.009 | 5.0051 | 5.0069 | 5.0251 | 5.0228 | 5.0267 | 5.0143 | 5.0181 | 5.0196 | 5.0187 | 5.0166 | 5.022 | 5.0161 | 0.015321 | 0.0030544
455 | 5.0402 | 5.053 | 4.9898 | 5.0062 | 5.023 | 5.0237 | 5.0281 | 5.0453 | 5.0418 | 5.0024 | 5.0063 | 5.0101 | 4.9926 | 4.9952 | 4.9967 | 5.0063 | 5.0038 | 5.0089 | 5.0088 | 5.0049 | 5.0067 | 5.0249 | 5.0225 | 5.0264 | 5.0141 | 5.0178 | 5.0194 | 5.0184 | 5.0163 | 5.0218 | 5.0158 | 0.015307 | 0.0030517
456 | 5.0396 | 5.0528 | 4.9896 | 5.006 | 5.0227 | 5.0235 | 5.0279 | 5.0451 | 5.0416 | 5.0021 | 5.0061 | 5.0099 | 4.9924 | 4.995 | 4.9964 | 5.006 | 5.0035 | 5.0086 | 5.0086 | 5.0047 | 5.0065 | 5.0246 | 5.0223 | 5.0262 | 5.0139 | 5.0176 | 5.0191 | 5.0182 | 5.0161 | 5.0215 | 5.0156 | 0.015292 | 0.003049
457 | 5.0391 | 5.0526 | 4.9893 | 5.0057 | 5.0225 | 5.0232 | 5.0277 | 5.0448 | 5.0413 | 5.0018 | 5.0059 | 5.0097 | 4.9921 | 4.9948 | 4.9962 | 5.0058 | 5.0033 | 5.0084 | 5.0083 | 5.0044 | 5.0062 | 5.0244 | 5.0221 | 5.026 | 5.0136 | 5.0173 | 5.0189 | 5.0179 | 5.0158 | 5.0213 | 5.0154 | 0.015278 | 0.0030463
458 | 5.0386 | 5.0524 | 4.9891 | 5.0055 | 5.0222 | 5.023 | 5.0274 | 5.0446 | 5.0411 | 5.0015 | 5.0057 | 5.0095 | 4.9919 | 4.9945 | 4.996 | 5.0055 | 5.003 | 5.0081 | 5.0081 | 5.0042 | 5.006 | 5.0241 | 5.0219 | 5.0257 | 5.0134 | 5.0171 | 5.0186 | 5.0177 | 5.0156 | 5.021 | 5.0151 | 0.015264 | 0.0030437
459 | 5.0381 | 5.0521 | 4.9888 | 5.0053 | 5.022 | 5.0227 | 5.0272 | 5.0443 | 5.0409 | 5.0012 | 5.0055 | 5.0093 | 4.9917 | 4.9943 | 4.9957 | 5.0053 | 5.0028 | 5.0079 | 5.0079 | 5.0039 | 5.0057 | 5.0239 | 5.0216 | 5.0255 | 5.0131 | 5.0169 | 5.0184 | 5.0174 | 5.0153 | 5.0208 | 5.0149 | 0.015251 | 0.0030411
460 | 5.0376 | 5.0519 | 4.9886 | 5.005 | 5.0218 | 5.0225 | 5.027 | 5.0441 | 5.0406 | 5.001 | 5.0053 | 5.0091 | 4.9915 | 4.9941 | 4.9955 | 5.0051 | 5.0026 | 5.0077 | 5.0077 | 5.0037 | 5.0055 | 5.0237 | 5.0214 | 5.0253 | 5.0129 | 5.0166 | 5.0182 | 5.0172 | 5.0151 | 5.0205 | 5.0146 | 0.015237 | 0.0030385
461 | 5.0371 | 5.0517 | 4.9883 | 5.0048 | 5.0215 | 5.0223 | 5.0267 | 5.0439 | 5.0404 | 5.0007 | 5.0052 | 5.0089 | 4.9913 | 4.9938 | 4.9953 | 5.0048 | 5.0023 | 5.0074 | 5.0074 | 5.0035 | 5.0053 | 5.0234 | 5.0212 | 5.0251 | 5.0127 | 5.0164 | 5.0179 | 5.017 | 5.0148 | 5.0203 | 5.0144 | 0.015224 | 0.003036
462 | 5.0365 | 5.0515 | 4.9881 | 5.0046 | 5.0213 | 5.022 | 5.0265 | 5.0436 | 5.0402 | 5.0004 | 5.005 | 5.0087 | 4.9911 | 4.9936 | 4.9951 | 5.0046 | 5.0021 | 5.0072 | 5.0072 | 5.0032 | 5.005 | 5.0232 | 5.021 | 5.0248 | 5.0125 | 5.0162 | 5.0177 | 5.0167 | 5.0146 | 5.02 | 5.0141 | 0.015211 | 0.0030335
463 | 5.036 | 5.0513 | 4.9878 | 5.0043 | 5.0211 | 5.0218 | 5.0263 | 5.0434 | 5.0399 | 5.0001 | 5.0048 | 5.0085 | 4.9908 | 4.9934 | 4.9948 | 5.0044 | 5.0018 | 5.007 | 5.007 | 5.003 | 5.0048 | 5.0229 | 5.0207 | 5.0246 | 5.0122 | 5.0159 | 5.0174 | 5.0165 | 5.0143 | 5.0198 | 5.0139 | 0.015197 | 0.0030311
464 | 5.0355 | 5.0511 | 4.9876 | 5.0041 | 5.0208 | 5.0216 | 5.026 | 5.0431 | 5.0397 | 4.9998 | 5.0046 | 5.0083 | 4.9906 | 4.9932 | 4.9946 | 5.0042 | 5.0016 | 5.0067 | 5.0068 | 5.0028 | 5.0045 | 5.0227 | 5.0205 | 5.0244 | 5.012 | 5.0157 | 5.0172 | 5.0162 | 5.0141 | 5.0195 | 5.0137 | 0.015185 | 0.0030287
465 | 5.035 | 5.0508 | 4.9873 | 5.0039 | 5.0206 | 5.0213 | 5.0258 | 5.0429 | 5.0395 | 4.9996 | 5.0044 | 5.0081 | 4.9904 | 4.9929 | 4.9944 | 5.0039 | 5.0013 | 5.0065 | 5.0066 | 5.0025 | 5.0043 | 5.0224 | 5.0203 | 5.0242 | 5.0118 | 5.0155 | 5.017 | 5.016 | 5.0139 | 5.0193 | 5.0134 | 0.015172 | 0.0030263
466 | 5.0345 | 5.0506 | 4.9871 | 5.0036 | 5.0204 | 5.0211 | 5.0256 | 5.0427 | 5.0393 | 4.9993 | 5.0042 | 5.0079 | 4.9902 | 4.9927 | 4.9942 | 5.0037 | 5.0011 | 5.0063 | 5.0064 | 5.0023 | 5.0041 | 5.0222 | 5.0201 | 5.0239 | 5.0115 | 5.0152 | 5.0167 | 5.0158 | 5.0136 | 5.0191 | 5.0132 | 0.015159 | 0.0030239
467 | 5.034 | 5.0504 | 4.9868 | 5.0034 | 5.0202 | 5.0209 | 5.0254 | 5.0425 | 5.039 | 4.999 | 5.004 | 5.0077 | 4.99 | 4.9925 | 4.994 | 5.0035 | 5.0009 | 5.006 | 5.0062 | 5.0021 | 5.0039 | 5.022 | 5.0199 | 5.0237 | 5.0113 | 5.015 | 5.0165 | 5.0155 | 5.0134 | 5.0188 | 5.0129 | 0.015147 | 0.0030216
468 | 5.0335 | 5.0502 | 4.9866 | 5.0032 | 5.0199 | 5.0206 | 5.0251 | 5.0422 | 5.0388 | 4.9987 | 5.0038 | 5.0075 | 4.9898 | 4.9923 | 4.9937 | 5.0033 | 5.0006 | 5.0058 | 5.006 | 5.0018 | 5.0036 | 5.0217 | 5.0196 | 5.0235 | 5.0111 | 5.0148 | 5.0163 | 5.0153 | 5.0132 | 5.0186 | 5.0127 | 0.015135 | 0.0030193
469 | 5.033 | 5.05 | 4.9863 | 5.0029 | 5.0197 | 5.0204 | 5.0249 | 5.042 | 5.0386 | 4.9985 | 5.0037 | 5.0074 | 4.9896 | 4.9921 | 4.9935 | 5.003 | 5.0004 | 5.0056 | 5.0057 | 5.0016 | 5.0034 | 5.0215 | 5.0194 | 5.0233 | 5.0109 | 5.0146 | 5.016 | 5.0151 | 5.0129 | 5.0184 | 5.0125 | 0.015123 | 0.003017
470 | 5.0325 | 5.0498 | 4.9861 | 5.0027 | 5.0195 | 5.0202 | 5.0247 | 5.0418 | 5.0384 | 4.9982 | 5.0035 | 5.0072 | 4.9894 | 4.9919 | 4.9933 | 5.0028 | 5.0002 | 5.0054 | 5.0055 | 5.0014 | 5.0032 | 5.0213 | 5.0192 | 5.0231 | 5.0107 | 5.0143 | 5.0158 | 5.0148 | 5.0127 | 5.0181 | 5.0123 | 0.015111 | 0.0030148
471 | 5.032 | 5.0496 | 4.9859 | 5.0025 | 5.0193 | 5.02 | 5.0245 | 5.0415 | 5.0382 | 4.9979 | 5.0033 | 5.007 | 4.9892 | 4.9916 | 4.9931 | 5.0026 | 4.9999 | 5.0051 | 5.0053 | 5.0012 | 5.0029 | 5.021 | 5.019 | 5.0229 | 5.0104 | 5.0141 | 5.0156 | 5.0146 | 5.0125 | 5.0179 | 5.012 | 0.015099 | 0.0030126
472 | 5.0315 | 5.0494 | 4.9856 | 5.0023 | 5.019 | 5.0197 | 5.0243 | 5.0413 | 5.0379 | 4.9977 | 5.0031 | 5.0068 | 4.989 | 4.9914 | 4.9929 | 5.0024 | 4.9997 | 5.0049 | 5.0051 | 5.001 | 5.0027 | 5.0208 | 5.0188 | 5.0227 | 5.0102 | 5.0139 | 5.0154 | 5.0144 | 5.0122 | 5.0177 | 5.0118 | 0.015088 | 0.0030105
473 | 5.031 | 5.0492 | 4.9854 | 5.0021 | 5.0188 | 5.0195 | 5.024 | 5.0411 | 5.0377 | 4.9974 | 5.003 | 5.0066 | 4.9888 | 4.9912 | 4.9927 | 5.0022 | 4.9995 | 5.0047 | 5.0049 | 5.0007 | 5.0025 | 5.0206 | 5.0186 | 5.0224 | 5.01 | 5.0137 | 5.0151 | 5.0141 | 5.012 | 5.0174 | 5.0116 | 0.015077 | 0.0030084
474 | 5.0305 | 5.049 | 4.9852 | 5.0018 | 5.0186 | 5.0193 | 5.0238 | 5.0409 | 5.0375 | 4.9971 | 5.0028 | 5.0064 | 4.9886 | 4.991 | 4.9925 | 5.0019 | 4.9993 | 5.0045 | 5.0047 | 5.0005 | 5.0023 | 5.0203 | 5.0184 | 5.0222 | 5.0098 | 5.0135 | 5.0149 | 5.0139 | 5.0118 | 5.0172 | 5.0113 | 0.015066 | 0.0030063
475 | 5.03 | 5.0488 | 4.9849 | 5.0016 | 5.0184 | 5.0191 | 5.0236 | 5.0407 | 5.0373 | 4.9969 | 5.0026 | 5.0062 | 4.9884 | 4.9908 | 4.9923 | 5.0017 | 4.999 | 5.0042 | 5.0045 | 5.0003 | 5.002 | 5.0201 | 5.0182 | 5.022 | 5.0096 | 5.0132 | 5.0147 | 5.0137 | 5.0116 | 5.017 | 5.0111 | 0.015055 | 0.0030043
476 | 5.0295 | 5.0486 | 4.9847 | 5.0014 | 5.0182 | 5.0189 | 5.0234 | 5.0404 | 5.0371 | 4.9966 | 5.0024 | 5.0061 | 4.9882 | 4.9906 | 4.9921 | 5.0015 | 4.9988 | 5.004 | 5.0043 | 5.0001 | 5.0018 | 5.0199 | 5.018 | 5.0218 | 5.0094 | 5.013 | 5.0145 | 5.0135 | 5.0113 | 5.0168 | 5.0109 | 0.015044 | 0.0030023
477 | 5.029 | 5.0484 | 4.9845 | 5.0012 | 5.018 | 5.0186 | 5.0232 | 5.0402 | 5.0369 | 4.9963 | 5.0023 | 5.0059 | 4.988 | 4.9904 | 4.9919 | 5.0013 | 4.9986 | 5.0038 | 5.0041 | 4.9999 | 5.0016 | 5.0197 | 5.0178 | 5.0216 | 5.0092 | 5.0128 | 5.0142 | 5.0132 | 5.0111 | 5.0165 | 5.0107 | 0.015033 | 0.0030003
478 | 5.0285 | 5.0482 | 4.9842 | 5.001 | 5.0177 | 5.0184 | 5.023 | 5.04 | 5.0367 | 4.9961 | 5.0021 | 5.0057 | 4.9878 | 4.9902 | 4.9917 | 5.0011 | 4.9984 | 5.0036 | 5.0039 | 4.9997 | 5.0014 | 5.0194 | 5.0176 | 5.0214 | 5.0089 | 5.0126 | 5.014 | 5.013 | 5.0109 | 5.0163 | 5.0104 | 0.015023 | 0.0029984
479 | 5.0281 | 5.048 | 4.984 | 5.0008 | 5.0175 | 5.0182 | 5.0228 | 5.0398 | 5.0364 | 4.9958 | 5.0019 | 5.0055 | 4.9876 | 4.99 | 4.9915 | 5.0009 | 4.9981 | 5.0034 | 5.0037 | 4.9994 | 5.0012 | 5.0192 | 5.0174 | 5.0212 | 5.0087 | 5.0124 | 5.0138 | 5.0128 | 5.0107 | 5.0161 | 5.0102 | 0.015013 | 0.0029965
480 | 5.0276 | 5.0478 | 4.9838 | 5.0005 | 5.0173 | 5.018 | 5.0225 | 5.0396 | 5.0362 | 4.9956 | 5.0017 | 5.0053 | 4.9874 | 4.9898 | 4.9913 | 5.0007 | 4.9979 | 5.0032 | 5.0035 | 4.9992 | 5.001 | 5.019 | 5.0172 | 5.021 | 5.0085 | 5.0122 | 5.0136 | 5.0126 | 5.0104 | 5.0159 | 5.01 | 0.015003 | 0.0029946
481 | 5.0271 | 5.0476 | 4.9835 | 5.0003 | 5.0171 | 5.0178 | 5.0223 | 5.0394 | 5.036 | 4.9953 | 5.0016 | 5.0052 | 4.9872 | 4.9896 | 4.9911 | 5.0005 | 4.9977 | 5.0029 | 5.0033 | 4.999 | 5.0007 | 5.0188 | 5.017 | 5.0208 | 5.0083 | 5.012 | 5.0134 | 5.0123 | 5.0102 | 5.0157 | 5.0098 | 0.014993 | 0.0029928
482 | 5.0266 | 5.0474 | 4.9833 | 5.0001 | 5.0169 | 5.0176 | 5.0221 | 5.0392 | 5.0358 | 4.9951 | 5.0014 | 5.005 | 4.987 | 4.9894 | 4.9909 | 5.0003 | 4.9975 | 5.0027 | 5.0031 | 4.9988 | 5.0005 | 5.0186 | 5.0168 | 5.0206 | 5.0081 | 5.0118 | 5.0132 | 5.0121 | 5.01 | 5.0154 | 5.0096 | 0.014983 | 0.002991
483 | 5.0261 | 5.0472 | 4.9831 | 4.9999 | 5.0167 | 5.0174 | 5.0219 | 5.0389 | 5.0356 | 4.9948 | 5.0012 | 5.0048 | 4.9869 | 4.9892 | 4.9907 | 5.0001 | 4.9973 | 5.0025 | 5.0029 | 4.9986 | 5.0003 | 5.0183 | 5.0166 | 5.0204 | 5.0079 | 5.0115 | 5.0129 | 5.0119 | 5.0098 | 5.0152 | 5.0094 | 0.014974 | 0.0029892
484 | 5.0256 | 5.047 | 4.9829 | 4.9997 | 5.0165 | 5.0171 | 5.0217 | 5.0387 | 5.0354 | 4.9945 | 5.0011 | 5.0046 | 4.9867 | 4.989 | 4.9905 | 4.9999 | 4.997 | 5.0023 | 5.0028 | 4.9984 | 5.0001 | 5.0181 | 5.0164 | 5.0202 | 5.0077 | 5.0113 | 5.0127 | 5.0117 | 5.0096 | 5.015 | 5.0091 | 0.014965 | 0.0029875
485 | 5.0251 | 5.0468 | 4.9826 | 4.9995 | 5.0163 | 5.0169 | 5.0215 | 5.0385 | 5.0352 | 4.9943 | 5.0009 | 5.0045 | 4.9865 | 4.9888 | 4.9903 | 4.9996 | 4.9968 | 5.0021 | 5.0026 | 4.9982 | 4.9999 | 5.0179 | 5.0162 | 5.02 | 5.0075 | 5.0111 | 5.0125 | 5.0115 | 5.0094 | 5.0148 | 5.0089 | 0.014956 | 0.0029858
486 | 5.0247 | 5.0466 | 4.9824 | 4.9993 | 5.0161 | 5.0167 | 5.0213 | 5.0383 | 5.035 | 4.994 | 5.0008 | 5.0043 | 4.9863 | 4.9886 | 4.9901 | 4.9994 | 4.9966 | 5.0019 | 5.0024 | 4.998 | 4.9997 | 5.0177 | 5.016 | 5.0198 | 5.0073 | 5.0109 | 5.0123 | 5.0113 | 5.0091 | 5.0146 | 5.0087 | 0.014947 | 0.0029841
487 | 5.0242 | 5.0464 | 4.9822 | 4.9991 | 5.0159 | 5.0165 | 5.0211 | 5.0381 | 5.0348 | 4.9938 | 5.0006 | 5.0041 | 4.9861 | 4.9884 | 4.9899 | 4.9992 | 4.9964 | 5.0017 | 5.0022 | 4.9978 | 4.9995 | 5.0175 | 5.0158 | 5.0196 | 5.0071 | 5.0107 | 5.0121 | 5.011 | 5.0089 | 5.0144 | 5.0085 | 0.014938 | 0.0029825
488 | 5.0237 | 5.0462 | 4.982 | 4.9989 | 5.0157 | 5.0163 | 5.0209 | 5.0379 | 5.0346 | 4.9935 | 5.0004 | 5.004 | 4.9859 | 4.9882 | 4.9897 | 4.999 | 4.9962 | 5.0015 | 5.002 | 4.9976 | 4.9993 | 5.0173 | 5.0156 | 5.0194 | 5.0069 | 5.0105 | 5.0119 | 5.0108 | 5.0087 | 5.0141 | 5.0083 | 0.014929 | 0.002981
489 | 5.0232 | 5.046 | 4.9818 | 4.9987 | 5.0155 | 5.0161 | 5.0207 | 5.0377 | 5.0344 | 4.9933 | 5.0003 | 5.0038 | 4.9857 | 4.988 | 4.9895 | 4.9988 | 4.996 | 5.0013 | 5.0018 | 4.9974 | 4.9991 | 5.0171 | 5.0154 | 5.0192 | 5.0067 | 5.0103 | 5.0117 | 5.0106 | 5.0085 | 5.0139 | 5.0081 | 0.014921 | 0.0029794
490 | 5.0228 | 5.0459 | 4.9816 | 4.9985 | 5.0153 | 5.0159 | 5.0205 | 5.0375 | 5.0342 | 4.993 | 5.0001 | 5.0036 | 4.9856 | 4.9878 | 4.9893 | 4.9986 | 4.9958 | 5.0011 | 5.0016 | 4.9972 | 4.9989 | 5.0169 | 5.0152 | 5.019 | 5.0065 | 5.0101 | 5.0115 | 5.0104 | 5.0083 | 5.0137 | 5.0079 | 0.014913 | 0.0029779
491 | 5.0223 | 5.0457 | 4.9813 | 4.9983 | 5.0151 | 5.0157 | 5.0203 | 5.0373 | 5.034 | 4.9928 | 5 | 5.0035 | 4.9854 | 4.9876 | 4.9891 | 4.9984 | 4.9956 | 5.0009 | 5.0014 | 4.997 | 4.9986 | 5.0166 | 5.015 | 5.0188 | 5.0063 | 5.0099 | 5.0113 | 5.0102 | 5.0081 | 5.0135 | 5.0077 | 0.014905 | 0.0029764
492 | 5.0218 | 5.0455 | 4.9811 | 4.9981 | 5.0149 | 5.0155 | 5.0201 | 5.0371 | 5.0338 | 4.9925 | 4.9998 | 5.0033 | 4.9852 | 4.9874 | 4.9889 | 4.9983 | 4.9954 | 5.0007 | 5.0013 | 4.9968 | 4.9984 | 5.0164 | 5.0148 | 5.0186 | 5.0061 | 5.0097 | 5.0111 | 5.01 | 5.0079 | 5.0133 | 5.0075 | 0.014897 | 0.002975
493 | 5.0213 | 5.0453 | 4.9809 | 4.9979 | 5.0147 | 5.0153 | 5.0199 | 5.0369 | 5.0336 | 4.9923 | 4.9996 | 5.0031 | 4.985 | 4.9872 | 4.9887 | 4.9981 | 4.9952 | 5.0005 | 5.0011 | 4.9966 | 4.9982 | 5.0162 | 5.0146 | 5.0184 | 5.0059 | 5.0095 | 5.0109 | 5.0098 | 5.0077 | 5.0131 | 5.0073 | 0.01489 | 0.0029736
494 | 5.0209 | 5.0451 | 4.9807 | 4.9977 | 5.0145 | 5.0151 | 5.0197 | 5.0367 | 5.0334 | 4.992 | 4.9995 | 5.003 | 4.9849 | 4.987 | 4.9885 | 4.9979 | 4.9949 | 5.0003 | 5.0009 | 4.9964 | 4.998 | 5.016 | 5.0144 | 5.0183 | 5.0057 | 5.0093 | 5.0107 | 5.0096 | 5.0075 | 5.0129 | 5.007 | 0.014882 | 0.0029722
495 | 5.0204 | 5.0449 | 4.9805 | 4.9975 | 5.0143 | 5.0149 | 5.0195 | 5.0365 | 5.0332 | 4.9918 | 4.9993 | 5.0028 | 4.9847 | 4.9868 | 4.9884 | 4.9977 | 4.9947 | 5.0001 | 5.0007 | 4.9962 | 4.9978 | 5.0158 | 5.0143 | 5.0181 | 5.0055 | 5.0091 | 5.0104 | 5.0094 | 5.0073 | 5.0127 | 5.0068 | 0.014875 | 0.0029709
496 | 5.0199 | 5.0448 | 4.9803 | 4.9973 | 5.0141 | 5.0147 | 5.0194 | 5.0363 | 5.033 | 4.9915 | 4.9992 | 5.0026 | 4.9845 | 4.9866 | 4.9882 | 4.9975 | 4.9945 | 4.9999 | 5.0005 | 4.996 | 4.9976 | 5.0156 | 5.0141 | 5.0179 | 5.0053 | 5.0089 | 5.0102 | 5.0092 | 5.0071 | 5.0125 | 5.0066 | 0.014868 | 0.0029696
497 | 5.0195 | 5.0446 | 4.9801 | 4.9971 | 5.0139 | 5.0145 | 5.0192 | 5.0361 | 5.0328 | 4.9913 | 4.999 | 5.0025 | 4.9843 | 4.9865 | 4.988 | 4.9973 | 4.9943 | 4.9997 | 5.0004 | 4.9958 | 4.9974 | 5.0154 | 5.0139 | 5.0177 | 5.0052 | 5.0087 | 5.01 | 5.009 | 5.0069 | 5.0123 | 5.0064 | 0.014861 | 0.0029684
498 | 5.019 | 5.0444 | 4.9799 | 4.9969 | 5.0137 | 5.0143 | 5.019 | 5.0359 | 5.0327 | 4.9911 | 4.9989 | 5.0023 | 4.9842 | 4.9863 | 4.9878 | 4.9971 | 4.9941 | 4.9995 | 5.0002 | 4.9956 | 4.9972 | 5.0152 | 5.0137 | 5.0175 | 5.005 | 5.0085 | 5.0098 | 5.0088 | 5.0066 | 5.0121 | 5.0062 | 0.014854 | 0.0029671
499 | 5.0185 | 5.0442 | 4.9796 | 4.9967 | 5.0135 | 5.0141 | 5.0188 | 5.0357 | 5.0325 | 4.9908 | 4.9987 | 5.0022 | 4.984 | 4.9861 | 4.9876 | 4.9969 | 4.9939 | 4.9993 | 5 | 4.9954 | 4.997 | 5.015 | 5.0135 | 5.0173 | 5.0048 | 5.0084 | 5.0096 | 5.0086 | 5.0064 | 5.0119 | 5.006 | 0.014848 | 0.002966
500 | 5.0181 | 5.044 | 4.9794 | 4.9965 | 5.0133 | 5.0139 | 5.0186 | 5.0355 | 5.0323 | 4.9906 | 4.9986 | 5.002 | 4.9838 | 4.9859 | 4.9875 | 4.9967 | 4.9937 | 4.9991 | 4.9998 | 4.9952 | 4.9969 | 5.0148 | 5.0133 | 5.0171 | 5.0046 | 5.0082 | 5.0095 | 5.0084 | 5.0062 | 5.0117 | 5.0058 | 0.014841 | 0.0029648
501 | 5.0176 | 5.0439 | 4.9792 | 4.9963 | 5.0131 | 5.0137 | 5.0184 | 5.0353 | 5.0321 | 4.9903 | 4.9984 | 5.0018 | 4.9836 | 4.9857 | 4.9873 | 4.9965 | 4.9935 | 4.9989 | 4.9997 | 4.995 | 4.9967 | 5.0146 | 5.0132 | 5.017 | 5.0044 | 5.008 | 5.0093 | 5.0082 | 5.006 | 5.0115 | 5.0056 | 0.014835 | 0.0029637
502 | 5.0172 | 5.0437 | 4.979 | 4.9961 | 5.0129 | 5.0135 | 5.0182 | 5.0351 | 5.0319 | 4.9901 | 4.9983 | 5.0017 | 4.9835 | 4.9855 | 4.9871 | 4.9963 | 4.9933 | 4.9987 | 4.9995 | 4.9948 | 4.9965 | 5.0144 | 5.013 | 5.0168 | 5.0042 | 5.0078 | 5.0091 | 5.008 | 5.0059 | 5.0113 | 5.0054 | 0.014829 | 0.0029626
503 | 5.0167 | 5.0435 | 4.9788 | 4.9959 | 5.0127 | 5.0133 | 5.018 | 5.0349 | 5.0317 | 4.9899 | 4.9982 | 5.0015 | 4.9833 | 4.9854 | 4.9869 | 4.9962 | 4.9931 | 4.9985 | 4.9993 | 4.9946 | 4.9963 | 5.0142 | 5.0128 | 5.0166 | 5.004 | 5.0076 | 5.0089 | 5.0078 | 5.0057 | 5.0111 | 5.0053 | 0.014824 | 0.0029616
504 | 5.0162 | 5.0433 | 4.9786 | 4.9957 | 5.0126 | 5.0132 | 5.0179 | 5.0347 | 5.0315 | 4.9896 | 4.998 | 5.0014 | 4.9831 | 4.9852 | 4.9867 | 4.996 | 4.9929 | 4.9983 | 4.9991 | 4.9945 | 4.9961 | 5.014 | 5.0126 | 5.0164 | 5.0038 | 5.0074 | 5.0087 | 5.0076 | 5.0055 | 5.0109 | 5.0051 | 0.014818 | 0.0029606
505 | 5.0158 | 5.0432 | 4.9784 | 4.9955 | 5.0124 | 5.013 | 5.0177 | 5.0346 | 5.0313 | 4.9894 | 4.9979 | 5.0012 | 4.983 | 4.985 | 4.9866 | 4.9958 | 4.9928 | 4.9981 | 4.999 | 4.9943 | 4.9959 | 5.0138 | 5.0124 | 5.0162 | 5.0037 | 5.0072 | 5.0085 | 5.0074 | 5.0053 | 5.0107 | 5.0049 | 0.014813 | 0.0029596
506 | 5.0153 | 5.043 | 4.9782 | 4.9954 | 5.0122 | 5.0128 | 5.0175 | 5.0344 | 5.0312 | 4.9891 | 4.9977 | 5.0011 | 4.9828 | 4.9848 | 4.9864 | 4.9956 | 4.9926 | 4.998 | 4.9988 | 4.9941 | 4.9957 | 5.0136 | 5.0123 | 5.0161 | 5.0035 | 5.007 | 5.0083 | 5.0072 | 5.0051 | 5.0105 | 5.0047 | 0.014807 | 0.0029587
507 | 5.0149 | 5.0428 | 4.978 | 4.9952 | 5.012 | 5.0126 | 5.0173 | 5.0342 | 5.031 | 4.9889 | 4.9976 | 5.0009 | 4.9826 | 4.9846 | 4.9862 | 4.9954 | 4.9924 | 4.9978 | 4.9986 | 4.9939 | 4.9955 | 5.0134 | 5.0121 | 5.0159 | 5.0033 | 5.0069 | 5.0081 | 5.007 | 5.0049 | 5.0103 | 5.0045 | 0.014802 | 0.0029578
508 | 5.0144 | 5.0427 | 4.9778 | 4.995 | 5.0118 | 5.0124 | 5.0171 | 5.034 | 5.0308 | 4.9887 | 4.9974 | 5.0008 | 4.9825 | 4.9845 | 4.986 | 4.9952 | 4.9922 | 4.9976 | 4.9985 | 4.9937 | 4.9953 | 5.0132 | 5.0119 | 5.0157 | 5.0031 | 5.0067 | 5.0079 | 5.0068 | 5.0047 | 5.0101 | 5.0043 | 0.014797 | 0.0029569
509 | 5.014 | 5.0425 | 4.9776 | 4.9948 | 5.0116 | 5.0122 | 5.0169 | 5.0338 | 5.0306 | 4.9884 | 4.9973 | 5.0006 | 4.9823 | 4.9843 | 4.9859 | 4.9951 | 4.992 | 4.9974 | 4.9983 | 4.9935 | 4.9951 | 5.013 | 5.0117 | 5.0155 | 5.0029 | 5.0065 | 5.0077 | 5.0066 | 5.0045 | 5.0099 | 5.0041 | 0.014793 | 0.0029561
510 | 5.0135 | 5.0423 | 4.9774 | 4.9946 | 5.0115 | 5.012 | 5.0168 | 5.0336 | 5.0304 | 4.9882 | 4.9972 | 5.0005 | 4.9822 | 4.9841 | 4.9857 | 4.9949 | 4.9918 | 4.9972 | 4.9981 | 4.9933 | 4.9949 | 5.0128 | 5.0116 | 5.0154 | 5.0028 | 5.0063 | 5.0075 | 5.0064 | 5.0043 | 5.0097 | 5.0039 | 0.014788 | 0.0029553
511 | 5.0131 | 5.0422 | 4.9772 | 4.9944 | 5.0113 | 5.0118 | 5.0166 | 5.0334 | 5.0303 | 4.988 | 4.997 | 5.0003 | 4.982 | 4.9839 | 4.9855 | 4.9947 | 4.9916 | 4.997 | 4.998 | 4.9932 | 4.9947 | 5.0127 | 5.0114 | 5.0152 | 5.0026 | 5.0061 | 5.0073 | 5.0062 | 5.0041 | 5.0095 | 5.0037 | 0.014784 | 0.0029546
512 | 5.0126 | 5.042 | 4.977 | 4.9942 | 5.0111 | 5.0117 | 5.0164 | 5.0333 | 5.0301 | 4.9877 | 4.9969 | 5.0002 | 4.9818 | 4.9838 | 4.9854 | 4.9945 | 4.9914 | 4.9968 | 4.9978 | 4.993 | 4.9946 | 5.0125 | 5.0112 | 5.015 | 5.0024 | 5.0059 | 5.0072 | 5.006 | 5.0039 | 5.0094 | 5.0035 | 0.01478 | 0.0029539
513 | 5.0122 | 5.0418 | 4.9768 | 4.9941 | 5.0109 | 5.0115 | 5.0162 | 5.0331 | 5.0299 | 4.9875 | 4.9967 | 5 | 4.9817 | 4.9836 | 4.9852 | 4.9943 | 4.9912 | 4.9967 | 4.9976 | 4.9928 | 4.9944 | 5.0123 | 5.0111 | 5.0149 | 5.0022 | 5.0058 | 5.007 | 5.0058 | 5.0037 | 5.0092 | 5.0033 | 0.014776 | 0.0029532
514 | 5.0117 | 5.0417 | 4.9766 | 4.9939 | 5.0107 | 5.0113 | 5.0161 | 5.0329 | 5.0297 | 4.9873 | 4.9966 | 4.9999 | 4.9815 | 4.9834 | 4.985 | 4.9942 | 4.991 | 4.9965 | 4.9975 | 4.9926 | 4.9942 | 5.0121 | 5.0109 | 5.0147 | 5.0021 | 5.0056 | 5.0068 | 5.0056 | 5.0035 | 5.009 | 5.0032 | 0.014772 | 0.0029525
515 | 5.0113 | 5.0415 | 4.9764 | 4.9937 | 5.0106 | 5.0111 | 5.0159 | 5.0327 | 5.0296 | 4.9871 | 4.9965 | 4.9997 | 4.9814 | 4.9833 | 4.9849 | 4.994 | 4.9909 | 4.9963 | 4.9973 | 4.9924 | 4.994 | 5.0119 | 5.0107 | 5.0145 | 5.0019 | 5.0054 | 5.0066 | 5.0054 | 5.0034 | 5.0088 | 5.003 | 0.014768 | 0.0029519
516 | 5.0108 | 5.0413 | 4.9762 | 4.9935 | 5.0104 | 5.0109 | 5.0157 | 5.0325 | 5.0294 | 4.9868 | 4.9963 | 4.9996 | 4.9812 | 4.9831 | 4.9847 | 4.9938 | 4.9907 | 4.9961 | 4.9972 | 4.9923 | 4.9938 | 5.0117 | 5.0106 | 5.0143 | 5.0017 | 5.0052 | 5.0064 | 5.0053 | 5.0032 | 5.0086 | 5.0028 | 0.014765 | 0.0029514
517 | 5.0104 | 5.0412 | 4.9761 | 4.9933 | 5.0102 | 5.0108 | 5.0155 | 5.0324 | 5.0292 | 4.9866 | 4.9962 | 4.9995 | 4.981 | 4.9829 | 4.9845 | 4.9936 | 4.9905 | 4.9959 | 4.997 | 4.9921 | 4.9936 | 5.0115 | 5.0104 | 5.0142 | 5.0015 | 5.0051 | 5.0062 | 5.0051 | 5.003 | 5.0084 | 5.0026 | 0.014762 | 0.0029508
518 | 5.01 | 5.041 | 4.9759 | 4.9932 | 5.01 | 5.0106 | 5.0154 | 5.0322 | 5.029 | 4.9864 | 4.9961 | 4.9993 | 4.9809 | 4.9828 | 4.9844 | 4.9935 | 4.9903 | 4.9958 | 4.9968 | 4.9919 | 4.9935 | 5.0113 | 5.0102 | 5.014 | 5.0014 | 5.0049 | 5.006 | 5.0049 | 5.0028 | 5.0082 | 5.0024 | 0.014759 | 0.0029503
519 | 5.0095 | 5.0408 | 4.9757 | 4.993 | 5.0099 | 5.0104 | 5.0152 | 5.032 | 5.0289 | 4.9861 | 4.9959 | 4.9992 | 4.9807 | 4.9826 | 4.9842 | 4.9933 | 4.9901 | 4.9956 | 4.9967 | 4.9917 | 4.9933 | 5.0112 | 5.0101 | 5.0138 | 5.0012 | 5.0047 | 5.0059 | 5.0047 | 5.0026 | 5.0081 | 5.0022 | 0.014756 | 0.0029498
520 | 5.0091 | 5.0407 | 4.9755 | 4.9928 | 5.0097 | 5.0102 | 5.015 | 5.0318 | 5.0287 | 4.9859 | 4.9958 | 4.999 | 4.9806 | 4.9824 | 4.984 | 4.9931 | 4.9899 | 4.9954 | 4.9965 | 4.9916 | 4.9931 | 5.011 | 5.0099 | 5.0137 | 5.001 | 5.0045 | 5.0057 | 5.0045 | 5.0024 | 5.0079 | 5.0021 | 0.014753 | 0.0029494
521 | 5.0086 | 5.0405 | 4.9753 | 4.9926 | 5.0095 | 5.0101 | 5.0149 | 5.0316 | 5.0285 | 4.9857 | 4.9957 | 4.9989 | 4.9804 | 4.9823 | 4.9839 | 4.993 | 4.9898 | 4.9952 | 4.9964 | 4.9914 | 4.9929 | 5.0108 | 5.0097 | 5.0135 | 5.0009 | 5.0044 | 5.0055 | 5.0043 | 5.0022 | 5.0077 | 5.0019 | 0.014751 | 0.002949
522 | 5.0082 | 5.0404 | 4.9751 | 4.9925 | 5.0093 | 5.0099 | 5.0147 | 5.0315 | 5.0284 | 4.9855 | 4.9955 | 4.9988 | 4.9803 | 4.9821 | 4.9837 | 4.9928 | 4.9896 | 4.9951 | 4.9962 | 4.9912 | 4.9927 | 5.0106 | 5.0096 | 5.0133 | 5.0007 | 5.0042 | 5.0053 | 5.0042 | 5.0021 | 5.0075 | 5.0017 | 0.014748 | 0.0029487
523 | 5.0078 | 5.0402 | 4.9749 | 4.9923 | 5.0092 | 5.0097 | 5.0145 | 5.0313 | 5.0282 | 4.9852 | 4.9954 | 4.9986 | 4.9801 | 4.9819 | 4.9836 | 4.9926 | 4.9894 | 4.9949 | 4.9961 | 4.9911 | 4.9926 | 5.0104 | 5.0094 | 5.0132 | 5.0005 | 5.004 | 5.0051 | 5.004 | 5.0019 | 5.0073 | 5.0015 | 0.014746 | 0.0029483
524 | 5.0073 | 5.0401 | 4.9747 | 4.9921 | 5.009 | 5.0095 | 5.0143 | 5.0311 | 5.028 | 4.985 | 4.9953 | 4.9985 | 4.98 | 4.9818 | 4.9834 | 4.9925 | 4.9892 | 4.9947 | 4.9959 | 4.9909 | 4.9924 | 5.0103 | 5.0092 | 5.013 | 5.0003 | 5.0039 | 5.005 | 5.0038 | 5.0017 | 5.0071 | 5.0013 | 0.014744 | 0.002948
525 | 5.0069 | 5.0399 | 4.9745 | 4.9919 | 5.0088 | 5.0094 | 5.0142 | 5.0309 | 5.0278 | 4.9848 | 4.9952 | 4.9983 | 4.9798 | 4.9816 | 4.9832 | 4.9923 | 4.989 | 4.9945 | 4.9957 | 4.9907 | 4.9922 | 5.0101 | 5.0091 | 5.0129 | 5.0002 | 5.0037 | 5.0048 | 5.0036 | 5.0015 | 5.007 | 5.0012 | 0.014742 | 0.0029478
526 | 5.0065 | 5.0397 | 4.9744 | 4.9918 | 5.0087 | 5.0092 | 5.014 | 5.0308 | 5.0277 | 4.9846 | 4.995 | 4.9982 | 4.9797 | 4.9814 | 4.9831 | 4.9921 | 4.9889 | 4.9944 | 4.9956 | 4.9905 | 4.992 | 5.0099 | 5.0089 | 5.0127 | 5 | 5.0035 | 5.0046 | 5.0034 | 5.0013 | 5.0068 | 5.001 | 0.014741 | 0.0029475
527 | 5.006 | 5.0396 | 4.9742 | 4.9916 | 5.0085 | 5.009 | 5.0139 | 5.0306 | 5.0275 | 4.9844 | 4.9949 | 4.9981 | 4.9795 | 4.9813 | 4.9829 | 4.992 | 4.9887 | 4.9942 | 4.9954 | 4.9904 | 4.9919 | 5.0097 | 5.0088 | 5.0125 | 4.9999 | 5.0034 | 5.0044 | 5.0032 | 5.0012 | 5.0066 | 5.0008 | 0.014739 | 0.0029473
528 | 5.0056 | 5.0394 | 4.974 | 4.9914 | 5.0083 | 5.0088 | 5.0137 | 5.0304 | 5.0274 | 4.9841 | 4.9948 | 4.9979 | 4.9794 | 4.9811 | 4.9828 | 4.9918 | 4.9885 | 4.994 | 4.9953 | 4.9902 | 4.9917 | 5.0095 | 5.0086 | 5.0124 | 4.9997 | 5.0032 | 5.0043 | 5.0031 | 5.001 | 5.0064 | 5.0006 | 0.014738 | 0.0029472
529 | 5.0052 | 5.0393 | 4.9738 | 4.9913 | 5.0082 | 5.0087 | 5.0135 | 5.0303 | 5.0272 | 4.9839 | 4.9947 | 4.9978 | 4.9792 | 4.981 | 4.9826 | 4.9916 | 4.9883 | 4.9939 | 4.9951 | 4.99 | 4.9915 | 5.0094 | 5.0084 | 5.0122 | 4.9995 | 5.003 | 5.0041 | 5.0029 | 5.0008 | 5.0063 | 5.0005 | 0.014737 | 0.0029471
530 | 5.0048 | 5.0391 | 4.9736 | 4.9911 | 5.008 | 5.0085 | 5.0134 | 5.0301 | 5.027 | 4.9837 | 4.9945 | 4.9977 | 4.9791 | 4.9808 | 4.9825 | 4.9915 | 4.9882 | 4.9937 | 4.995 | 4.9899 | 4.9913 | 5.0092 | 5.0083 | 5.0121 | 4.9994 | 5.0029 | 5.0039 | 5.0027 | 5.0006 | 5.0061 | 5.0003 | 0.014736 | 0.002947
531 | 5.0043 | 5.039 | 4.9734 | 4.9909 | 5.0078 | 5.0083 | 5.0132 | 5.0299 | 5.0269 | 4.9835 | 4.9944 | 4.9975 | 4.979 | 4.9806 | 4.9823 | 4.9913 | 4.988 | 4.9935 | 4.9948 | 4.9897 | 4.9912 | 5.009 | 5.0081 | 5.0119 | 4.9992 | 5.0027 | 5.0037 | 5.0025 | 5.0005 | 5.0059 | 5.0001 | 0.014735 | 0.0029469
532 | 5.0039 | 5.0388 | 4.9733 | 4.9907 | 5.0077 | 5.0082 | 5.013 | 5.0298 | 5.0267 | 4.9833 | 4.9943 | 4.9974 | 4.9788 | 4.9805 | 4.9821 | 4.9911 | 4.9878 | 4.9933 | 4.9947 | 4.9895 | 4.991 | 5.0088 | 5.008 | 5.0117 | 4.999 | 5.0025 | 5.0036 | 5.0024 | 5.0003 | 5.0057 | 4.9999 | 0.014734 | 0.0029469
533 | 5.0035 | 5.0387 | 4.9731 | 4.9906 | 5.0075 | 5.008 | 5.0129 | 5.0296 | 5.0265 | 4.9831 | 4.9942 | 4.9973 | 4.9787 | 4.9803 | 4.982 | 4.991 | 4.9876 | 4.9932 | 4.9945 | 4.9894 | 4.9908 | 5.0087 | 5.0078 | 5.0116 | 4.9989 | 5.0024 | 5.0034 | 5.0022 | 5.0001 | 5.0056 | 4.9998 | 0.014734 | 0.0029469
534 | 5.0031 | 5.0385 | 4.9729 | 4.9904 | 5.0073 | 5.0078 | 5.0127 | 5.0294 | 5.0264 | 4.9828 | 4.994 | 4.9971 | 4.9785 | 4.9802 | 4.9818 | 4.9908 | 4.9875 | 4.993 | 4.9944 | 4.9892 | 4.9907 | 5.0085 | 5.0077 | 5.0114 | 4.9987 | 5.0022 | 5.0032 | 5.002 | 4.9999 | 5.0054 | 4.9996 | 0.014734 | 0.002947
535 | 5.0026 | 5.0384 | 4.9727 | 4.9902 | 5.0072 | 5.0077 | 5.0126 | 5.0293 | 5.0262 | 4.9826 | 4.9939 | 4.997 | 4.9784 | 4.98 | 4.9817 | 4.9906 | 4.9873 | 4.9928 | 4.9942 | 4.9891 | 4.9905 | 5.0083 | 5.0075 | 5.0113 | 4.9986 | 5.002 | 5.0031 | 5.0018 | 4.9998 | 5.0052 | 4.9994 | 0.014733 | 0.002947
536 | 5.0022 | 5.0382 | 4.9726 | 4.9901 | 5.007 | 5.0075 | 5.0124 | 5.0291 | 5.0261 | 4.9824 | 4.9938 | 4.9969 | 4.9782 | 4.9799 | 4.9815 | 4.9905 | 4.9871 | 4.9927 | 4.9941 | 4.9889 | 4.9903 | 5.0081 | 5.0073 | 5.0111 | 4.9984 | 5.0019 | 5.0029 | 5.0017 | 4.9996 | 5.0051 | 4.9993 | 0.014734 | 0.0029471
537 | 5.0018 | 5.0381 | 4.9724 | 4.9899 | 5.0068 | 5.0073 | 5.0122 | 5.0289 | 5.0259 | 4.9822 | 4.9937 | 4.9967 | 4.9781 | 4.9797 | 4.9814 | 4.9903 | 4.987 | 4.9925 | 4.994 | 4.9887 | 4.9902 | 5.008 | 5.0072 | 5.011 | 4.9982 | 5.0017 | 5.0027 | 5.0015 | 4.9994 | 5.0049 | 4.9991 | 0.014734 | 0.0029473
538 | 5.0014 | 5.0379 | 4.9722 | 4.9898 | 5.0067 | 5.0072 | 5.0121 | 5.0288 | 5.0257 | 4.982 | 4.9936 | 4.9966 | 4.978 | 4.9795 | 4.9812 | 4.9902 | 4.9868 | 4.9924 | 4.9938 | 4.9886 | 4.99 | 5.0078 | 5.007 | 5.0108 | 4.9981 | 5.0016 | 5.0026 | 5.0013 | 4.9993 | 5.0047 | 4.9989 | 0.014734 | 0.0029475
539 | 5.001 | 5.0378 | 4.972 | 4.9896 | 5.0065 | 5.007 | 5.0119 | 5.0286 | 5.0256 | 4.9818 | 4.9934 | 4.9965 | 4.9778 | 4.9794 | 4.9811 | 4.99 | 4.9866 | 4.9922 | 4.9937 | 4.9884 | 4.9898 | 5.0076 | 5.0069 | 5.0107 | 4.9979 | 5.0014 | 5.0024 | 5.0012 | 4.9991 | 5.0045 | 4.9987 | 0.014735 | 0.0029477
540 | 5.0005 | 5.0376 | 4.9719 | 4.9894 | 5.0064 | 5.0069 | 5.0118 | 5.0284 | 5.0254 | 4.9816 | 4.9933 | 4.9964 | 4.9777 | 4.9792 | 4.9809 | 4.9899 | 4.9865 | 4.992 | 4.9935 | 4.9883 | 4.9897 | 5.0075 | 5.0067 | 5.0105 | 4.9978 | 5.0012 | 5.0022 | 5.001 | 4.9989 | 5.0044 | 4.9986 | 0.014735 | 0.0029479
541 | 5.0001 | 5.0375 | 4.9717 | 4.9893 | 5.0062 | 5.0067 | 5.0116 | 5.0283 | 5.0253 | 4.9813 | 4.9932 | 4.9962 | 4.9775 | 4.9791 | 4.9808 | 4.9897 | 4.9863 | 4.9919 | 4.9934 | 4.9881 | 4.9895 | 5.0073 | 5.0066 | 5.0104 | 4.9976 | 5.0011 | 5.0021 | 5.0008 | 4.9987 | 5.0042 | 4.9984 | 0.014736 | 0.0029482
542 | 4.9997 | 5.0373 | 4.9715 | 4.9891 | 5.0061 | 5.0065 | 5.0115 | 5.0281 | 5.0251 | 4.9811 | 4.9931 | 4.9961 | 4.9774 | 4.9789 | 4.9806 | 4.9895 | 4.9861 | 4.9917 | 4.9932 | 4.9879 | 4.9893 | 5.0071 | 5.0064 | 5.0102 | 4.9975 | 5.0009 | 5.0019 | 5.0006 | 4.9986 | 5.004 | 4.9982 | 0.014737 | 0.0029485
543 | 4.9993 | 5.0372 | 4.9713 | 4.9889 | 5.0059 | 5.0064 | 5.0113 | 5.028 | 5.025 | 4.9809 | 4.993 | 4.996 | 4.9773 | 4.9788 | 4.9805 | 4.9894 | 4.986 | 4.9915 | 4.9931 | 4.9878 | 4.9892 | 5.007 | 5.0063 | 5.0101 | 4.9973 | 5.0008 | 5.0017 | 5.0005 | 4.9984 | 5.0039 | 4.9981 | 0.014739 | 0.0029488
544 | 4.9989 | 5.0371 | 4.9712 | 4.9888 | 5.0057 | 5.0062 | 5.0112 | 5.0278 | 5.0248 | 4.9807 | 4.9929 | 4.9959 | 4.9771 | 4.9786 | 4.9804 | 4.9892 | 4.9858 | 4.9914 | 4.993 | 4.9876 | 4.989 | 5.0068 | 5.0061 | 5.0099 | 4.9972 | 5.0006 | 5.0016 | 5.0003 | 4.9982 | 5.0037 | 4.9979 | 0.01474 | 0.0029492
545 | 4.9985 | 5.0369 | 4.971 | 4.9886 | 5.0056 | 5.0061 | 5.011 | 5.0276 | 5.0247 | 4.9805 | 4.9927 | 4.9957 | 4.977 | 4.9785 | 4.9802 | 4.9891 | 4.9856 | 4.9912 | 4.9928 | 4.9875 | 4.9888 | 5.0066 | 5.006 | 5.0098 | 4.997 | 5.0005 | 5.0014 | 5.0001 | 4.9981 | 5.0036 | 4.9978 | 0.014742 | 0.0029496
546 | 4.9981 | 5.0368 | 4.9708 | 4.9885 | 5.0054 | 5.0059 | 5.0109 | 5.0275 | 5.0245 | 4.9803 | 4.9926 | 4.9956 | 4.9769 | 4.9783 | 4.9801 | 4.9889 | 4.9855 | 4.9911 | 4.9927 | 4.9873 | 4.9887 | 5.0065 | 5.0058 | 5.0096 | 4.9968 | 5.0003 | 5.0012 | 5 | 4.9979 | 5.0034 | 4.9976 | 0.014743 | 0.0029501
547 | 4.9977 | 5.0366 | 4.9706 | 4.9883 | 5.0053 | 5.0057 | 5.0107 | 5.0273 | 5.0243 | 4.9801 | 4.9925 | 4.9955 | 4.9767 | 4.9782 | 4.9799 | 4.9888 | 4.9853 | 4.9909 | 4.9925 | 4.9872 | 4.9885 | 5.0063 | 5.0057 | 5.0095 | 4.9967 | 5.0001 | 5.0011 | 4.9998 | 4.9978 | 5.0032 | 4.9974 | 0.014745 | 0.0029505
548 | 4.9972 | 5.0365 | 4.9705 | 4.9881 | 5.0051 | 5.0056 | 5.0106 | 5.0272 | 5.0242 | 4.9799 | 4.9924 | 4.9954 | 4.9766 | 4.978 | 4.9798 | 4.9886 | 4.9851 | 4.9908 | 4.9924 | 4.987 | 4.9883 | 5.0061 | 5.0055 | 5.0093 | 4.9965 | 5 | 5.0009 | 4.9996 | 4.9976 | 5.0031 | 4.9973 | 0.014747 | 0.002951
549 | 4.9968 | 5.0363 | 4.9703 | 4.988 | 5.005 | 5.0054 | 5.0104 | 5.027 | 5.024 | 4.9797 | 4.9923 | 4.9952 | 4.9765 | 4.9779 | 4.9796 | 4.9885 | 4.985 | 4.9906 | 4.9923 | 4.9868 | 4.9882 | 5.006 | 5.0054 | 5.0092 | 4.9964 | 4.9998 | 5.0007 | 4.9995 | 4.9974 | 5.0029 | 4.9971 | 0.014749 | 0.0029516
550 | 4.9964 | 5.0362 | 4.9701 | 4.9878 | 5.0048 | 5.0053 | 5.0103 | 5.0269 | 5.0239 | 4.9795 | 4.9922 | 4.9951 | 4.9763 | 4.9777 | 4.9795 | 4.9883 | 4.9848 | 4.9904 | 4.9921 | 4.9867 | 4.988 | 5.0058 | 5.0053 | 5.009 | 4.9962 | 4.9997 | 5.0006 | 4.9993 | 4.9973 | 5.0027 | 4.9969 | 0.014752 | 0.0029521
551 | 4.996 | 5.0361 | 4.97 | 4.9877 | 5.0047 | 5.0051 | 5.0101 | 5.0267 | 5.0237 | 4.9793 | 4.9921 | 4.995 | 4.9762 | 4.9776 | 4.9794 | 4.9882 | 4.9847 | 4.9903 | 4.992 | 4.9865 | 4.9879 | 5.0057 | 5.0051 | 5.0089 | 4.9961 | 4.9995 | 5.0004 | 4.9992 | 4.9971 | 5.0026 | 4.9968 | 0.014754 | 0.0029527
552 | 4.9956 | 5.0359 | 4.9698 | 4.9875 | 5.0045 | 5.005 | 5.01 | 5.0265 | 5.0236 | 4.9791 | 4.992 | 4.9949 | 4.9761 | 4.9775 | 4.9792 | 4.988 | 4.9845 | 4.9901 | 4.9919 | 4.9864 | 4.9877 | 5.0055 | 5.005 | 5.0087 | 4.9959 | 4.9994 | 5.0003 | 4.999 | 4.9969 | 5.0024 | 4.9966 | 0.014757 | 0.0029533
553 | 4.9952 | 5.0358 | 4.9696 | 4.9874 | 5.0044 | 5.0048 | 5.0098 | 5.0264 | 5.0234 | 4.9789 | 4.9919 | 4.9948 | 4.9759 | 4.9773 | 4.9791 | 4.9879 | 4.9843 | 4.99 | 4.9917 | 4.9862 | 4.9876 | 5.0053 | 5.0048 | 5.0086 | 4.9958 | 4.9992 | 5.0001 | 4.9988 | 4.9968 | 5.0023 | 4.9965 | 0.01476 | 0.002954
554 | 4.9948 | 5.0356 | 4.9695 | 4.9872 | 5.0042 | 5.0047 | 5.0097 | 5.0262 | 5.0233 | 4.9786 | 4.9917 | 4.9946 | 4.9758 | 4.9772 | 4.9789 | 4.9877 | 4.9842 | 4.9898 | 4.9916 | 4.9861 | 4.9874 | 5.0052 | 5.0047 | 5.0085 | 4.9956 | 4.9991 | 5 | 4.9987 | 4.9966 | 5.0021 | 4.9963 | 0.014762 | 0.0029547
555 | 4.9944 | 5.0355 | 4.9693 | 4.9871 | 5.0041 | 5.0045 | 5.0095 | 5.0261 | 5.0232 | 4.9784 | 4.9916 | 4.9945 | 4.9757 | 4.977 | 4.9788 | 4.9876 | 4.984 | 4.9897 | 4.9914 | 4.9859 | 4.9872 | 5.005 | 5.0045 | 5.0083 | 4.9955 | 4.9989 | 4.9998 | 4.9985 | 4.9965 | 5.0019 | 4.9962 | 0.014766 | 0.0029554
556 | 4.994 | 5.0354 | 4.9691 | 4.9869 | 5.0039 | 5.0044 | 5.0094 | 5.0259 | 5.023 | 4.9782 | 4.9915 | 4.9944 | 4.9755 | 4.9769 | 4.9787 | 4.9874 | 4.9839 | 4.9895 | 4.9913 | 4.9858 | 4.9871 | 5.0049 | 5.0044 | 5.0082 | 4.9954 | 4.9988 | 4.9996 | 4.9983 | 4.9963 | 5.0018 | 4.996 | 0.014769 | 0.0029561
557 | 4.9936 | 5.0352 | 4.969 | 4.9867 | 5.0038 | 5.0042 | 5.0092 | 5.0258 | 5.0229 | 4.978 | 4.9914 | 4.9943 | 4.9754 | 4.9767 | 4.9785 | 4.9873 | 4.9837 | 4.9894 | 4.9912 | 4.9856 | 4.9869 | 5.0047 | 5.0043 | 5.008 | 4.9952 | 4.9986 | 4.9995 | 4.9982 | 4.9961 | 5.0016 | 4.9958 | 0.014772 | 0.0029569
558 | 4.9932 | 5.0351 | 4.9688 | 4.9866 | 5.0036 | 5.004 | 5.0091 | 5.0256 | 5.0227 | 4.9778 | 4.9913 | 4.9942 | 4.9753 | 4.9766 | 4.9784 | 4.9871 | 4.9836 | 4.9892 | 4.991 | 4.9855 | 4.9868 | 5.0045 | 5.0041 | 5.0079 | 4.9951 | 4.9985 | 4.9993 | 4.998 | 4.996 | 5.0015 | 4.9957 | 0.014776 | 0.0029577
559 | 4.9928 | 5.035 | 4.9686 | 4.9864 | 5.0035 | 5.0039 | 5.0089 | 5.0255 | 5.0226 | 4.9776 | 4.9912 | 4.9941 | 4.9752 | 4.9765 | 4.9783 | 4.987 | 4.9834 | 4.9891 | 4.9909 | 4.9853 | 4.9866 | 5.0044 | 5.004 | 5.0077 | 4.9949 | 4.9983 | 4.9992 | 4.9979 | 4.9958 | 5.0013 | 4.9955 | 0.01478 | 0.0029585
560 | 4.9924 | 5.0348 | 4.9685 | 4.9863 | 5.0033 | 5.0037 | 5.0088 | 5.0253 | 5.0224 | 4.9774 | 4.9911 | 4.9939 | 4.975 | 4.9763 | 4.9781 | 4.9868 | 4.9832 | 4.9889 | 4.9908 | 4.9852 | 4.9865 | 5.0042 | 5.0038 | 5.0076 | 4.9948 | 4.9982 | 4.999 | 4.9977 | 4.9957 | 5.0012 | 4.9954 | 0.014783 | 0.0029594
561 | 4.992 | 5.0347 | 4.9683 | 4.9861 | 5.0032 | 5.0036 | 5.0087 | 5.0252 | 5.0223 | 4.9772 | 4.991 | 4.9938 | 4.9749 | 4.9762 | 4.978 | 4.9867 | 4.9831 | 4.9888 | 4.9907 | 4.9851 | 4.9863 | 5.0041 | 5.0037 | 5.0075 | 4.9946 | 4.9981 | 4.9989 | 4.9975 | 4.9955 | 5.001 | 4.9952 | 0.014787 | 0.0029603
562 | 4.9916 | 5.0346 | 4.9682 | 4.986 | 5.003 | 5.0034 | 5.0085 | 5.025 | 5.0221 | 4.977 | 4.9909 | 4.9937 | 4.9748 | 4.976 | 4.9778 | 4.9866 | 4.9829 | 4.9886 | 4.9905 | 4.9849 | 4.9861 | 5.0039 | 5.0036 | 5.0073 | 4.9945 | 4.9979 | 4.9987 | 4.9974 | 4.9954 | 5.0009 | 4.9951 | 0.014792 | 0.0029612
563 | 4.9913 | 5.0344 | 4.968 | 4.9858 | 5.0029 | 5.0033 | 5.0084 | 5.0249 | 5.022 | 4.9768 | 4.9908 | 4.9936 | 4.9747 | 4.9759 | 4.9777 | 4.9864 | 4.9828 | 4.9885 | 4.9904 | 4.9848 | 4.986 | 5.0038 | 5.0034 | 5.0072 | 4.9943 | 4.9978 | 4.9986 | 4.9972 | 4.9952 | 5.0007 | 4.9949 | 0.014796 | 0.0029622
564 | 4.9909 | 5.0343 | 4.9678 | 4.9857 | 5.0027 | 5.0032 | 5.0082 | 5.0247 | 5.0218 | 4.9766 | 4.9907 | 4.9935 | 4.9745 | 4.9758 | 4.9776 | 4.9863 | 4.9826 | 4.9883 | 4.9903 | 4.9846 | 4.9858 | 5.0036 | 5.0033 | 5.0071 | 4.9942 | 4.9976 | 4.9984 | 4.9971 | 4.995 | 5.0005 | 4.9948 | 0.0148 | 0.0029632
565 | 4.9905 | 5.0342 | 4.9677 | 4.9855 | 5.0026 | 5.003 | 5.0081 | 5.0246 | 5.0217 | 4.9764 | 4.9906 | 4.9934 | 4.9744 | 4.9756 | 4.9774 | 4.9861 | 4.9825 | 4.9882 | 4.9901 | 4.9845 | 4.9857 | 5.0035 | 5.0031 | 5.0069 | 4.9941 | 4.9975 | 4.9982 | 4.9969 | 4.9949 | 5.0004 | 4.9946 | 0.014805 | 0.0029642
566 | 4.9901 | 5.034 | 4.9675 | 4.9854 | 5.0025 | 5.0029 | 5.0079 | 5.0244 | 5.0216 | 4.9762 | 4.9905 | 4.9933 | 4.9743 | 4.9755 | 4.9773 | 4.986 | 4.9823 | 4.988 | 4.99 | 4.9843 | 4.9855 | 5.0033 | 5.003 | 5.0068 | 4.9939 | 4.9973 | 4.9981 | 4.9968 | 4.9947 | 5.0002 | 4.9945 | 0.01481 | 0.0029652
567 | 4.9897 | 5.0339 | 4.9674 | 4.9852 | 5.0023 | 5.0027 | 5.0078 | 5.0243 | 5.0214 | 4.9761 | 4.9904 | 4.9931 | 4.9742 | 4.9753 | 4.9772 | 4.9858 | 4.9822 | 4.9879 | 4.9899 | 4.9842 | 4.9854 | 5.0031 | 5.0029 | 5.0066 | 4.9938 | 4.9972 | 4.9979 | 4.9966 | 4.9946 | 5.0001 | 4.9943 | 0.014814 | 0.0029663
568 | 4.9893 | 5.0338 | 4.9672 | 4.9851 | 5.0022 | 5.0026 | 5.0077 | 5.0242 | 5.0213 | 4.9759 | 4.9903 | 4.993 | 4.974 | 4.9752 | 4.9771 | 4.9857 | 4.982 | 4.9877 | 4.9898 | 4.984 | 4.9852 | 5.003 | 5.0027 | 5.0065 | 4.9936 | 4.9971 | 4.9978 | 4.9965 | 4.9944 | 4.9999 | 4.9942 | 0.014819 | 0.0029673
569 | 4.9889 | 5.0336 | 4.967 | 4.985 | 5.002 | 5.0024 | 5.0075 | 5.024 | 5.0211 | 4.9757 | 4.9902 | 4.9929 | 4.9739 | 4.9751 | 4.9769 | 4.9856 | 4.9819 | 4.9876 | 4.9896 | 4.9839 | 4.9851 | 5.0028 | 5.0026 | 5.0064 | 4.9935 | 4.9969 | 4.9976 | 4.9963 | 4.9943 | 4.9998 | 4.994 | 0.014825 | 0.0029685
570 | 4.9885 | 5.0335 | 4.9669 | 4.9848 | 5.0019 | 5.0023 | 5.0074 | 5.0239 | 5.021 | 4.9755 | 4.9901 | 4.9928 | 4.9738 | 4.9749 | 4.9768 | 4.9854 | 4.9817 | 4.9874 | 4.9895 | 4.9838 | 4.9849 | 5.0027 | 5.0025 | 5.0062 | 4.9934 | 4.9968 | 4.9975 | 4.9961 | 4.9941 | 4.9996 | 4.9939 | 0.01483 | 0.0029696
571 | 4.9882 | 5.0334 | 4.9667 | 4.9847 | 5.0018 | 5.0021 | 5.0073 | 5.0237 | 5.0209 | 4.9753 | 4.99 | 4.9927 | 4.9737 | 4.9748 | 4.9767 | 4.9853 | 4.9816 | 4.9873 | 4.9894 | 4.9836 | 4.9848 | 5.0025 | 5.0023 | 5.0061 | 4.9932 | 4.9966 | 4.9973 | 4.996 | 4.994 | 4.9995 | 4.9937 | 0.014835 | 0.0029708
572 | 4.9878 | 5.0333 | 4.9666 | 4.9845 | 5.0016 | 5.002 | 5.0071 | 5.0236 | 5.0207 | 4.9751 | 4.9899 | 4.9926 | 4.9736 | 4.9747 | 4.9765 | 4.9851 | 4.9814 | 4.9871 | 4.9892 | 4.9835 | 4.9846 | 5.0024 | 5.0022 | 5.006 | 4.9931 | 4.9965 | 4.9972 | 4.9958 | 4.9938 | 4.9993 | 4.9936 | 0.014841 | 0.002972
573 | 4.9874 | 5.0331 | 4.9664 | 4.9844 | 5.0015 | 5.0018 | 5.007 | 5.0234 | 5.0206 | 4.9749 | 4.9898 | 4.9925 | 4.9734 | 4.9745 | 4.9764 | 4.985 | 4.9813 | 4.987 | 4.9891 | 4.9833 | 4.9845 | 5.0022 | 5.0021 | 5.0058 | 4.9929 | 4.9964 | 4.997 | 4.9957 | 4.9937 | 4.9992 | 4.9934 | 0.014846 | 0.0029732
574 | 4.987 | 5.033 | 4.9663 | 4.9842 | 5.0013 | 5.0017 | 5.0068 | 5.0233 | 5.0205 | 4.9747 | 4.9897 | 4.9924 | 4.9733 | 4.9744 | 4.9763 | 4.9849 | 4.9811 | 4.9869 | 4.989 | 4.9832 | 4.9843 | 5.0021 | 5.0019 | 5.0057 | 4.9928 | 4.9962 | 4.9969 | 4.9955 | 4.9935 | 4.9991 | 4.9933 | 0.014852 | 0.0029745
575 | 4.9866 | 5.0329 | 4.9661 | 4.9841 | 5.0012 | 5.0016 | 5.0067 | 5.0231 | 5.0203 | 4.9745 | 4.9896 | 4.9923 | 4.9732 | 4.9743 | 4.9761 | 4.9847 | 4.981 | 4.9867 | 4.9889 | 4.9831 | 4.9842 | 5.0019 | 5.0018 | 5.0056 | 4.9927 | 4.9961 | 4.9967 | 4.9954 | 4.9934 | 4.9989 | 4.9931 | 0.014858 | 0.0029757
576 | 4.9863 | 5.0327 | 4.966 | 4.9839 | 5.0011 | 5.0014 | 5.0066 | 5.023 | 5.0202 | 4.9743 | 4.9895 | 4.9922 | 4.9731 | 4.9741 | 4.976 | 4.9846 | 4.9808 | 4.9866 | 4.9888 | 4.9829 | 4.984 | 5.0018 | 5.0017 | 5.0054 | 4.9925 | 4.9959 | 4.9966 | 4.9952 | 4.9932 | 4.9988 | 4.993 | 0.014864 | 0.0029771
577 | 4.9859 | 5.0326 | 4.9658 | 4.9838 | 5.0009 | 5.0013 | 5.0064 | 5.0229 | 5.02 | 4.9741 | 4.9894 | 4.992 | 4.973 | 4.974 | 4.9759 | 4.9845 | 4.9807 | 4.9864 | 4.9886 | 4.9828 | 4.9839 | 5.0017 | 5.0015 | 5.0053 | 4.9924 | 4.9958 | 4.9965 | 4.9951 | 4.9931 | 4.9986 | 4.9928 | 0.014871 | 0.0029784
578 | 4.9855 | 5.0325 | 4.9657 | 4.9837 | 5.0008 | 5.0011 | 5.0063 | 5.0227 | 5.0199 | 4.9739 | 4.9893 | 4.9919 | 4.9729 | 4.9739 | 4.9758 | 4.9843 | 4.9805 | 4.9863 | 4.9885 | 4.9826 | 4.9838 | 5.0015 | 5.0014 | 5.0052 | 4.9923 | 4.9957 | 4.9963 | 4.9949 | 4.9929 | 4.9985 | 4.9927 | 0.014877 | 0.0029797
579 | 4.9851 | 5.0324 | 4.9655 | 4.9835 | 5.0006 | 5.001 | 5.0062 | 5.0226 | 5.0198 | 4.9737 | 4.9892 | 4.9918 | 4.9727 | 4.9737 | 4.9756 | 4.9842 | 4.9804 | 4.9861 | 4.9884 | 4.9825 | 4.9836 | 5.0014 | 5.0013 | 5.0051 | 4.9921 | 4.9955 | 4.9962 | 4.9948 | 4.9928 | 4.9983 | 4.9925 | 0.014883 | 0.0029811
580 | 4.9847 | 5.0322 | 4.9653 | 4.9834 | 5.0005 | 5.0009 | 5.006 | 5.0224 | 5.0196 | 4.9735 | 4.9891 | 4.9917 | 4.9726 | 4.9736 | 4.9755 | 4.984 | 4.9802 | 4.986 | 4.9883 | 4.9824 | 4.9835 | 5.0012 | 5.0011 | 5.0049 | 4.992 | 4.9954 | 4.996 | 4.9946 | 4.9926 | 4.9982 | 4.9924 | 0.01489 | 0.0029825
581 | 4.9844 | 5.0321 | 4.9652 | 4.9832 | 5.0004 | 5.0007 | 5.0059 | 5.0223 | 5.0195 | 4.9734 | 4.989 | 4.9916 | 4.9725 | 4.9735 | 4.9754 | 4.9839 | 4.9801 | 4.9859 | 4.9881 | 4.9822 | 4.9833 | 5.0011 | 5.001 | 5.0048 | 4.9919 | 4.9953 | 4.9959 | 4.9945 | 4.9925 | 4.998 | 4.9923 | 0.014897 | 0.002984
582 | 4.984 | 5.032 | 4.965 | 4.9831 | 5.0002 | 5.0006 | 5.0058 | 5.0222 | 5.0194 | 4.9732 | 4.9889 | 4.9915 | 4.9724 | 4.9733 | 4.9753 | 4.9838 | 4.98 | 4.9857 | 4.988 | 4.9821 | 4.9832 | 5.0009 | 5.0009 | 5.0047 | 4.9917 | 4.9951 | 4.9957 | 4.9943 | 4.9923 | 4.9979 | 4.9921 | 0.014904 | 0.0029854
583 | 4.9836 | 5.0319 | 4.9649 | 4.9829 | 5.0001 | 5.0004 | 5.0056 | 5.022 | 5.0192 | 4.973 | 4.9888 | 4.9914 | 4.9723 | 4.9732 | 4.9751 | 4.9836 | 4.9798 | 4.9856 | 4.9879 | 4.982 | 4.983 | 5.0008 | 5.0008 | 5.0045 | 4.9916 | 4.995 | 4.9956 | 4.9942 | 4.9922 | 4.9977 | 4.992 | 0.014911 | 0.0029869
584 | 4.9833 | 5.0317 | 4.9647 | 4.9828 | 5 | 5.0003 | 5.0055 | 5.0219 | 5.0191 | 4.9728 | 4.9887 | 4.9913 | 4.9722 | 4.9731 | 4.975 | 4.9835 | 4.9797 | 4.9854 | 4.9878 | 4.9818 | 4.9829 | 5.0006 | 5.0006 | 5.0044 | 4.9915 | 4.9949 | 4.9954 | 4.994 | 4.992 | 4.9976 | 4.9918 | 0.014918 | 0.0029884
585 | 4.9829 | 5.0316 | 4.9646 | 4.9827 | 4.9998 | 5.0002 | 5.0054 | 5.0218 | 5.019 | 4.9726 | 4.9886 | 4.9912 | 4.972 | 4.973 | 4.9749 | 4.9834 | 4.9795 | 4.9853 | 4.9877 | 4.9817 | 4.9827 | 5.0005 | 5.0005 | 5.0043 | 4.9913 | 4.9947 | 4.9953 | 4.9939 | 4.9919 | 4.9975 | 4.9917 | 0.014925 | 0.00299
586 | 4.9825 | 5.0315 | 4.9644 | 4.9825 | 4.9997 | 5 | 5.0053 | 5.0216 | 5.0188 | 4.9724 | 4.9885 | 4.9911 | 4.9719 | 4.9728 | 4.9748 | 4.9832 | 4.9794 | 4.9852 | 4.9875 | 4.9816 | 4.9826 | 5.0003 | 5.0004 | 5.0042 | 4.9912 | 4.9946 | 4.9951 | 4.9938 | 4.9917 | 4.9973 | 4.9915 | 0.014932 | 0.0029915
587 | 4.9821 | 5.0314 | 4.9643 | 4.9824 | 4.9996 | 4.9999 | 5.0051 | 5.0215 | 5.0187 | 4.9722 | 4.9884 | 4.991 | 4.9718 | 4.9727 | 4.9747 | 4.9831 | 4.9792 | 4.985 | 4.9874 | 4.9814 | 4.9825 | 5.0002 | 5.0002 | 5.004 | 4.9911 | 4.9945 | 4.995 | 4.9936 | 4.9916 | 4.9972 | 4.9914 | 0.01494 | 0.0029931
588 | 4.9818 | 5.0313 | 4.9641 | 4.9823 | 4.9994 | 4.9998 | 5.005 | 5.0213 | 5.0186 | 4.972 | 4.9883 | 4.9909 | 4.9717 | 4.9726 | 4.9745 | 4.983 | 4.9791 | 4.9849 | 4.9873 | 4.9813 | 4.9823 | 5.0001 | 5.0001 | 5.0039 | 4.9909 | 4.9943 | 4.9949 | 4.9935 | 4.9915 | 4.997 | 4.9913 | 0.014947 | 0.0029947
589 | 4.9814 | 5.0311 | 4.964 | 4.9821 | 4.9993 | 4.9996 | 5.0049 | 5.0212 | 5.0184 | 4.9719 | 4.9883 | 4.9908 | 4.9716 | 4.9724 | 4.9744 | 4.9828 | 4.979 | 4.9848 | 4.9872 | 4.9812 | 4.9822 | 4.9999 | 5 | 5.0038 | 4.9908 | 4.9942 | 4.9947 | 4.9933 | 4.9913 | 4.9969 | 4.9911 | 0.014955 | 0.0029963
590 | 4.981 | 5.031 | 4.9639 | 4.982 | 4.9992 | 4.9995 | 5.0047 | 5.0211 | 5.0183 | 4.9717 | 4.9882 | 4.9907 | 4.9715 | 4.9723 | 4.9743 | 4.9827 | 4.9788 | 4.9846 | 4.9871 | 4.981 | 4.982 | 4.9998 | 4.9999 | 5.0037 | 4.9907 | 4.9941 | 4.9946 | 4.9932 | 4.9912 | 4.9968 | 4.991 | 0.014963 | 0.002998
591 | 4.9807 | 5.0309 | 4.9637 | 4.9818 | 4.999 | 4.9994 | 5.0046 | 5.0209 | 5.0182 | 4.9715 | 4.9881 | 4.9906 | 4.9714 | 4.9722 | 4.9742 | 4.9826 | 4.9787 | 4.9845 | 4.987 | 4.9809 | 4.9819 | 4.9996 | 4.9997 | 5.0035 | 4.9906 | 4.9939 | 4.9944 | 4.993 | 4.991 | 4.9966 | 4.9908 | 0.014971 | 0.0029997
592 | 4.9803 | 5.0308 | 4.9636 | 4.9817 | 4.9989 | 4.9992 | 5.0045 | 5.0208 | 5.0181 | 4.9713 | 4.988 | 4.9905 | 4.9713 | 4.9721 | 4.974 | 4.9825 | 4.9785 | 4.9843 | 4.9868 | 4.9808 | 4.9817 | 4.9995 | 4.9996 | 5.0034 | 4.9904 | 4.9938 | 4.9943 | 4.9929 | 4.9909 | 4.9965 | 4.9907 | 0.014979 | 0.0030014
593 | 4.98 | 5.0307 | 4.9634 | 4.9816 | 4.9988 | 4.9991 | 5.0044 | 5.0207 | 5.0179 | 4.9711 | 4.9879 | 4.9904 | 4.9712 | 4.9719 | 4.9739 | 4.9823 | 4.9784 | 4.9842 | 4.9867 | 4.9806 | 4.9816 | 4.9994 | 4.9995 | 5.0033 | 4.9903 | 4.9937 | 4.9942 | 4.9927 | 4.9907 | 4.9963 | 4.9906 | 0.014987 | 0.0030031
594 | 4.9796 | 5.0305 | 4.9633 | 4.9814 | 4.9986 | 4.999 | 5.0042 | 5.0205 | 5.0178 | 4.9709 | 4.9878 | 4.9903 | 4.971 | 4.9718 | 4.9738 | 4.9822 | 4.9783 | 4.9841 | 4.9866 | 4.9805 | 4.9815 | 4.9992 | 4.9994 | 5.0032 | 4.9902 | 4.9935 | 4.994 | 4.9926 | 4.9906 | 4.9962 | 4.9904 | 0.014995 | 0.0030048
595 | 4.9792 | 5.0304 | 4.9631 | 4.9813 | 4.9985 | 4.9988 | 5.0041 | 5.0204 | 5.0177 | 4.9707 | 4.9877 | 4.9902 | 4.9709 | 4.9717 | 4.9737 | 4.9821 | 4.9781 | 4.9839 | 4.9865 | 4.9804 | 4.9813 | 4.9991 | 4.9992 | 5.003 | 4.99 | 4.9934 | 4.9939 | 4.9925 | 4.9905 | 4.9961 | 4.9903 | 0.015004 | 0.0030066
596 | 4.9789 | 5.0303 | 4.963 | 4.9812 | 4.9984 | 4.9987 | 5.004 | 5.0203 | 5.0175 | 4.9706 | 4.9876 | 4.9901 | 4.9708 | 4.9716 | 4.9736 | 4.9819 | 4.978 | 4.9838 | 4.9864 | 4.9802 | 4.9812 | 4.9989 | 4.9991 | 5.0029 | 4.9899 | 4.9933 | 4.9937 | 4.9923 | 4.9903 | 4.9959 | 4.9901 | 0.015012 | 0.0030084
597 | 4.9785 | 5.0302 | 4.9628 | 4.981 | 4.9983 | 4.9986 | 5.0039 | 5.0201 | 5.0174 | 4.9704 | 4.9875 | 4.99 | 4.9707 | 4.9714 | 4.9735 | 4.9818 | 4.9779 | 4.9837 | 4.9863 | 4.9801 | 4.9811 | 4.9988 | 4.999 | 5.0028 | 4.9898 | 4.9932 | 4.9936 | 4.9922 | 4.9902 | 4.9958 | 4.99 | 0.015021 | 0.0030102
598 | 4.9782 | 5.0301 | 4.9627 | 4.9809 | 4.9981 | 4.9984 | 5.0037 | 5.02 | 5.0173 | 4.9702 | 4.9874 | 4.9899 | 4.9706 | 4.9713 | 4.9733 | 4.9817 | 4.9777 | 4.9835 | 4.9862 | 4.98 | 4.9809 | 4.9987 | 4.9989 | 5.0027 | 4.9897 | 4.993 | 4.9935 | 4.992 | 4.99 | 4.9956 | 4.9899 | 0.01503 | 0.003012
599 | 4.9778 | 5.0299 | 4.9625 | 4.9808 | 4.998 | 4.9983 | 5.0036 | 5.0199 | 5.0172 | 4.97 | 4.9874 | 4.9898 | 4.9705 | 4.9712 | 4.9732 | 4.9816 | 4.9776 | 4.9834 | 4.986 | 4.9799 | 4.9808 | 4.9985 | 4.9987 | 5.0025 | 4.9895 | 4.9929 | 4.9933 | 4.9919 | 4.9899 | 4.9955 | 4.9897 | 0.015038 | 0.0030139
600 | 4.9774 | 5.0298 | 4.9624 | 4.9806 | 4.9979 | 4.9982 | 5.0035 | 5.0197 | 5.017 | 4.9698 | 4.9873 | 4.9897 | 4.9704 | 4.9711 | 4.9731 | 4.9814 | 4.9774 | 4.9833 | 4.9859 | 4.9797 | 4.9806 | 4.9984 | 4.9986 | 5.0024 | 4.9894 | 4.9928 | 4.9932 | 4.9918 | 4.9898 | 4.9954 | 4.9896 | 0.015047 | 0.0030157
601 | 4.9771 | 5.0297 | 4.9623 | 4.9805 | 4.9977 | 4.998 | 5.0034 | 5.0196 | 5.0169 | 4.9697 | 4.9872 | 4.9896 | 4.9703 | 4.9709 | 4.973 | 4.9813 | 4.9773 | 4.9831 | 4.9858 | 4.9796 | 4.9805 | 4.9982 | 4.9985 | 5.0023 | 4.9893 | 4.9926 | 4.9931 | 4.9916 | 4.9896 | 4.9952 | 4.9895 | 0.015056 | 0.0030176
602 | 4.9767 | 5.0296 | 4.9621 | 4.9804 | 4.9976 | 4.9979 | 5.0032 | 5.0195 | 5.0168 | 4.9695 | 4.9871 | 4.9895 | 4.9702 | 4.9708 | 4.9729 | 4.9812 | 4.9772 | 4.983 | 4.9857 | 4.9795 | 4.9804 | 4.9981 | 4.9984 | 5.0022 | 4.9892 | 4.9925 | 4.9929 | 4.9915 | 4.9895 | 4.9951 | 4.9893 | 0.015066 | 0.0030196
603 | 4.9764 | 5.0295 | 4.962 | 4.9802 | 4.9975 | 4.9978 | 5.0031 | 5.0194 | 5.0167 | 4.9693 | 4.987 | 4.9894 | 4.9701 | 4.9707 | 4.9728 | 4.981 | 4.977 | 4.9829 | 4.9856 | 4.9793 | 4.9802 | 4.998 | 4.9983 | 5.0021 | 4.989 | 4.9924 | 4.9928 | 4.9913 | 4.9893 | 4.995 | 4.9892 | 0.015075 | 0.0030215
604 | 4.976 | 5.0294 | 4.9618 | 4.9801 | 4.9974 | 4.9976 | 5.003 | 5.0192 | 5.0165 | 4.9691 | 4.9869 | 4.9893 | 4.97 | 4.9706 | 4.9726 | 4.9809 | 4.9769 | 4.9827 | 4.9855 | 4.9792 | 4.9801 | 4.9978 | 4.9981 | 5.0019 | 4.9889 | 4.9923 | 4.9926 | 4.9912 | 4.9892 | 4.9948 | 4.9891 | 0.015084 | 0.0030235
605 | 4.9757 | 5.0292 | 4.9617 | 4.98 | 4.9972 | 4.9975 | 5.0029 | 5.0191 | 5.0164 | 4.9689 | 4.9868 | 4.9892 | 4.9699 | 4.9705 | 4.9725 | 4.9808 | 4.9768 | 4.9826 | 4.9854 | 4.9791 | 4.98 | 4.9977 | 4.998 | 5.0018 | 4.9888 | 4.9921 | 4.9925 | 4.9911 | 4.9891 | 4.9947 | 4.9889 | 0.015094 | 0.0030254
606 | 4.9753 | 5.0291 | 4.9615 | 4.9798 | 4.9971 | 4.9974 | 5.0027 | 5.019 | 5.0163 | 4.9688 | 4.9868 | 4.9891 | 4.9697 | 4.9703 | 4.9724 | 4.9807 | 4.9766 | 4.9825 | 4.9853 | 4.979 | 4.9798 | 4.9976 | 4.9979 | 5.0017 | 4.9887 | 4.992 | 4.9924 | 4.9909 | 4.9889 | 4.9946 | 4.9888 | 0.015103 | 0.0030274
607 | 4.975 | 5.029 | 4.9614 | 4.9797 | 4.997 | 4.9973 | 5.0026 | 5.0188 | 5.0162 | 4.9686 | 4.9867 | 4.989 | 4.9696 | 4.9702 | 4.9723 | 4.9805 | 4.9765 | 4.9824 | 4.9851 | 4.9788 | 4.9797 | 4.9974 | 4.9978 | 5.0016 | 4.9885 | 4.9919 | 4.9922 | 4.9908 | 4.9888 | 4.9944 | 4.9887 | 0.015113 | 0.0030294
608 | 4.9746 | 5.0289 | 4.9613 | 4.9796 | 4.9969 | 4.9971 | 5.0025 | 5.0187 | 5.016 | 4.9684 | 4.9866 | 4.9889 | 4.9695 | 4.9701 | 4.9722 | 4.9804 | 4.9764 | 4.9822 | 4.985 | 4.9787 | 4.9795 | 4.9973 | 4.9977 | 5.0015 | 4.9884 | 4.9918 | 4.9921 | 4.9906 | 4.9887 | 4.9943 | 4.9885 | 0.015123 | 0.0030315
609 | 4.9743 | 5.0288 | 4.9611 | 4.9794 | 4.9967 | 4.997 | 5.0024 | 5.0186 | 5.0159 | 4.9682 | 4.9865 | 4.9888 | 4.9694 | 4.97 | 4.9721 | 4.9803 | 4.9762 | 4.9821 | 4.9849 | 4.9786 | 4.9794 | 4.9972 | 4.9975 | 5.0013 | 4.9883 | 4.9916 | 4.992 | 4.9905 | 4.9885 | 4.9942 | 4.9884 | 0.015132 | 0.0030335
610 | 4.9739 | 5.0287 | 4.961 | 4.9793 | 4.9966 | 4.9969 | 5.0023 | 5.0185 | 5.0158 | 4.968 | 4.9864 | 4.9887 | 4.9693 | 4.9699 | 4.972 | 4.9802 | 4.9761 | 4.982 | 4.9848 | 4.9785 | 4.9793 | 4.997 | 4.9974 | 5.0012 | 4.9882 | 4.9915 | 4.9918 | 4.9904 | 4.9884 | 4.994 | 4.9883 | 0.015142 | 0.0030356
611 | 4.9736 | 5.0286 | 4.9608 | 4.9792 | 4.9965 | 4.9967 | 5.0021 | 5.0183 | 5.0157 | 4.9679 | 4.9863 | 4.9886 | 4.9692 | 4.9697 | 4.9718 | 4.9801 | 4.976 | 4.9818 | 4.9847 | 4.9783 | 4.9791 | 4.9969 | 4.9973 | 5.0011 | 4.988 | 4.9914 | 4.9917 | 4.9902 | 4.9882 | 4.9939 | 4.9881 | 0.015152 | 0.0030377
612 | 4.9732 | 5.0284 | 4.9607 | 4.979 | 4.9964 | 4.9966 | 5.002 | 5.0182 | 5.0155 | 4.9677 | 4.9862 | 4.9885 | 4.9691 | 4.9696 | 4.9717 | 4.9799 | 4.9758 | 4.9817 | 4.9846 | 4.9782 | 4.979 | 4.9968 | 4.9972 | 5.001 | 4.9879 | 4.9913 | 4.9916 | 4.9901 | 4.9881 | 4.9938 | 4.988 | 0.015163 | 0.0030398
613 | 4.9729 | 5.0283 | 4.9606 | 4.9789 | 4.9962 | 4.9965 | 5.0019 | 5.0181 | 5.0154 | 4.9675 | 4.9862 | 4.9884 | 4.969 | 4.9695 | 4.9716 | 4.9798 | 4.9757 | 4.9816 | 4.9845 | 4.9781 | 4.9789 | 4.9966 | 4.9971 | 5.0009 | 4.9878 | 4.9912 | 4.9914 | 4.99 | 4.988 | 4.9936 | 4.9879 | 0.015173 | 0.0030419
614 | 4.9726 | 5.0282 | 4.9604 | 4.9788 | 4.9961 | 4.9964 | 5.0018 | 5.0179 | 5.0153 | 4.9673 | 4.9861 | 4.9883 | 4.9689 | 4.9694 | 4.9715 | 4.9797 | 4.9756 | 4.9815 | 4.9844 | 4.978 | 4.9787 | 4.9965 | 4.9969 | 5.0008 | 4.9877 | 4.991 | 4.9913 | 4.9898 | 4.9878 | 4.9935 | 4.9877 | 0.015183 | 0.0030441
615 | 4.9722 | 5.0281 | 4.9603 | 4.9787 | 4.996 | 4.9962 | 5.0017 | 5.0178 | 5.0152 | 4.9672 | 4.986 | 4.9882 | 4.9688 | 4.9693 | 4.9714 | 4.9796 | 4.9754 | 4.9813 | 4.9843 | 4.9779 | 4.9786 | 4.9964 | 4.9968 | 5.0006 | 4.9876 | 4.9909 | 4.9912 | 4.9897 | 4.9877 | 4.9934 | 4.9876 | 0.015194 | 0.0030463
616 | 4.9719 | 5.028 | 4.9601 | 4.9785 | 4.9959 | 4.9961 | 5.0015 | 5.0177 | 5.0151 | 4.967 | 4.9859 | 4.9881 | 4.9687 | 4.9691 | 4.9713 | 4.9794 | 4.9753 | 4.9812 | 4.9842 | 4.9777 | 4.9785 | 4.9962 | 4.9967 | 5.0005 | 4.9874 | 4.9908 | 4.991 | 4.9895 | 4.9876 | 4.9933 | 4.9875 | 0.015204 | 0.0030484
617 | 4.9715 | 5.0279 | 4.96 | 4.9784 | 4.9958 | 4.996 | 5.0014 | 5.0176 | 5.0149 | 4.9668 | 4.9858 | 4.9881 | 4.9686 | 4.969 | 4.9712 | 4.9793 | 4.9752 | 4.9811 | 4.9841 | 4.9776 | 4.9783 | 4.9961 | 4.9966 | 5.0004 | 4.9873 | 4.9907 | 4.9909 | 4.9894 | 4.9874 | 4.9931 | 4.9874 | 0.015215 | 0.0030506
618 | 4.9712 | 5.0278 | 4.9599 | 4.9783 | 4.9956 | 4.9959 | 5.0013 | 5.0174 | 5.0148 | 4.9666 | 4.9857 | 4.988 | 4.9685 | 4.9689 | 4.9711 | 4.9792 | 4.9751 | 4.9809 | 4.984 | 4.9775 | 4.9782 | 4.996 | 4.9965 | 5.0003 | 4.9872 | 4.9905 | 4.9908 | 4.9893 | 4.9873 | 4.993 | 4.9872 | 0.015225 | 0.0030529
619 | 4.9709 | 5.0277 | 4.9597 | 4.9781 | 4.9955 | 4.9957 | 5.0012 | 5.0173 | 5.0147 | 4.9665 | 4.9857 | 4.9879 | 4.9684 | 4.9688 | 4.971 | 4.9791 | 4.9749 | 4.9808 | 4.9838 | 4.9774 | 4.9781 | 4.9958 | 4.9964 | 5.0002 | 4.9871 | 4.9904 | 4.9906 | 4.9891 | 4.9872 | 4.9929 | 4.9871 | 0.015236 | 0.0030551
620 | 4.9705 | 5.0275 | 4.9596 | 4.978 | 4.9954 | 4.9956 | 5.0011 | 5.0172 | 5.0146 | 4.9663 | 4.9856 | 4.9878 | 4.9683 | 4.9687 | 4.9708 | 4.979 | 4.9748 | 4.9807 | 4.9837 | 4.9772 | 4.978 | 4.9957 | 4.9962 | 5.0001 | 4.987 | 4.9903 | 4.9905 | 4.989 | 4.987 | 4.9927 | 4.987 | 0.015247 | 0.0030573
621 | 4.9702 | 5.0274 | 4.9595 | 4.9779 | 4.9953 | 4.9955 | 5.001 | 5.0171 | 5.0145 | 4.9661 | 4.9855 | 4.9877 | 4.9682 | 4.9686 | 4.9707 | 4.9788 | 4.9747 | 4.9806 | 4.9836 | 4.9771 | 4.9778 | 4.9956 | 4.9961 | 5 | 4.9868 | 4.9902 | 4.9904 | 4.9889 | 4.9869 | 4.9926 | 4.9868 | 0.015258 | 0.0030596
622 | 4.9698 | 5.0273 | 4.9593 | 4.9778 | 4.9952 | 4.9954 | 5.0008 | 5.0169 | 5.0143 | 4.9659 | 4.9854 | 4.9876 | 4.9681 | 4.9684 | 4.9706 | 4.9787 | 4.9745 | 4.9804 | 4.9835 | 4.977 | 4.9777 | 4.9954 | 4.996 | 4.9998 | 4.9867 | 4.9901 | 4.9902 | 4.9887 | 4.9868 | 4.9925 | 4.9867 | 0.015269 | 0.0030619
623 | 4.9695 | 5.0272 | 4.9592 | 4.9776 | 4.995 | 4.9952 | 5.0007 | 5.0168 | 5.0142 | 4.9658 | 4.9853 | 4.9875 | 4.968 | 4.9683 | 4.9705 | 4.9786 | 4.9744 | 4.9803 | 4.9834 | 4.9769 | 4.9776 | 4.9953 | 4.9959 | 4.9997 | 4.9866 | 4.99 | 4.9901 | 4.9886 | 4.9866 | 4.9924 | 4.9866 | 0.01528 | 0.0030642
624 | 4.9692 | 5.0271 | 4.959 | 4.9775 | 4.9949 | 4.9951 | 5.0006 | 5.0167 | 5.0141 | 4.9656 | 4.9853 | 4.9874 | 4.9679 | 4.9682 | 4.9704 | 4.9785 | 4.9743 | 4.9802 | 4.9833 | 4.9768 | 4.9774 | 4.9952 | 4.9958 | 4.9996 | 4.9865 | 4.9898 | 4.99 | 4.9885 | 4.9865 | 4.9922 | 4.9865 | 0.015291 | 0.0030665
625 | 4.9688 | 5.027 | 4.9589 | 4.9774 | 4.9948 | 4.995 | 5.0005 | 5.0166 | 5.014 | 4.9654 | 4.9852 | 4.9873 | 4.9678 | 4.9681 | 4.9703 | 4.9784 | 4.9742 | 4.9801 | 4.9832 | 4.9766 | 4.9773 | 4.9951 | 4.9957 | 4.9995 | 4.9864 | 4.9897 | 4.9899 | 4.9883 | 4.9864 | 4.9921 | 4.9863 | 0.015302 | 0.0030689
626 | 4.9685 | 5.0269 | 4.9588 | 4.9773 | 4.9947 | 4.9949 | 5.0004 | 5.0165 | 5.0139 | 4.9653 | 4.9851 | 4.9872 | 4.9677 | 4.968 | 4.9702 | 4.9782 | 4.974 | 4.9799 | 4.9831 | 4.9765 | 4.9772 | 4.9949 | 4.9955 | 4.9994 | 4.9863 | 4.9896 | 4.9897 | 4.9882 | 4.9862 | 4.992 | 4.9862 | 0.015314 | 0.0030712
627 | 4.9682 | 5.0268 | 4.9586 | 4.9771 | 4.9946 | 4.9947 | 5.0003 | 5.0163 | 5.0138 | 4.9651 | 4.985 | 4.9871 | 4.9676 | 4.9679 | 4.9701 | 4.9781 | 4.9739 | 4.9798 | 4.983 | 4.9764 | 4.977 | 4.9948 | 4.9954 | 4.9993 | 4.9861 | 4.9895 | 4.9896 | 4.9881 | 4.9861 | 4.9918 | 4.9861 | 0.015325 | 0.0030736
628 | 4.9678 | 5.0267 | 4.9585 | 4.977 | 4.9944 | 4.9946 | 5.0002 | 5.0162 | 5.0136 | 4.9649 | 4.9849 | 4.987 | 4.9675 | 4.9677 | 4.97 | 4.978 | 4.9738 | 4.9797 | 4.9829 | 4.9763 | 4.9769 | 4.9947 | 4.9953 | 4.9992 | 4.986 | 4.9894 | 4.9895 | 4.9879 | 4.986 | 4.9917 | 4.9859 | 0.015337 | 0.003076
629 | 4.9675 | 5.0266 | 4.9584 | 4.9769 | 4.9943 | 4.9945 | 5 | 5.0161 | 5.0135 | 4.9647 | 4.9849 | 4.987 | 4.9674 | 4.9676 | 4.9699 | 4.9779 | 4.9736 | 4.9796 | 4.9828 | 4.9762 | 4.9768 | 4.9945 | 4.9952 | 4.9991 | 4.9859 | 4.9892 | 4.9893 | 4.9878 | 4.9859 | 4.9916 | 4.9858 | 0.015348 | 0.0030784
630 | 4.9672 | 5.0265 | 4.9582 | 4.9768 | 4.9942 | 4.9944 | 4.9999 | 5.016 | 5.0134 | 4.9646 | 4.9848 | 4.9869 | 4.9673 | 4.9675 | 4.9698 | 4.9778 | 4.9735 | 4.9794 | 4.9827 | 4.976 | 4.9767 | 4.9944 | 4.9951 | 4.9989 | 4.9858 | 4.9891 | 4.9892 | 4.9877 | 4.9857 | 4.9915 | 4.9857 | 0.01536 | 0.0030808
631 | 4.9669 | 5.0263 | 4.9581 | 4.9766 | 4.9941 | 4.9943 | 4.9998 | 5.0158 | 5.0133 | 4.9644 | 4.9847 | 4.9868 | 4.9672 | 4.9674 | 4.9697 | 4.9777 | 4.9734 | 4.9793 | 4.9826 | 4.9759 | 4.9765 | 4.9943 | 4.995 | 4.9988 | 4.9857 | 4.989 | 4.9891 | 4.9875 | 4.9856 | 4.9913 | 4.9856 | 0.015371 | 0.0030832
632 | 4.9665 | 5.0262 | 4.958 | 4.9765 | 4.994 | 4.9941 | 4.9997 | 5.0157 | 5.0132 | 4.9642 | 4.9846 | 4.9867 | 4.9671 | 4.9673 | 4.9695 | 4.9775 | 4.9733 | 4.9792 | 4.9825 | 4.9758 | 4.9764 | 4.9942 | 4.9949 | 4.9987 | 4.9856 | 4.9889 | 4.989 | 4.9874 | 4.9855 | 4.9912 | 4.9854 | 0.015383 | 0.0030856
633 | 4.9662 | 5.0261 | 4.9578 | 4.9764 | 4.9938 | 4.994 | 4.9996 | 5.0156 | 5.0131 | 4.9641 | 4.9846 | 4.9866 | 4.967 | 4.9672 | 4.9694 | 4.9774 | 4.9731 | 4.9791 | 4.9824 | 4.9757 | 4.9763 | 4.994 | 4.9947 | 4.9986 | 4.9854 | 4.9888 | 4.9888 | 4.9873 | 4.9853 | 4.9911 | 4.9853 | 0.015395 | 0.0030881
634 | 4.9659 | 5.026 | 4.9577 | 4.9763 | 4.9937 | 4.9939 | 4.9995 | 5.0155 | 5.0129 | 4.9639 | 4.9845 | 4.9865 | 4.9669 | 4.9671 | 4.9693 | 4.9773 | 4.973 | 4.979 | 4.9823 | 4.9756 | 4.9761 | 4.9939 | 4.9946 | 4.9985 | 4.9853 | 4.9887 | 4.9887 | 4.9872 | 4.9852 | 4.991 | 4.9852 | 0.015407 | 0.0030905
635 | 4.9656 | 5.0259 | 4.9576 | 4.9761 | 4.9936 | 4.9938 | 4.9994 | 5.0154 | 5.0128 | 4.9637 | 4.9844 | 4.9864 | 4.9668 | 4.9669 | 4.9692 | 4.9772 | 4.9729 | 4.9788 | 4.9822 | 4.9755 | 4.976 | 4.9938 | 4.9945 | 4.9984 | 4.9852 | 4.9885 | 4.9886 | 4.987 | 4.9851 | 4.9908 | 4.9851 | 0.015419 | 0.003093
636 | 4.9652 | 5.0258 | 4.9574 | 4.976 | 4.9935 | 4.9937 | 4.9992 | 5.0152 | 5.0127 | 4.9636 | 4.9843 | 4.9863 | 4.9667 | 4.9668 | 4.9691 | 4.9771 | 4.9728 | 4.9787 | 4.9821 | 4.9753 | 4.9759 | 4.9937 | 4.9944 | 4.9983 | 4.9851 | 4.9884 | 4.9884 | 4.9869 | 4.9849 | 4.9907 | 4.985 | 0.015431 | 0.0030955
637 | 4.9649 | 5.0257 | 4.9573 | 4.9759 | 4.9934 | 4.9935 | 4.9991 | 5.0151 | 5.0126 | 4.9634 | 4.9843 | 4.9862 | 4.9666 | 4.9667 | 4.969 | 4.977 | 4.9726 | 4.9786 | 4.982 | 4.9752 | 4.9758 | 4.9935 | 4.9943 | 4.9982 | 4.985 | 4.9883 | 4.9883 | 4.9868 | 4.9848 | 4.9906 | 4.9848 | 0.015443 | 0.003098
638 | 4.9646 | 5.0256 | 4.9572 | 4.9758 | 4.9933 | 4.9934 | 4.999 | 5.015 | 5.0125 | 4.9632 | 4.9842 | 4.9862 | 4.9665 | 4.9666 | 4.9689 | 4.9768 | 4.9725 | 4.9785 | 4.9819 | 4.9751 | 4.9756 | 4.9934 | 4.9942 | 4.9981 | 4.9849 | 4.9882 | 4.9882 | 4.9866 | 4.9847 | 4.9905 | 4.9847 | 0.015455 | 0.0031006
639 | 4.9643 | 5.0255 | 4.957 | 4.9756 | 4.9932 | 4.9933 | 4.9989 | 5.0149 | 5.0124 | 4.9631 | 4.9841 | 4.9861 | 4.9664 | 4.9665 | 4.9688 | 4.9767 | 4.9724 | 4.9783 | 4.9818 | 4.975 | 4.9755 | 4.9933 | 4.9941 | 4.998 | 4.9848 | 4.9881 | 4.9881 | 4.9865 | 4.9846 | 4.9903 | 4.9846 | 0.015468 | 0.0031031
640 | 4.9639 | 5.0254 | 4.9569 | 4.9755 | 4.993 | 4.9932 | 4.9988 | 5.0148 | 5.0122 | 4.9629 | 4.984 | 4.986 | 4.9663 | 4.9664 | 4.9687 | 4.9766 | 4.9723 | 4.9782 | 4.9817 | 4.9749 | 4.9754 | 4.9932 | 4.994 | 4.9978 | 4.9846 | 4.988 | 4.9879 | 4.9864 | 4.9844 | 4.9902 | 4.9845 | 0.01548 | 0.0031056
641 | 4.9636 | 5.0253 | 4.9568 | 4.9754 | 4.9929 | 4.9931 | 4.9987 | 5.0147 | 5.0121 | 4.9627 | 4.9839 | 4.9859 | 4.9663 | 4.9663 | 4.9686 | 4.9765 | 4.9722 | 4.9781 | 4.9816 | 4.9748 | 4.9752 | 4.993 | 4.9938 | 4.9977 | 4.9845 | 4.9879 | 4.9878 | 4.9862 | 4.9843 | 4.9901 | 4.9843 | 0.015492 | 0.0031082
642 | 4.9633 | 5.0252 | 4.9567 | 4.9753 | 4.9928 | 4.9929 | 4.9986 | 5.0145 | 5.012 | 4.9626 | 4.9839 | 4.9858 | 4.9662 | 4.9662 | 4.9685 | 4.9764 | 4.972 | 4.978 | 4.9815 | 4.9746 | 4.9751 | 4.9929 | 4.9937 | 4.9976 | 4.9844 | 4.9877 | 4.9877 | 4.9861 | 4.9842 | 4.99 | 4.9842 | 0.015505 | 0.0031108
643 | 4.963 | 5.0251 | 4.9565 | 4.9752 | 4.9927 | 4.9928 | 4.9985 | 5.0144 | 5.0119 | 4.9624 | 4.9838 | 4.9857 | 4.9661 | 4.9661 | 4.9684 | 4.9763 | 4.9719 | 4.9779 | 4.9814 | 4.9745 | 4.975 | 4.9928 | 4.9936 | 4.9975 | 4.9843 | 4.9876 | 4.9876 | 4.986 | 4.9841 | 4.9899 | 4.9841 | 0.015517 | 0.0031134
644 | 4.9627 | 5.025 | 4.9564 | 4.975 | 4.9926 | 4.9927 | 4.9984 | 5.0143 | 5.0118 | 4.9622 | 4.9837 | 4.9856 | 4.966 | 4.9659 | 4.9683 | 4.9762 | 4.9718 | 4.9777 | 4.9813 | 4.9744 | 4.9749 | 4.9927 | 4.9935 | 4.9974 | 4.9842 | 4.9875 | 4.9874 | 4.9859 | 4.9839 | 4.9897 | 4.984 | 0.01553 | 0.003116
645 | 4.9624 | 5.0249 | 4.9563 | 4.9749 | 4.9925 | 4.9926 | 4.9983 | 5.0142 | 5.0117 | 4.9621 | 4.9836 | 4.9855 | 4.9659 | 4.9658 | 4.9682 | 4.976 | 4.9717 | 4.9776 | 4.9812 | 4.9743 | 4.9747 | 4.9925 | 4.9934 | 4.9973 | 4.9841 | 4.9874 | 4.9873 | 4.9857 | 4.9838 | 4.9896 | 4.9838 | 0.015542 | 0.0031186
646 | 4.962 | 5.0248 | 4.9561 | 4.9748 | 4.9924 | 4.9925 | 4.9981 | 5.0141 | 5.0116 | 4.9619 | 4.9836 | 4.9855 | 4.9658 | 4.9657 | 4.9681 | 4.9759 | 4.9715 | 4.9775 | 4.9811 | 4.9742 | 4.9746 | 4.9924 | 4.9933 | 4.9972 | 4.984 | 4.9873 | 4.9872 | 4.9856 | 4.9837 | 4.9895 | 4.9837 | 0.015555 | 0.0031212
647 | 4.9617 | 5.0247 | 4.956 | 4.9747 | 4.9922 | 4.9924 | 4.998 | 5.0139 | 5.0115 | 4.9617 | 4.9835 | 4.9854 | 4.9657 | 4.9656 | 4.968 | 4.9758 | 4.9714 | 4.9774 | 4.981 | 4.9741 | 4.9745 | 4.9923 | 4.9932 | 4.9971 | 4.9839 | 4.9872 | 4.9871 | 4.9855 | 4.9835 | 4.9894 | 4.9836 | 0.015568 | 0.0031238
648 | 4.9614 | 5.0246 | 4.9559 | 4.9746 | 4.9921 | 4.9922 | 4.9979 | 5.0138 | 5.0113 | 4.9616 | 4.9834 | 4.9853 | 4.9656 | 4.9655 | 4.9679 | 4.9757 | 4.9713 | 4.9773 | 4.9809 | 4.974 | 4.9744 | 4.9922 | 4.9931 | 4.997 | 4.9838 | 4.9871 | 4.9869 | 4.9854 | 4.9834 | 4.9893 | 4.9835 | 0.015581 | 0.0031265
649 | 4.9611 | 5.0245 | 4.9557 | 4.9744 | 4.992 | 4.9921 | 4.9978 | 5.0137 | 5.0112 | 4.9614 | 4.9834 | 4.9852 | 4.9655 | 4.9654 | 4.9678 | 4.9756 | 4.9712 | 4.9772 | 4.9808 | 4.9739 | 4.9742 | 4.992 | 4.993 | 4.9969 | 4.9836 | 4.987 | 4.9868 | 4.9852 | 4.9833 | 4.9891 | 4.9834 | 0.015594 | 0.0031291
650 | 4.9608 | 5.0243 | 4.9556 | 4.9743 | 4.9919 | 4.992 | 4.9977 | 5.0136 | 5.0111 | 4.9612 | 4.9833 | 4.9851 | 4.9654 | 4.9653 | 4.9677 | 4.9755 | 4.9711 | 4.977 | 4.9807 | 4.9737 | 4.9741 | 4.9919 | 4.9929 | 4.9968 | 4.9835 | 4.9868 | 4.9867 | 4.9851 | 4.9832 | 4.989 | 4.9832 | 0.015606 | 0.0031318
651 | 4.9605 | 5.0242 | 4.9555 | 4.9742 | 4.9918 | 4.9919 | 4.9976 | 5.0135 | 5.011 | 4.9611 | 4.9832 | 4.985 | 4.9653 | 4.9652 | 4.9676 | 4.9754 | 4.9709 | 4.9769 | 4.9806 | 4.9736 | 4.974 | 4.9918 | 4.9927 | 4.9967 | 4.9834 | 4.9867 | 4.9866 | 4.985 | 4.983 | 4.9889 | 4.9831 | 0.015619 | 0.0031345
652 | 4.9602 | 5.0241 | 4.9554 | 4.9741 | 4.9917 | 4.9918 | 4.9975 | 5.0134 | 5.0109 | 4.9609 | 4.9831 | 4.9849 | 4.9652 | 4.9651 | 4.9675 | 4.9753 | 4.9708 | 4.9768 | 4.9805 | 4.9735 | 4.9739 | 4.9917 | 4.9926 | 4.9966 | 4.9833 | 4.9866 | 4.9864 | 4.9848 | 4.9829 | 4.9888 | 4.983 | 0.015632 | 0.0031371
653 | 4.9599 | 5.024 | 4.9552 | 4.974 | 4.9916 | 4.9917 | 4.9974 | 5.0133 | 5.0108 | 4.9608 | 4.9831 | 4.9849 | 4.9651 | 4.9649 | 4.9674 | 4.9751 | 4.9707 | 4.9767 | 4.9804 | 4.9734 | 4.9737 | 4.9916 | 4.9925 | 4.9965 | 4.9832 | 4.9865 | 4.9863 | 4.9847 | 4.9828 | 4.9887 | 4.9829 | 0.015645 | 0.0031398
654 | 4.9595 | 5.0239 | 4.9551 | 4.9738 | 4.9915 | 4.9915 | 4.9973 | 5.0131 | 5.0107 | 4.9606 | 4.983 | 4.9848 | 4.965 | 4.9648 | 4.9673 | 4.975 | 4.9706 | 4.9766 | 4.9803 | 4.9733 | 4.9736 | 4.9914 | 4.9924 | 4.9963 | 4.9831 | 4.9864 | 4.9862 | 4.9846 | 4.9827 | 4.9885 | 4.9828 | 0.015659 | 0.0031425
655 | 4.9592 | 5.0238 | 4.955 | 4.9737 | 4.9913 | 4.9914 | 4.9972 | 5.013 | 5.0106 | 4.9604 | 4.9829 | 4.9847 | 4.9649 | 4.9647 | 4.9672 | 4.9749 | 4.9705 | 4.9764 | 4.9802 | 4.9732 | 4.9735 | 4.9913 | 4.9923 | 4.9962 | 4.983 | 4.9863 | 4.9861 | 4.9845 | 4.9825 | 4.9884 | 4.9827 | 0.015672 | 0.0031453
656 | 4.9589 | 5.0237 | 4.9549 | 4.9736 | 4.9912 | 4.9913 | 4.9971 | 5.0129 | 5.0105 | 4.9603 | 4.9828 | 4.9846 | 4.9648 | 4.9646 | 4.9671 | 4.9748 | 4.9703 | 4.9763 | 4.9801 | 4.9731 | 4.9734 | 4.9912 | 4.9922 | 4.9961 | 4.9829 | 4.9862 | 4.9859 | 4.9843 | 4.9824 | 4.9883 | 4.9825 | 0.015685 | 0.003148
657 | 4.9586 | 5.0236 | 4.9547 | 4.9735 | 4.9911 | 4.9912 | 4.997 | 5.0128 | 5.0103 | 4.9601 | 4.9828 | 4.9845 | 4.9648 | 4.9645 | 4.967 | 4.9747 | 4.9702 | 4.9762 | 4.98 | 4.973 | 4.9733 | 4.9911 | 4.9921 | 4.996 | 4.9828 | 4.9861 | 4.9858 | 4.9842 | 4.9823 | 4.9882 | 4.9824 | 0.015698 | 0.0031507
658 | 4.9583 | 5.0235 | 4.9546 | 4.9734 | 4.991 | 4.9911 | 4.9968 | 5.0127 | 5.0102 | 4.9599 | 4.9827 | 4.9844 | 4.9647 | 4.9644 | 4.9669 | 4.9746 | 4.9701 | 4.9761 | 4.9799 | 4.9728 | 4.9731 | 4.991 | 4.992 | 4.9959 | 4.9826 | 4.986 | 4.9857 | 4.9841 | 4.9822 | 4.9881 | 4.9823 | 0.015712 | 0.0031535
659 | 4.958 | 5.0234 | 4.9545 | 4.9733 | 4.9909 | 4.991 | 4.9967 | 5.0126 | 5.0101 | 4.9598 | 4.9826 | 4.9843 | 4.9646 | 4.9643 | 4.9668 | 4.9745 | 4.97 | 4.976 | 4.9798 | 4.9727 | 4.973 | 4.9908 | 4.9919 | 4.9958 | 4.9825 | 4.9859 | 4.9856 | 4.984 | 4.982 | 4.9879 | 4.9822 | 0.015725 | 0.0031562
660 | 4.9577 | 5.0233 | 4.9543 | 4.9731 | 4.9908 | 4.9909 | 4.9966 | 5.0124 | 5.01 | 4.9596 | 4.9826 | 4.9843 | 4.9645 | 4.9642 | 4.9667 | 4.9744 | 4.9699 | 4.9759 | 4.9797 | 4.9726 | 4.9729 | 4.9907 | 4.9918 | 4.9957 | 4.9824 | 4.9857 | 4.9855 | 4.9838 | 4.9819 | 4.9878 | 4.9821 | 0.015738 | 0.003159
661 | 4.9574 | 5.0232 | 4.9542 | 4.973 | 4.9907 | 4.9907 | 4.9965 | 5.0123 | 5.0099 | 4.9595 | 4.9825 | 4.9842 | 4.9644 | 4.9641 | 4.9666 | 4.9743 | 4.9697 | 4.9757 | 4.9796 | 4.9725 | 4.9728 | 4.9906 | 4.9917 | 4.9956 | 4.9823 | 4.9856 | 4.9853 | 4.9837 | 4.9818 | 4.9877 | 4.9819 | 0.015752 | 0.0031618
662 | 4.9571 | 5.0231 | 4.9541 | 4.9729 | 4.9906 | 4.9906 | 4.9964 | 5.0122 | 5.0098 | 4.9593 | 4.9824 | 4.9841 | 4.9643 | 4.964 | 4.9665 | 4.9741 | 4.9696 | 4.9756 | 4.9795 | 4.9724 | 4.9726 | 4.9905 | 4.9916 | 4.9955 | 4.9822 | 4.9855 | 4.9852 | 4.9836 | 4.9817 | 4.9876 | 4.9818 | 0.015765 | 0.0031645
663 | 4.9568 | 5.023 | 4.954 | 4.9728 | 4.9905 | 4.9905 | 4.9963 | 5.0121 | 5.0097 | 4.9591 | 4.9823 | 4.984 | 4.9642 | 4.9639 | 4.9664 | 4.974 | 4.9695 | 4.9755 | 4.9794 | 4.9723 | 4.9725 | 4.9904 | 4.9915 | 4.9954 | 4.9821 | 4.9854 | 4.9851 | 4.9835 | 4.9816 | 4.9875 | 4.9817 | 0.015779 | 0.0031673
664 | 4.9565 | 5.0229 | 4.9538 | 4.9727 | 4.9903 | 4.9904 | 4.9962 | 5.012 | 5.0096 | 4.959 | 4.9823 | 4.9839 | 4.9641 | 4.9638 | 4.9663 | 4.9739 | 4.9694 | 4.9754 | 4.9793 | 4.9722 | 4.9724 | 4.9902 | 4.9913 | 4.9953 | 4.982 | 4.9853 | 4.985 | 4.9833 | 4.9814 | 4.9874 | 4.9816 | 0.015792 | 0.0031701
665 | 4.9562 | 5.0228 | 4.9537 | 4.9725 | 4.9902 | 4.9903 | 4.9961 | 5.0119 | 5.0095 | 4.9588 | 4.9822 | 4.9838 | 4.964 | 4.9637 | 4.9662 | 4.9738 | 4.9693 | 4.9753 | 4.9792 | 4.9721 | 4.9723 | 4.9901 | 4.9912 | 4.9952 | 4.9819 | 4.9852 | 4.9849 | 4.9832 | 4.9813 | 4.9872 | 4.9815 | 0.015806 | 0.0031729
666 | 4.9559 | 5.0227 | 4.9536 | 4.9724 | 4.9901 | 4.9902 | 4.996 | 5.0118 | 5.0094 | 4.9587 | 4.9821 | 4.9838 | 4.9639 | 4.9635 | 4.9661 | 4.9737 | 4.9691 | 4.9752 | 4.9791 | 4.972 | 4.9722 | 4.99 | 4.9911 | 4.9951 | 4.9818 | 4.9851 | 4.9847 | 4.9831 | 4.9812 | 4.9871 | 4.9814 | 0.01582 | 0.0031758
667 | 4.9556 | 5.0226 | 4.9535 | 4.9723 | 4.99 | 4.9901 | 4.9959 | 5.0117 | 5.0093 | 4.9585 | 4.9821 | 4.9837 | 4.9639 | 4.9634 | 4.966 | 4.9736 | 4.969 | 4.9751 | 4.9791 | 4.9718 | 4.972 | 4.9899 | 4.991 | 4.995 | 4.9817 | 4.985 | 4.9846 | 4.983 | 4.9811 | 4.987 | 4.9812 | 0.015833 | 0.0031786
668 | 4.9553 | 5.0225 | 4.9533 | 4.9722 | 4.9899 | 4.99 | 4.9958 | 5.0115 | 5.0091 | 4.9583 | 4.982 | 4.9836 | 4.9638 | 4.9633 | 4.9659 | 4.9735 | 4.9689 | 4.9749 | 4.979 | 4.9717 | 4.9719 | 4.9898 | 4.9909 | 4.9949 | 4.9816 | 4.9849 | 4.9845 | 4.9829 | 4.9809 | 4.9869 | 4.9811 | 0.015847 | 0.0031814
669 | 4.955 | 5.0224 | 4.9532 | 4.9721 | 4.9898 | 4.9898 | 4.9957 | 5.0114 | 5.009 | 4.9582 | 4.9819 | 4.9835 | 4.9637 | 4.9632 | 4.9658 | 4.9734 | 4.9688 | 4.9748 | 4.9789 | 4.9716 | 4.9718 | 4.9896 | 4.9908 | 4.9948 | 4.9815 | 4.9848 | 4.9844 | 4.9827 | 4.9808 | 4.9868 | 4.981 | 0.015861 | 0.0031842
670 | 4.9547 | 5.0223 | 4.9531 | 4.972 | 4.9897 | 4.9897 | 4.9956 | 5.0113 | 5.0089 | 4.958 | 4.9818 | 4.9834 | 4.9636 | 4.9631 | 4.9657 | 4.9733 | 4.9687 | 4.9747 | 4.9788 | 4.9715 | 4.9717 | 4.9895 | 4.9907 | 4.9947 | 4.9814 | 4.9847 | 4.9843 | 4.9826 | 4.9807 | 4.9867 | 4.9809 | 0.015875 | 0.0031871
671 | 4.9544 | 5.0222 | 4.953 | 4.9719 | 4.9896 | 4.9896 | 4.9955 | 5.0112 | 5.0088 | 4.9579 | 4.9818 | 4.9834 | 4.9635 | 4.963 | 4.9656 | 4.9732 | 4.9686 | 4.9746 | 4.9787 | 4.9714 | 4.9716 | 4.9894 | 4.9906 | 4.9946 | 4.9812 | 4.9846 | 4.9841 | 4.9825 | 4.9806 | 4.9865 | 4.9808 | 0.015888 | 0.0031899
672 | 4.9541 | 5.0221 | 4.9528 | 4.9717 | 4.9895 | 4.9895 | 4.9954 | 5.0111 | 5.0087 | 4.9577 | 4.9817 | 4.9833 | 4.9634 | 4.9629 | 4.9655 | 4.9731 | 4.9684 | 4.9745 | 4.9786 | 4.9713 | 4.9714 | 4.9893 | 4.9905 | 4.9945 | 4.9811 | 4.9845 | 4.984 | 4.9824 | 4.9804 | 4.9864 | 4.9807 | 0.015902 | 0.0031928
673 | 4.9538 | 5.022 | 4.9527 | 4.9716 | 4.9894 | 4.9894 | 4.9953 | 5.011 | 5.0086 | 4.9576 | 4.9816 | 4.9832 | 4.9633 | 4.9628 | 4.9654 | 4.973 | 4.9683 | 4.9744 | 4.9785 | 4.9712 | 4.9713 | 4.9892 | 4.9904 | 4.9944 | 4.981 | 4.9843 | 4.9839 | 4.9822 | 4.9803 | 4.9863 | 4.9805 | 0.015916 | 0.0031957
674 | 4.9535 | 5.0219 | 4.9526 | 4.9715 | 4.9893 | 4.9893 | 4.9952 | 5.0109 | 5.0085 | 4.9574 | 4.9816 | 4.9831 | 4.9632 | 4.9627 | 4.9653 | 4.9728 | 4.9682 | 4.9743 | 4.9784 | 4.9711 | 4.9712 | 4.9891 | 4.9903 | 4.9943 | 4.9809 | 4.9842 | 4.9838 | 4.9821 | 4.9802 | 4.9862 | 4.9804 | 0.01593 | 0.0031985
675 | 4.9533 | 5.0218 | 4.9525 | 4.9714 | 4.9891 | 4.9892 | 4.9951 | 5.0108 | 5.0084 | 4.9572 | 4.9815 | 4.983 | 4.9631 | 4.9626 | 4.9652 | 4.9727 | 4.9681 | 4.9741 | 4.9783 | 4.971 | 4.9711 | 4.9889 | 4.9902 | 4.9942 | 4.9808 | 4.9841 | 4.9837 | 4.982 | 4.9801 | 4.9861 | 4.9803 | 0.015944 | 0.0032014
676 | 4.953 | 5.0217 | 4.9523 | 4.9713 | 4.989 | 4.9891 | 4.995 | 5.0107 | 5.0083 | 4.9571 | 4.9814 | 4.9829 | 4.9631 | 4.9625 | 4.9651 | 4.9726 | 4.968 | 4.974 | 4.9782 | 4.9709 | 4.9709 | 4.9888 | 4.9901 | 4.9941 | 4.9807 | 4.984 | 4.9835 | 4.9819 | 4.98 | 4.986 | 4.9802 | 0.015958 | 0.0032043
677 | 4.9527 | 5.0216 | 4.9522 | 4.9712 | 4.9889 | 4.989 | 4.9949 | 5.0105 | 5.0082 | 4.9569 | 4.9814 | 4.9829 | 4.963 | 4.9624 | 4.965 | 4.9725 | 4.9679 | 4.9739 | 4.9781 | 4.9708 | 4.9708 | 4.9887 | 4.99 | 4.994 | 4.9806 | 4.9839 | 4.9834 | 4.9817 | 4.9798 | 4.9859 | 4.9801 | 0.015972 | 0.0032072
678 | 4.9524 | 5.0215 | 4.9521 | 4.971 | 4.9888 | 4.9888 | 4.9948 | 5.0104 | 5.0081 | 4.9568 | 4.9813 | 4.9828 | 4.9629 | 4.9623 | 4.9649 | 4.9724 | 4.9678 | 4.9738 | 4.978 | 4.9707 | 4.9707 | 4.9886 | 4.9899 | 4.9939 | 4.9805 | 4.9838 | 4.9833 | 4.9816 | 4.9797 | 4.9857 | 4.98 | 0.015986 | 0.0032101
679 | 4.9521 | 5.0214 | 4.952 | 4.9709 | 4.9887 | 4.9887 | 4.9947 | 5.0103 | 5.008 | 4.9566 | 4.9812 | 4.9827 | 4.9628 | 4.9622 | 4.9648 | 4.9723 | 4.9676 | 4.9737 | 4.9779 | 4.9705 | 4.9706 | 4.9885 | 4.9898 | 4.9938 | 4.9804 | 4.9837 | 4.9832 | 4.9815 | 4.9796 | 4.9856 | 4.9799 | 0.016 | 0.003213
680 | 4.9518 | 5.0213 | 4.9519 | 4.9708 | 4.9886 | 4.9886 | 4.9945 | 5.0102 | 5.0079 | 4.9565 | 4.9812 | 4.9826 | 4.9627 | 4.9621 | 4.9647 | 4.9722 | 4.9675 | 4.9736 | 4.9778 | 4.9704 | 4.9705 | 4.9884 | 4.9897 | 4.9937 | 4.9803 | 4.9836 | 4.9831 | 4.9814 | 4.9795 | 4.9855 | 4.9798 | 0.016014 | 0.0032159
681 | 4.9515 | 5.0212 | 4.9517 | 4.9707 | 4.9885 | 4.9885 | 4.9944 | 5.0101 | 5.0078 | 4.9563 | 4.9811 | 4.9825 | 4.9626 | 4.962 | 4.9646 | 4.9721 | 4.9674 | 4.9735 | 4.9777 | 4.9703 | 4.9704 | 4.9882 | 4.9896 | 4.9936 | 4.9802 | 4.9835 | 4.9829 | 4.9813 | 4.9794 | 4.9854 | 4.9796 | 0.016028 | 0.0032188
682 | 4.9512 | 5.0211 | 4.9516 | 4.9706 | 4.9884 | 4.9884 | 4.9943 | 5.01 | 5.0076 | 4.9562 | 4.981 | 4.9825 | 4.9625 | 4.9619 | 4.9645 | 4.972 | 4.9673 | 4.9733 | 4.9777 | 4.9702 | 4.9702 | 4.9881 | 4.9894 | 4.9935 | 4.9801 | 4.9834 | 4.9828 | 4.9811 | 4.9792 | 4.9853 | 4.9795 | 0.016043 | 0.0032217
683 | 4.951 | 5.021 | 4.9515 | 4.9705 | 4.9883 | 4.9883 | 4.9942 | 5.0099 | 5.0075 | 4.956 | 4.981 | 4.9824 | 4.9624 | 4.9618 | 4.9644 | 4.9719 | 4.9672 | 4.9732 | 4.9776 | 4.9701 | 4.9701 | 4.988 | 4.9893 | 4.9934 | 4.98 | 4.9833 | 4.9827 | 4.981 | 4.9791 | 4.9852 | 4.9794 | 0.016057 | 0.0032246
684 | 4.9507 | 5.0209 | 4.9514 | 4.9704 | 4.9882 | 4.9882 | 4.9941 | 5.0098 | 5.0074 | 4.9558 | 4.9809 | 4.9823 | 4.9624 | 4.9617 | 4.9643 | 4.9718 | 4.9671 | 4.9731 | 4.9775 | 4.97 | 4.97 | 4.9879 | 4.9892 | 4.9933 | 4.9799 | 4.9832 | 4.9826 | 4.9809 | 4.979 | 4.9851 | 4.9793 | 0.016071 | 0.0032276
685 | 4.9504 | 5.0209 | 4.9512 | 4.9702 | 4.9881 | 4.9881 | 4.994 | 5.0097 | 5.0073 | 4.9557 | 4.9808 | 4.9822 | 4.9623 | 4.9615 | 4.9642 | 4.9717 | 4.967 | 4.973 | 4.9774 | 4.9699 | 4.9699 | 4.9878 | 4.9891 | 4.9932 | 4.9798 | 4.9831 | 4.9825 | 4.9808 | 4.9789 | 4.985 | 4.9792 | 0.016085 | 0.0032305
686 | 4.9501 | 5.0208 | 4.9511 | 4.9701 | 4.988 | 4.988 | 4.9939 | 5.0096 | 5.0072 | 4.9555 | 4.9808 | 4.9821 | 4.9622 | 4.9614 | 4.9642 | 4.9716 | 4.9668 | 4.9729 | 4.9773 | 4.9698 | 4.9698 | 4.9877 | 4.989 | 4.9931 | 4.9797 | 4.983 | 4.9823 | 4.9807 | 4.9788 | 4.9848 | 4.9791 | 0.016099 | 0.0032334
687 | 4.9498 | 5.0207 | 4.951 | 4.97 | 4.9879 | 4.9879 | 4.9938 | 5.0094 | 5.0071 | 4.9554 | 4.9807 | 4.9821 | 4.9621 | 4.9613 | 4.9641 | 4.9715 | 4.9667 | 4.9728 | 4.9772 | 4.9697 | 4.9696 | 4.9876 | 4.9889 | 4.993 | 4.9796 | 4.9829 | 4.9822 | 4.9805 | 4.9786 | 4.9847 | 4.979 | 0.016114 | 0.0032364
688 | 4.9495 | 5.0206 | 4.9509 | 4.9699 | 4.9878 | 4.9877 | 4.9937 | 5.0093 | 5.007 | 4.9552 | 4.9806 | 4.982 | 4.962 | 4.9612 | 4.964 | 4.9714 | 4.9666 | 4.9727 | 4.9771 | 4.9696 | 4.9695 | 4.9874 | 4.9888 | 4.9929 | 4.9795 | 4.9828 | 4.9821 | 4.9804 | 4.9785 | 4.9846 | 4.9788 | 0.016128 | 0.0032393
689 | 4.9493 | 5.0205 | 4.9508 | 4.9698 | 4.9877 | 4.9876 | 4.9936 | 5.0092 | 5.0069 | 4.9551 | 4.9806 | 4.9819 | 4.9619 | 4.9611 | 4.9639 | 4.9713 | 4.9665 | 4.9726 | 4.977 | 4.9695 | 4.9694 | 4.9873 | 4.9887 | 4.9928 | 4.9794 | 4.9827 | 4.982 | 4.9803 | 4.9784 | 4.9845 | 4.9787 | 0.016142 | 0.0032423
690 | 4.949 | 5.0204 | 4.9506 | 4.9697 | 4.9876 | 4.9875 | 4.9935 | 5.0091 | 5.0068 | 4.9549 | 4.9805 | 4.9818 | 4.9618 | 4.961 | 4.9638 | 4.9711 | 4.9664 | 4.9725 | 4.9769 | 4.9694 | 4.9693 | 4.9872 | 4.9886 | 4.9927 | 4.9793 | 4.9826 | 4.9819 | 4.9802 | 4.9783 | 4.9844 | 4.9786 | 0.016157 | 0.0032452
691 | 4.9487 | 5.0203 | 4.9505 | 4.9696 | 4.9874 | 4.9874 | 4.9934 | 5.009 | 5.0067 | 4.9548 | 4.9804 | 4.9817 | 4.9618 | 4.9609 | 4.9637 | 4.971 | 4.9663 | 4.9723 | 4.9768 | 4.9693 | 4.9692 | 4.9871 | 4.9885 | 4.9926 | 4.9792 | 4.9825 | 4.9818 | 4.9801 | 4.9782 | 4.9843 | 4.9785 | 0.016171 | 0.0032482
692 | 4.9484 | 5.0202 | 4.9504 | 4.9695 | 4.9873 | 4.9873 | 4.9933 | 5.0089 | 5.0066 | 4.9546 | 4.9804 | 4.9817 | 4.9617 | 4.9608 | 4.9636 | 4.9709 | 4.9662 | 4.9722 | 4.9767 | 4.9692 | 4.969 | 4.987 | 4.9884 | 4.9925 | 4.979 | 4.9824 | 4.9816 | 4.9799 | 4.978 | 4.9842 | 4.9784 | 0.016185 | 0.0032511
693 | 4.9481 | 5.0201 | 4.9503 | 4.9693 | 4.9872 | 4.9872 | 4.9932 | 5.0088 | 5.0065 | 4.9545 | 4.9803 | 4.9816 | 4.9616 | 4.9607 | 4.9635 | 4.9708 | 4.966 | 4.9721 | 4.9767 | 4.9691 | 4.9689 | 4.9869 | 4.9883 | 4.9924 | 4.9789 | 4.9822 | 4.9815 | 4.9798 | 4.9779 | 4.9841 | 4.9783 | 0.0162 | 0.0032541
694 | 4.9479 | 5.02 | 4.9501 | 4.9692 | 4.9871 | 4.9871 | 4.9931 | 5.0087 | 5.0064 | 4.9543 | 4.9802 | 4.9815 | 4.9615 | 4.9606 | 4.9634 | 4.9707 | 4.9659 | 4.972 | 4.9766 | 4.969 | 4.9688 | 4.9868 | 4.9882 | 4.9923 | 4.9788 | 4.9821 | 4.9814 | 4.9797 | 4.9778 | 4.9839 | 4.9782 | 0.016214 | 0.0032571
695 | 4.9476 | 5.0199 | 4.95 | 4.9691 | 4.987 | 4.987 | 4.993 | 5.0086 | 5.0063 | 4.9542 | 4.9802 | 4.9814 | 4.9614 | 4.9605 | 4.9633 | 4.9706 | 4.9658 | 4.9719 | 4.9765 | 4.9689 | 4.9687 | 4.9866 | 4.9881 | 4.9922 | 4.9787 | 4.982 | 4.9813 | 4.9796 | 4.9777 | 4.9838 | 4.9781 | 0.016229 | 0.00326
696 | 4.9473 | 5.0198 | 4.9499 | 4.969 | 4.9869 | 4.9869 | 4.9929 | 5.0085 | 5.0062 | 4.954 | 4.9801 | 4.9814 | 4.9613 | 4.9604 | 4.9632 | 4.9705 | 4.9657 | 4.9718 | 4.9764 | 4.9687 | 4.9686 | 4.9865 | 4.988 | 4.9921 | 4.9786 | 4.9819 | 4.9812 | 4.9795 | 4.9776 | 4.9837 | 4.978 | 0.016243 | 0.003263
697 | 4.947 | 5.0197 | 4.9498 | 4.9689 | 4.9868 | 4.9868 | 4.9928 | 5.0084 | 5.0061 | 4.9539 | 4.98 | 4.9813 | 4.9612 | 4.9603 | 4.9631 | 4.9704 | 4.9656 | 4.9717 | 4.9763 | 4.9686 | 4.9685 | 4.9864 | 4.9879 | 4.992 | 4.9785 | 4.9818 | 4.9811 | 4.9793 | 4.9775 | 4.9836 | 4.9778 | 0.016258 | 0.003266
698 | 4.9468 | 5.0196 | 4.9497 | 4.9688 | 4.9867 | 4.9867 | 4.9927 | 5.0083 | 5.006 | 4.9537 | 4.98 | 4.9812 | 4.9612 | 4.9602 | 4.963 | 4.9703 | 4.9655 | 4.9716 | 4.9762 | 4.9685 | 4.9683 | 4.9863 | 4.9878 | 4.9919 | 4.9784 | 4.9817 | 4.9809 | 4.9792 | 4.9773 | 4.9835 | 4.9777 | 0.016272 | 0.003269
699 | 4.9465 | 5.0195 | 4.9495 | 4.9687 | 4.9866 | 4.9866 | 4.9926 | 5.0082 | 5.0059 | 4.9536 | 4.9799 | 4.9811 | 4.9611 | 4.9601 | 4.9629 | 4.9702 | 4.9654 | 4.9715 | 4.9761 | 4.9684 | 4.9682 | 4.9862 | 4.9877 | 4.9918 | 4.9783 | 4.9816 | 4.9808 | 4.9791 | 4.9772 | 4.9834 | 4.9776 | 0.016287 | 0.003272
700 | 4.9462 | 5.0194 | 4.9494 | 4.9686 | 4.9865 | 4.9864 | 4.9925 | 5.008 | 5.0058 | 4.9534 | 4.9798 | 4.981 | 4.961 | 4.96 | 4.9628 | 4.9701 | 4.9653 | 4.9713 | 4.976 | 4.9683 | 4.9681 | 4.9861 | 4.9876 | 4.9917 | 4.9782 | 4.9815 | 4.9807 | 4.979 | 4.9771 | 4.9833 | 4.9775 | 0.016301 | 0.0032749
701 | 4.9459 | 5.0193 | 4.9493 | 4.9684 | 4.9864 | 4.9863 | 4.9924 | 5.0079 | 5.0057 | 4.9532 | 4.9798 | 4.981 | 4.9609 | 4.9599 | 4.9627 | 4.97 | 4.9651 | 4.9712 | 4.9759 | 4.9682 | 4.968 | 4.986 | 4.9875 | 4.9916 | 4.9781 | 4.9814 | 4.9806 | 4.9789 | 4.977 | 4.9832 | 4.9774 | 0.016316 | 0.0032779
702 | 4.9457 | 5.0192 | 4.9492 | 4.9683 | 4.9863 | 4.9862 | 4.9923 | 5.0078 | 5.0056 | 4.9531 | 4.9797 | 4.9809 | 4.9608 | 4.9598 | 4.9627 | 4.9699 | 4.965 | 4.9711 | 4.9758 | 4.9681 | 4.9679 | 4.9859 | 4.9874 | 4.9915 | 4.978 | 4.9813 | 4.9805 | 4.9787 | 4.9769 | 4.9831 | 4.9773 | 0.01633 | 0.0032809
703 | 4.9454 | 5.0191 | 4.9491 | 4.9682 | 4.9862 | 4.9861 | 4.9922 | 5.0077 | 5.0055 | 4.9529 | 4.9796 | 4.9808 | 4.9607 | 4.9597 | 4.9626 | 4.9698 | 4.9649 | 4.971 | 4.9758 | 4.968 | 4.9678 | 4.9857 | 4.9873 | 4.9914 | 4.9779 | 4.9812 | 4.9804 | 4.9786 | 4.9767 | 4.9829 | 4.9772 | 0.016345 | 0.0032839
704 | 4.9451 | 5.019 | 4.9489 | 4.9681 | 4.9861 | 4.986 | 4.9921 | 5.0076 | 5.0054 | 4.9528 | 4.9796 | 4.9807 | 4.9606 | 4.9596 | 4.9625 | 4.9697 | 4.9648 | 4.9709 | 4.9757 | 4.9679 | 4.9676 | 4.9856 | 4.9872 | 4.9913 | 4.9778 | 4.9811 | 4.9802 | 4.9785 | 4.9766 | 4.9828 | 4.9771 | 0.016359 | 0.0032869
705 | 4.9449 | 5.0189 | 4.9488 | 4.968 | 4.986 | 4.9859 | 4.992 | 5.0075 | 5.0053 | 4.9526 | 4.9795 | 4.9806 | 4.9606 | 4.9595 | 4.9624 | 4.9696 | 4.9647 | 4.9708 | 4.9756 | 4.9678 | 4.9675 | 4.9855 | 4.9871 | 4.9912 | 4.9777 | 4.981 | 4.9801 | 4.9784 | 4.9765 | 4.9827 | 4.977 | 0.016374 | 0.0032899
706 | 4.9446 | 5.0188 | 4.9487 | 4.9679 | 4.9859 | 4.9858 | 4.9919 | 5.0074 | 5.0052 | 4.9525 | 4.9794 | 4.9806 | 4.9605 | 4.9594 | 4.9623 | 4.9695 | 4.9646 | 4.9707 | 4.9755 | 4.9677 | 4.9674 | 4.9854 | 4.987 | 4.9911 | 4.9776 | 4.9809 | 4.98 | 4.9783 | 4.9764 | 4.9826 | 4.9769 | 0.016388 | 0.0032929
707 | 4.9443 | 5.0187 | 4.9486 | 4.9678 | 4.9858 | 4.9857 | 4.9918 | 5.0073 | 5.0051 | 4.9524 | 4.9794 | 4.9805 | 4.9604 | 4.9593 | 4.9622 | 4.9694 | 4.9645 | 4.9706 | 4.9754 | 4.9676 | 4.9673 | 4.9853 | 4.9869 | 4.991 | 4.9775 | 4.9808 | 4.9799 | 4.9782 | 4.9763 | 4.9825 | 4.9767 | 0.016403 | 0.0032959
708 | 4.9441 | 5.0186 | 4.9485 | 4.9677 | 4.9857 | 4.9856 | 4.9917 | 5.0072 | 5.0049 | 4.9522 | 4.9793 | 4.9804 | 4.9603 | 4.9592 | 4.9621 | 4.9693 | 4.9644 | 4.9705 | 4.9753 | 4.9675 | 4.9672 | 4.9852 | 4.9868 | 4.9909 | 4.9774 | 4.9807 | 4.9798 | 4.978 | 4.9762 | 4.9824 | 4.9766 | 0.016417 | 0.0032989
709 | 4.9438 | 5.0186 | 4.9483 | 4.9676 | 4.9856 | 4.9855 | 4.9916 | 5.0071 | 5.0048 | 4.9521 | 4.9792 | 4.9803 | 4.9602 | 4.9591 | 4.962 | 4.9692 | 4.9643 | 4.9704 | 4.9752 | 4.9674 | 4.9671 | 4.9851 | 4.9867 | 4.9908 | 4.9773 | 4.9806 | 4.9797 | 4.9779 | 4.976 | 4.9823 | 4.9765 | 0.016432 | 0.0033019
710 | 4.9435 | 5.0185 | 4.9482 | 4.9674 | 4.9855 | 4.9854 | 4.9915 | 5.007 | 5.0047 | 4.9519 | 4.9792 | 4.9803 | 4.9601 | 4.959 | 4.9619 | 4.9691 | 4.9641 | 4.9703 | 4.9751 | 4.9673 | 4.9669 | 4.985 | 4.9866 | 4.9907 | 4.9772 | 4.9805 | 4.9796 | 4.9778 | 4.9759 | 4.9822 | 4.9764 | 0.016447 | 0.0033049
711 | 4.9433 | 5.0184 | 4.9481 | 4.9673 | 4.9854 | 4.9853 | 4.9915 | 5.0069 | 5.0046 | 4.9518 | 4.9791 | 4.9802 | 4.9601 | 4.9589 | 4.9618 | 4.969 | 4.964 | 4.9701 | 4.9751 | 4.9672 | 4.9668 | 4.9848 | 4.9865 | 4.9906 | 4.9771 | 4.9804 | 4.9794 | 4.9777 | 4.9758 | 4.9821 | 4.9763 | 0.016461 | 0.0033079
712 | 4.943 | 5.0183 | 4.948 | 4.9672 | 4.9853 | 4.9852 | 4.9914 | 5.0068 | 5.0045 | 4.9516 | 4.9791 | 4.9801 | 4.96 | 4.9588 | 4.9617 | 4.9689 | 4.9639 | 4.97 | 4.975 | 4.9671 | 4.9667 | 4.9847 | 4.9864 | 4.9905 | 4.977 | 4.9803 | 4.9793 | 4.9776 | 4.9757 | 4.982 | 4.9762 | 0.016476 | 0.0033109
713 | 4.9427 | 5.0182 | 4.9479 | 4.9671 | 4.9852 | 4.9851 | 4.9913 | 5.0067 | 5.0044 | 4.9515 | 4.979 | 4.98 | 4.9599 | 4.9587 | 4.9616 | 4.9688 | 4.9638 | 4.9699 | 4.9749 | 4.967 | 4.9666 | 4.9846 | 4.9863 | 4.9904 | 4.9769 | 4.9802 | 4.9792 | 4.9774 | 4.9756 | 4.9819 | 4.9761 | 0.01649 | 0.0033139
714 | 4.9425 | 5.0181 | 4.9478 | 4.967 | 4.9851 | 4.985 | 4.9912 | 5.0066 | 5.0043 | 4.9513 | 4.9789 | 4.9799 | 4.9598 | 4.9586 | 4.9615 | 4.9687 | 4.9637 | 4.9698 | 4.9748 | 4.9669 | 4.9665 | 4.9845 | 4.9862 | 4.9903 | 4.9768 | 4.9801 | 4.9791 | 4.9773 | 4.9755 | 4.9817 | 4.976 | 0.016505 | 0.0033169
715 | 4.9422 | 5.018 | 4.9476 | 4.9669 | 4.985 | 4.9848 | 4.9911 | 5.0065 | 5.0042 | 4.9512 | 4.9789 | 4.9799 | 4.9597 | 4.9585 | 4.9615 | 4.9686 | 4.9636 | 4.9697 | 4.9747 | 4.9668 | 4.9664 | 4.9844 | 4.9861 | 4.9902 | 4.9767 | 4.98 | 4.979 | 4.9772 | 4.9753 | 4.9816 | 4.9759 | 0.01652 | 0.0033199
716 | 4.942 | 5.0179 | 4.9475 | 4.9668 | 4.9849 | 4.9847 | 4.991 | 5.0063 | 5.0041 | 4.951 | 4.9788 | 4.9798 | 4.9596 | 4.9584 | 4.9614 | 4.9685 | 4.9635 | 4.9696 | 4.9746 | 4.9667 | 4.9662 | 4.9843 | 4.986 | 4.9901 | 4.9766 | 4.9799 | 4.9789 | 4.9771 | 4.9752 | 4.9815 | 4.9758 | 0.016534 | 0.003323
717 | 4.9417 | 5.0178 | 4.9474 | 4.9667 | 4.9848 | 4.9846 | 4.9909 | 5.0062 | 5.004 | 4.9509 | 4.9787 | 4.9797 | 4.9596 | 4.9583 | 4.9613 | 4.9684 | 4.9634 | 4.9695 | 4.9745 | 4.9666 | 4.9661 | 4.9842 | 4.9859 | 4.9901 | 4.9765 | 4.9798 | 4.9788 | 4.977 | 4.9751 | 4.9814 | 4.9757 | 0.016549 | 0.003326
718 | 4.9414 | 5.0177 | 4.9473 | 4.9666 | 4.9847 | 4.9845 | 4.9908 | 5.0061 | 5.0039 | 4.9507 | 4.9787 | 4.9796 | 4.9595 | 4.9582 | 4.9612 | 4.9683 | 4.9633 | 4.9694 | 4.9744 | 4.9665 | 4.966 | 4.9841 | 4.9858 | 4.99 | 4.9764 | 4.9797 | 4.9786 | 4.9769 | 4.975 | 4.9813 | 4.9755 | 0.016563 | 0.003329
719 | 4.9412 | 5.0176 | 4.9472 | 4.9665 | 4.9846 | 4.9844 | 4.9907 | 5.006 | 5.0038 | 4.9506 | 4.9786 | 4.9796 | 4.9594 | 4.9581 | 4.9611 | 4.9682 | 4.9632 | 4.9693 | 4.9744 | 4.9664 | 4.9659 | 4.984 | 4.9857 | 4.9899 | 4.9763 | 4.9796 | 4.9785 | 4.9767 | 4.9749 | 4.9812 | 4.9754 | 0.016578 | 0.003332
720 | 4.9409 | 5.0175 | 4.947 | 4.9663 | 4.9845 | 4.9843 | 4.9906 | 5.0059 | 5.0037 | 4.9504 | 4.9785 | 4.9795 | 4.9593 | 4.958 | 4.961 | 4.968 | 4.963 | 4.9692 | 4.9743 | 4.9663 | 4.9658 | 4.9839 | 4.9856 | 4.9898 | 4.9762 | 4.9795 | 4.9784 | 4.9766 | 4.9748 | 4.9811 | 4.9753 | 0.016593 | 0.003335
721 | 4.9407 | 5.0174 | 4.9469 | 4.9662 | 4.9844 | 4.9842 | 4.9905 | 5.0058 | 5.0036 | 4.9503 | 4.9785 | 4.9794 | 4.9592 | 4.9579 | 4.9609 | 4.9679 | 4.9629 | 4.9691 | 4.9742 | 4.9662 | 4.9657 | 4.9837 | 4.9855 | 4.9897 | 4.9761 | 4.9794 | 4.9783 | 4.9765 | 4.9746 | 4.981 | 4.9752 | 0.016607 | 0.003338
722 | 4.9404 | 5.0173 | 4.9468 | 4.9661 | 4.9843 | 4.9841 | 4.9904 | 5.0057 | 5.0035 | 4.9501 | 4.9784 | 4.9793 | 4.9591 | 4.9578 | 4.9608 | 4.9678 | 4.9628 | 4.9689 | 4.9741 | 4.9661 | 4.9656 | 4.9836 | 4.9854 | 4.9896 | 4.976 | 4.9793 | 4.9782 | 4.9764 | 4.9745 | 4.9809 | 4.9751 | 0.016622 | 0.003341
723 | 4.9402 | 5.0172 | 4.9467 | 4.966 | 4.9842 | 4.984 | 4.9903 | 5.0056 | 5.0034 | 4.95 | 4.9784 | 4.9793 | 4.9591 | 4.9577 | 4.9607 | 4.9677 | 4.9627 | 4.9688 | 4.974 | 4.966 | 4.9654 | 4.9835 | 4.9853 | 4.9895 | 4.9759 | 4.9792 | 4.9781 | 4.9763 | 4.9744 | 4.9808 | 4.975 | 0.016637 | 0.003344
724 | 4.9399 | 5.0171 | 4.9466 | 4.9659 | 4.9841 | 4.9839 | 4.9902 | 5.0055 | 5.0033 | 4.9498 | 4.9783 | 4.9792 | 4.959 | 4.9576 | 4.9606 | 4.9676 | 4.9626 | 4.9687 | 4.9739 | 4.9659 | 4.9653 | 4.9834 | 4.9852 | 4.9894 | 4.9758 | 4.9791 | 4.978 | 4.9762 | 4.9743 | 4.9807 | 4.9749 | 0.016651 | 0.003347
725 | 4.9396 | 5.0171 | 4.9465 | 4.9658 | 4.9839 | 4.9838 | 4.9901 | 5.0054 | 5.0032 | 4.9497 | 4.9782 | 4.9791 | 4.9589 | 4.9575 | 4.9605 | 4.9675 | 4.9625 | 4.9686 | 4.9738 | 4.9658 | 4.9652 | 4.9833 | 4.9851 | 4.9893 | 4.9757 | 4.979 | 4.9778 | 4.976 | 4.9742 | 4.9806 | 4.9748 | 0.016666 | 0.00335
726 | 4.9394 | 5.017 | 4.9463 | 4.9657 | 4.9838 | 4.9837 | 4.99 | 5.0053 | 5.0031 | 4.9496 | 4.9782 | 4.979 | 4.9588 | 4.9574 | 4.9605 | 4.9674 | 4.9624 | 4.9685 | 4.9737 | 4.9657 | 4.9651 | 4.9832 | 4.985 | 4.9892 | 4.9756 | 4.9789 | 4.9777 | 4.9759 | 4.9741 | 4.9805 | 4.9747 | 0.01668 | 0.003353
727 | 4.9391 | 5.0169 | 4.9462 | 4.9656 | 4.9837 | 4.9836 | 4.9899 | 5.0052 | 5.003 | 4.9494 | 4.9781 | 4.9789 | 4.9587 | 4.9573 | 4.9604 | 4.9673 | 4.9623 | 4.9684 | 4.9737 | 4.9656 | 4.965 | 4.9831 | 4.9849 | 4.9891 | 4.9755 | 4.9788 | 4.9776 | 4.9758 | 4.9739 | 4.9803 | 4.9746 | 0.016695 | 0.003356
728 | 4.9389 | 5.0168 | 4.9461 | 4.9655 | 4.9836 | 4.9835 | 4.9898 | 5.0051 | 5.0029 | 4.9493 | 4.978 | 4.9789 | 4.9587 | 4.9572 | 4.9603 | 4.9672 | 4.9622 | 4.9683 | 4.9736 | 4.9655 | 4.9649 | 4.983 | 4.9848 | 4.989 | 4.9754 | 4.9787 | 4.9775 | 4.9757 | 4.9738 | 4.9802 | 4.9745 | 0.01671 | 0.0033591
729 | 4.9386 | 5.0167 | 4.946 | 4.9654 | 4.9835 | 4.9834 | 4.9897 | 5.005 | 5.0028 | 4.9491 | 4.978 | 4.9788 | 4.9586 | 4.9571 | 4.9602 | 4.9671 | 4.9621 | 4.9682 | 4.9735 | 4.9654 | 4.9647 | 4.9829 | 4.9847 | 4.9889 | 4.9753 | 4.9786 | 4.9774 | 4.9756 | 4.9737 | 4.9801 | 4.9744 | 0.016724 | 0.0033621
730 | 4.9384 | 5.0166 | 4.9459 | 4.9653 | 4.9834 | 4.9833 | 4.9896 | 5.0049 | 5.0027 | 4.949 | 4.9779 | 4.9787 | 4.9585 | 4.957 | 4.9601 | 4.967 | 4.962 | 4.9681 | 4.9734 | 4.9653 | 4.9646 | 4.9828 | 4.9846 | 4.9888 | 4.9752 | 4.9785 | 4.9773 | 4.9755 | 4.9736 | 4.98 | 4.9743 | 0.016739 | 0.0033651
731 | 4.9381 | 5.0165 | 4.9458 | 4.9651 | 4.9833 | 4.9832 | 4.9895 | 5.0048 | 5.0026 | 4.9488 | 4.9778 | 4.9786 | 4.9584 | 4.9569 | 4.96 | 4.9669 | 4.9619 | 4.968 | 4.9733 | 4.9652 | 4.9645 | 4.9826 | 4.9845 | 4.9887 | 4.9751 | 4.9784 | 4.9772 | 4.9753 | 4.9735 | 4.9799 | 4.9742 | 0.016753 | 0.0033681
732 | 4.9379 | 5.0164 | 4.9456 | 4.965 | 4.9832 | 4.9831 | 4.9894 | 5.0047 | 5.0025 | 4.9487 | 4.9778 | 4.9786 | 4.9583 | 4.9568 | 4.9599 | 4.9668 | 4.9617 | 4.9679 | 4.9732 | 4.9651 | 4.9644 | 4.9825 | 4.9844 | 4.9886 | 4.975 | 4.9783 | 4.977 | 4.9752 | 4.9734 | 4.9798 | 4.974 | 0.016768 | 0.0033711
733 | 4.9376 | 5.0163 | 4.9455 | 4.9649 | 4.9831 | 4.983 | 4.9893 | 5.0046 | 5.0024 | 4.9485 | 4.9777 | 4.9785 | 4.9582 | 4.9567 | 4.9598 | 4.9667 | 4.9616 | 4.9678 | 4.9731 | 4.965 | 4.9643 | 4.9824 | 4.9843 | 4.9885 | 4.9749 | 4.9782 | 4.9769 | 4.9751 | 4.9732 | 4.9797 | 4.9739 | 0.016782 | 0.0033741
734 | 4.9374 | 5.0162 | 4.9454 | 4.9648 | 4.983 | 4.9829 | 4.9892 | 5.0045 | 5.0023 | 4.9484 | 4.9777 | 4.9784 | 4.9582 | 4.9566 | 4.9597 | 4.9666 | 4.9615 | 4.9677 | 4.9731 | 4.9649 | 4.9642 | 4.9823 | 4.9842 | 4.9884 | 4.9748 | 4.9781 | 4.9768 | 4.975 | 4.9731 | 4.9796 | 4.9738 | 0.016797 | 0.0033771
735 | 4.9371 | 5.0161 | 4.9453 | 4.9647 | 4.9829 | 4.9828 | 4.9891 | 5.0044 | 5.0022 | 4.9483 | 4.9776 | 4.9783 | 4.9581 | 4.9565 | 4.9596 | 4.9665 | 4.9614 | 4.9676 | 4.973 | 4.9648 | 4.9641 | 4.9822 | 4.9841 | 4.9883 | 4.9747 | 4.978 | 4.9767 | 4.9749 | 4.973 | 4.9795 | 4.9737 | 0.016811 | 0.00338
736 | 4.9369 | 5.016 | 4.9452 | 4.9646 | 4.9828 | 4.9827 | 4.989 | 5.0043 | 5.0021 | 4.9481 | 4.9775 | 4.9783 | 4.958 | 4.9564 | 4.9596 | 4.9664 | 4.9613 | 4.9674 | 4.9729 | 4.9647 | 4.9639 | 4.9821 | 4.984 | 4.9883 | 4.9746 | 4.9779 | 4.9766 | 4.9748 | 4.9729 | 4.9794 | 4.9736 | 0.016826 | 0.003383
737 | 4.9366 | 5.0159 | 4.9451 | 4.9645 | 4.9827 | 4.9825 | 4.9889 | 5.0042 | 5.002 | 4.948 | 4.9775 | 4.9782 | 4.9579 | 4.9563 | 4.9595 | 4.9663 | 4.9612 | 4.9673 | 4.9728 | 4.9646 | 4.9638 | 4.982 | 4.9839 | 4.9882 | 4.9745 | 4.9778 | 4.9765 | 4.9747 | 4.9728 | 4.9793 | 4.9735 | 0.01684 | 0.003386
738 | 4.9364 | 5.0158 | 4.9449 | 4.9644 | 4.9826 | 4.9824 | 4.9888 | 5.0041 | 5.0019 | 4.9478 | 4.9774 | 4.9781 | 4.9578 | 4.9562 | 4.9594 | 4.9662 | 4.9611 | 4.9672 | 4.9727 | 4.9645 | 4.9637 | 4.9819 | 4.9838 | 4.9881 | 4.9744 | 4.9777 | 4.9764 | 4.9745 | 4.9727 | 4.9792 | 4.9734 | 0.016855 | 0.003389
739 | 4.9362 | 5.0158 | 4.9448 | 4.9643 | 4.9825 | 4.9823 | 4.9888 | 5.004 | 5.0018 | 4.9477 | 4.9774 | 4.978 | 4.9578 | 4.9561 | 4.9593 | 4.9661 | 4.961 | 4.9671 | 4.9726 | 4.9644 | 4.9636 | 4.9818 | 4.9837 | 4.988 | 4.9743 | 4.9776 | 4.9763 | 4.9744 | 4.9726 | 4.9791 | 4.9733 | 0.016869 | 0.003392
740 | 4.9359 | 5.0157 | 4.9447 | 4.9642 | 4.9824 | 4.9822 | 4.9887 | 5.0038 | 5.0017 | 4.9475 | 4.9773 | 4.978 | 4.9577 | 4.956 | 4.9592 | 4.966 | 4.9609 | 4.967 | 4.9725 | 4.9643 | 4.9635 | 4.9817 | 4.9836 | 4.9879 | 4.9742 | 4.9775 | 4.9761 | 4.9743 | 4.9724 | 4.979 | 4.9732 | 0.016884 | 0.003395
741 | 4.9357 | 5.0156 | 4.9446 | 4.9641 | 4.9823 | 4.9821 | 4.9886 | 5.0037 | 5.0016 | 4.9474 | 4.9772 | 4.9779 | 4.9576 | 4.9559 | 4.9591 | 4.9659 | 4.9608 | 4.9669 | 4.9725 | 4.9642 | 4.9634 | 4.9816 | 4.9835 | 4.9878 | 4.9741 | 4.9774 | 4.976 | 4.9742 | 4.9723 | 4.9789 | 4.9731 | 0.016898 | 0.003398
742 | 4.9354 | 5.0155 | 4.9445 | 4.964 | 4.9822 | 4.982 | 4.9885 | 5.0036 | 5.0015 | 4.9472 | 4.9772 | 4.9778 | 4.9575 | 4.9558 | 4.959 | 4.9658 | 4.9607 | 4.9668 | 4.9724 | 4.9641 | 4.9633 | 4.9814 | 4.9834 | 4.9877 | 4.974 | 4.9773 | 4.9759 | 4.9741 | 4.9722 | 4.9788 | 4.973 | 0.016913 | 0.003401
743 | 4.9352 | 5.0154 | 4.9444 | 4.9638 | 4.9821 | 4.9819 | 4.9884 | 5.0035 | 5.0014 | 4.9471 | 4.9771 | 4.9777 | 4.9574 | 4.9557 | 4.9589 | 4.9657 | 4.9606 | 4.9667 | 4.9723 | 4.964 | 4.9631 | 4.9813 | 4.9833 | 4.9876 | 4.9739 | 4.9772 | 4.9758 | 4.974 | 4.9721 | 4.9786 | 4.9729 | 0.016927 | 0.0034039
744 | 4.9349 | 5.0153 | 4.9442 | 4.9637 | 4.982 | 4.9818 | 4.9883 | 5.0034 | 5.0013 | 4.947 | 4.977 | 4.9777 | 4.9573 | 4.9556 | 4.9588 | 4.9656 | 4.9604 | 4.9666 | 4.9722 | 4.9639 | 4.963 | 4.9812 | 4.9832 | 4.9875 | 4.9738 | 4.9771 | 4.9757 | 4.9738 | 4.972 | 4.9785 | 4.9728 | 0.016942 | 0.0034069
745 | 4.9347 | 5.0152 | 4.9441 | 4.9636 | 4.9819 | 4.9817 | 4.9882 | 5.0033 | 5.0012 | 4.9468 | 4.977 | 4.9776 | 4.9573 | 4.9555 | 4.9588 | 4.9655 | 4.9603 | 4.9665 | 4.9721 | 4.9638 | 4.9629 | 4.9811 | 4.9831 | 4.9874 | 4.9737 | 4.977 | 4.9756 | 4.9737 | 4.9719 | 4.9784 | 4.9727 | 0.016956 | 0.0034099
746 | 4.9345 | 5.0151 | 4.944 | 4.9635 | 4.9818 | 4.9816 | 4.9881 | 5.0032 | 5.0011 | 4.9467 | 4.9769 | 4.9775 | 4.9572 | 4.9554 | 4.9587 | 4.9654 | 4.9602 | 4.9664 | 4.972 | 4.9637 | 4.9628 | 4.981 | 4.983 | 4.9873 | 4.9736 | 4.9769 | 4.9755 | 4.9736 | 4.9717 | 4.9783 | 4.9726 | 0.016971 | 0.0034129
747 | 4.9342 | 5.015 | 4.9439 | 4.9634 | 4.9817 | 4.9815 | 4.988 | 5.0031 | 5.001 | 4.9465 | 4.9769 | 4.9774 | 4.9571 | 4.9553 | 4.9586 | 4.9653 | 4.9601 | 4.9663 | 4.9719 | 4.9636 | 4.9627 | 4.9809 | 4.9829 | 4.9872 | 4.9735 | 4.9768 | 4.9754 | 4.9735 | 4.9716 | 4.9782 | 4.9725 | 0.016985 | 0.0034158
748 | 4.934 | 5.0149 | 4.9438 | 4.9633 | 4.9816 | 4.9814 | 4.9879 | 5.003 | 5.0009 | 4.9464 | 4.9768 | 4.9774 | 4.957 | 4.9552 | 4.9585 | 4.9652 | 4.96 | 4.9662 | 4.9719 | 4.9635 | 4.9626 | 4.9808 | 4.9828 | 4.9871 | 4.9734 | 4.9768 | 4.9752 | 4.9734 | 4.9715 | 4.9781 | 4.9724 | 0.016999 | 0.0034188
749 | 4.9338 | 5.0148 | 4.9437 | 4.9632 | 4.9815 | 4.9813 | 4.9878 | 5.0029 | 5.0008 | 4.9463 | 4.9767 | 4.9773 | 4.9569 | 4.9551 | 4.9584 | 4.9651 | 4.9599 | 4.9661 | 4.9718 | 4.9634 | 4.9625 | 4.9807 | 4.9827 | 4.987 | 4.9733 | 4.9767 | 4.9751 | 4.9733 | 4.9714 | 4.978 | 4.9722 | 0.017014 | 0.0034217
750 | 4.9335 | 5.0147 | 4.9435 | 4.9631 | 4.9814 | 4.9812 | 4.9877 | 5.0028 | 5.0007 | 4.9461 | 4.9767 | 4.9772 | 4.9569 | 4.955 | 4.9583 | 4.965 | 4.9598 | 4.966 | 4.9717 | 4.9633 | 4.9623 | 4.9806 | 4.9826 | 4.9869 | 4.9732 | 4.9766 | 4.975 | 4.9732 | 4.9713 | 4.9779 | 4.9721 | 0.017028 | 0.0034247
751 | 4.9333 | 5.0147 | 4.9434 | 4.963 | 4.9813 | 4.9811 | 4.9876 | 5.0027 | 5.0006 | 4.946 | 4.9766 | 4.9771 | 4.9568 | 4.9549 | 4.9582 | 4.9649 | 4.9597 | 4.9658 | 4.9716 | 4.9632 | 4.9622 | 4.9805 | 4.9825 | 4.9868 | 4.9732 | 4.9765 | 4.9749 | 4.973 | 4.9712 | 4.9778 | 4.972 | 0.017042 | 0.0034277
752 | 4.933 | 5.0146 | 4.9433 | 4.9629 | 4.9812 | 4.981 | 4.9875 | 5.0026 | 5.0005 | 4.9458 | 4.9765 | 4.9771 | 4.9567 | 4.9548 | 4.9581 | 4.9648 | 4.9596 | 4.9657 | 4.9715 | 4.9631 | 4.9621 | 4.9804 | 4.9824 | 4.9868 | 4.9731 | 4.9764 | 4.9748 | 4.9729 | 4.9711 | 4.9777 | 4.9719 | 0.017057 | 0.0034306
753 | 4.9328 | 5.0145 | 4.9432 | 4.9628 | 4.9811 | 4.9809 | 4.9874 | 5.0025 | 5.0004 | 4.9457 | 4.9765 | 4.977 | 4.9566 | 4.9547 | 4.958 | 4.9647 | 4.9595 | 4.9656 | 4.9714 | 4.963 | 4.962 | 4.9803 | 4.9823 | 4.9867 | 4.973 | 4.9763 | 4.9747 | 4.9728 | 4.9709 | 4.9776 | 4.9718 | 0.017071 | 0.0034336
754 | 4.9326 | 5.0144 | 4.9431 | 4.9627 | 4.981 | 4.9808 | 4.9873 | 5.0024 | 5.0003 | 4.9456 | 4.9764 | 4.9769 | 4.9565 | 4.9546 | 4.958 | 4.9646 | 4.9594 | 4.9655 | 4.9714 | 4.9629 | 4.9619 | 4.9802 | 4.9822 | 4.9866 | 4.9729 | 4.9762 | 4.9746 | 4.9727 | 4.9708 | 4.9775 | 4.9717 | 0.017085 | 0.0034365
755 | 4.9323 | 5.0143 | 4.943 | 4.9626 | 4.981 | 4.9807 | 4.9872 | 5.0023 | 5.0002 | 4.9454 | 4.9764 | 4.9768 | 4.9565 | 4.9545 | 4.9579 | 4.9645 | 4.9593 | 4.9654 | 4.9713 | 4.9628 | 4.9618 | 4.98 | 4.9821 | 4.9865 | 4.9728 | 4.9761 | 4.9745 | 4.9726 | 4.9707 | 4.9774 | 4.9716 | 0.0171 | 0.0034394
756 | 4.9321 | 5.0142 | 4.9428 | 4.9625 | 4.9809 | 4.9806 | 4.9871 | 5.0022 | 5.0001 | 4.9453 | 4.9763 | 4.9768 | 4.9564 | 4.9544 | 4.9578 | 4.9644 | 4.9592 | 4.9653 | 4.9712 | 4.9627 | 4.9617 | 4.9799 | 4.982 | 4.9864 | 4.9727 | 4.976 | 4.9743 | 4.9725 | 4.9706 | 4.9773 | 4.9715 | 0.017114 | 0.0034424
757 | 4.9319 | 5.0141 | 4.9427 | 4.9623 | 4.9808 | 4.9805 | 4.9871 | 5.0021 | 5 | 4.9451 | 4.9762 | 4.9767 | 4.9563 | 4.9543 | 4.9577 | 4.9643 | 4.9591 | 4.9652 | 4.9711 | 4.9626 | 4.9616 | 4.9798 | 4.9819 | 4.9863 | 4.9726 | 4.9759 | 4.9742 | 4.9724 | 4.9705 | 4.9772 | 4.9714 | 0.017128 | 0.0034453
758 | 4.9316 | 5.014 | 4.9426 | 4.9622 | 4.9807 | 4.9804 | 4.987 | 5.002 | 4.9999 | 4.945 | 4.9762 | 4.9766 | 4.9562 | 4.9542 | 4.9576 | 4.9642 | 4.959 | 4.9651 | 4.971 | 4.9625 | 4.9614 | 4.9797 | 4.9818 | 4.9862 | 4.9725 | 4.9758 | 4.9741 | 4.9722 | 4.9704 | 4.9771 | 4.9713 | 0.017142 | 0.0034482
759 | 4.9314 | 5.0139 | 4.9425 | 4.9621 | 4.9806 | 4.9803 | 4.9869 | 5.0019 | 4.9998 | 4.9449 | 4.9761 | 4.9765 | 4.9561 | 4.9541 | 4.9575 | 4.9641 | 4.9588 | 4.965 | 4.9709 | 4.9624 | 4.9613 | 4.9796 | 4.9817 | 4.9861 | 4.9724 | 4.9757 | 4.974 | 4.9721 | 4.9703 | 4.977 | 4.9712 | 0.017156 | 0.0034512
760 | 4.9312 | 5.0138 | 4.9424 | 4.962 | 4.9805 | 4.9802 | 4.9868 | 5.0018 | 4.9997 | 4.9447 | 4.9761 | 4.9765 | 4.9561 | 4.954 | 4.9574 | 4.964 | 4.9587 | 4.9649 | 4.9708 | 4.9623 | 4.9612 | 4.9795 | 4.9816 | 4.986 | 4.9723 | 4.9756 | 4.9739 | 4.972 | 4.9701 | 4.9769 | 4.9711 | 0.017171 | 0.0034541
761 | 4.931 | 5.0137 | 4.9423 | 4.9619 | 4.9804 | 4.9801 | 4.9867 | 5.0017 | 4.9996 | 4.9446 | 4.976 | 4.9764 | 4.956 | 4.9539 | 4.9573 | 4.9639 | 4.9586 | 4.9648 | 4.9708 | 4.9622 | 4.9611 | 4.9794 | 4.9815 | 4.9859 | 4.9722 | 4.9755 | 4.9738 | 4.9719 | 4.97 | 4.9768 | 4.971 | 0.017185 | 0.003457
762 | 4.9307 | 5.0136 | 4.9422 | 4.9618 | 4.9803 | 4.98 | 4.9866 | 5.0016 | 4.9995 | 4.9444 | 4.9759 | 4.9763 | 4.9559 | 4.9538 | 4.9573 | 4.9638 | 4.9585 | 4.9647 | 4.9707 | 4.9621 | 4.961 | 4.9793 | 4.9814 | 4.9858 | 4.9721 | 4.9754 | 4.9737 | 4.9718 | 4.9699 | 4.9767 | 4.9709 | 0.017199 | 0.0034599
763 | 4.9305 | 5.0136 | 4.942 | 4.9617 | 4.9802 | 4.9799 | 4.9865 | 5.0015 | 4.9994 | 4.9443 | 4.9759 | 4.9762 | 4.9558 | 4.9537 | 4.9572 | 4.9637 | 4.9584 | 4.9646 | 4.9706 | 4.962 | 4.9609 | 4.9792 | 4.9813 | 4.9857 | 4.972 | 4.9753 | 4.9736 | 4.9717 | 4.9698 | 4.9766 | 4.9708 | 0.017213 | 0.0034628
764 | 4.9303 | 5.0135 | 4.9419 | 4.9616 | 4.9801 | 4.9798 | 4.9864 | 5.0014 | 4.9993 | 4.9442 | 4.9758 | 4.9762 | 4.9557 | 4.9536 | 4.9571 | 4.9636 | 4.9583 | 4.9645 | 4.9705 | 4.9619 | 4.9608 | 4.9791 | 4.9812 | 4.9856 | 4.9719 | 4.9752 | 4.9734 | 4.9716 | 4.9697 | 4.9765 | 4.9707 | 0.017227 | 0.0034657
765 | 4.93 | 5.0134 | 4.9418 | 4.9615 | 4.98 | 4.9797 | 4.9863 | 5.0013 | 4.9993 | 4.944 | 4.9758 | 4.9761 | 4.9557 | 4.9535 | 4.957 | 4.9635 | 4.9582 | 4.9644 | 4.9704 | 4.9618 | 4.9606 | 4.979 | 4.9811 | 4.9856 | 4.9718 | 4.9751 | 4.9733 | 4.9714 | 4.9696 | 4.9763 | 4.9706 | 0.017241 | 0.0034686
766 | 4.9298 | 5.0133 | 4.9417 | 4.9614 | 4.9799 | 4.9796 | 4.9862 | 5.0012 | 4.9992 | 4.9439 | 4.9757 | 4.976 | 4.9556 | 4.9534 | 4.9569 | 4.9635 | 4.9581 | 4.9643 | 4.9703 | 4.9617 | 4.9605 | 4.9789 | 4.981 | 4.9855 | 4.9717 | 4.975 | 4.9732 | 4.9713 | 4.9695 | 4.9762 | 4.9705 | 0.017255 | 0.0034715
767 | 4.9296 | 5.0132 | 4.9416 | 4.9613 | 4.9798 | 4.9795 | 4.9861 | 5.0011 | 4.9991 | 4.9437 | 4.9756 | 4.9759 | 4.9555 | 4.9533 | 4.9568 | 4.9634 | 4.958 | 4.9642 | 4.9703 | 4.9616 | 4.9604 | 4.9788 | 4.9809 | 4.9854 | 4.9716 | 4.9749 | 4.9731 | 4.9712 | 4.9693 | 4.9761 | 4.9704 | 0.017269 | 0.0034744
768 | 4.9294 | 5.0131 | 4.9415 | 4.9612 | 4.9797 | 4.9794 | 4.986 | 5.001 | 4.999 | 4.9436 | 4.9756 | 4.9759 | 4.9554 | 4.9532 | 4.9567 | 4.9633 | 4.9579 | 4.9641 | 4.9702 | 4.9615 | 4.9603 | 4.9787 | 4.9808 | 4.9853 | 4.9715 | 4.9748 | 4.973 | 4.9711 | 4.9692 | 4.976 | 4.9703 | 0.017283 | 0.0034773
769 | 4.9291 | 5.013 | 4.9414 | 4.9611 | 4.9796 | 4.9793 | 4.9859 | 5.0009 | 4.9989 | 4.9435 | 4.9755 | 4.9758 | 4.9553 | 4.9531 | 4.9566 | 4.9632 | 4.9578 | 4.9639 | 4.9701 | 4.9614 | 4.9602 | 4.9786 | 4.9807 | 4.9852 | 4.9714 | 4.9747 | 4.9729 | 4.971 | 4.9691 | 4.9759 | 4.9702 | 0.017297 | 0.0034802
770 | 4.9289 | 5.0129 | 4.9412 | 4.961 | 4.9795 | 4.9792 | 4.9858 | 5.0008 | 4.9988 | 4.9433 | 4.9754 | 4.9757 | 4.9553 | 4.953 | 4.9565 | 4.9631 | 4.9577 | 4.9638 | 4.97 | 4.9613 | 4.9601 | 4.9784 | 4.9806 | 4.9851 | 4.9713 | 4.9746 | 4.9728 | 4.9709 | 4.969 | 4.9758 | 4.9701 | 0.017311 | 0.0034831
771 | 4.9287 | 5.0128 | 4.9411 | 4.9608 | 4.9794 | 4.9791 | 4.9857 | 5.0007 | 4.9987 | 4.9432 | 4.9754 | 4.9756 | 4.9552 | 4.9529 | 4.9565 | 4.963 | 4.9576 | 4.9637 | 4.9699 | 4.9612 | 4.96 | 4.9783 | 4.9805 | 4.985 | 4.9712 | 4.9745 | 4.9727 | 4.9708 | 4.9689 | 4.9757 | 4.97 | 0.017325 | 0.0034859
772 | 4.9285 | 5.0127 | 4.941 | 4.9607 | 4.9793 | 4.979 | 4.9856 | 5.0006 | 4.9986 | 4.9431 | 4.9753 | 4.9756 | 4.9551 | 4.9528 | 4.9564 | 4.9629 | 4.9575 | 4.9636 | 4.9698 | 4.9611 | 4.9599 | 4.9782 | 4.9804 | 4.9849 | 4.9711 | 4.9744 | 4.9725 | 4.9706 | 4.9688 | 4.9756 | 4.9699 | 0.017339 | 0.0034888
773 | 4.9282 | 5.0126 | 4.9409 | 4.9606 | 4.9792 | 4.9789 | 4.9856 | 5.0005 | 4.9985 | 4.9429 | 4.9753 | 4.9755 | 4.955 | 4.9527 | 4.9563 | 4.9628 | 4.9574 | 4.9635 | 4.9698 | 4.961 | 4.9597 | 4.9781 | 4.9803 | 4.9848 | 4.971 | 4.9743 | 4.9724 | 4.9705 | 4.9687 | 4.9755 | 4.9697 | 0.017353 | 0.0034917
774 | 4.928 | 5.0126 | 4.9408 | 4.9605 | 4.9791 | 4.9788 | 4.9855 | 5.0004 | 4.9984 | 4.9428 | 4.9752 | 4.9754 | 4.9549 | 4.9526 | 4.9562 | 4.9627 | 4.9573 | 4.9634 | 4.9697 | 4.9609 | 4.9596 | 4.978 | 4.9802 | 4.9847 | 4.9709 | 4.9742 | 4.9723 | 4.9704 | 4.9685 | 4.9754 | 4.9696 | 0.017366 | 0.0034945
775 | 4.9278 | 5.0125 | 4.9407 | 4.9604 | 4.979 | 4.9787 | 4.9854 | 5.0003 | 4.9983 | 4.9426 | 4.9751 | 4.9753 | 4.9549 | 4.9525 | 4.9561 | 4.9626 | 4.9571 | 4.9633 | 4.9696 | 4.9608 | 4.9595 | 4.9779 | 4.9801 | 4.9846 | 4.9708 | 4.9741 | 4.9722 | 4.9703 | 4.9684 | 4.9753 | 4.9695 | 0.01738 | 0.0034974
776 | 4.9276 | 5.0124 | 4.9405 | 4.9603 | 4.9789 | 4.9786 | 4.9853 | 5.0002 | 4.9982 | 4.9425 | 4.9751 | 4.9753 | 4.9548 | 4.9524 | 4.956 | 4.9625 | 4.957 | 4.9632 | 4.9695 | 4.9607 | 4.9594 | 4.9778 | 4.98 | 4.9845 | 4.9707 | 4.974 | 4.9721 | 4.9702 | 4.9683 | 4.9752 | 4.9694 | 0.017394 | 0.0035002
777 | 4.9274 | 5.0123 | 4.9404 | 4.9602 | 4.9788 | 4.9785 | 4.9852 | 5.0001 | 4.9981 | 4.9424 | 4.975 | 4.9752 | 4.9547 | 4.9523 | 4.9559 | 4.9624 | 4.9569 | 4.9631 | 4.9694 | 4.9606 | 4.9593 | 4.9777 | 4.9799 | 4.9844 | 4.9706 | 4.974 | 4.972 | 4.9701 | 4.9682 | 4.9751 | 4.9693 | 0.017408 | 0.0035031
778 | 4.9271 | 5.0122 | 4.9403 | 4.9601 | 4.9787 | 4.9783 | 4.9851 | 5 | 4.998 | 4.9422 | 4.975 | 4.9751 | 4.9546 | 4.9522 | 4.9558 | 4.9623 | 4.9568 | 4.963 | 4.9693 | 4.9605 | 4.9592 | 4.9776 | 4.9798 | 4.9844 | 4.9705 | 4.9739 | 4.9719 | 4.97 | 4.9681 | 4.975 | 4.9692 | 0.017422 | 0.0035059
779 | 4.9269 | 5.0121 | 4.9402 | 4.96 | 4.9786 | 4.9782 | 4.985 | 4.9999 | 4.9979 | 4.9421 | 4.9749 | 4.975 | 4.9545 | 4.9521 | 4.9558 | 4.9622 | 4.9567 | 4.9629 | 4.9693 | 4.9604 | 4.9591 | 4.9775 | 4.9797 | 4.9843 | 4.9704 | 4.9738 | 4.9718 | 4.9698 | 4.968 | 4.9749 | 4.9691 | 0.017435 | 0.0035087
780 | 4.9267 | 5.012 | 4.9401 | 4.9599 | 4.9785 | 4.9781 | 4.9849 | 4.9998 | 4.9978 | 4.942 | 4.9748 | 4.975 | 4.9545 | 4.952 | 4.9557 | 4.9621 | 4.9566 | 4.9628 | 4.9692 | 4.9603 | 4.9589 | 4.9774 | 4.9796 | 4.9842 | 4.9703 | 4.9737 | 4.9717 | 4.9697 | 4.9679 | 4.9748 | 4.969 | 0.017449 | 0.0035115
781 | 4.9265 | 5.0119 | 4.94 | 4.9598 | 4.9784 | 4.978 | 4.9848 | 4.9997 | 4.9977 | 4.9418 | 4.9748 | 4.9749 | 4.9544 | 4.9519 | 4.9556 | 4.962 | 4.9565 | 4.9627 | 4.9691 | 4.9602 | 4.9588 | 4.9773 | 4.9795 | 4.9841 | 4.9702 | 4.9736 | 4.9715 | 4.9696 | 4.9677 | 4.9747 | 4.9689 | 0.017463 | 0.0035144
782 | 4.9263 | 5.0118 | 4.9399 | 4.9597 | 4.9783 | 4.9779 | 4.9847 | 4.9996 | 4.9976 | 4.9417 | 4.9747 | 4.9748 | 4.9543 | 4.9518 | 4.9555 | 4.9619 | 4.9564 | 4.9626 | 4.969 | 4.9601 | 4.9587 | 4.9772 | 4.9794 | 4.984 | 4.9702 | 4.9735 | 4.9714 | 4.9695 | 4.9676 | 4.9746 | 4.9688 | 0.017476 | 0.0035172
783 | 4.9261 | 5.0117 | 4.9397 | 4.9596 | 4.9782 | 4.9778 | 4.9846 | 4.9995 | 4.9975 | 4.9415 | 4.9747 | 4.9747 | 4.9542 | 4.9517 | 4.9554 | 4.9618 | 4.9563 | 4.9625 | 4.9689 | 4.96 | 4.9586 | 4.9771 | 4.9793 | 4.9839 | 4.9701 | 4.9734 | 4.9713 | 4.9694 | 4.9675 | 4.9745 | 4.9687 | 0.01749 | 0.00352
784 | 4.9258 | 5.0116 | 4.9396 | 4.9595 | 4.9781 | 4.9777 | 4.9845 | 4.9994 | 4.9974 | 4.9414 | 4.9746 | 4.9747 | 4.9541 | 4.9516 | 4.9553 | 4.9617 | 4.9562 | 4.9624 | 4.9688 | 4.9599 | 4.9585 | 4.977 | 4.9792 | 4.9838 | 4.97 | 4.9733 | 4.9712 | 4.9693 | 4.9674 | 4.9744 | 4.9686 | 0.017503 | 0.0035228
785 | 4.9256 | 5.0116 | 4.9395 | 4.9594 | 4.978 | 4.9776 | 4.9844 | 4.9993 | 4.9973 | 4.9413 | 4.9745 | 4.9746 | 4.9541 | 4.9515 | 4.9552 | 4.9616 | 4.9561 | 4.9623 | 4.9688 | 4.9598 | 4.9584 | 4.9769 | 4.9791 | 4.9837 | 4.9699 | 4.9732 | 4.9711 | 4.9692 | 4.9673 | 4.9743 | 4.9685 | 0.017517 | 0.0035256
786 | 4.9254 | 5.0115 | 4.9394 | 4.9593 | 4.9779 | 4.9775 | 4.9844 | 4.9992 | 4.9972 | 4.9411 | 4.9745 | 4.9745 | 4.954 | 4.9514 | 4.9551 | 4.9615 | 4.956 | 4.9622 | 4.9687 | 4.9597 | 4.9583 | 4.9768 | 4.979 | 4.9836 | 4.9698 | 4.9731 | 4.971 | 4.969 | 4.9672 | 4.9742 | 4.9684 | 0.01753 | 0.0035284
787 | 4.9252 | 5.0114 | 4.9393 | 4.9591 | 4.9778 | 4.9774 | 4.9843 | 4.9991 | 4.9971 | 4.941 | 4.9744 | 4.9744 | 4.9539 | 4.9513 | 4.9551 | 4.9614 | 4.9559 | 4.9621 | 4.9686 | 4.9596 | 4.9582 | 4.9766 | 4.9789 | 4.9835 | 4.9697 | 4.973 | 4.9709 | 4.9689 | 4.9671 | 4.9741 | 4.9683 | 0.017544 | 0.0035312
788 | 4.925 | 5.0113 | 4.9392 | 4.959 | 4.9777 | 4.9773 | 4.9842 | 4.999 | 4.997 | 4.9409 | 4.9744 | 4.9744 | 4.9538 | 4.9512 | 4.955 | 4.9613 | 4.9558 | 4.9619 | 4.9685 | 4.9595 | 4.958 | 4.9765 | 4.9788 | 4.9834 | 4.9696 | 4.9729 | 4.9708 | 4.9688 | 4.9669 | 4.974 | 4.9682 | 0.017557 | 0.0035339
789 | 4.9248 | 5.0112 | 4.9391 | 4.9589 | 4.9776 | 4.9772 | 4.9841 | 4.9989 | 4.9969 | 4.9407 | 4.9743 | 4.9743 | 4.9537 | 4.9511 | 4.9549 | 4.9612 | 4.9557 | 4.9618 | 4.9684 | 4.9594 | 4.9579 | 4.9764 | 4.9787 | 4.9833 | 4.9695 | 4.9728 | 4.9707 | 4.9687 | 4.9668 | 4.9739 | 4.9681 | 0.017571 | 0.0035367
790 | 4.9246 | 5.0111 | 4.9389 | 4.9588 | 4.9775 | 4.9771 | 4.984 | 4.9988 | 4.9968 | 4.9406 | 4.9742 | 4.9742 | 4.9537 | 4.951 | 4.9548 | 4.9611 | 4.9556 | 4.9617 | 4.9683 | 4.9593 | 4.9578 | 4.9763 | 4.9786 | 4.9833 | 4.9694 | 4.9727 | 4.9705 | 4.9686 | 4.9667 | 4.9738 | 4.968 | 0.017584 | 0.0035395
791 | 4.9244 | 5.011 | 4.9388 | 4.9587 | 4.9774 | 4.977 | 4.9839 | 4.9987 | 4.9967 | 4.9405 | 4.9742 | 4.9741 | 4.9536 | 4.9509 | 4.9547 | 4.961 | 4.9555 | 4.9616 | 4.9683 | 4.9592 | 4.9577 | 4.9762 | 4.9785 | 4.9832 | 4.9693 | 4.9726 | 4.9704 | 4.9685 | 4.9666 | 4.9737 | 4.9679 | 0.017597 | 0.0035422
792 | 4.9242 | 5.0109 | 4.9387 | 4.9586 | 4.9773 | 4.9769 | 4.9838 | 4.9986 | 4.9966 | 4.9403 | 4.9741 | 4.9741 | 4.9535 | 4.9508 | 4.9546 | 4.9609 | 4.9554 | 4.9615 | 4.9682 | 4.9591 | 4.9576 | 4.9761 | 4.9784 | 4.9831 | 4.9692 | 4.9725 | 4.9703 | 4.9684 | 4.9665 | 4.9736 | 4.9678 | 0.017611 | 0.003545
793 | 4.9239 | 5.0108 | 4.9386 | 4.9585 | 4.9772 | 4.9768 | 4.9837 | 4.9985 | 4.9965 | 4.9402 | 4.9741 | 4.974 | 4.9534 | 4.9507 | 4.9545 | 4.9608 | 4.9553 | 4.9614 | 4.9681 | 4.959 | 4.9575 | 4.976 | 4.9783 | 4.983 | 4.9691 | 4.9724 | 4.9702 | 4.9682 | 4.9664 | 4.9735 | 4.9677 | 0.017624 | 0.0035477
794 | 4.9237 | 5.0107 | 4.9385 | 4.9584 | 4.9771 | 4.9767 | 4.9836 | 4.9984 | 4.9964 | 4.9401 | 4.974 | 4.9739 | 4.9533 | 4.9506 | 4.9544 | 4.9607 | 4.9551 | 4.9613 | 4.968 | 4.9589 | 4.9574 | 4.9759 | 4.9782 | 4.9829 | 4.969 | 4.9723 | 4.9701 | 4.9681 | 4.9663 | 4.9734 | 4.9676 | 0.017637 | 0.0035505
795 | 4.9235 | 5.0107 | 4.9384 | 4.9583 | 4.977 | 4.9766 | 4.9835 | 4.9983 | 4.9963 | 4.9399 | 4.9739 | 4.9738 | 4.9533 | 4.9505 | 4.9544 | 4.9606 | 4.955 | 4.9612 | 4.9679 | 4.9588 | 4.9572 | 4.9758 | 4.9781 | 4.9828 | 4.9689 | 4.9722 | 4.97 | 4.968 | 4.9661 | 4.9733 | 4.9675 | 0.017651 | 0.0035532
796 | 4.9233 | 5.0106 | 4.9383 | 4.9582 | 4.9769 | 4.9765 | 4.9834 | 4.9982 | 4.9962 | 4.9398 | 4.9739 | 4.9738 | 4.9532 | 4.9504 | 4.9543 | 4.9605 | 4.9549 | 4.9611 | 4.9678 | 4.9587 | 4.9571 | 4.9757 | 4.978 | 4.9827 | 4.9688 | 4.9721 | 4.9699 | 4.9679 | 4.966 | 4.9732 | 4.9674 | 0.017664 | 0.003556
797 | 4.9231 | 5.0105 | 4.9381 | 4.9581 | 4.9768 | 4.9764 | 4.9833 | 4.9981 | 4.9961 | 4.9397 | 4.9738 | 4.9737 | 4.9531 | 4.9503 | 4.9542 | 4.9604 | 4.9548 | 4.961 | 4.9678 | 4.9586 | 4.957 | 4.9756 | 4.9779 | 4.9826 | 4.9687 | 4.972 | 4.9698 | 4.9678 | 4.9659 | 4.9731 | 4.9673 | 0.017677 | 0.0035587
798 | 4.9229 | 5.0104 | 4.938 | 4.958 | 4.9767 | 4.9763 | 4.9832 | 4.998 | 4.996 | 4.9395 | 4.9738 | 4.9736 | 4.953 | 4.9502 | 4.9541 | 4.9603 | 4.9547 | 4.9609 | 4.9677 | 4.9585 | 4.9569 | 4.9755 | 4.9778 | 4.9825 | 4.9686 | 4.9719 | 4.9696 | 4.9677 | 4.9658 | 4.9729 | 4.9672 | 0.01769 | 0.0035614
799 | 4.9227 | 5.0103 | 4.9379 | 4.9579 | 4.9766 | 4.9762 | 4.9831 | 4.9979 | 4.9959 | 4.9394 | 4.9737 | 4.9735 | 4.9529 | 4.9501 | 4.954 | 4.9602 | 4.9546 | 4.9608 | 4.9676 | 4.9584 | 4.9568 | 4.9754 | 4.9777 | 4.9824 | 4.9685 | 4.9719 | 4.9695 | 4.9676 | 4.9657 | 4.9728 | 4.9671 | 0.017703 | 0.0035641
800 | 4.9225 | 5.0102 | 4.9378 | 4.9578 | 4.9765 | 4.9761 | 4.9831 | 4.9978 | 4.9958 | 4.9393 | 4.9736 | 4.9735 | 4.9529 | 4.95 | 4.9539 | 4.9601 | 4.9545 | 4.9607 | 4.9675 | 4.9583 | 4.9567 | 4.9753 | 4.9776 | 4.9823 | 4.9684 | 4.9718 | 4.9694 | 4.9675 | 4.9656 | 4.9727 | 4.967 | 0.017716 | 0.0035668
801 | 4.9223 | 5.0101 | 4.9377 | 4.9577 | 4.9764 | 4.976 | 4.983 | 4.9977 | 4.9957 | 4.9391 | 4.9736 | 4.9734 | 4.9528 | 4.9499 | 4.9538 | 4.96 | 4.9544 | 4.9606 | 4.9674 | 4.9582 | 4.9566 | 4.9752 | 4.9775 | 4.9823 | 4.9683 | 4.9717 | 4.9693 | 4.9673 | 4.9655 | 4.9726 | 4.9669 | 0.017729 | 0.0035695
802 | 4.9221 | 5.01 | 4.9376 | 4.9576 | 4.9763 | 4.9759 | 4.9829 | 4.9976 | 4.9956 | 4.939 | 4.9735 | 4.9733 | 4.9527 | 4.9498 | 4.9537 | 4.9599 | 4.9543 | 4.9605 | 4.9674 | 4.9581 | 4.9565 | 4.9751 | 4.9774 | 4.9822 | 4.9682 | 4.9716 | 4.9692 | 4.9672 | 4.9653 | 4.9725 | 4.9668 | 0.017742 | 0.0035722
803 | 4.9219 | 5.0099 | 4.9375 | 4.9575 | 4.9762 | 4.9758 | 4.9828 | 4.9975 | 4.9955 | 4.9389 | 4.9735 | 4.9732 | 4.9526 | 4.9497 | 4.9537 | 4.9598 | 4.9542 | 4.9604 | 4.9673 | 4.958 | 4.9563 | 4.975 | 4.9773 | 4.9821 | 4.9681 | 4.9715 | 4.9691 | 4.9671 | 4.9652 | 4.9724 | 4.9667 | 0.017755 | 0.0035749
804 | 4.9217 | 5.0098 | 4.9373 | 4.9573 | 4.9761 | 4.9757 | 4.9827 | 4.9974 | 4.9954 | 4.9387 | 4.9734 | 4.9732 | 4.9525 | 4.9496 | 4.9536 | 4.9597 | 4.9541 | 4.9603 | 4.9672 | 4.9579 | 4.9562 | 4.9749 | 4.9772 | 4.982 | 4.968 | 4.9714 | 4.969 | 4.967 | 4.9651 | 4.9723 | 4.9666 | 0.017768 | 0.0035776
805 | 4.9215 | 5.0097 | 4.9372 | 4.9572 | 4.9761 | 4.9756 | 4.9826 | 4.9973 | 4.9953 | 4.9386 | 4.9733 | 4.9731 | 4.9525 | 4.9495 | 4.9535 | 4.9596 | 4.954 | 4.9602 | 4.9671 | 4.9578 | 4.9561 | 4.9747 | 4.9771 | 4.9819 | 4.9679 | 4.9713 | 4.9689 | 4.9669 | 4.965 | 4.9722 | 4.9665 | 0.017781 | 0.0035802
806 | 4.9213 | 5.0097 | 4.9371 | 4.9571 | 4.976 | 4.9755 | 4.9825 | 4.9972 | 4.9952 | 4.9385 | 4.9733 | 4.973 | 4.9524 | 4.9494 | 4.9534 | 4.9595 | 4.9539 | 4.9601 | 4.967 | 4.9577 | 4.956 | 4.9746 | 4.977 | 4.9818 | 4.9679 | 4.9712 | 4.9688 | 4.9668 | 4.9649 | 4.9721 | 4.9664 | 0.017794 | 0.0035829
807 | 4.9211 | 5.0096 | 4.937 | 4.957 | 4.9759 | 4.9754 | 4.9824 | 4.9971 | 4.9951 | 4.9383 | 4.9732 | 4.9729 | 4.9523 | 4.9493 | 4.9533 | 4.9594 | 4.9538 | 4.96 | 4.9669 | 4.9576 | 4.9559 | 4.9745 | 4.9769 | 4.9817 | 4.9678 | 4.9711 | 4.9686 | 4.9667 | 4.9648 | 4.972 | 4.9663 | 0.017807 | 0.0035855
808 | 4.9209 | 5.0095 | 4.9369 | 4.9569 | 4.9758 | 4.9753 | 4.9823 | 4.997 | 4.995 | 4.9382 | 4.9732 | 4.9729 | 4.9522 | 4.9492 | 4.9532 | 4.9593 | 4.9537 | 4.9599 | 4.9669 | 4.9575 | 4.9558 | 4.9744 | 4.9768 | 4.9816 | 4.9677 | 4.971 | 4.9685 | 4.9665 | 4.9647 | 4.9719 | 4.9662 | 0.01782 | 0.0035882
809 | 4.9207 | 5.0094 | 4.9368 | 4.9568 | 4.9757 | 4.9752 | 4.9822 | 4.9969 | 4.9949 | 4.9381 | 4.9731 | 4.9728 | 4.9521 | 4.9491 | 4.9531 | 4.9592 | 4.9536 | 4.9597 | 4.9668 | 4.9575 | 4.9557 | 4.9743 | 4.9767 | 4.9815 | 4.9676 | 4.9709 | 4.9684 | 4.9664 | 4.9646 | 4.9718 | 4.9661 | 0.017832 | 0.0035908
810 | 4.9205 | 5.0093 | 4.9367 | 4.9567 | 4.9756 | 4.9751 | 4.9821 | 4.9968 | 4.9949 | 4.9379 | 4.973 | 4.9727 | 4.9521 | 4.949 | 4.953 | 4.9591 | 4.9535 | 4.9596 | 4.9667 | 4.9574 | 4.9556 | 4.9742 | 4.9766 | 4.9814 | 4.9675 | 4.9708 | 4.9683 | 4.9663 | 4.9644 | 4.9717 | 4.966 | 0.017845 | 0.0035935
811 | 4.9203 | 5.0092 | 4.9365 | 4.9566 | 4.9755 | 4.975 | 4.982 | 4.9967 | 4.9948 | 4.9378 | 4.973 | 4.9726 | 4.952 | 4.9489 | 4.953 | 4.959 | 4.9534 | 4.9595 | 4.9666 | 4.9573 | 4.9554 | 4.9741 | 4.9765 | 4.9813 | 4.9674 | 4.9707 | 4.9682 | 4.9662 | 4.9643 | 4.9716 | 4.9659 | 0.017858 | 0.0035961
812 | 4.9201 | 5.0091 | 4.9364 | 4.9565 | 4.9754 | 4.9749 | 4.982 | 4.9966 | 4.9947 | 4.9377 | 4.9729 | 4.9726 | 4.9519 | 4.9488 | 4.9529 | 4.9589 | 4.9533 | 4.9594 | 4.9665 | 4.9572 | 4.9553 | 4.974 | 4.9764 | 4.9813 | 4.9673 | 4.9706 | 4.9681 | 4.9661 | 4.9642 | 4.9715 | 4.9658 | 0.01787 | 0.0035987
813 | 4.9199 | 5.009 | 4.9363 | 4.9564 | 4.9753 | 4.9748 | 4.9819 | 4.9965 | 4.9946 | 4.9375 | 4.9728 | 4.9725 | 4.9518 | 4.9487 | 4.9528 | 4.9588 | 4.9532 | 4.9593 | 4.9664 | 4.9571 | 4.9552 | 4.9739 | 4.9763 | 4.9812 | 4.9672 | 4.9705 | 4.968 | 4.966 | 4.9641 | 4.9714 | 4.9656 | 0.017883 | 0.0036013
814 | 4.9197 | 5.0089 | 4.9362 | 4.9563 | 4.9752 | 4.9747 | 4.9818 | 4.9964 | 4.9945 | 4.9374 | 4.9728 | 4.9724 | 4.9517 | 4.9486 | 4.9527 | 4.9587 | 4.9531 | 4.9592 | 4.9664 | 4.957 | 4.9551 | 4.9738 | 4.9762 | 4.9811 | 4.9671 | 4.9704 | 4.9679 | 4.9659 | 4.964 | 4.9713 | 4.9655 | 0.017896 | 0.0036039
815 | 4.9195 | 5.0088 | 4.9361 | 4.9562 | 4.9751 | 4.9746 | 4.9817 | 4.9963 | 4.9944 | 4.9373 | 4.9727 | 4.9723 | 4.9517 | 4.9485 | 4.9526 | 4.9586 | 4.9529 | 4.9591 | 4.9663 | 4.9569 | 4.955 | 4.9737 | 4.9762 | 4.981 | 4.967 | 4.9703 | 4.9678 | 4.9657 | 4.9639 | 4.9712 | 4.9654 | 0.017908 | 0.0036065
816 | 4.9193 | 5.0088 | 4.936 | 4.9561 | 4.975 | 4.9745 | 4.9816 | 4.9962 | 4.9943 | 4.9371 | 4.9727 | 4.9723 | 4.9516 | 4.9484 | 4.9525 | 4.9586 | 4.9528 | 4.959 | 4.9662 | 4.9568 | 4.9549 | 4.9736 | 4.9761 | 4.9809 | 4.9669 | 4.9702 | 4.9676 | 4.9656 | 4.9638 | 4.9711 | 4.9653 | 0.017921 | 0.0036091
817 | 4.9191 | 5.0087 | 4.9359 | 4.956 | 4.9749 | 4.9744 | 4.9815 | 4.9961 | 4.9942 | 4.937 | 4.9726 | 4.9722 | 4.9515 | 4.9483 | 4.9524 | 4.9585 | 4.9527 | 4.9589 | 4.9661 | 4.9567 | 4.9548 | 4.9735 | 4.976 | 4.9808 | 4.9668 | 4.9701 | 4.9675 | 4.9655 | 4.9636 | 4.971 | 4.9652 | 0.017933 | 0.0036117
818 | 4.9189 | 5.0086 | 4.9357 | 4.9559 | 4.9748 | 4.9743 | 4.9814 | 4.996 | 4.9941 | 4.9369 | 4.9725 | 4.9721 | 4.9514 | 4.9483 | 4.9524 | 4.9584 | 4.9526 | 4.9588 | 4.966 | 4.9566 | 4.9546 | 4.9734 | 4.9759 | 4.9807 | 4.9667 | 4.9701 | 4.9674 | 4.9654 | 4.9635 | 4.9709 | 4.9651 | 0.017945 | 0.0036143
819 | 4.9187 | 5.0085 | 4.9356 | 4.9558 | 4.9747 | 4.9742 | 4.9813 | 4.9959 | 4.994 | 4.9367 | 4.9725 | 4.972 | 4.9514 | 4.9482 | 4.9523 | 4.9583 | 4.9525 | 4.9587 | 4.9659 | 4.9565 | 4.9545 | 4.9733 | 4.9758 | 4.9806 | 4.9666 | 4.97 | 4.9673 | 4.9653 | 4.9634 | 4.9708 | 4.965 | 0.017958 | 0.0036169
820 | 4.9185 | 5.0084 | 4.9355 | 4.9557 | 4.9746 | 4.9741 | 4.9812 | 4.9958 | 4.9939 | 4.9366 | 4.9724 | 4.972 | 4.9513 | 4.9481 | 4.9522 | 4.9582 | 4.9524 | 4.9586 | 4.9659 | 4.9564 | 4.9544 | 4.9732 | 4.9757 | 4.9805 | 4.9665 | 4.9699 | 4.9672 | 4.9652 | 4.9633 | 4.9707 | 4.9649 | 0.01797 | 0.0036194
821 | 4.9183 | 5.0083 | 4.9354 | 4.9555 | 4.9745 | 4.974 | 4.9811 | 4.9957 | 4.9938 | 4.9365 | 4.9724 | 4.9719 | 4.9512 | 4.948 | 4.9521 | 4.9581 | 4.9523 | 4.9585 | 4.9658 | 4.9563 | 4.9543 | 4.9731 | 4.9756 | 4.9804 | 4.9664 | 4.9698 | 4.9671 | 4.9651 | 4.9632 | 4.9706 | 4.9648 | 0.017982 | 0.003622
822 | 4.9181 | 5.0082 | 4.9353 | 4.9554 | 4.9744 | 4.9739 | 4.981 | 4.9956 | 4.9937 | 4.9364 | 4.9723 | 4.9718 | 4.9511 | 4.9479 | 4.952 | 4.958 | 4.9522 | 4.9584 | 4.9657 | 4.9562 | 4.9542 | 4.973 | 4.9755 | 4.9803 | 4.9663 | 4.9697 | 4.967 | 4.965 | 4.9631 | 4.9705 | 4.9647 | 0.017995 | 0.0036245
823 | 4.9179 | 5.0081 | 4.9352 | 4.9553 | 4.9743 | 4.9738 | 4.9809 | 4.9955 | 4.9936 | 4.9362 | 4.9722 | 4.9717 | 4.951 | 4.9478 | 4.9519 | 4.9579 | 4.9521 | 4.9583 | 4.9656 | 4.9561 | 4.9541 | 4.9729 | 4.9754 | 4.9803 | 4.9662 | 4.9696 | 4.9669 | 4.9648 | 4.963 | 4.9704 | 4.9646 | 0.018007 | 0.0036271
824 | 4.9177 | 5.008 | 4.9351 | 4.9552 | 4.9742 | 4.9737 | 4.9809 | 4.9954 | 4.9935 | 4.9361 | 4.9722 | 4.9717 | 4.951 | 4.9477 | 4.9518 | 4.9578 | 4.952 | 4.9582 | 4.9655 | 4.956 | 4.954 | 4.9728 | 4.9753 | 4.9802 | 4.9661 | 4.9695 | 4.9668 | 4.9647 | 4.9628 | 4.9703 | 4.9645 | 0.018019 | 0.0036296
825 | 4.9176 | 5.0079 | 4.9349 | 4.9551 | 4.9741 | 4.9736 | 4.9808 | 4.9953 | 4.9934 | 4.936 | 4.9721 | 4.9716 | 4.9509 | 4.9476 | 4.9517 | 4.9577 | 4.9519 | 4.9581 | 4.9655 | 4.9559 | 4.9539 | 4.9727 | 4.9752 | 4.9801 | 4.966 | 4.9694 | 4.9666 | 4.9646 | 4.9627 | 4.9702 | 4.9644 | 0.018031 | 0.0036321
826 | 4.9174 | 5.0078 | 4.9348 | 4.955 | 4.974 | 4.9735 | 4.9807 | 4.9952 | 4.9933 | 4.9358 | 4.9721 | 4.9715 | 4.9508 | 4.9475 | 4.9517 | 4.9576 | 4.9518 | 4.958 | 4.9654 | 4.9558 | 4.9537 | 4.9725 | 4.9751 | 4.98 | 4.9659 | 4.9693 | 4.9665 | 4.9645 | 4.9626 | 4.9701 | 4.9643 | 0.018043 | 0.0036346
827 | 4.9172 | 5.0078 | 4.9347 | 4.9549 | 4.9739 | 4.9734 | 4.9806 | 4.9951 | 4.9932 | 4.9357 | 4.972 | 4.9714 | 4.9507 | 4.9474 | 4.9516 | 4.9575 | 4.9517 | 4.9579 | 4.9653 | 4.9557 | 4.9536 | 4.9724 | 4.975 | 4.9799 | 4.9658 | 4.9692 | 4.9664 | 4.9644 | 4.9625 | 4.97 | 4.9642 | 0.018056 | 0.0036371
828 | 4.917 | 5.0077 | 4.9346 | 4.9548 | 4.9738 | 4.9733 | 4.9805 | 4.995 | 4.9931 | 4.9356 | 4.9719 | 4.9714 | 4.9506 | 4.9473 | 4.9515 | 4.9574 | 4.9516 | 4.9578 | 4.9652 | 4.9556 | 4.9535 | 4.9723 | 4.9749 | 4.9798 | 4.9658 | 4.9691 | 4.9663 | 4.9643 | 4.9624 | 4.9699 | 4.9641 | 0.018068 | 0.0036396
829 | 4.9168 | 5.0076 | 4.9345 | 4.9547 | 4.9737 | 4.9732 | 4.9804 | 4.9949 | 4.993 | 4.9354 | 4.9719 | 4.9713 | 4.9506 | 4.9472 | 4.9514 | 4.9573 | 4.9515 | 4.9577 | 4.9651 | 4.9555 | 4.9534 | 4.9722 | 4.9748 | 4.9797 | 4.9657 | 4.969 | 4.9662 | 4.9642 | 4.9623 | 4.9698 | 4.964 | 0.01808 | 0.0036421
830 | 4.9166 | 5.0075 | 4.9344 | 4.9546 | 4.9736 | 4.9731 | 4.9803 | 4.9948 | 4.9929 | 4.9353 | 4.9718 | 4.9712 | 4.9505 | 4.9471 | 4.9513 | 4.9572 | 4.9514 | 4.9575 | 4.965 | 4.9554 | 4.9533 | 4.9721 | 4.9747 | 4.9796 | 4.9656 | 4.9689 | 4.9661 | 4.964 | 4.9622 | 4.9697 | 4.9639 | 0.018092 | 0.0036446
831 | 4.9164 | 5.0074 | 4.9343 | 4.9545 | 4.9735 | 4.973 | 4.9802 | 4.9947 | 4.9928 | 4.9352 | 4.9718 | 4.9711 | 4.9504 | 4.947 | 4.9512 | 4.9571 | 4.9513 | 4.9574 | 4.965 | 4.9553 | 4.9532 | 4.972 | 4.9746 | 4.9795 | 4.9655 | 4.9688 | 4.966 | 4.9639 | 4.962 | 4.9696 | 4.9638 | 0.018104 | 0.0036471
832 | 4.9162 | 5.0073 | 4.9341 | 4.9544 | 4.9734 | 4.9729 | 4.9801 | 4.9946 | 4.9927 | 4.9351 | 4.9717 | 4.971 | 4.9503 | 4.9469 | 4.9511 | 4.957 | 4.9512 | 4.9573 | 4.9649 | 4.9552 | 4.9531 | 4.9719 | 4.9745 | 4.9794 | 4.9654 | 4.9687 | 4.9659 | 4.9638 | 4.9619 | 4.9695 | 4.9637 | 0.018115 | 0.0036496
833 | 4.9161 | 5.0072 | 4.934 | 4.9543 | 4.9733 | 4.9728 | 4.98 | 4.9945 | 4.9926 | 4.9349 | 4.9716 | 4.971 | 4.9502 | 4.9468 | 4.951 | 4.9569 | 4.9511 | 4.9572 | 4.9648 | 4.9551 | 4.9529 | 4.9718 | 4.9744 | 4.9793 | 4.9653 | 4.9686 | 4.9657 | 4.9637 | 4.9618 | 4.9694 | 4.9636 | 0.018127 | 0.003652
834 | 4.9159 | 5.0071 | 4.9339 | 4.9542 | 4.9732 | 4.9727 | 4.9799 | 4.9944 | 4.9925 | 4.9348 | 4.9716 | 4.9709 | 4.9502 | 4.9467 | 4.951 | 4.9568 | 4.951 | 4.9571 | 4.9647 | 4.955 | 4.9528 | 4.9717 | 4.9743 | 4.9793 | 4.9652 | 4.9685 | 4.9656 | 4.9636 | 4.9617 | 4.9693 | 4.9635 | 0.018139 | 0.0036545
835 | 4.9157 | 5.007 | 4.9338 | 4.9541 | 4.9731 | 4.9726 | 4.9798 | 4.9943 | 4.9924 | 4.9347 | 4.9715 | 4.9708 | 4.9501 | 4.9466 | 4.9509 | 4.9567 | 4.9509 | 4.957 | 4.9646 | 4.9549 | 4.9527 | 4.9716 | 4.9742 | 4.9792 | 4.9651 | 4.9684 | 4.9655 | 4.9635 | 4.9616 | 4.9692 | 4.9634 | 0.018151 | 0.0036569
836 | 4.9155 | 5.0069 | 4.9337 | 4.954 | 4.973 | 4.9725 | 4.9797 | 4.9942 | 4.9923 | 4.9345 | 4.9715 | 4.9707 | 4.95 | 4.9465 | 4.9508 | 4.9566 | 4.9508 | 4.9569 | 4.9645 | 4.9548 | 4.9526 | 4.9715 | 4.9741 | 4.9791 | 4.965 | 4.9683 | 4.9654 | 4.9634 | 4.9615 | 4.9691 | 4.9633 | 0.018163 | 0.0036594
837 | 4.9153 | 5.0068 | 4.9336 | 4.9539 | 4.9729 | 4.9724 | 4.9797 | 4.9941 | 4.9922 | 4.9344 | 4.9714 | 4.9707 | 4.9499 | 4.9464 | 4.9507 | 4.9565 | 4.9507 | 4.9568 | 4.9645 | 4.9547 | 4.9525 | 4.9714 | 4.974 | 4.979 | 4.9649 | 4.9683 | 4.9653 | 4.9632 | 4.9613 | 4.969 | 4.9632 | 0.018174 | 0.0036618
838 | 4.9151 | 5.0068 | 4.9335 | 4.9537 | 4.9728 | 4.9723 | 4.9796 | 4.994 | 4.9921 | 4.9343 | 4.9713 | 4.9706 | 4.9498 | 4.9463 | 4.9506 | 4.9564 | 4.9505 | 4.9567 | 4.9644 | 4.9546 | 4.9524 | 4.9713 | 4.9739 | 4.9789 | 4.9648 | 4.9682 | 4.9652 | 4.9631 | 4.9612 | 4.9689 | 4.9631 | 0.018186 | 0.0036642
839 | 4.9149 | 5.0067 | 4.9333 | 4.9536 | 4.9727 | 4.9722 | 4.9795 | 4.9939 | 4.992 | 4.9341 | 4.9713 | 4.9705 | 4.9498 | 4.9462 | 4.9505 | 4.9563 | 4.9504 | 4.9566 | 4.9643 | 4.9545 | 4.9523 | 4.9712 | 4.9738 | 4.9788 | 4.9647 | 4.9681 | 4.9651 | 4.963 | 4.9611 | 4.9688 | 4.963 | 0.018197 | 0.0036666
840 | 4.9148 | 5.0066 | 4.9332 | 4.9535 | 4.9726 | 4.9721 | 4.9794 | 4.9938 | 4.9919 | 4.934 | 4.9712 | 4.9704 | 4.9497 | 4.9461 | 4.9504 | 4.9562 | 4.9503 | 4.9565 | 4.9642 | 4.9544 | 4.9522 | 4.9711 | 4.9737 | 4.9787 | 4.9646 | 4.968 | 4.965 | 4.9629 | 4.961 | 4.9687 | 4.9629 | 0.018209 | 0.003669
841 | 4.9146 | 5.0065 | 4.9331 | 4.9534 | 4.9725 | 4.972 | 4.9793 | 4.9937 | 4.9918 | 4.9339 | 4.9712 | 4.9704 | 4.9496 | 4.946 | 4.9503 | 4.9561 | 4.9502 | 4.9564 | 4.9641 | 4.9543 | 4.952 | 4.971 | 4.9736 | 4.9786 | 4.9645 | 4.9679 | 4.9649 | 4.9628 | 4.9609 | 4.9686 | 4.9628 | 0.018221 | 0.0036714
842 | 4.9144 | 5.0064 | 4.933 | 4.9533 | 4.9725 | 4.9719 | 4.9792 | 4.9936 | 4.9917 | 4.9338 | 4.9711 | 4.9703 | 4.9495 | 4.9459 | 4.9503 | 4.956 | 4.9501 | 4.9563 | 4.964 | 4.9542 | 4.9519 | 4.9709 | 4.9735 | 4.9785 | 4.9644 | 4.9678 | 4.9647 | 4.9627 | 4.9608 | 4.9685 | 4.9627 | 0.018232 | 0.0036738
843 | 4.9142 | 5.0063 | 4.9329 | 4.9532 | 4.9724 | 4.9718 | 4.9791 | 4.9935 | 4.9916 | 4.9336 | 4.971 | 4.9702 | 4.9494 | 4.9458 | 4.9502 | 4.9559 | 4.95 | 4.9562 | 4.964 | 4.9541 | 4.9518 | 4.9708 | 4.9734 | 4.9784 | 4.9643 | 4.9677 | 4.9646 | 4.9626 | 4.9607 | 4.9684 | 4.9626 | 0.018243 | 0.0036762
844 | 4.914 | 5.0062 | 4.9328 | 4.9531 | 4.9723 | 4.9717 | 4.979 | 4.9934 | 4.9915 | 4.9335 | 4.971 | 4.9701 | 4.9494 | 4.9457 | 4.9501 | 4.9558 | 4.9499 | 4.9561 | 4.9639 | 4.954 | 4.9517 | 4.9707 | 4.9733 | 4.9783 | 4.9642 | 4.9676 | 4.9645 | 4.9624 | 4.9605 | 4.9683 | 4.9625 | 0.018255 | 0.0036786
845 | 4.9139 | 5.0061 | 4.9327 | 4.953 | 4.9722 | 4.9716 | 4.9789 | 4.9933 | 4.9914 | 4.9334 | 4.9709 | 4.9701 | 4.9493 | 4.9456 | 4.95 | 4.9557 | 4.9498 | 4.956 | 4.9638 | 4.9539 | 4.9516 | 4.9706 | 4.9732 | 4.9783 | 4.9641 | 4.9675 | 4.9644 | 4.9623 | 4.9604 | 4.9682 | 4.9624 | 0.018266 | 0.0036809
846 | 4.9137 | 5.006 | 4.9325 | 4.9529 | 4.9721 | 4.9715 | 4.9788 | 4.9932 | 4.9913 | 4.9333 | 4.9709 | 4.97 | 4.9492 | 4.9455 | 4.9499 | 4.9556 | 4.9497 | 4.9559 | 4.9637 | 4.9538 | 4.9515 | 4.9705 | 4.9731 | 4.9782 | 4.964 | 4.9674 | 4.9643 | 4.9622 | 4.9603 | 4.9681 | 4.9623 | 0.018278 | 0.0036833
847 | 4.9135 | 5.0059 | 4.9324 | 4.9528 | 4.972 | 4.9714 | 4.9787 | 4.9931 | 4.9912 | 4.9331 | 4.9708 | 4.9699 | 4.9491 | 4.9454 | 4.9498 | 4.9555 | 4.9496 | 4.9558 | 4.9636 | 4.9537 | 4.9514 | 4.9703 | 4.973 | 4.9781 | 4.9639 | 4.9673 | 4.9642 | 4.9621 | 4.9602 | 4.968 | 4.9622 | 0.018289 | 0.0036856
848 | 4.9133 | 5.0058 | 4.9323 | 4.9527 | 4.9719 | 4.9713 | 4.9786 | 4.993 | 4.9911 | 4.933 | 4.9707 | 4.9698 | 4.949 | 4.9453 | 4.9497 | 4.9554 | 4.9495 | 4.9557 | 4.9636 | 4.9536 | 4.9512 | 4.9702 | 4.9729 | 4.978 | 4.9638 | 4.9672 | 4.9641 | 4.962 | 4.9601 | 4.9679 | 4.9621 | 0.0183 | 0.003688
849 | 4.9132 | 5.0058 | 4.9322 | 4.9526 | 4.9718 | 4.9712 | 4.9786 | 4.9929 | 4.991 | 4.9329 | 4.9707 | 4.9698 | 4.949 | 4.9452 | 4.9496 | 4.9553 | 4.9494 | 4.9556 | 4.9635 | 4.9535 | 4.9511 | 4.9701 | 4.9728 | 4.9779 | 4.9637 | 4.9671 | 4.964 | 4.9619 | 4.96 | 4.9678 | 4.962 | 0.018311 | 0.0036903
850 | 4.913 | 5.0057 | 4.9321 | 4.9525 | 4.9717 | 4.9711 | 4.9785 | 4.9928 | 4.9909 | 4.9327 | 4.9706 | 4.9697 | 4.9489 | 4.9451 | 4.9496 | 4.9552 | 4.9493 | 4.9554 | 4.9634 | 4.9535 | 4.951 | 4.97 | 4.9727 | 4.9778 | 4.9637 | 4.967 | 4.9638 | 4.9618 | 4.9599 | 4.9677 | 4.9619 | 0.018322 | 0.0036926
851 | 4.9128 | 5.0056 | 4.932 | 4.9524 | 4.9716 | 4.971 | 4.9784 | 4.9927 | 4.9909 | 4.9326 | 4.9705 | 4.9696 | 4.9488 | 4.945 | 4.9495 | 4.9551 | 4.9492 | 4.9553 | 4.9633 | 4.9534 | 4.9509 | 4.9699 | 4.9726 | 4.9777 | 4.9636 | 4.9669 | 4.9637 | 4.9616 | 4.9597 | 4.9676 | 4.9618 | 0.018334 | 0.0036949
852 | 4.9126 | 5.0055 | 4.9319 | 4.9523 | 4.9715 | 4.9709 | 4.9783 | 4.9926 | 4.9908 | 4.9325 | 4.9705 | 4.9695 | 4.9487 | 4.9449 | 4.9494 | 4.955 | 4.9491 | 4.9552 | 4.9632 | 4.9533 | 4.9508 | 4.9698 | 4.9725 | 4.9776 | 4.9635 | 4.9668 | 4.9636 | 4.9615 | 4.9596 | 4.9675 | 4.9617 | 0.018345 | 0.0036973
853 | 4.9124 | 5.0054 | 4.9317 | 4.9522 | 4.9714 | 4.9708 | 4.9782 | 4.9925 | 4.9907 | 4.9324 | 4.9704 | 4.9695 | 4.9486 | 4.9448 | 4.9493 | 4.9549 | 4.949 | 4.9551 | 4.9631 | 4.9532 | 4.9507 | 4.9697 | 4.9724 | 4.9775 | 4.9634 | 4.9667 | 4.9635 | 4.9614 | 4.9595 | 4.9674 | 4.9616 | 0.018356 | 0.0036996
854 | 4.9123 | 5.0053 | 4.9316 | 4.9521 | 4.9713 | 4.9707 | 4.9781 | 4.9924 | 4.9906 | 4.9322 | 4.9704 | 4.9694 | 4.9486 | 4.9447 | 4.9492 | 4.9548 | 4.9489 | 4.955 | 4.9631 | 4.9531 | 4.9506 | 4.9696 | 4.9723 | 4.9774 | 4.9633 | 4.9666 | 4.9634 | 4.9613 | 4.9594 | 4.9673 | 4.9615 | 0.018367 | 0.0037018
855 | 4.9121 | 5.0052 | 4.9315 | 4.9519 | 4.9712 | 4.9706 | 4.978 | 4.9923 | 4.9905 | 4.9321 | 4.9703 | 4.9693 | 4.9485 | 4.9446 | 4.9491 | 4.9547 | 4.9488 | 4.9549 | 4.963 | 4.953 | 4.9504 | 4.9695 | 4.9722 | 4.9773 | 4.9632 | 4.9665 | 4.9633 | 4.9612 | 4.9593 | 4.9672 | 4.9614 | 0.018378 | 0.0037041
856 | 4.9119 | 5.0051 | 4.9314 | 4.9518 | 4.9711 | 4.9705 | 4.9779 | 4.9922 | 4.9904 | 4.932 | 4.9702 | 4.9692 | 4.9484 | 4.9445 | 4.949 | 4.9546 | 4.9487 | 4.9548 | 4.9629 | 4.9529 | 4.9503 | 4.9694 | 4.9721 | 4.9773 | 4.9631 | 4.9665 | 4.9632 | 4.9611 | 4.9592 | 4.967 | 4.9613 | 0.018388 | 0.0037064
857 | 4.9118 | 5.005 | 4.9313 | 4.9517 | 4.971 | 4.9704 | 4.9778 | 4.9921 | 4.9903 | 4.9319 | 4.9702 | 4.9691 | 4.9483 | 4.9444 | 4.9489 | 4.9545 | 4.9486 | 4.9547 | 4.9628 | 4.9528 | 4.9502 | 4.9693 | 4.972 | 4.9772 | 4.963 | 4.9664 | 4.9631 | 4.961 | 4.959 | 4.9669 | 4.9612 | 0.018399 | 0.0037087
858 | 4.9116 | 5.0049 | 4.9312 | 4.9516 | 4.9709 | 4.9703 | 4.9777 | 4.992 | 4.9902 | 4.9317 | 4.9701 | 4.9691 | 4.9482 | 4.9443 | 4.9489 | 4.9544 | 4.9485 | 4.9546 | 4.9627 | 4.9527 | 4.9501 | 4.9692 | 4.9719 | 4.9771 | 4.9629 | 4.9663 | 4.9629 | 4.9609 | 4.9589 | 4.9668 | 4.9611 | 0.01841 | 0.0037109
859 | 4.9114 | 5.0048 | 4.9311 | 4.9515 | 4.9708 | 4.9702 | 4.9776 | 4.9919 | 4.9901 | 4.9316 | 4.9701 | 4.969 | 4.9482 | 4.9442 | 4.9488 | 4.9543 | 4.9484 | 4.9545 | 4.9626 | 4.9526 | 4.95 | 4.9691 | 4.9718 | 4.977 | 4.9628 | 4.9662 | 4.9628 | 4.9607 | 4.9588 | 4.9667 | 4.961 | 0.018421 | 0.0037132
860 | 4.9112 | 5.0047 | 4.9309 | 4.9514 | 4.9707 | 4.9701 | 4.9775 | 4.9918 | 4.99 | 4.9315 | 4.97 | 4.9689 | 4.9481 | 4.9441 | 4.9487 | 4.9543 | 4.9482 | 4.9544 | 4.9626 | 4.9525 | 4.9499 | 4.969 | 4.9717 | 4.9769 | 4.9627 | 4.9661 | 4.9627 | 4.9606 | 4.9587 | 4.9666 | 4.9609 | 0.018432 | 0.0037154
861 | 4.9111 | 5.0047 | 4.9308 | 4.9513 | 4.9706 | 4.97 | 4.9775 | 4.9917 | 4.9899 | 4.9313 | 4.9699 | 4.9688 | 4.948 | 4.944 | 4.9486 | 4.9542 | 4.9481 | 4.9543 | 4.9625 | 4.9524 | 4.9498 | 4.9689 | 4.9716 | 4.9768 | 4.9626 | 4.966 | 4.9626 | 4.9605 | 4.9586 | 4.9665 | 4.9608 | 0.018442 | 0.0037176
862 | 4.9109 | 5.0046 | 4.9307 | 4.9512 | 4.9705 | 4.9699 | 4.9774 | 4.9916 | 4.9898 | 4.9312 | 4.9699 | 4.9688 | 4.9479 | 4.9439 | 4.9485 | 4.9541 | 4.948 | 4.9542 | 4.9624 | 4.9523 | 4.9496 | 4.9688 | 4.9715 | 4.9767 | 4.9625 | 4.9659 | 4.9625 | 4.9604 | 4.9585 | 4.9664 | 4.9607 | 0.018453 | 0.0037199
863 | 4.9107 | 5.0045 | 4.9306 | 4.9511 | 4.9704 | 4.9698 | 4.9773 | 4.9915 | 4.9897 | 4.9311 | 4.9698 | 4.9687 | 4.9478 | 4.9438 | 4.9484 | 4.954 | 4.9479 | 4.9541 | 4.9623 | 4.9522 | 4.9495 | 4.9687 | 4.9714 | 4.9766 | 4.9624 | 4.9658 | 4.9624 | 4.9603 | 4.9584 | 4.9663 | 4.9606 | 0.018464 | 0.0037221
864 | 4.9106 | 5.0044 | 4.9305 | 4.951 | 4.9703 | 4.9697 | 4.9772 | 4.9914 | 4.9896 | 4.931 | 4.9698 | 4.9686 | 4.9478 | 4.9437 | 4.9483 | 4.9539 | 4.9478 | 4.954 | 4.9622 | 4.9521 | 4.9494 | 4.9686 | 4.9713 | 4.9765 | 4.9623 | 4.9657 | 4.9623 | 4.9602 | 4.9582 | 4.9662 | 4.9605 | 0.018474 | 0.0037243
865 | 4.9104 | 5.0043 | 4.9304 | 4.9509 | 4.9702 | 4.9696 | 4.9771 | 4.9913 | 4.9895 | 4.9308 | 4.9697 | 4.9685 | 4.9477 | 4.9436 | 4.9482 | 4.9538 | 4.9477 | 4.9539 | 4.9621 | 4.952 | 4.9493 | 4.9685 | 4.9712 | 4.9764 | 4.9622 | 4.9656 | 4.9622 | 4.96 | 4.9581 | 4.9661 | 4.9604 | 0.018485 | 0.0037265
866 | 4.9102 | 5.0042 | 4.9303 | 4.9508 | 4.9701 | 4.9695 | 4.977 | 4.9912 | 4.9894 | 4.9307 | 4.9696 | 4.9685 | 4.9476 | 4.9435 | 4.9482 | 4.9537 | 4.9476 | 4.9538 | 4.9621 | 4.9519 | 4.9492 | 4.9684 | 4.9711 | 4.9763 | 4.9621 | 4.9655 | 4.962 | 4.9599 | 4.958 | 4.966 | 4.9603 | 0.018495 | 0.0037287
867 | 4.9101 | 5.0041 | 4.9301 | 4.9507 | 4.97 | 4.9694 | 4.9769 | 4.9911 | 4.9893 | 4.9306 | 4.9696 | 4.9684 | 4.9475 | 4.9434 | 4.9481 | 4.9536 | 4.9475 | 4.9537 | 4.962 | 4.9518 | 4.9491 | 4.9683 | 4.971 | 4.9762 | 4.962 | 4.9654 | 4.9619 | 4.9598 | 4.9579 | 4.9659 | 4.9602 | 0.018506 | 0.0037308
868 | 4.9099 | 5.004 | 4.93 | 4.9506 | 4.9699 | 4.9693 | 4.9768 | 4.991 | 4.9892 | 4.9305 | 4.9695 | 4.9683 | 4.9474 | 4.9433 | 4.948 | 4.9535 | 4.9474 | 4.9535 | 4.9619 | 4.9517 | 4.949 | 4.9681 | 4.9709 | 4.9762 | 4.9619 | 4.9653 | 4.9618 | 4.9597 | 4.9578 | 4.9658 | 4.9601 | 0.018516 | 0.003733
869 | 4.9097 | 5.0039 | 4.9299 | 4.9505 | 4.9698 | 4.9692 | 4.9767 | 4.9909 | 4.9891 | 4.9303 | 4.9694 | 4.9682 | 4.9473 | 4.9432 | 4.9479 | 4.9534 | 4.9473 | 4.9534 | 4.9618 | 4.9516 | 4.9488 | 4.968 | 4.9708 | 4.9761 | 4.9618 | 4.9652 | 4.9617 | 4.9596 | 4.9577 | 4.9657 | 4.96 | 0.018526 | 0.0037352
870 | 4.9096 | 5.0038 | 4.9298 | 4.9504 | 4.9697 | 4.9691 | 4.9766 | 4.9908 | 4.989 | 4.9302 | 4.9694 | 4.9682 | 4.9473 | 4.9431 | 4.9478 | 4.9533 | 4.9472 | 4.9533 | 4.9617 | 4.9515 | 4.9487 | 4.9679 | 4.9707 | 4.976 | 4.9617 | 4.9651 | 4.9616 | 4.9595 | 4.9575 | 4.9656 | 4.9599 | 0.018537 | 0.0037373
871 | 4.9094 | 5.0037 | 4.9297 | 4.9502 | 4.9696 | 4.969 | 4.9765 | 4.9907 | 4.9889 | 4.9301 | 4.9693 | 4.9681 | 4.9472 | 4.943 | 4.9477 | 4.9532 | 4.9471 | 4.9532 | 4.9617 | 4.9514 | 4.9486 | 4.9678 | 4.9706 | 4.9759 | 4.9616 | 4.965 | 4.9615 | 4.9594 | 4.9574 | 4.9655 | 4.9598 | 0.018547 | 0.0037395
872 | 4.9092 | 5.0036 | 4.9296 | 4.9501 | 4.9695 | 4.9689 | 4.9764 | 4.9906 | 4.9888 | 4.93 | 4.9693 | 4.968 | 4.9471 | 4.9429 | 4.9476 | 4.9531 | 4.947 | 4.9531 | 4.9616 | 4.9513 | 4.9485 | 4.9677 | 4.9705 | 4.9758 | 4.9615 | 4.9649 | 4.9614 | 4.9592 | 4.9573 | 4.9654 | 4.9597 | 0.018557 | 0.0037416
873 | 4.9091 | 5.0036 | 4.9294 | 4.95 | 4.9694 | 4.9688 | 4.9763 | 4.9905 | 4.9887 | 4.9298 | 4.9692 | 4.9679 | 4.947 | 4.9428 | 4.9475 | 4.953 | 4.9469 | 4.953 | 4.9615 | 4.9512 | 4.9484 | 4.9676 | 4.9704 | 4.9757 | 4.9615 | 4.9648 | 4.9613 | 4.9591 | 4.9572 | 4.9653 | 4.9596 | 0.018567 | 0.0037437
874 | 4.9089 | 5.0035 | 4.9293 | 4.9499 | 4.9693 | 4.9686 | 4.9763 | 4.9904 | 4.9886 | 4.9297 | 4.9691 | 4.9678 | 4.9469 | 4.9427 | 4.9474 | 4.9529 | 4.9468 | 4.9529 | 4.9614 | 4.9511 | 4.9483 | 4.9675 | 4.9703 | 4.9756 | 4.9614 | 4.9647 | 4.9611 | 4.959 | 4.9571 | 4.9652 | 4.9595 | 0.018577 | 0.0037458
875 | 4.9087 | 5.0034 | 4.9292 | 4.9498 | 4.9692 | 4.9685 | 4.9762 | 4.9903 | 4.9885 | 4.9296 | 4.9691 | 4.9678 | 4.9469 | 4.9425 | 4.9474 | 4.9528 | 4.9467 | 4.9528 | 4.9613 | 4.951 | 4.9482 | 4.9674 | 4.9702 | 4.9755 | 4.9613 | 4.9647 | 4.961 | 4.9589 | 4.957 | 4.9651 | 4.9594 | 0.018587 | 0.0037479
876 | 4.9086 | 5.0033 | 4.9291 | 4.9497 | 4.9691 | 4.9684 | 4.9761 | 4.9902 | 4.9884 | 4.9295 | 4.969 | 4.9677 | 4.9468 | 4.9424 | 4.9473 | 4.9527 | 4.9466 | 4.9527 | 4.9612 | 4.9509 | 4.948 | 4.9673 | 4.9701 | 4.9754 | 4.9612 | 4.9646 | 4.9609 | 4.9588 | 4.9568 | 4.965 | 4.9593 | 0.018597 | 0.00375
877 | 4.9084 | 5.0032 | 4.929 | 4.9496 | 4.969 | 4.9683 | 4.976 | 4.9901 | 4.9883 | 4.9293 | 4.969 | 4.9676 | 4.9467 | 4.9423 | 4.9472 | 4.9526 | 4.9465 | 4.9526 | 4.9612 | 4.9508 | 4.9479 | 4.9672 | 4.97 | 4.9753 | 4.9611 | 4.9645 | 4.9608 | 4.9587 | 4.9567 | 4.9649 | 4.9592 | 0.018607 | 0.0037521
878 | 4.9082 | 5.0031 | 4.9289 | 4.9495 | 4.9689 | 4.9682 | 4.9759 | 4.99 | 4.9882 | 4.9292 | 4.9689 | 4.9675 | 4.9466 | 4.9422 | 4.9471 | 4.9525 | 4.9464 | 4.9525 | 4.9611 | 4.9507 | 4.9478 | 4.9671 | 4.9699 | 4.9752 | 4.961 | 4.9644 | 4.9607 | 4.9586 | 4.9566 | 4.9648 | 4.9591 | 0.018617 | 0.0037542
879 | 4.9081 | 5.003 | 4.9288 | 4.9494 | 4.9688 | 4.9681 | 4.9758 | 4.9899 | 4.9881 | 4.9291 | 4.9688 | 4.9675 | 4.9465 | 4.9421 | 4.947 | 4.9524 | 4.9463 | 4.9524 | 4.961 | 4.9506 | 4.9477 | 4.967 | 4.9698 | 4.9752 | 4.9609 | 4.9643 | 4.9606 | 4.9584 | 4.9565 | 4.9647 | 4.959 | 0.018627 | 0.0037563
880 | 4.9079 | 5.0029 | 4.9286 | 4.9493 | 4.9687 | 4.968 | 4.9757 | 4.9898 | 4.988 | 4.929 | 4.9688 | 4.9674 | 4.9465 | 4.942 | 4.9469 | 4.9523 | 4.9462 | 4.9523 | 4.9609 | 4.9505 | 4.9476 | 4.9669 | 4.9697 | 4.9751 | 4.9608 | 4.9642 | 4.9605 | 4.9583 | 4.9564 | 4.9646 | 4.9589 | 0.018637 | 0.0037584
881 | 4.9078 | 5.0028 | 4.9285 | 4.9492 | 4.9686 | 4.9679 | 4.9756 | 4.9897 | 4.9879 | 4.9288 | 4.9687 | 4.9673 | 4.9464 | 4.9419 | 4.9468 | 4.9522 | 4.946 | 4.9522 | 4.9608 | 4.9504 | 4.9475 | 4.9668 | 4.9696 | 4.975 | 4.9607 | 4.9641 | 4.9604 | 4.9582 | 4.9563 | 4.9645 | 4.9588 | 0.018647 | 0.0037604
882 | 4.9076 | 5.0027 | 4.9284 | 4.9491 | 4.9685 | 4.9678 | 4.9755 | 4.9896 | 4.9878 | 4.9287 | 4.9686 | 4.9672 | 4.9463 | 4.9418 | 4.9467 | 4.9521 | 4.9459 | 4.9521 | 4.9607 | 4.9503 | 4.9473 | 4.9667 | 4.9694 | 4.9749 | 4.9606 | 4.964 | 4.9602 | 4.9581 | 4.9562 | 4.9644 | 4.9587 | 0.018657 | 0.0037625
883 | 4.9074 | 5.0026 | 4.9283 | 4.949 | 4.9684 | 4.9677 | 4.9754 | 4.9895 | 4.9877 | 4.9286 | 4.9686 | 4.9671 | 4.9462 | 4.9417 | 4.9467 | 4.952 | 4.9458 | 4.952 | 4.9607 | 4.9502 | 4.9472 | 4.9666 | 4.9693 | 4.9748 | 4.9605 | 4.9639 | 4.9601 | 4.958 | 4.956 | 4.9643 | 4.9586 | 0.018666 | 0.0037645
884 | 4.9073 | 5.0025 | 4.9282 | 4.9489 | 4.9684 | 4.9676 | 4.9753 | 4.9894 | 4.9876 | 4.9285 | 4.9685 | 4.9671 | 4.9461 | 4.9416 | 4.9466 | 4.9519 | 4.9457 | 4.9519 | 4.9606 | 4.9501 | 4.9471 | 4.9665 | 4.9692 | 4.9747 | 4.9604 | 4.9638 | 4.96 | 4.9579 | 4.9559 | 4.9642 | 4.9584 | 0.018676 | 0.0037665
885 | 4.9071 | 5.0025 | 4.9281 | 4.9488 | 4.9683 | 4.9675 | 4.9752 | 4.9893 | 4.9875 | 4.9283 | 4.9685 | 4.967 | 4.9461 | 4.9415 | 4.9465 | 4.9518 | 4.9456 | 4.9517 | 4.9605 | 4.95 | 4.947 | 4.9664 | 4.9691 | 4.9746 | 4.9603 | 4.9637 | 4.9599 | 4.9578 | 4.9558 | 4.9641 | 4.9583 | 0.018686 | 0.0037686
886 | 4.907 | 5.0024 | 4.9279 | 4.9487 | 4.9682 | 4.9674 | 4.9751 | 4.9892 | 4.9874 | 4.9282 | 4.9684 | 4.9669 | 4.946 | 4.9414 | 4.9464 | 4.9517 | 4.9455 | 4.9516 | 4.9604 | 4.9499 | 4.9469 | 4.9663 | 4.969 | 4.9745 | 4.9602 | 4.9636 | 4.9598 | 4.9576 | 4.9557 | 4.964 | 4.9582 | 0.018695 | 0.0037706
887 | 4.9068 | 5.0023 | 4.9278 | 4.9485 | 4.9681 | 4.9673 | 4.9751 | 4.9891 | 4.9873 | 4.9281 | 4.9683 | 4.9668 | 4.9459 | 4.9413 | 4.9463 | 4.9516 | 4.9454 | 4.9515 | 4.9603 | 4.9498 | 4.9468 | 4.9662 | 4.9689 | 4.9744 | 4.9601 | 4.9635 | 4.9597 | 4.9575 | 4.9556 | 4.9639 | 4.9581 | 0.018705 | 0.0037726
888 | 4.9066 | 5.0022 | 4.9277 | 4.9484 | 4.968 | 4.9672 | 4.975 | 4.9889 | 4.9872 | 4.928 | 4.9683 | 4.9668 | 4.9458 | 4.9412 | 4.9462 | 4.9515 | 4.9453 | 4.9514 | 4.9602 | 4.9497 | 4.9467 | 4.9661 | 4.9688 | 4.9743 | 4.96 | 4.9634 | 4.9596 | 4.9574 | 4.9555 | 4.9638 | 4.958 | 0.018714 | 0.0037746
889 | 4.9065 | 5.0021 | 4.9276 | 4.9483 | 4.9679 | 4.9671 | 4.9749 | 4.9888 | 4.9871 | 4.9279 | 4.9682 | 4.9667 | 4.9457 | 4.9411 | 4.9461 | 4.9514 | 4.9452 | 4.9513 | 4.9602 | 4.9496 | 4.9465 | 4.9659 | 4.9687 | 4.9742 | 4.9599 | 4.9633 | 4.9594 | 4.9573 | 4.9553 | 4.9637 | 4.9579 | 0.018724 | 0.0037765
890 | 4.9063 | 5.002 | 4.9275 | 4.9482 | 4.9678 | 4.967 | 4.9748 | 4.9887 | 4.987 | 4.9277 | 4.9681 | 4.9666 | 4.9456 | 4.941 | 4.946 | 4.9513 | 4.9451 | 4.9512 | 4.9601 | 4.9495 | 4.9464 | 4.9658 | 4.9686 | 4.9741 | 4.9598 | 4.9632 | 4.9593 | 4.9572 | 4.9552 | 4.9636 | 4.9578 | 0.018733 | 0.0037785
891 | 4.9062 | 5.0019 | 4.9274 | 4.9481 | 4.9677 | 4.9669 | 4.9747 | 4.9886 | 4.9869 | 4.9276 | 4.9681 | 4.9665 | 4.9456 | 4.9409 | 4.9459 | 4.9512 | 4.945 | 4.9511 | 4.96 | 4.9494 | 4.9463 | 4.9657 | 4.9685 | 4.9741 | 4.9597 | 4.9631 | 4.9592 | 4.9571 | 4.9551 | 4.9635 | 4.9577 | 0.018743 | 0.0037805
892 | 4.906 | 5.0018 | 4.9273 | 4.948 | 4.9676 | 4.9668 | 4.9746 | 4.9885 | 4.9868 | 4.9275 | 4.968 | 4.9664 | 4.9455 | 4.9408 | 4.9459 | 4.9511 | 4.9449 | 4.951 | 4.9599 | 4.9493 | 4.9462 | 4.9656 | 4.9684 | 4.974 | 4.9596 | 4.963 | 4.9591 | 4.9569 | 4.955 | 4.9634 | 4.9576 | 0.018752 | 0.0037825
893 | 4.9059 | 5.0017 | 4.9271 | 4.9479 | 4.9675 | 4.9667 | 4.9745 | 4.9884 | 4.9867 | 4.9274 | 4.968 | 4.9664 | 4.9454 | 4.9407 | 4.9458 | 4.951 | 4.9448 | 4.9509 | 4.9598 | 4.9492 | 4.9461 | 4.9655 | 4.9683 | 4.9739 | 4.9595 | 4.9629 | 4.959 | 4.9568 | 4.9549 | 4.9633 | 4.9575 | 0.018761 | 0.0037844
894 | 4.9057 | 5.0016 | 4.927 | 4.9478 | 4.9674 | 4.9666 | 4.9744 | 4.9883 | 4.9866 | 4.9272 | 4.9679 | 4.9663 | 4.9453 | 4.9406 | 4.9457 | 4.9509 | 4.9447 | 4.9508 | 4.9597 | 4.9491 | 4.946 | 4.9654 | 4.9682 | 4.9738 | 4.9594 | 4.9628 | 4.9589 | 4.9567 | 4.9548 | 4.9632 | 4.9574 | 0.018771 | 0.0037863
895 | 4.9056 | 5.0015 | 4.9269 | 4.9477 | 4.9673 | 4.9665 | 4.9743 | 4.9882 | 4.9865 | 4.9271 | 4.9678 | 4.9662 | 4.9452 | 4.9405 | 4.9456 | 4.9508 | 4.9446 | 4.9507 | 4.9597 | 4.949 | 4.9458 | 4.9653 | 4.9681 | 4.9737 | 4.9593 | 4.9628 | 4.9588 | 4.9566 | 4.9546 | 4.9631 | 4.9573 | 0.01878 | 0.0037883
896 | 4.9054 | 5.0014 | 4.9268 | 4.9476 | 4.9672 | 4.9664 | 4.9742 | 4.9881 | 4.9864 | 4.927 | 4.9678 | 4.9661 | 4.9452 | 4.9404 | 4.9455 | 4.9507 | 4.9445 | 4.9506 | 4.9596 | 4.9489 | 4.9457 | 4.9652 | 4.968 | 4.9736 | 4.9592 | 4.9627 | 4.9587 | 4.9565 | 4.9545 | 4.963 | 4.9572 | 0.018789 | 0.0037902
897 | 4.9053 | 5.0013 | 4.9267 | 4.9475 | 4.9671 | 4.9663 | 4.9741 | 4.988 | 4.9863 | 4.9269 | 4.9677 | 4.9661 | 4.9451 | 4.9403 | 4.9454 | 4.9506 | 4.9444 | 4.9505 | 4.9595 | 4.9488 | 4.9456 | 4.9651 | 4.9679 | 4.9735 | 4.9591 | 4.9626 | 4.9585 | 4.9564 | 4.9544 | 4.9629 | 4.9571 | 0.018798 | 0.0037921
898 | 4.9051 | 5.0013 | 4.9266 | 4.9474 | 4.967 | 4.9662 | 4.974 | 4.9879 | 4.9862 | 4.9267 | 4.9677 | 4.966 | 4.945 | 4.9402 | 4.9453 | 4.9505 | 4.9443 | 4.9504 | 4.9594 | 4.9487 | 4.9455 | 4.965 | 4.9678 | 4.9734 | 4.9591 | 4.9625 | 4.9584 | 4.9563 | 4.9543 | 4.9628 | 4.957 | 0.018807 | 0.003794
899 | 4.905 | 5.0012 | 4.9264 | 4.9473 | 4.9669 | 4.9661 | 4.9739 | 4.9878 | 4.9861 | 4.9266 | 4.9676 | 4.9659 | 4.9449 | 4.9401 | 4.9452 | 4.9504 | 4.9442 | 4.9503 | 4.9593 | 4.9486 | 4.9454 | 4.9649 | 4.9677 | 4.9733 | 4.959 | 4.9624 | 4.9583 | 4.9561 | 4.9542 | 4.9627 | 4.9569 | 0.018816 | 0.0037959
900 | 4.9048 | 5.0011 | 4.9263 | 4.9472 | 4.9668 | 4.966 | 4.9739 | 4.9877 | 4.986 | 4.9265 | 4.9675 | 4.9658 | 4.9448 | 4.94 | 4.9451 | 4.9503 | 4.9441 | 4.9502 | 4.9592 | 4.9485 | 4.9453 | 4.9648 | 4.9676 | 4.9732 | 4.9589 | 4.9623 | 4.9582 | 4.956 | 4.9541 | 4.9626 | 4.9568 | 0.018825 | 0.0037978
901 | 4.9047 | 5.001 | 4.9262 | 4.947 | 4.9667 | 4.9659 | 4.9738 | 4.9876 | 4.9859 | 4.9264 | 4.9675 | 4.9657 | 4.9447 | 4.9399 | 4.9451 | 4.9502 | 4.9439 | 4.95 | 4.9592 | 4.9484 | 4.9452 | 4.9647 | 4.9675 | 4.9731 | 4.9588 | 4.9622 | 4.9581 | 4.9559 | 4.9539 | 4.9625 | 4.9567 | 0.018834 | 0.0037997
902 | 4.9045 | 5.0009 | 4.9261 | 4.9469 | 4.9666 | 4.9658 | 4.9737 | 4.9875 | 4.9858 | 4.9263 | 4.9674 | 4.9657 | 4.9447 | 4.9398 | 4.945 | 4.9501 | 4.9438 | 4.9499 | 4.9591 | 4.9483 | 4.945 | 4.9646 | 4.9674 | 4.973 | 4.9587 | 4.9621 | 4.958 | 4.9558 | 4.9538 | 4.9624 | 4.9566 | 0.018843 | 0.0038016
903 | 4.9043 | 5.0008 | 4.926 | 4.9468 | 4.9665 | 4.9657 | 4.9736 | 4.9874 | 4.9857 | 4.9261 | 4.9673 | 4.9656 | 4.9446 | 4.9397 | 4.9449 | 4.95 | 4.9437 | 4.9498 | 4.959 | 4.9482 | 4.9449 | 4.9645 | 4.9673 | 4.973 | 4.9586 | 4.962 | 4.9579 | 4.9557 | 4.9537 | 4.9623 | 4.9565 | 0.018852 | 0.0038034
904 | 4.9042 | 5.0007 | 4.9259 | 4.9467 | 4.9664 | 4.9656 | 4.9735 | 4.9873 | 4.9856 | 4.926 | 4.9673 | 4.9655 | 4.9445 | 4.9396 | 4.9448 | 4.9499 | 4.9436 | 4.9497 | 4.9589 | 4.9481 | 4.9448 | 4.9644 | 4.9672 | 4.9729 | 4.9585 | 4.9619 | 4.9577 | 4.9556 | 4.9536 | 4.9622 | 4.9564 | 0.018861 | 0.0038053
905 | 4.9041 | 5.0006 | 4.9257 | 4.9466 | 4.9663 | 4.9655 | 4.9734 | 4.9872 | 4.9855 | 4.9259 | 4.9672 | 4.9654 | 4.9444 | 4.9395 | 4.9447 | 4.9498 | 4.9435 | 4.9496 | 4.9588 | 4.948 | 4.9447 | 4.9643 | 4.9671 | 4.9728 | 4.9584 | 4.9618 | 4.9576 | 4.9554 | 4.9535 | 4.9621 | 4.9563 | 0.018869 | 0.0038071
906 | 4.9039 | 5.0005 | 4.9256 | 4.9465 | 4.9662 | 4.9654 | 4.9733 | 4.9871 | 4.9854 | 4.9258 | 4.9672 | 4.9653 | 4.9443 | 4.9394 | 4.9446 | 4.9497 | 4.9434 | 4.9495 | 4.9587 | 4.948 | 4.9446 | 4.9642 | 4.967 | 4.9727 | 4.9583 | 4.9617 | 4.9575 | 4.9553 | 4.9533 | 4.962 | 4.9562 | 0.018878 | 0.003809
907 | 4.9038 | 5.0004 | 4.9255 | 4.9464 | 4.9661 | 4.9653 | 4.9732 | 4.987 | 4.9853 | 4.9257 | 4.9671 | 4.9653 | 4.9443 | 4.9393 | 4.9445 | 4.9496 | 4.9433 | 4.9494 | 4.9587 | 4.9479 | 4.9445 | 4.9641 | 4.9669 | 4.9726 | 4.9582 | 4.9616 | 4.9574 | 4.9552 | 4.9532 | 4.9619 | 4.9561 | 0.018887 | 0.0038108
908 | 4.9036 | 5.0003 | 4.9254 | 4.9463 | 4.966 | 4.9652 | 4.9731 | 4.9869 | 4.9852 | 4.9255 | 4.967 | 4.9652 | 4.9442 | 4.9392 | 4.9444 | 4.9495 | 4.9432 | 4.9493 | 4.9586 | 4.9478 | 4.9443 | 4.9639 | 4.9668 | 4.9725 | 4.9581 | 4.9615 | 4.9573 | 4.9551 | 4.9531 | 4.9618 | 4.956 | 0.018895 | 0.0038126
909 | 4.9035 | 5.0002 | 4.9253 | 4.9462 | 4.9659 | 4.9651 | 4.973 | 4.9868 | 4.9851 | 4.9254 | 4.967 | 4.9651 | 4.9441 | 4.9391 | 4.9443 | 4.9494 | 4.9431 | 4.9492 | 4.9585 | 4.9477 | 4.9442 | 4.9638 | 4.9667 | 4.9724 | 4.958 | 4.9614 | 4.9572 | 4.955 | 4.953 | 4.9617 | 4.9559 | 0.018904 | 0.0038144
910 | 4.9033 | 5.0001 | 4.9252 | 4.9461 | 4.9658 | 4.965 | 4.9729 | 4.9867 | 4.985 | 4.9253 | 4.9669 | 4.965 | 4.944 | 4.939 | 4.9443 | 4.9493 | 4.943 | 4.9491 | 4.9584 | 4.9476 | 4.9441 | 4.9637 | 4.9666 | 4.9723 | 4.9579 | 4.9613 | 4.9571 | 4.9549 | 4.9529 | 4.9616 | 4.9558 | 0.018912 | 0.0038162
911 | 4.9032 | 5 | 4.9251 | 4.946 | 4.9657 | 4.9649 | 4.9728 | 4.9866 | 4.9849 | 4.9252 | 4.9668 | 4.965 | 4.9439 | 4.9389 | 4.9442 | 4.9492 | 4.9429 | 4.949 | 4.9583 | 4.9475 | 4.944 | 4.9636 | 4.9665 | 4.9722 | 4.9578 | 4.9612 | 4.9569 | 4.9548 | 4.9528 | 4.9615 | 4.9557 | 0.018921 | 0.003818
912 | 4.903 | 5 | 4.9249 | 4.9459 | 4.9656 | 4.9648 | 4.9727 | 4.9865 | 4.9848 | 4.925 | 4.9668 | 4.9649 | 4.9438 | 4.9388 | 4.9441 | 4.9491 | 4.9428 | 4.9489 | 4.9582 | 4.9474 | 4.9439 | 4.9635 | 4.9664 | 4.9721 | 4.9577 | 4.9611 | 4.9568 | 4.9546 | 4.9526 | 4.9614 | 4.9556 | 0.018929 | 0.0038198
913 | 4.9029 | 4.9999 | 4.9248 | 4.9458 | 4.9655 | 4.9647 | 4.9726 | 4.9864 | 4.9847 | 4.9249 | 4.9667 | 4.9648 | 4.9438 | 4.9387 | 4.944 | 4.949 | 4.9427 | 4.9488 | 4.9581 | 4.9473 | 4.9438 | 4.9634 | 4.9663 | 4.972 | 4.9576 | 4.961 | 4.9567 | 4.9545 | 4.9525 | 4.9613 | 4.9555 | 0.018938 | 0.0038216
914 | 4.9027 | 4.9998 | 4.9247 | 4.9457 | 4.9654 | 4.9646 | 4.9726 | 4.9863 | 4.9846 | 4.9248 | 4.9666 | 4.9647 | 4.9437 | 4.9386 | 4.9439 | 4.9489 | 4.9426 | 4.9487 | 4.9581 | 4.9472 | 4.9436 | 4.9633 | 4.9662 | 4.9719 | 4.9575 | 4.9609 | 4.9566 | 4.9544 | 4.9524 | 4.9612 | 4.9554 | 0.018946 | 0.0038233
915 | 4.9026 | 4.9997 | 4.9246 | 4.9455 | 4.9653 | 4.9645 | 4.9725 | 4.9862 | 4.9845 | 4.9247 | 4.9666 | 4.9646 | 4.9436 | 4.9385 | 4.9438 | 4.9488 | 4.9425 | 4.9485 | 4.958 | 4.9471 | 4.9435 | 4.9632 | 4.9661 | 4.9718 | 4.9574 | 4.9608 | 4.9565 | 4.9543 | 4.9523 | 4.9611 | 4.9553 | 0.018954 | 0.0038251
916 | 4.9024 | 4.9996 | 4.9245 | 4.9454 | 4.9652 | 4.9644 | 4.9724 | 4.9861 | 4.9844 | 4.9246 | 4.9665 | 4.9646 | 4.9435 | 4.9384 | 4.9437 | 4.9487 | 4.9424 | 4.9484 | 4.9579 | 4.947 | 4.9434 | 4.9631 | 4.966 | 4.9718 | 4.9573 | 4.9607 | 4.9564 | 4.9542 | 4.9522 | 4.961 | 4.9552 | 0.018963 | 0.0038268
917 | 4.9023 | 4.9995 | 4.9244 | 4.9453 | 4.9651 | 4.9643 | 4.9723 | 4.986 | 4.9843 | 4.9244 | 4.9665 | 4.9645 | 4.9434 | 4.9383 | 4.9436 | 4.9486 | 4.9423 | 4.9483 | 4.9578 | 4.9469 | 4.9433 | 4.963 | 4.9659 | 4.9717 | 4.9572 | 4.9606 | 4.9563 | 4.9541 | 4.9521 | 4.9609 | 4.9551 | 0.018971 | 0.0038286
918 | 4.9022 | 4.9994 | 4.9242 | 4.9452 | 4.965 | 4.9642 | 4.9722 | 4.9859 | 4.9842 | 4.9243 | 4.9664 | 4.9644 | 4.9433 | 4.9382 | 4.9435 | 4.9485 | 4.9422 | 4.9482 | 4.9577 | 4.9468 | 4.9432 | 4.9629 | 4.9658 | 4.9716 | 4.9571 | 4.9606 | 4.9561 | 4.9539 | 4.9519 | 4.9607 | 4.955 | 0.018979 | 0.0038303
919 | 4.902 | 4.9993 | 4.9241 | 4.9451 | 4.9649 | 4.9641 | 4.9721 | 4.9858 | 4.9841 | 4.9242 | 4.9663 | 4.9643 | 4.9433 | 4.9381 | 4.9434 | 4.9484 | 4.942 | 4.9481 | 4.9576 | 4.9467 | 4.9431 | 4.9628 | 4.9657 | 4.9715 | 4.957 | 4.9605 | 4.956 | 4.9538 | 4.9518 | 4.9606 | 4.9549 | 0.018987 | 0.003832
920 | 4.9019 | 4.9992 | 4.924 | 4.945 | 4.9648 | 4.964 | 4.972 | 4.9857 | 4.984 | 4.9241 | 4.9663 | 4.9642 | 4.9432 | 4.9379 | 4.9434 | 4.9483 | 4.9419 | 4.948 | 4.9576 | 4.9466 | 4.9429 | 4.9627 | 4.9656 | 4.9714 | 4.9569 | 4.9604 | 4.9559 | 4.9537 | 4.9517 | 4.9605 | 4.9548 | 0.018995 | 0.0038337
921 | 4.9017 | 4.9991 | 4.9239 | 4.9449 | 4.9647 | 4.9639 | 4.9719 | 4.9856 | 4.9839 | 4.924 | 4.9662 | 4.9642 | 4.9431 | 4.9378 | 4.9433 | 4.9482 | 4.9418 | 4.9479 | 4.9575 | 4.9465 | 4.9428 | 4.9626 | 4.9655 | 4.9713 | 4.9568 | 4.9603 | 4.9558 | 4.9536 | 4.9516 | 4.9604 | 4.9547 | 0.019003 | 0.0038354
922 | 4.9016 | 4.999 | 4.9238 | 4.9448 | 4.9646 | 4.9638 | 4.9718 | 4.9855 | 4.9838 | 4.9238 | 4.9661 | 4.9641 | 4.943 | 4.9377 | 4.9432 | 4.9481 | 4.9417 | 4.9478 | 4.9574 | 4.9464 | 4.9427 | 4.9625 | 4.9654 | 4.9712 | 4.9567 | 4.9602 | 4.9557 | 4.9535 | 4.9515 | 4.9603 | 4.9546 | 0.019011 | 0.0038371
923 | 4.9015 | 4.9989 | 4.9237 | 4.9447 | 4.9645 | 4.9637 | 4.9717 | 4.9854 | 4.9837 | 4.9237 | 4.9661 | 4.964 | 4.9429 | 4.9376 | 4.9431 | 4.948 | 4.9416 | 4.9477 | 4.9573 | 4.9463 | 4.9426 | 4.9624 | 4.9653 | 4.9711 | 4.9566 | 4.9601 | 4.9556 | 4.9534 | 4.9513 | 4.9602 | 4.9545 | 0.019019 | 0.0038388
924 | 4.9013 | 4.9988 | 4.9235 | 4.9446 | 4.9644 | 4.9635 | 4.9716 | 4.9853 | 4.9836 | 4.9236 | 4.966 | 4.9639 | 4.9429 | 4.9375 | 4.943 | 4.9479 | 4.9415 | 4.9476 | 4.9572 | 4.9462 | 4.9425 | 4.9623 | 4.9652 | 4.971 | 4.9565 | 4.96 | 4.9555 | 4.9532 | 4.9512 | 4.9601 | 4.9544 | 0.019027 | 0.0038405
925 | 4.9012 | 4.9987 | 4.9234 | 4.9445 | 4.9643 | 4.9634 | 4.9715 | 4.9852 | 4.9835 | 4.9235 | 4.966 | 4.9638 | 4.9428 | 4.9374 | 4.9429 | 4.9478 | 4.9414 | 4.9475 | 4.9571 | 4.9461 | 4.9424 | 4.9622 | 4.965 | 4.9709 | 4.9564 | 4.9599 | 4.9553 | 4.9531 | 4.9511 | 4.96 | 4.9543 | 0.019035 | 0.0038421
926 | 4.901 | 4.9986 | 4.9233 | 4.9444 | 4.9642 | 4.9633 | 4.9714 | 4.9851 | 4.9834 | 4.9234 | 4.9659 | 4.9638 | 4.9427 | 4.9373 | 4.9428 | 4.9477 | 4.9413 | 4.9474 | 4.9571 | 4.946 | 4.9422 | 4.962 | 4.9649 | 4.9708 | 4.9563 | 4.9598 | 4.9552 | 4.953 | 4.951 | 4.9599 | 4.9542 | 0.019043 | 0.0038438
927 | 4.9009 | 4.9985 | 4.9232 | 4.9443 | 4.9641 | 4.9632 | 4.9713 | 4.985 | 4.9833 | 4.9232 | 4.9658 | 4.9637 | 4.9426 | 4.9372 | 4.9427 | 4.9476 | 4.9412 | 4.9473 | 4.957 | 4.9459 | 4.9421 | 4.9619 | 4.9648 | 4.9707 | 4.9562 | 4.9597 | 4.9551 | 4.9529 | 4.9509 | 4.9598 | 4.9541 | 0.019051 | 0.0038454
928 | 4.9008 | 4.9985 | 4.9231 | 4.9441 | 4.964 | 4.9631 | 4.9712 | 4.9849 | 4.9832 | 4.9231 | 4.9658 | 4.9636 | 4.9425 | 4.9371 | 4.9426 | 4.9475 | 4.9411 | 4.9472 | 4.9569 | 4.9458 | 4.942 | 4.9618 | 4.9647 | 4.9706 | 4.9561 | 4.9596 | 4.955 | 4.9528 | 4.9508 | 4.9597 | 4.954 | 0.019058 | 0.0038471
929 | 4.9006 | 4.9984 | 4.923 | 4.944 | 4.9639 | 4.963 | 4.9712 | 4.9848 | 4.9831 | 4.923 | 4.9657 | 4.9635 | 4.9424 | 4.937 | 4.9425 | 4.9474 | 4.941 | 4.947 | 4.9568 | 4.9457 | 4.9419 | 4.9617 | 4.9646 | 4.9705 | 4.956 | 4.9595 | 4.9549 | 4.9527 | 4.9506 | 4.9596 | 4.9539 | 0.019066 | 0.0038487
930 | 4.9005 | 4.9983 | 4.9228 | 4.9439 | 4.9638 | 4.9629 | 4.9711 | 4.9847 | 4.983 | 4.9229 | 4.9656 | 4.9634 | 4.9424 | 4.9369 | 4.9425 | 4.9473 | 4.9409 | 4.9469 | 4.9567 | 4.9456 | 4.9418 | 4.9616 | 4.9645 | 4.9705 | 4.956 | 4.9594 | 4.9548 | 4.9525 | 4.9505 | 4.9595 | 4.9538 | 0.019074 | 0.0038503
931 | 4.9003 | 4.9982 | 4.9227 | 4.9438 | 4.9637 | 4.9628 | 4.971 | 4.9846 | 4.9829 | 4.9228 | 4.9656 | 4.9634 | 4.9423 | 4.9368 | 4.9424 | 4.9472 | 4.9408 | 4.9468 | 4.9566 | 4.9455 | 4.9417 | 4.9615 | 4.9644 | 4.9704 | 4.9559 | 4.9593 | 4.9547 | 4.9524 | 4.9504 | 4.9594 | 4.9537 | 0.019081 | 0.003852
932 | 4.9002 | 4.9981 | 4.9226 | 4.9437 | 4.9636 | 4.9627 | 4.9709 | 4.9845 | 4.9828 | 4.9226 | 4.9655 | 4.9633 | 4.9422 | 4.9367 | 4.9423 | 4.9471 | 4.9407 | 4.9467 | 4.9566 | 4.9454 | 4.9415 | 4.9614 | 4.9643 | 4.9703 | 4.9558 | 4.9592 | 4.9545 | 4.9523 | 4.9503 | 4.9593 | 4.9536 | 0.019089 | 0.0038536
933 | 4.9001 | 4.998 | 4.9225 | 4.9436 | 4.9635 | 4.9626 | 4.9708 | 4.9844 | 4.9827 | 4.9225 | 4.9654 | 4.9632 | 4.9421 | 4.9366 | 4.9422 | 4.947 | 4.9406 | 4.9466 | 4.9565 | 4.9453 | 4.9414 | 4.9613 | 4.9642 | 4.9702 | 4.9557 | 4.9591 | 4.9544 | 4.9522 | 4.9502 | 4.9592 | 4.9535 | 0.019096 | 0.0038552
934 | 4.8999 | 4.9979 | 4.9224 | 4.9435 | 4.9634 | 4.9625 | 4.9707 | 4.9843 | 4.9826 | 4.9224 | 4.9654 | 4.9631 | 4.942 | 4.9365 | 4.9421 | 4.9469 | 4.9405 | 4.9465 | 4.9564 | 4.9452 | 4.9413 | 4.9612 | 4.9641 | 4.9701 | 4.9556 | 4.959 | 4.9543 | 4.9521 | 4.95 | 4.9591 | 4.9534 | 0.019104 | 0.0038567
935 | 4.8998 | 4.9978 | 4.9222 | 4.9434 | 4.9633 | 4.9624 | 4.9706 | 4.9842 | 4.9825 | 4.9223 | 4.9653 | 4.963 | 4.9419 | 4.9364 | 4.942 | 4.9468 | 4.9404 | 4.9464 | 4.9563 | 4.9451 | 4.9412 | 4.9611 | 4.964 | 4.97 | 4.9555 | 4.9589 | 4.9542 | 4.952 | 4.9499 | 4.959 | 4.9533 | 0.019111 | 0.0038583
936 | 4.8997 | 4.9977 | 4.9221 | 4.9433 | 4.9632 | 4.9623 | 4.9705 | 4.9841 | 4.9824 | 4.9222 | 4.9652 | 4.963 | 4.9419 | 4.9363 | 4.9419 | 4.9467 | 4.9402 | 4.9463 | 4.9562 | 4.945 | 4.9411 | 4.961 | 4.9639 | 4.9699 | 4.9554 | 4.9588 | 4.9541 | 4.9518 | 4.9498 | 4.9589 | 4.9532 | 0.019119 | 0.0038599
937 | 4.8995 | 4.9976 | 4.922 | 4.9432 | 4.9631 | 4.9622 | 4.9704 | 4.984 | 4.9823 | 4.9221 | 4.9652 | 4.9629 | 4.9418 | 4.9362 | 4.9418 | 4.9466 | 4.9401 | 4.9462 | 4.9561 | 4.9449 | 4.941 | 4.9609 | 4.9638 | 4.9698 | 4.9553 | 4.9587 | 4.954 | 4.9517 | 4.9497 | 4.9588 | 4.9531 | 0.019126 | 0.0038615
938 | 4.8994 | 4.9975 | 4.9219 | 4.9431 | 4.963 | 4.9621 | 4.9703 | 4.9839 | 4.9822 | 4.9219 | 4.9651 | 4.9628 | 4.9417 | 4.9361 | 4.9417 | 4.9465 | 4.94 | 4.9461 | 4.956 | 4.9448 | 4.9408 | 4.9608 | 4.9637 | 4.9697 | 4.9552 | 4.9586 | 4.9538 | 4.9516 | 4.9496 | 4.9587 | 4.953 | 0.019133 | 0.003863
939 | 4.8993 | 4.9974 | 4.9218 | 4.943 | 4.9629 | 4.962 | 4.9702 | 4.9838 | 4.9821 | 4.9218 | 4.9651 | 4.9627 | 4.9416 | 4.936 | 4.9416 | 4.9464 | 4.9399 | 4.946 | 4.956 | 4.9447 | 4.9407 | 4.9607 | 4.9636 | 4.9696 | 4.9551 | 4.9585 | 4.9537 | 4.9515 | 4.9494 | 4.9586 | 4.9529 | 0.019141 | 0.0038646
940 | 4.8991 | 4.9973 | 4.9217 | 4.9428 | 4.9628 | 4.9619 | 4.9701 | 4.9837 | 4.982 | 4.9217 | 4.965 | 4.9626 | 4.9415 | 4.9359 | 4.9416 | 4.9463 | 4.9398 | 4.9459 | 4.9559 | 4.9446 | 4.9406 | 4.9606 | 4.9635 | 4.9695 | 4.955 | 4.9584 | 4.9536 | 4.9514 | 4.9493 | 4.9585 | 4.9528 | 0.019148 | 0.0038661
941 | 4.899 | 4.9972 | 4.9215 | 4.9427 | 4.9627 | 4.9618 | 4.97 | 4.9836 | 4.9819 | 4.9216 | 4.9649 | 4.9626 | 4.9414 | 4.9358 | 4.9415 | 4.9462 | 4.9397 | 4.9457 | 4.9558 | 4.9445 | 4.9405 | 4.9605 | 4.9634 | 4.9694 | 4.9549 | 4.9583 | 4.9535 | 4.9513 | 4.9492 | 4.9584 | 4.9527 | 0.019155 | 0.0038676
942 | 4.8989 | 4.9971 | 4.9214 | 4.9426 | 4.9626 | 4.9617 | 4.9699 | 4.9835 | 4.9818 | 4.9215 | 4.9649 | 4.9625 | 4.9414 | 4.9357 | 4.9414 | 4.9461 | 4.9396 | 4.9456 | 4.9557 | 4.9444 | 4.9404 | 4.9604 | 4.9633 | 4.9693 | 4.9548 | 4.9582 | 4.9534 | 4.9511 | 4.9491 | 4.9583 | 4.9526 | 0.019162 | 0.0038691
943 | 4.8987 | 4.997 | 4.9213 | 4.9425 | 4.9625 | 4.9616 | 4.9698 | 4.9834 | 4.9817 | 4.9213 | 4.9648 | 4.9624 | 4.9413 | 4.9356 | 4.9413 | 4.946 | 4.9395 | 4.9455 | 4.9556 | 4.9443 | 4.9402 | 4.9602 | 4.9632 | 4.9692 | 4.9547 | 4.9581 | 4.9533 | 4.951 | 4.949 | 4.9582 | 4.9524 | 0.019169 | 0.0038706
944 | 4.8986 | 4.9969 | 4.9212 | 4.9424 | 4.9624 | 4.9615 | 4.9697 | 4.9833 | 4.9816 | 4.9212 | 4.9647 | 4.9623 | 4.9412 | 4.9354 | 4.9412 | 4.9459 | 4.9394 | 4.9454 | 4.9555 | 4.9442 | 4.9401 | 4.9601 | 4.9631 | 4.9692 | 4.9546 | 4.9581 | 4.9532 | 4.9509 | 4.9489 | 4.9581 | 4.9523 | 0.019176 | 0.0038721
945 | 4.8985 | 4.9968 | 4.9211 | 4.9423 | 4.9623 | 4.9614 | 4.9697 | 4.9832 | 4.9815 | 4.9211 | 4.9647 | 4.9622 | 4.9411 | 4.9353 | 4.9411 | 4.9458 | 4.9393 | 4.9453 | 4.9554 | 4.9441 | 4.94 | 4.96 | 4.963 | 4.9691 | 4.9545 | 4.958 | 4.953 | 4.9508 | 4.9487 | 4.958 | 4.9522 | 0.019183 | 0.0038736
946 | 4.8983 | 4.9968 | 4.921 | 4.9422 | 4.9622 | 4.9613 | 4.9696 | 4.9831 | 4.9814 | 4.921 | 4.9646 | 4.9622 | 4.941 | 4.9352 | 4.941 | 4.9457 | 4.9392 | 4.9452 | 4.9554 | 4.944 | 4.9399 | 4.9599 | 4.9629 | 4.969 | 4.9544 | 4.9579 | 4.9529 | 4.9507 | 4.9486 | 4.9579 | 4.9521 | 0.01919 | 0.0038751
947 | 4.8982 | 4.9967 | 4.9208 | 4.9421 | 4.9621 | 4.9612 | 4.9695 | 4.983 | 4.9813 | 4.9209 | 4.9645 | 4.9621 | 4.9409 | 4.9351 | 4.9409 | 4.9456 | 4.9391 | 4.9451 | 4.9553 | 4.9439 | 4.9398 | 4.9598 | 4.9628 | 4.9689 | 4.9543 | 4.9578 | 4.9528 | 4.9506 | 4.9485 | 4.9578 | 4.952 | 0.019197 | 0.0038766
948 | 4.8981 | 4.9966 | 4.9207 | 4.942 | 4.962 | 4.9611 | 4.9694 | 4.9829 | 4.9812 | 4.9208 | 4.9645 | 4.962 | 4.9408 | 4.935 | 4.9408 | 4.9455 | 4.939 | 4.945 | 4.9552 | 4.9438 | 4.9397 | 4.9597 | 4.9627 | 4.9688 | 4.9542 | 4.9577 | 4.9527 | 4.9504 | 4.9484 | 4.9577 | 4.9519 | 0.019204 | 0.0038781
949 | 4.898 | 4.9965 | 4.9206 | 4.9419 | 4.9619 | 4.961 | 4.9693 | 4.9828 | 4.9811 | 4.9206 | 4.9644 | 4.9619 | 4.9408 | 4.9349 | 4.9407 | 4.9454 | 4.9389 | 4.9449 | 4.9551 | 4.9437 | 4.9395 | 4.9596 | 4.9625 | 4.9687 | 4.9541 | 4.9576 | 4.9526 | 4.9503 | 4.9483 | 4.9576 | 4.9518 | 0.019211 | 0.0038795
950 | 4.8978 | 4.9964 | 4.9205 | 4.9418 | 4.9618 | 4.9609 | 4.9692 | 4.9827 | 4.981 | 4.9205 | 4.9643 | 4.9618 | 4.9407 | 4.9348 | 4.9406 | 4.9453 | 4.9388 | 4.9448 | 4.955 | 4.9436 | 4.9394 | 4.9595 | 4.9624 | 4.9686 | 4.954 | 4.9575 | 4.9525 | 4.9502 | 4.9481 | 4.9575 | 4.9517 | 0.019218 | 0.003881
951 | 4.8977 | 4.9963 | 4.9204 | 4.9417 | 4.9617 | 4.9608 | 4.9691 | 4.9826 | 4.9809 | 4.9204 | 4.9643 | 4.9618 | 4.9406 | 4.9347 | 4.9406 | 4.9452 | 4.9386 | 4.9447 | 4.9549 | 4.9435 | 4.9393 | 4.9594 | 4.9623 | 4.9685 | 4.9539 | 4.9574 | 4.9523 | 4.9501 | 4.948 | 4.9574 | 4.9516 | 0.019224 | 0.0038824
952 | 4.8976 | 4.9962 | 4.9202 | 4.9415 | 4.9616 | 4.9607 | 4.969 | 4.9825 | 4.9808 | 4.9203 | 4.9642 | 4.9617 | 4.9405 | 4.9346 | 4.9405 | 4.9451 | 4.9385 | 4.9446 | 4.9549 | 4.9434 | 4.9392 | 4.9593 | 4.9622 | 4.9684 | 4.9538 | 4.9573 | 4.9522 | 4.95 | 4.9479 | 4.9573 | 4.9515 | 0.019231 | 0.0038838
953 | 4.8974 | 4.9961 | 4.9201 | 4.9414 | 4.9615 | 4.9606 | 4.9689 | 4.9823 | 4.9807 | 4.9202 | 4.9642 | 4.9616 | 4.9404 | 4.9345 | 4.9404 | 4.945 | 4.9384 | 4.9444 | 4.9548 | 4.9433 | 4.9391 | 4.9592 | 4.9621 | 4.9683 | 4.9537 | 4.9572 | 4.9521 | 4.9498 | 4.9478 | 4.9572 | 4.9514 | 0.019238 | 0.0038853
954 | 4.8973 | 4.996 | 4.92 | 4.9413 | 4.9614 | 4.9604 | 4.9688 | 4.9822 | 4.9806 | 4.92 | 4.9641 | 4.9615 | 4.9403 | 4.9344 | 4.9403 | 4.9449 | 4.9383 | 4.9443 | 4.9547 | 4.9432 | 4.9389 | 4.9591 | 4.962 | 4.9682 | 4.9536 | 4.9571 | 4.952 | 4.9497 | 4.9477 | 4.9571 | 4.9513 | 0.019244 | 0.0038867
955 | 4.8972 | 4.9959 | 4.9199 | 4.9412 | 4.9613 | 4.9603 | 4.9687 | 4.9821 | 4.9805 | 4.9199 | 4.964 | 4.9614 | 4.9403 | 4.9343 | 4.9402 | 4.9448 | 4.9382 | 4.9442 | 4.9546 | 4.9431 | 4.9388 | 4.959 | 4.9619 | 4.9681 | 4.9535 | 4.957 | 4.9519 | 4.9496 | 4.9475 | 4.957 | 4.9512 | 0.019251 | 0.0038881
956 | 4.8971 | 4.9958 | 4.9198 | 4.9411 | 4.9612 | 4.9602 | 4.9686 | 4.982 | 4.9804 | 4.9198 | 4.964 | 4.9613 | 4.9402 | 4.9342 | 4.9401 | 4.9447 | 4.9381 | 4.9441 | 4.9545 | 4.943 | 4.9387 | 4.9589 | 4.9618 | 4.968 | 4.9534 | 4.9569 | 4.9518 | 4.9495 | 4.9474 | 4.9569 | 4.9511 | 0.019257 | 0.0038895
957 | 4.8969 | 4.9957 | 4.9197 | 4.941 | 4.9611 | 4.9601 | 4.9685 | 4.9819 | 4.9803 | 4.9197 | 4.9639 | 4.9613 | 4.9401 | 4.9341 | 4.94 | 4.9446 | 4.938 | 4.944 | 4.9544 | 4.9429 | 4.9386 | 4.9588 | 4.9617 | 4.9679 | 4.9533 | 4.9568 | 4.9516 | 4.9494 | 4.9473 | 4.9568 | 4.951 | 0.019264 | 0.0038909
958 | 4.8968 | 4.9956 | 4.9195 | 4.9409 | 4.961 | 4.96 | 4.9684 | 4.9818 | 4.9802 | 4.9196 | 4.9638 | 4.9612 | 4.94 | 4.934 | 4.9399 | 4.9445 | 4.9379 | 4.9439 | 4.9543 | 4.9428 | 4.9385 | 4.9587 | 4.9616 | 4.9678 | 4.9532 | 4.9567 | 4.9515 | 4.9493 | 4.9472 | 4.9567 | 4.9509 | 0.01927 | 0.0038923
959 | 4.8967 | 4.9955 | 4.9194 | 4.9408 | 4.9609 | 4.9599 | 4.9683 | 4.9817 | 4.9801 | 4.9195 | 4.9638 | 4.9611 | 4.9399 | 4.9339 | 4.9398 | 4.9444 | 4.9378 | 4.9438 | 4.9543 | 4.9427 | 4.9384 | 4.9585 | 4.9615 | 4.9678 | 4.9531 | 4.9566 | 4.9514 | 4.9491 | 4.9471 | 4.9565 | 4.9508 | 0.019277 | 0.0038936
960 | 4.8966 | 4.9954 | 4.9193 | 4.9407 | 4.9608 | 4.9598 | 4.9682 | 4.9816 | 4.98 | 4.9193 | 4.9637 | 4.961 | 4.9398 | 4.9338 | 4.9397 | 4.9443 | 4.9377 | 4.9437 | 4.9542 | 4.9426 | 4.9382 | 4.9584 | 4.9614 | 4.9677 | 4.953 | 4.9565 | 4.9513 | 4.949 | 4.9469 | 4.9564 | 4.9507 | 0.019283 | 0.003895
961 | 4.8964 | 4.9953 | 4.9192 | 4.9406 | 4.9607 | 4.9597 | 4.9681 | 4.9815 | 4.9799 | 4.9192 | 4.9636 | 4.9609 | 4.9398 | 4.9337 | 4.9396 | 4.9442 | 4.9376 | 4.9436 | 4.9541 | 4.9425 | 4.9381 | 4.9583 | 4.9613 | 4.9676 | 4.9529 | 4.9564 | 4.9512 | 4.9489 | 4.9468 | 4.9563 | 4.9506 | 0.019289 | 0.0038964
962 | 4.8963 | 4.9952 | 4.9191 | 4.9405 | 4.9606 | 4.9596 | 4.9681 | 4.9814 | 4.9798 | 4.9191 | 4.9636 | 4.9609 | 4.9397 | 4.9336 | 4.9395 | 4.9441 | 4.9375 | 4.9435 | 4.954 | 4.9424 | 4.938 | 4.9582 | 4.9612 | 4.9675 | 4.9528 | 4.9563 | 4.9511 | 4.9488 | 4.9467 | 4.9562 | 4.9505 | 0.019296 | 0.0038977
963 | 4.8962 | 4.9951 | 4.9189 | 4.9404 | 4.9605 | 4.9595 | 4.968 | 4.9813 | 4.9797 | 4.919 | 4.9635 | 4.9608 | 4.9396 | 4.9334 | 4.9395 | 4.944 | 4.9374 | 4.9434 | 4.9539 | 4.9423 | 4.9379 | 4.9581 | 4.9611 | 4.9674 | 4.9527 | 4.9562 | 4.9509 | 4.9487 | 4.9466 | 4.9561 | 4.9504 | 0.019302 | 0.003899
964 | 4.8961 | 4.995 | 4.9188 | 4.9402 | 4.9604 | 4.9594 | 4.9679 | 4.9812 | 4.9796 | 4.9189 | 4.9634 | 4.9607 | 4.9395 | 4.9333 | 4.9394 | 4.9439 | 4.9373 | 4.9432 | 4.9538 | 4.9422 | 4.9378 | 4.958 | 4.961 | 4.9673 | 4.9526 | 4.9561 | 4.9508 | 4.9486 | 4.9465 | 4.956 | 4.9503 | 0.019308 | 0.0039004
965 | 4.8959 | 4.9949 | 4.9187 | 4.9401 | 4.9603 | 4.9593 | 4.9678 | 4.9811 | 4.9795 | 4.9188 | 4.9634 | 4.9606 | 4.9394 | 4.9332 | 4.9393 | 4.9438 | 4.9372 | 4.9431 | 4.9537 | 4.9421 | 4.9376 | 4.9579 | 4.9609 | 4.9672 | 4.9525 | 4.956 | 4.9507 | 4.9484 | 4.9463 | 4.9559 | 4.9502 | 0.019314 | 0.0039017
966 | 4.8958 | 4.9949 | 4.9186 | 4.94 | 4.9602 | 4.9592 | 4.9677 | 4.981 | 4.9794 | 4.9187 | 4.9633 | 4.9605 | 4.9393 | 4.9331 | 4.9392 | 4.9437 | 4.937 | 4.943 | 4.9537 | 4.942 | 4.9375 | 4.9578 | 4.9608 | 4.9671 | 4.9524 | 4.9559 | 4.9506 | 4.9483 | 4.9462 | 4.9558 | 4.9501 | 0.01932 | 0.003903
967 | 4.8957 | 4.9948 | 4.9185 | 4.9399 | 4.9601 | 4.9591 | 4.9676 | 4.9809 | 4.9793 | 4.9185 | 4.9632 | 4.9604 | 4.9392 | 4.933 | 4.9391 | 4.9436 | 4.9369 | 4.9429 | 4.9536 | 4.9419 | 4.9374 | 4.9577 | 4.9607 | 4.967 | 4.9523 | 4.9558 | 4.9505 | 4.9482 | 4.9461 | 4.9557 | 4.95 | 0.019326 | 0.0039043
968 | 4.8956 | 4.9947 | 4.9184 | 4.9398 | 4.96 | 4.959 | 4.9675 | 4.9808 | 4.9792 | 4.9184 | 4.9632 | 4.9604 | 4.9392 | 4.9329 | 4.939 | 4.9435 | 4.9368 | 4.9428 | 4.9535 | 4.9418 | 4.9373 | 4.9576 | 4.9605 | 4.9669 | 4.9522 | 4.9557 | 4.9504 | 4.9481 | 4.946 | 4.9556 | 4.9499 | 0.019332 | 0.0039056
969 | 4.8955 | 4.9946 | 4.9182 | 4.9397 | 4.9599 | 4.9589 | 4.9674 | 4.9807 | 4.9791 | 4.9183 | 4.9631 | 4.9603 | 4.9391 | 4.9328 | 4.9389 | 4.9434 | 4.9367 | 4.9427 | 4.9534 | 4.9416 | 4.9372 | 4.9575 | 4.9604 | 4.9668 | 4.9521 | 4.9556 | 4.9502 | 4.948 | 4.9458 | 4.9555 | 4.9498 | 0.019338 | 0.0039069
970 | 4.8953 | 4.9945 | 4.9181 | 4.9396 | 4.9598 | 4.9588 | 4.9673 | 4.9806 | 4.979 | 4.9182 | 4.963 | 4.9602 | 4.939 | 4.9327 | 4.9388 | 4.9433 | 4.9366 | 4.9426 | 4.9533 | 4.9415 | 4.937 | 4.9574 | 4.9603 | 4.9667 | 4.952 | 4.9555 | 4.9501 | 4.9478 | 4.9457 | 4.9554 | 4.9497 | 0.019344 | 0.0039082
971 | 4.8952 | 4.9944 | 4.918 | 4.9395 | 4.9597 | 4.9587 | 4.9672 | 4.9805 | 4.9789 | 4.9181 | 4.963 | 4.9601 | 4.9389 | 4.9326 | 4.9387 | 4.9432 | 4.9365 | 4.9425 | 4.9532 | 4.9414 | 4.9369 | 4.9573 | 4.9602 | 4.9666 | 4.9519 | 4.9554 | 4.95 | 4.9477 | 4.9456 | 4.9553 | 4.9496 | 0.01935 | 0.0039095
972 | 4.8951 | 4.9943 | 4.9179 | 4.9394 | 4.9596 | 4.9586 | 4.9671 | 4.9804 | 4.9788 | 4.918 | 4.9629 | 4.96 | 4.9388 | 4.9325 | 4.9386 | 4.9431 | 4.9364 | 4.9424 | 4.9531 | 4.9413 | 4.9368 | 4.9572 | 4.9601 | 4.9665 | 4.9518 | 4.9553 | 4.9499 | 4.9476 | 4.9455 | 4.9552 | 4.9495 | 0.019356 | 0.0039107
973 | 4.895 | 4.9942 | 4.9178 | 4.9393 | 4.9595 | 4.9585 | 4.967 | 4.9803 | 4.9787 | 4.9178 | 4.9628 | 4.96 | 4.9387 | 4.9324 | 4.9385 | 4.943 | 4.9363 | 4.9423 | 4.9531 | 4.9412 | 4.9367 | 4.957 | 4.96 | 4.9664 | 4.9517 | 4.9552 | 4.9498 | 4.9475 | 4.9454 | 4.9551 | 4.9494 | 0.019362 | 0.003912
974 | 4.8949 | 4.9941 | 4.9176 | 4.9391 | 4.9594 | 4.9584 | 4.9669 | 4.9802 | 4.9786 | 4.9177 | 4.9628 | 4.9599 | 4.9386 | 4.9323 | 4.9384 | 4.9429 | 4.9362 | 4.9422 | 4.953 | 4.9411 | 4.9366 | 4.9569 | 4.9599 | 4.9663 | 4.9516 | 4.9551 | 4.9497 | 4.9474 | 4.9452 | 4.955 | 4.9493 | 0.019368 | 0.0039132
975 | 4.8947 | 4.994 | 4.9175 | 4.939 | 4.9593 | 4.9583 | 4.9668 | 4.9801 | 4.9785 | 4.9176 | 4.9627 | 4.9598 | 4.9386 | 4.9322 | 4.9384 | 4.9428 | 4.9361 | 4.942 | 4.9529 | 4.941 | 4.9365 | 4.9568 | 4.9598 | 4.9662 | 4.9515 | 4.955 | 4.9495 | 4.9472 | 4.9451 | 4.9549 | 4.9492 | 0.019373 | 0.0039145
976 | 4.8946 | 4.9939 | 4.9174 | 4.9389 | 4.9592 | 4.9582 | 4.9667 | 4.98 | 4.9784 | 4.9175 | 4.9626 | 4.9597 | 4.9385 | 4.9321 | 4.9383 | 4.9427 | 4.936 | 4.9419 | 4.9528 | 4.9409 | 4.9363 | 4.9567 | 4.9597 | 4.9662 | 4.9514 | 4.9549 | 4.9494 | 4.9471 | 4.945 | 4.9548 | 4.9491 | 0.019379 | 0.0039157
977 | 4.8945 | 4.9938 | 4.9173 | 4.9388 | 4.9591 | 4.958 | 4.9666 | 4.9799 | 4.9783 | 4.9174 | 4.9626 | 4.9596 | 4.9384 | 4.932 | 4.9382 | 4.9426 | 4.9359 | 4.9418 | 4.9527 | 4.9408 | 4.9362 | 4.9566 | 4.9596 | 4.9661 | 4.9513 | 4.9548 | 4.9493 | 4.947 | 4.9449 | 4.9547 | 4.949 | 0.019385 | 0.0039169
978 | 4.8944 | 4.9937 | 4.9172 | 4.9387 | 4.959 | 4.9579 | 4.9665 | 4.9798 | 4.9782 | 4.9173 | 4.9625 | 4.9595 | 4.9383 | 4.9319 | 4.9381 | 4.9425 | 4.9358 | 4.9417 | 4.9526 | 4.9407 | 4.9361 | 4.9565 | 4.9595 | 4.966 | 4.9512 | 4.9547 | 4.9492 | 4.9469 | 4.9448 | 4.9546 | 4.9489 | 0.01939 | 0.0039181
979 | 4.8943 | 4.9936 | 4.917 | 4.9386 | 4.9589 | 4.9578 | 4.9664 | 4.9797 | 4.9781 | 4.9171 | 4.9624 | 4.9595 | 4.9382 | 4.9317 | 4.938 | 4.9424 | 4.9356 | 4.9416 | 4.9525 | 4.9406 | 4.936 | 4.9564 | 4.9594 | 4.9659 | 4.9511 | 4.9547 | 4.9491 | 4.9468 | 4.9446 | 4.9545 | 4.9488 | 0.019396 | 0.0039193
980 | 4.8942 | 4.9935 | 4.9169 | 4.9385 | 4.9588 | 4.9577 | 4.9663 | 4.9796 | 4.978 | 4.917 | 4.9624 | 4.9594 | 4.9381 | 4.9316 | 4.9379 | 4.9423 | 4.9355 | 4.9415 | 4.9525 | 4.9405 | 4.9359 | 4.9563 | 4.9593 | 4.9658 | 4.951 | 4.9546 | 4.949 | 4.9467 | 4.9445 | 4.9544 | 4.9487 | 0.019401 | 0.0039205
981 | 4.894 | 4.9934 | 4.9168 | 4.9384 | 4.9587 | 4.9576 | 4.9662 | 4.9795 | 4.9779 | 4.9169 | 4.9623 | 4.9593 | 4.938 | 4.9315 | 4.9378 | 4.9422 | 4.9354 | 4.9414 | 4.9524 | 4.9404 | 4.9357 | 4.9562 | 4.9592 | 4.9657 | 4.9509 | 4.9545 | 4.9488 | 4.9465 | 4.9444 | 4.9543 | 4.9486 | 0.019407 | 0.0039217
982 | 4.8939 | 4.9933 | 4.9167 | 4.9383 | 4.9586 | 4.9575 | 4.9662 | 4.9794 | 4.9778 | 4.9168 | 4.9623 | 4.9592 | 4.938 | 4.9314 | 4.9377 | 4.9421 | 4.9353 | 4.9413 | 4.9523 | 4.9403 | 4.9356 | 4.9561 | 4.9591 | 4.9656 | 4.9508 | 4.9544 | 4.9487 | 4.9464 | 4.9443 | 4.9542 | 4.9484 | 0.019412 | 0.0039229
983 | 4.8938 | 4.9932 | 4.9166 | 4.9382 | 4.9585 | 4.9574 | 4.9661 | 4.9793 | 4.9777 | 4.9167 | 4.9622 | 4.9591 | 4.9379 | 4.9313 | 4.9376 | 4.942 | 4.9352 | 4.9412 | 4.9522 | 4.9402 | 4.9355 | 4.956 | 4.959 | 4.9655 | 4.9507 | 4.9543 | 4.9486 | 4.9463 | 4.9442 | 4.9541 | 4.9483 | 0.019418 | 0.0039241
984 | 4.8937 | 4.9931 | 4.9164 | 4.9381 | 4.9584 | 4.9573 | 4.966 | 4.9792 | 4.9776 | 4.9166 | 4.9621 | 4.959 | 4.9378 | 4.9312 | 4.9375 | 4.9419 | 4.9351 | 4.9411 | 4.9521 | 4.9401 | 4.9354 | 4.9559 | 4.9588 | 4.9654 | 4.9506 | 4.9542 | 4.9485 | 4.9462 | 4.944 | 4.954 | 4.9482 | 0.019423 | 0.0039252
985 | 4.8936 | 4.993 | 4.9163 | 4.9379 | 4.9583 | 4.9572 | 4.9659 | 4.979 | 4.9775 | 4.9165 | 4.9621 | 4.959 | 4.9377 | 4.9311 | 4.9374 | 4.9418 | 4.935 | 4.9409 | 4.952 | 4.94 | 4.9353 | 4.9558 | 4.9587 | 4.9653 | 4.9505 | 4.9541 | 4.9484 | 4.9461 | 4.9439 | 4.9539 | 4.9481 | 0.019428 | 0.0039264
986 | 4.8935 | 4.9929 | 4.9162 | 4.9378 | 4.9582 | 4.9571 | 4.9658 | 4.9789 | 4.9774 | 4.9163 | 4.962 | 4.9589 | 4.9376 | 4.931 | 4.9373 | 4.9417 | 4.9349 | 4.9408 | 4.9519 | 4.9399 | 4.9351 | 4.9557 | 4.9586 | 4.9652 | 4.9504 | 4.954 | 4.9483 | 4.9459 | 4.9438 | 4.9538 | 4.948 | 0.019434 | 0.0039276
987 | 4.8933 | 4.9928 | 4.9161 | 4.9377 | 4.9581 | 4.957 | 4.9657 | 4.9788 | 4.9773 | 4.9162 | 4.9619 | 4.9588 | 4.9375 | 4.9309 | 4.9372 | 4.9416 | 4.9348 | 4.9407 | 4.9519 | 4.9398 | 4.935 | 4.9555 | 4.9585 | 4.9651 | 4.9504 | 4.9539 | 4.9481 | 4.9458 | 4.9437 | 4.9537 | 4.9479 | 0.019439 | 0.0039287
988 | 4.8932 | 4.9927 | 4.916 | 4.9376 | 4.958 | 4.9569 | 4.9656 | 4.9787 | 4.9772 | 4.9161 | 4.9619 | 4.9587 | 4.9374 | 4.9308 | 4.9372 | 4.9415 | 4.9347 | 4.9406 | 4.9518 | 4.9397 | 4.9349 | 4.9554 | 4.9584 | 4.965 | 4.9503 | 4.9538 | 4.948 | 4.9457 | 4.9435 | 4.9536 | 4.9478 | 0.019444 | 0.0039298
989 | 4.8931 | 4.9926 | 4.9158 | 4.9375 | 4.9579 | 4.9568 | 4.9655 | 4.9786 | 4.9771 | 4.916 | 4.9618 | 4.9586 | 4.9374 | 4.9307 | 4.9371 | 4.9414 | 4.9346 | 4.9405 | 4.9517 | 4.9396 | 4.9348 | 4.9553 | 4.9583 | 4.9649 | 4.9502 | 4.9537 | 4.9479 | 4.9456 | 4.9434 | 4.9534 | 4.9477 | 0.019449 | 0.003931
990 | 4.893 | 4.9925 | 4.9157 | 4.9374 | 4.9578 | 4.9567 | 4.9654 | 4.9785 | 4.977 | 4.9159 | 4.9617 | 4.9585 | 4.9373 | 4.9306 | 4.937 | 4.9413 | 4.9345 | 4.9404 | 4.9516 | 4.9395 | 4.9347 | 4.9552 | 4.9582 | 4.9648 | 4.9501 | 4.9536 | 4.9478 | 4.9455 | 4.9433 | 4.9533 | 4.9476 | 0.019454 | 0.0039321
[31 empty] | 5.1324 | 0.016882 | 0.0033063

## Run1
MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR
File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Test_Run 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Test_Run 02 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run1\3 01 001.RAW
Cursor Location: |  |  | (179, 194) |  |  | Cursor Location: |  |  | (179, 194) |  |  | Cursor Location: |  |  | (179, 194)
Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15)
Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev
1 | 40.014 | 3.6892 | 7.7669 | 5.1374 | 0.34065 | 1 | 40.001 | 3.6889 | 7.7703 | 5.1695 | 0.32072 | 1 | 40.005 | 3.689 | 7.7319 | 5.1081 | 0.33037
2 | 80.028 | 4.3824 | 7.4368 |  |  | 2 | 80.001 | 4.382 | 7.4347 |  |  | 2 | 80.009 | 4.3821 | 7.399
3 | 120.042 | 4.7878 | 7.2347 |  |  | 3 | 120.002 | 4.7875 | 7.2405 |  |  | 3 | 120.014 | 4.7876 | 7.2057
4 | 160.057 | 5.0755 | 7.0923 |  |  | 4 | 160.003 | 5.0752 | 7.1008 |  |  | 4 | 160.019 | 5.0753 | 7.0664
5 | 200.071 | 5.2987 | 6.9828 |  |  | 5 | 200.003 | 5.2983 | 6.991 |  |  | 5 | 200.024 | 5.2984 | 6.9569
6 | 240.085 | 5.481 | 6.8938 |  |  | 6 | 240.004 | 5.4807 | 6.9003 |  |  | 6 | 240.028 | 5.4808 | 6.8664
7 | 280.099 | 5.6351 | 6.8185 |  |  | 7 | 280.005 | 5.6348 | 6.823 |  |  | 7 | 280.033 | 5.6349 | 6.7893
8 | 320.113 | 5.7687 | 6.7531 |  |  | 8 | 320.005 | 5.7683 | 6.7557 |  |  | 8 | 320.038 | 5.7684 | 6.7221
9 | 360.127 | 5.8865 | 6.6951 |  |  | 9 | 360.006 | 5.8861 | 6.6961 |  |  | 9 | 360.042 | 5.8862 | 6.6626
10 | 400.141 | 5.9918 | 6.643 |  |  | 10 | 400.007 | 5.9915 | 6.6426 |  |  | 10 | 400.047 | 5.9916 | 6.6091
11 | 440.155 | 6.0871 | 6.5956 |  |  | 11 | 440.007 | 6.0868 | 6.594 |  |  | 11 | 440.052 | 6.0869 | 6.5606
12 | 480.17 | 6.1741 | 6.552 |  |  | 12 | 480.008 | 6.1738 | 6.5497 |  |  | 12 | 480.057 | 6.1739 | 6.5163
13 | 520.184 | 6.2542 | 6.5116 |  |  | 13 | 520.008 | 6.2538 | 6.5088 |  |  | 13 | 520.061 | 6.2539 | 6.4754
14 | 560.198 | 6.3283 | 6.474 |  |  | 14 | 560.009 | 6.328 | 6.4709 |  |  | 14 | 560.066 | 6.3281 | 6.4376
15 | 600.212 | 6.3973 | 6.4388 |  |  | 15 | 600.01 | 6.3969 | 6.4357 |  |  | 15 | 600.071 | 6.397 | 6.4023
16 | 640.226 | 6.4618 | 6.4057 |  |  | 16 | 640.01 | 6.4615 | 6.4027 |  |  | 16 | 640.076 | 6.4616 | 6.3693
17 | 680.24 | 6.5224 | 6.3744 |  |  | 17 | 680.011 | 6.5221 | 6.3716 |  |  | 17 | 680.08 | 6.5222 | 6.3382
18 | 720.254 | 6.5796 | 6.3448 |  |  | 18 | 720.012 | 6.5793 | 6.3424 |  |  | 18 | 720.085 | 6.5794 | 6.309
19 | 760.268 | 6.6337 | 6.3167 |  |  | 19 | 760.012 | 6.6333 | 6.3148 |  |  | 19 | 760.09 | 6.6334 | 6.2813
20 | 800.283 | 6.685 | 6.2899 |  |  | 20 | 800.013 | 6.6846 | 6.2885 |  |  | 20 | 800.094 | 6.6847 | 6.255
21 | 840.297 | 6.7338 | 6.2643 |  |  | 21 | 840.014 | 6.7334 | 6.2636 |  |  | 21 | 840.099 | 6.7335 | 6.23
22 | 880.311 | 6.7803 | 6.2399 |  |  | 22 | 880.014 | 6.7799 | 6.2398 |  |  | 22 | 880.104 | 6.78 | 6.2062
23 | 920.325 | 6.8247 | 6.2165 |  |  | 23 | 920.015 | 6.8244 | 6.2171 |  |  | 23 | 920.109 | 6.8245 | 6.1834
24 | 960.339 | 6.8673 | 6.194 |  |  | 24 | 960.016 | 6.8669 | 6.1954 |  |  | 24 | 960.113 | 6.8671 | 6.1616
25 | 1000.353 | 6.9081 | 6.1724 |  |  | 25 | 1000.016 | 6.9078 | 6.1745 |  |  | 25 | 1000.118 | 6.9079 | 6.1407
26 | 1040.367 | 6.9473 | 6.1517 |  |  | 26 | 1040.017 | 6.947 | 6.1545 |  |  | 26 | 1040.123 | 6.9471 | 6.1206
27 | 1080.381 | 6.9851 | 6.1316 |  |  | 27 | 1080.018 | 6.9847 | 6.1353 |  |  | 27 | 1080.127 | 6.9848 | 6.1013
28 | 1120.396 | 7.0214 | 6.1124 |  |  | 28 | 1120.018 | 7.0211 | 6.1168 |  |  | 28 | 1120.132 | 7.0212 | 6.0827
29 | 1160.41 | 7.0565 | 6.0937 |  |  | 29 | 1160.019 | 7.0562 | 6.0989 |  |  | 29 | 1160.137 | 7.0563 | 6.0647
30 | 1200.424 | 7.0904 | 6.0758 |  |  | 30 | 1200.02 | 7.0901 | 6.0817 |  |  | 30 | 1200.142 | 7.0902 | 6.0474
31 | 1240.438 | 7.1232 | 6.0584 |  |  | 31 | 1240.02 | 7.1229 | 6.065 |  |  | 31 | 1240.146 | 7.123 | 6.0306
32 | 1280.452 | 7.155 | 6.0415 |  |  | 32 | 1280.021 | 7.1546 | 6.0489 |  |  | 32 | 1280.151 | 7.1547 | 6.0144
33 | 1320.466 | 7.1857 | 6.0252 |  |  | 33 | 1320.022 | 7.1854 | 6.0332 |  |  | 33 | 1320.156 | 7.1855 | 5.9987
34 | 1360.48 | 7.2156 | 6.0094 |  |  | 34 | 1360.022 | 7.2153 | 6.0181 |  |  | 34 | 1360.161 | 7.2154 | 5.9834
35 | 1400.494 | 7.2446 | 5.9941 |  |  | 35 | 1400.023 | 7.2442 | 6.0034 |  |  | 35 | 1400.165 | 7.2443 | 5.9686
36 | 1440.509 | 7.2728 | 5.9792 |  |  | 36 | 1440.024 | 7.2724 | 5.9891 |  |  | 36 | 1440.17 | 7.2725 | 5.9543
37 | 1480.523 | 7.3002 | 5.9648 |  |  | 37 | 1480.024 | 7.2998 | 5.9753 |  |  | 37 | 1480.175 | 7.2999 | 5.9403
38 | 1520.537 | 7.3268 | 5.9508 |  |  | 38 | 1520.025 | 7.3265 | 5.9618 |  |  | 38 | 1520.179 | 7.3266 | 5.9267
39 | 1560.551 | 7.3528 | 5.9371 |  |  | 39 | 1560.025 | 7.3525 | 5.9487 |  |  | 39 | 1560.184 | 7.3526 | 5.9135
40 | 1600.565 | 7.3781 | 5.9239 |  |  | 40 | 1600.026 | 7.3778 | 5.9359 |  |  | 40 | 1600.189 | 7.3779 | 5.9006
41 | 1640.579 | 7.4028 | 5.911 |  |  | 41 | 1640.027 | 7.4025 | 5.9234 |  |  | 41 | 1640.194 | 7.4026 | 5.888
42 | 1680.593 | 7.4269 | 5.8984 |  |  | 42 | 1680.027 | 7.4266 | 5.9113 |  |  | 42 | 1680.198 | 7.4267 | 5.8758
43 | 1720.608 | 7.4504 | 5.8862 |  |  | 43 | 1720.028 | 7.4501 | 5.8995 |  |  | 43 | 1720.203 | 7.4502 | 5.8638
44 | 1760.622 | 7.4734 | 5.8743 |  |  | 44 | 1760.029 | 7.4731 | 5.8879 |  |  | 44 | 1760.208 | 7.4732 | 5.8522
45 | 1800.636 | 7.4959 | 5.8627 |  |  | 45 | 1800.029 | 7.4956 | 5.8766 |  |  | 45 | 1800.212 | 7.4957 | 5.8408
46 | 1840.65 | 7.5179 | 5.8514 |  |  | 46 | 1840.03 | 7.5175 | 5.8656 |  |  | 46 | 1840.217 | 7.5176 | 5.8296
47 | 1880.664 | 7.5394 | 5.8403 |  |  | 47 | 1880.031 | 7.539 | 5.8548 |  |  | 47 | 1880.222 | 7.5391 | 5.8188
48 | 1920.678 | 7.5604 | 5.8296 |  |  | 48 | 1920.031 | 7.5601 | 5.8443 |  |  | 48 | 1920.227 | 7.5602 | 5.8081
49 | 1960.692 | 7.5811 | 5.819 |  |  | 49 | 1960.032 | 7.5807 | 5.834 |  |  | 49 | 1960.231 | 7.5808 | 5.7977
50 | 2000.706 | 7.6013 | 5.8088 |  |  | 50 | 2000.033 | 7.6009 | 5.8239 |  |  | 50 | 2000.236 | 7.601 | 5.7875
51 | 2040.721 | 7.6211 | 5.7988 |  |  | 51 | 2040.033 | 7.6207 | 5.814 |  |  | 51 | 2040.241 | 7.6208 | 5.7775
52 | 2080.735 | 7.6405 | 5.789 |  |  | 52 | 2080.034 | 7.6401 | 5.8044 |  |  | 52 | 2080.246 | 7.6402 | 5.7677
53 | 2120.749 | 7.6595 | 5.7794 |  |  | 53 | 2120.035 | 7.6592 | 5.7949 |  |  | 53 | 2120.25 | 7.6593 | 5.7581
54 | 2160.763 | 7.6782 | 5.77 |  |  | 54 | 2160.035 | 7.6779 | 5.7856 |  |  | 54 | 2160.255 | 7.678 | 5.7487
55 | 2200.777 | 7.6966 | 5.7609 |  |  | 55 | 2200.036 | 7.6962 | 5.7765 |  |  | 55 | 2200.26 | 7.6963 | 5.7395
56 | 2240.791 | 7.7146 | 5.752 |  |  | 56 | 2240.037 | 7.7142 | 5.7676 |  |  | 56 | 2240.264 | 7.7143 | 5.7305
57 | 2280.805 | 7.7323 | 5.7432 |  |  | 57 | 2280.037 | 7.7319 | 5.7588 |  |  | 57 | 2280.269 | 7.732 | 5.7216
58 | 2320.819 | 7.7497 | 5.7346 |  |  | 58 | 2320.038 | 7.7493 | 5.7503 |  |  | 58 | 2320.274 | 7.7494 | 5.7129
59 | 2360.834 | 7.7668 | 5.7263 |  |  | 59 | 2360.039 | 7.7664 | 5.7418 |  |  | 59 | 2360.279 | 7.7665 | 5.7044
60 | 2400.848 | 7.7836 | 5.7181 |  |  | 60 | 2400.039 | 7.7832 | 5.7336 |  |  | 60 | 2400.283 | 7.7833 | 5.696
61 | 2440.862 | 7.8001 | 5.71 |  |  | 61 | 2440.04 | 7.7998 | 5.7254 |  |  | 61 | 2440.288 | 7.7999 | 5.6877
62 | 2480.876 | 7.8164 | 5.7022 |  |  | 62 | 2480.04 | 7.816 | 5.7175 |  |  | 62 | 2480.293 | 7.8161 | 5.6797
63 | 2520.89 | 7.8324 | 5.6945 |  |  | 63 | 2520.041 | 7.832 | 5.7096 |  |  | 63 | 2520.297 | 7.8321 | 5.6717
64 | 2560.904 | 7.8481 | 5.6869 |  |  | 64 | 2560.042 | 7.8478 | 5.7019 |  |  | 64 | 2560.302 | 7.8479 | 5.6639
65 | 2600.918 | 7.8636 | 5.6795 |  |  | 65 | 2600.042 | 7.8633 | 5.6944 |  |  | 65 | 2600.307 | 7.8634 | 5.6562
66 | 2640.932 | 7.8789 | 5.6723 |  |  | 66 | 2640.043 | 7.8786 | 5.6869 |  |  | 66 | 2640.312 | 7.8787 | 5.6486
67 | 2680.947 | 7.8939 | 5.6651 |  |  | 67 | 2680.044 | 7.8936 | 5.6796 |  |  | 67 | 2680.316 | 7.8937 | 5.6412
68 | 2720.961 | 7.9087 | 5.6582 |  |  | 68 | 2720.044 | 7.9084 | 5.6724 |  |  | 68 | 2720.321 | 7.9085 | 5.6339
69 | 2760.975 | 7.9233 | 5.6513 |  |  | 69 | 2760.045 | 7.923 | 5.6654 |  |  | 69 | 2760.326 | 7.9231 | 5.6267
70 | 2800.989 | 7.9377 | 5.6446 |  |  | 70 | 2800.046 | 7.9374 | 5.6584 |  |  | 70 | 2800.331 | 7.9375 | 5.6196
71 | 2841.003 | 7.9519 | 5.638 |  |  | 71 | 2840.046 | 7.9516 | 5.6515 |  |  | 71 | 2840.335 | 7.9517 | 5.6127
72 | 2881.017 | 7.9659 | 5.6316 |  |  | 72 | 2880.047 | 7.9656 | 5.6448 |  |  | 72 | 2880.34 | 7.9657 | 5.6058
73 | 2921.031 | 7.9797 | 5.6252 |  |  | 73 | 2920.048 | 7.9794 | 5.6382 |  |  | 73 | 2920.345 | 7.9795 | 5.5991
74 | 2961.046 | 7.9933 | 5.619 |  |  | 74 | 2960.048 | 7.993 | 5.6316 |  |  | 74 | 2960.349 | 7.9931 | 5.5924
75 | 3001.06 | 8.0067 | 5.6129 |  |  | 75 | 3000.049 | 8.0064 | 5.6252 |  |  | 75 | 3000.354 | 8.0065 | 5.5859
76 | 3041.074 | 8.02 | 5.6069 |  |  | 76 | 3040.05 | 8.0196 | 5.6189 |  |  | 76 | 3040.359 | 8.0197 | 5.5794
77 | 3081.088 | 8.033 | 5.601 |  |  | 77 | 3080.05 | 8.0327 | 5.6127 |  |  | 77 | 3080.364 | 8.0328 | 5.5731
78 | 3121.102 | 8.0459 | 5.5952 |  |  | 78 | 3120.051 | 8.0456 | 5.6065 |  |  | 78 | 3120.368 | 8.0457 | 5.5668
79 | 3161.116 | 8.0587 | 5.5895 |  |  | 79 | 3160.052 | 8.0583 | 5.6005 |  |  | 79 | 3160.373 | 8.0584 | 5.5607
80 | 3201.13 | 8.0713 | 5.5839 |  |  | 80 | 3200.052 | 8.0709 | 5.5945 |  |  | 80 | 3200.378 | 8.071 | 5.5546
81 | 3241.144 | 8.0837 | 5.5784 |  |  | 81 | 3240.053 | 8.0833 | 5.5886 |  |  | 81 | 3240.382 | 8.0834 | 5.5486
82 | 3281.159 | 8.096 | 5.573 |  |  | 82 | 3280.054 | 8.0956 | 5.5828 |  |  | 82 | 3280.387 | 8.0957 | 5.5427
83 | 3321.173 | 8.1081 | 5.5677 |  |  | 83 | 3320.054 | 8.1077 | 5.5771 |  |  | 83 | 3320.392 | 8.1078 | 5.5369
84 | 3361.187 | 8.12 | 5.5625 |  |  | 84 | 3360.055 | 8.1197 | 5.5715 |  |  | 84 | 3360.397 | 8.1198 | 5.5311
85 | 3401.201 | 8.1319 | 5.5573 |  |  | 85 | 3400.056 | 8.1315 | 5.566 |  |  | 85 | 3400.401 | 8.1316 | 5.5255
86 | 3441.215 | 8.1436 | 5.5523 |  |  | 86 | 3440.056 | 8.1432 | 5.5605 |  |  | 86 | 3440.406 | 8.1433 | 5.5199
87 | 3481.229 | 8.1551 | 5.5473 |  |  | 87 | 3480.057 | 8.1548 | 5.5551 |  |  | 87 | 3480.411 | 8.1549 | 5.5144
88 | 3521.243 | 8.1666 | 5.5424 |  |  | 88 | 3520.057 | 8.1662 | 5.5498 |  |  | 88 | 3520.416 | 8.1663 | 5.509
89 | 3561.257 | 8.1779 | 5.5376 |  |  | 89 | 3560.058 | 8.1775 | 5.5445 |  |  | 89 | 3560.42 | 8.1776 | 5.5036
90 | 3601.272 | 8.189 | 5.5329 |  |  | 90 | 3600.059 | 8.1887 | 5.5394 |  |  | 90 | 3600.425 | 8.1888 | 5.4983
91 | 3641.286 | 8.2001 | 5.5282 |  |  | 91 | 3640.059 | 8.1998 | 5.5343 |  |  | 91 | 3640.43 | 8.1999 | 5.4931
92 | 3681.3 | 8.211 | 5.5236 |  |  | 92 | 3680.06 | 8.2107 | 5.5292 |  |  | 92 | 3680.434 | 8.2108 | 5.488
93 | 3721.314 | 8.2218 | 5.5191 |  |  | 93 | 3720.061 | 8.2215 | 5.5243 |  |  | 93 | 3720.439 | 8.2216 | 5.4829
94 | 3761.328 | 8.2325 | 5.5146 |  |  | 94 | 3760.061 | 8.2322 | 5.5193 |  |  | 94 | 3760.444 | 8.2323 | 5.4779
95 | 3801.342 | 8.2431 | 5.5103 |  |  | 95 | 3800.062 | 8.2428 | 5.5145 |  |  | 95 | 3800.449 | 8.2429 | 5.4729
96 | 3841.356 | 8.2536 | 5.5059 |  |  | 96 | 3840.063 | 8.2532 | 5.5097 |  |  | 96 | 3840.453 | 8.2533 | 5.468
97 | 3881.37 | 8.2639 | 5.5017 |  |  | 97 | 3880.063 | 8.2636 | 5.505 |  |  | 97 | 3880.458 | 8.2637 | 5.4632
98 | 3921.385 | 8.2742 | 5.4975 |  |  | 98 | 3920.064 | 8.2739 | 5.5004 |  |  | 98 | 3920.463 | 8.274 | 5.4584
99 | 3961.399 | 8.2844 | 5.4934 |  |  | 99 | 3960.065 | 8.284 | 5.4958 |  |  | 99 | 3960.467 | 8.2841 | 5.4537
100 | 4001.413 | 8.2944 | 5.4893 |  |  | 100 | 4000.065 | 8.2941 | 5.4912 |  |  | 100 | 4000.472 | 8.2942 | 5.4491
101 | 4041.427 | 8.3044 | 5.4853 |  |  | 101 | 4040.066 | 8.304 | 5.4868 |  |  | 101 | 4040.477 | 8.3041 | 5.4445
102 | 4081.441 | 8.3142 | 5.4814 |  |  | 102 | 4080.067 | 8.3139 | 5.4823 |  |  | 102 | 4080.482 | 8.314 | 5.44
103 | 4121.455 | 8.324 | 5.4775 |  |  | 103 | 4120.067 | 8.3236 | 5.478 |  |  | 103 | 4120.486 | 8.3237 | 5.4355
104 | 4161.469 | 8.3336 | 5.4737 |  |  | 104 | 4160.068 | 8.3333 | 5.4737 |  |  | 104 | 4160.491 | 8.3334 | 5.4311
105 | 4201.483 | 8.3432 | 5.4699 |  |  | 105 | 4200.069 | 8.3429 | 5.4694 |  |  | 105 | 4200.496 | 8.343 | 5.4267
106 | 4241.498 | 8.3527 | 5.4661 |  |  | 106 | 4240.069 | 8.3523 | 5.4652 |  |  | 106 | 4240.5 | 8.3524 | 5.4224
107 | 4281.512 | 8.3621 | 5.4625 |  |  | 107 | 4280.07 | 8.3617 | 5.461 |  |  | 107 | 4280.505 | 8.3618 | 5.4181
108 | 4321.526 | 8.3714 | 5.4588 |  |  | 108 | 4320.071 | 8.371 | 5.4569 |  |  | 108 | 4320.51 | 8.3711 | 5.4139
109 | 4361.54 | 8.3806 | 5.4553 |  |  | 109 | 4360.071 | 8.3802 | 5.4529 |  |  | 109 | 4360.515 | 8.3803 | 5.4098
110 | 4401.554 | 8.3897 | 5.4517 |  |  | 110 | 4400.072 | 8.3894 | 5.4489 |  |  | 110 | 4400.519 | 8.3895 | 5.4057
111 | 4441.568 | 8.3988 | 5.4483 |  |  | 111 | 4440.072 | 8.3984 | 5.4449 |  |  | 111 | 4440.524 | 8.3985 | 5.4016
112 | 4481.582 | 8.4077 | 5.4448 |  |  | 112 | 4480.073 | 8.4074 | 5.441 |  |  | 112 | 4480.529 | 8.4075 | 5.3976
113 | 4521.597 | 8.4166 | 5.4414 |  |  | 113 | 4520.074 | 8.4163 | 5.4371 |  |  | 113 | 4520.534 | 8.4164 | 5.3936
114 | 4561.611 | 8.4254 | 5.4381 |  |  | 114 | 4560.074 | 8.4251 | 5.4333 |  |  | 114 | 4560.538 | 8.4252 | 5.3897
115 | 4601.625 | 8.4342 | 5.4348 |  |  | 115 | 4600.075 | 8.4338 | 5.4295 |  |  | 115 | 4600.543 | 8.4339 | 5.3858
116 | 4641.639 | 8.4428 | 5.4315 |  |  | 116 | 4640.076 | 8.4425 | 5.4258 |  |  | 116 | 4640.548 | 8.4426 | 5.382
117 | 4681.653 | 8.4514 | 5.4283 |  |  | 117 | 4680.076 | 8.4511 | 5.4221 |  |  | 117 | 4680.552 | 8.4512 | 5.3782
118 | 4721.667 | 8.4599 | 5.4251 |  |  | 118 | 4720.077 | 8.4596 | 5.4185 |  |  | 118 | 4720.557 | 8.4597 | 5.3744
119 | 4761.681 | 8.4684 | 5.422 |  |  | 119 | 4760.078 | 8.468 | 5.4149 |  |  | 119 | 4760.562 | 8.4681 | 5.3707
120 | 4801.695 | 8.4767 | 5.4189 |  |  | 120 | 4800.078 | 8.4764 | 5.4113 |  |  | 120 | 4800.567 | 8.4765 | 5.3671
121 | 4841.71 | 8.485 | 5.4159 |  |  | 121 | 4840.079 | 8.4847 | 5.4078 |  |  | 121 | 4840.571 | 8.4848 | 5.3635
122 | 4881.724 | 8.4933 | 5.4128 |  |  | 122 | 4880.08 | 8.4929 | 5.4043 |  |  | 122 | 4880.576 | 8.493 | 5.3599
123 | 4921.738 | 8.5014 | 5.4098 |  |  | 123 | 4920.08 | 8.5011 | 5.4009 |  |  | 123 | 4920.581 | 8.5012 | 5.3563
124 | 4961.752 | 8.5095 | 5.4069 |  |  | 124 | 4960.081 | 8.5092 | 5.3975 |  |  | 124 | 4960.585 | 8.5093 | 5.3528
125 | 5001.766 | 8.5175 | 5.404 |  |  | 125 | 5000.082 | 8.5172 | 5.3941 |  |  | 125 | 5000.59 | 8.5173 | 5.3494
126 | 5041.78 | 8.5255 | 5.4011 |  |  | 126 | 5040.082 | 8.5252 | 5.3908 |  |  | 126 | 5040.595 | 8.5253 | 5.3459
127 | 5081.794 | 8.5334 | 5.3983 |  |  | 127 | 5080.083 | 8.5331 | 5.3875 |  |  | 127 | 5080.6 | 8.5332 | 5.3426
128 | 5121.808 | 8.5413 | 5.3955 |  |  | 128 | 5120.084 | 8.5409 | 5.3842 |  |  | 128 | 5120.604 | 8.541 | 5.3392
129 | 5161.823 | 8.549 | 5.3927 |  |  | 129 | 5160.084 | 8.5487 | 5.381 |  |  | 129 | 5160.609 | 8.5488 | 5.3359
130 | 5201.837 | 8.5568 | 5.3899 |  |  | 130 | 5200.085 | 8.5564 | 5.3778 |  |  | 130 | 5200.614 | 8.5565 | 5.3326
131 | 5241.851 | 8.5644 | 5.3872 |  |  | 131 | 5240.086 | 8.5641 | 5.3747 |  |  | 131 | 5240.619 | 8.5642 | 5.3294
132 | 5281.865 | 8.572 | 5.3846 |  |  | 132 | 5280.086 | 8.5717 | 5.3716 |  |  | 132 | 5280.623 | 8.5718 | 5.3262
133 | 5321.879 | 8.5796 | 5.3819 |  |  | 133 | 5320.087 | 8.5792 | 5.3685 |  |  | 133 | 5320.628 | 8.5793 | 5.323
134 | 5361.893 | 8.5871 | 5.3793 |  |  | 134 | 5360.088 | 8.5867 | 5.3654 |  |  | 134 | 5360.633 | 8.5868 | 5.3198
135 | 5401.907 | 8.5945 | 5.3767 |  |  | 135 | 5400.088 | 8.5942 | 5.3624 |  |  | 135 | 5400.637 | 8.5943 | 5.3167
136 | 5441.921 | 8.6019 | 5.3741 |  |  | 136 | 5440.089 | 8.6016 | 5.3595 |  |  | 136 | 5440.642 | 8.6017 | 5.3137
137 | 5481.936 | 8.6092 | 5.3716 |  |  | 137 | 5480.089 | 8.6089 | 5.3565 |  |  | 137 | 5480.647 | 8.609 | 5.3106
138 | 5521.95 | 8.6165 | 5.3691 |  |  | 138 | 5520.09 | 8.6161 | 5.3536 |  |  | 138 | 5520.652 | 8.6163 | 5.3076
139 | 5561.964 | 8.6237 | 5.3666 |  |  | 139 | 5560.091 | 8.6234 | 5.3507 |  |  | 139 | 5560.656 | 8.6235 | 5.3046
140 | 5601.978 | 8.6309 | 5.3642 |  |  | 140 | 5600.091 | 8.6305 | 5.3478 |  |  | 140 | 5600.661 | 8.6306 | 5.3017
141 | 5641.992 | 8.638 | 5.3617 |  |  | 141 | 5640.092 | 8.6377 | 5.345 |  |  | 141 | 5640.666 | 8.6378 | 5.2988
142 | 5682.006 | 8.6451 | 5.3593 |  |  | 142 | 5680.093 | 8.6447 | 5.3422 |  |  | 142 | 5680.67 | 8.6448 | 5.2959
143 | 5722.02 | 8.6521 | 5.357 |  |  | 143 | 5720.093 | 8.6517 | 5.3395 |  |  | 143 | 5720.675 | 8.6518 | 5.293
144 | 5762.035 | 8.659 | 5.3546 |  |  | 144 | 5760.094 | 8.6587 | 5.3367 |  |  | 144 | 5760.68 | 8.6588 | 5.2902
145 | 5802.049 | 8.666 | 5.3523 |  |  | 145 | 5800.095 | 8.6656 | 5.334 |  |  | 145 | 5800.685 | 8.6657 | 5.2874
146 | 5842.063 | 8.6728 | 5.35 |  |  | 146 | 5840.095 | 8.6725 | 5.3313 |  |  | 146 | 5840.689 | 8.6726 | 5.2846
147 | 5882.077 | 8.6797 | 5.3477 |  |  | 147 | 5880.096 | 8.6793 | 5.3287 |  |  | 147 | 5880.694 | 8.6794 | 5.2819
148 | 5922.091 | 8.6864 | 5.3455 |  |  | 148 | 5920.097 | 8.6861 | 5.3261 |  |  | 148 | 5920.699 | 8.6862 | 5.2792
149 | 5962.105 | 8.6932 | 5.3432 |  |  | 149 | 5960.097 | 8.6928 | 5.3235 |  |  | 149 | 5960.704 | 8.6929 | 5.2765
150 | 6002.119 | 8.6999 | 5.341 |  |  | 150 | 6000.098 | 8.6995 | 5.3209 |  |  | 150 | 6000.708 | 8.6996 | 5.2738
151 | 6042.133 | 8.7065 | 5.3388 |  |  | 151 | 6040.099 | 8.7062 | 5.3183 |  |  | 151 | 6040.713 | 8.7063 | 5.2712
152 | 6082.148 | 8.7131 | 5.3367 |  |  | 152 | 6080.099 | 8.7128 | 5.3158 |  |  | 152 | 6080.718 | 8.7129 | 5.2686
153 | 6122.162 | 8.7197 | 5.3345 |  |  | 153 | 6120.1 | 8.7193 | 5.3133 |  |  | 153 | 6120.722 | 8.7194 | 5.266
154 | 6162.176 | 8.7262 | 5.3324 |  |  | 154 | 6160.101 | 8.7258 | 5.3109 |  |  | 154 | 6160.727 | 8.726 | 5.2634
155 | 6202.19 | 8.7327 | 5.3303 |  |  | 155 | 6200.101 | 8.7323 | 5.3084 |  |  | 155 | 6200.732 | 8.7324 | 5.2609
156 | 6242.204 | 8.7391 | 5.3282 |  |  | 156 | 6240.102 | 8.7388 | 5.306 |  |  | 156 | 6240.737 | 8.7389 | 5.2584
157 | 6282.218 | 8.7455 | 5.3261 |  |  | 157 | 6280.103 | 8.7451 | 5.3036 |  |  | 157 | 6280.741 | 8.7452 | 5.2559
158 | 6322.232 | 8.7518 | 5.3241 |  |  | 158 | 6320.103 | 8.7515 | 5.3012 |  |  | 158 | 6320.746 | 8.7516 | 5.2535
159 | 6362.246 | 8.7581 | 5.3221 |  |  | 159 | 6360.104 | 8.7578 | 5.2989 |  |  | 159 | 6360.751 | 8.7579 | 5.251
160 | 6402.261 | 8.7644 | 5.3201 |  |  | 160 | 6400.104 | 8.7641 | 5.2966 |  |  | 160 | 6400.755 | 8.7642 | 5.2486
161 | 6442.275 | 8.7706 | 5.3181 |  |  | 161 | 6440.105 | 8.7703 | 5.2943 |  |  | 161 | 6440.76 | 8.7704 | 5.2462
162 | 6482.289 | 8.7768 | 5.3161 |  |  | 162 | 6480.106 | 8.7765 | 5.292 |  |  | 162 | 6480.765 | 8.7766 | 5.2439
163 | 6522.303 | 8.783 | 5.3142 |  |  | 163 | 6520.106 | 8.7826 | 5.2898 |  |  | 163 | 6520.77 | 8.7827 | 5.2415
164 | 6562.317 | 8.7891 | 5.3122 |  |  | 164 | 6560.107 | 8.7888 | 5.2875 |  |  | 164 | 6560.774 | 8.7889 | 5.2392
165 | 6602.331 | 8.7952 | 5.3103 |  |  | 165 | 6600.108 | 8.7948 | 5.2853 |  |  | 165 | 6600.779 | 8.7949 | 5.2369
166 | 6642.345 | 8.8012 | 5.3084 |  |  | 166 | 6640.108 | 8.8009 | 5.2831 |  |  | 166 | 6640.784 | 8.801 | 5.2347
167 | 6682.359 | 8.8072 | 5.3065 |  |  | 167 | 6680.109 | 8.8069 | 5.281 |  |  | 167 | 6680.789 | 8.807 | 5.2324
168 | 6722.374 | 8.8132 | 5.3047 |  |  | 168 | 6720.11 | 8.8129 | 5.2788 |  |  | 168 | 6720.793 | 8.813 | 5.2302
169 | 6762.388 | 8.8191 | 5.3028 |  |  | 169 | 6760.11 | 8.8188 | 5.2767 |  |  | 169 | 6760.798 | 8.8189 | 5.228
170 | 6802.402 | 8.825 | 5.301 |  |  | 170 | 6800.111 | 8.8247 | 5.2746 |  |  | 170 | 6800.803 | 8.8248 | 5.2258
171 | 6842.416 | 8.8309 | 5.2992 |  |  | 171 | 6840.112 | 8.8306 | 5.2725 |  |  | 171 | 6840.807 | 8.8307 | 5.2236
172 | 6882.43 | 8.8367 | 5.2974 |  |  | 172 | 6880.112 | 8.8364 | 5.2704 |  |  | 172 | 6880.812 | 8.8365 | 5.2215
173 | 6922.444 | 8.8425 | 5.2956 |  |  | 173 | 6920.113 | 8.8422 | 5.2684 |  |  | 173 | 6920.817 | 8.8423 | 5.2194
174 | 6962.458 | 8.8483 | 5.2938 |  |  | 174 | 6960.114 | 8.848 | 5.2664 |  |  | 174 | 6960.822 | 8.8481 | 5.2173
175 | 7002.472 | 8.854 | 5.292 |  |  | 175 | 7000.114 | 8.8537 | 5.2644 |  |  | 175 | 7000.826 | 8.8538 | 5.2152
176 | 7042.487 | 8.8597 | 5.2903 |  |  | 176 | 7040.115 | 8.8594 | 5.2624 |  |  | 176 | 7040.831 | 8.8595 | 5.2131
177 | 7082.501 | 8.8654 | 5.2886 |  |  | 177 | 7080.116 | 8.865 | 5.2604 |  |  | 177 | 7080.836 | 8.8651 | 5.2111
178 | 7122.515 | 8.871 | 5.2869 |  |  | 178 | 7120.116 | 8.8707 | 5.2585 |  |  | 178 | 7120.84 | 8.8708 | 5.2091
179 | 7162.529 | 8.8766 | 5.2852 |  |  | 179 | 7160.117 | 8.8763 | 5.2566 |  |  | 179 | 7160.845 | 8.8764 | 5.207
180 | 7202.543 | 8.8822 | 5.2835 |  |  | 180 | 7200.118 | 8.8819 | 5.2546 |  |  | 180 | 7200.85 | 8.882 | 5.2051
181 | 7242.557 | 8.8877 | 5.2818 |  |  | 181 | 7240.118 | 8.8874 | 5.2528 |  |  | 181 | 7240.855 | 8.8875 | 5.2031
182 | 7282.571 | 8.8932 | 5.2801 |  |  | 182 | 7280.119 | 8.8929 | 5.2509 |  |  | 182 | 7280.859 | 8.893 | 5.2011
183 | 7322.586 | 8.8987 | 5.2785 |  |  | 183 | 7320.12 | 8.8984 | 5.249 |  |  | 183 | 7320.864 | 8.8985 | 5.1992
184 | 7362.6 | 8.9042 | 5.2769 |  |  | 184 | 7360.12 | 8.9038 | 5.2472 |  |  | 184 | 7360.869 | 8.9039 | 5.1973
185 | 7402.614 | 8.9096 | 5.2752 |  |  | 185 | 7400.121 | 8.9093 | 5.2454 |  |  | 185 | 7400.874 | 8.9094 | 5.1954
186 | 7442.628 | 8.915 | 5.2736 |  |  | 186 | 7440.121 | 8.9146 | 5.2436 |  |  | 186 | 7440.878 | 8.9147 | 5.1935
187 | 7482.642 | 8.9203 | 5.272 |  |  | 187 | 7480.122 | 8.92 | 5.2418 |  |  | 187 | 7480.883 | 8.9201 | 5.1917
188 | 7522.656 | 8.9257 | 5.2705 |  |  | 188 | 7520.123 | 8.9253 | 5.24 |  |  | 188 | 7520.888 | 8.9254 | 5.1898
189 | 7562.67 | 8.931 | 5.2689 |  |  | 189 | 7560.123 | 8.9306 | 5.2383 |  |  | 189 | 7560.892 | 8.9307 | 5.188
190 | 7602.684 | 8.9363 | 5.2673 |  |  | 190 | 7600.124 | 8.9359 | 5.2365 |  |  | 190 | 7600.897 | 8.936 | 5.1862
191 | 7642.699 | 8.9415 | 5.2658 |  |  | 191 | 7640.125 | 8.9412 | 5.2348 |  |  | 191 | 7640.902 | 8.9413 | 5.1844
192 | 7682.713 | 8.9467 | 5.2643 |  |  | 192 | 7680.125 | 8.9464 | 5.2331 |  |  | 192 | 7680.907 | 8.9465 | 5.1826
193 | 7722.727 | 8.9519 | 5.2627 |  |  | 193 | 7720.126 | 8.9516 | 5.2314 |  |  | 193 | 7720.911 | 8.9517 | 5.1808
194 | 7762.741 | 8.9571 | 5.2612 |  |  | 194 | 7760.127 | 8.9568 | 5.2297 |  |  | 194 | 7760.916 | 8.9569 | 5.1791
195 | 7802.755 | 8.9622 | 5.2597 |  |  | 195 | 7800.127 | 8.9619 | 5.2281 |  |  | 195 | 7800.921 | 8.962 | 5.1774
196 | 7842.769 | 8.9673 | 5.2582 |  |  | 196 | 7840.128 | 8.967 | 5.2264 |  |  | 196 | 7840.925 | 8.9671 | 5.1756
197 | 7882.783 | 8.9724 | 5.2567 |  |  | 197 | 7880.129 | 8.9721 | 5.2248 |  |  | 197 | 7880.93 | 8.9722 | 5.1739
198 | 7922.797 | 8.9775 | 5.2553 |  |  | 198 | 7920.129 | 8.9772 | 5.2232 |  |  | 198 | 7920.935 | 8.9773 | 5.1723
199 | 7962.812 | 8.9825 | 5.2538 |  |  | 199 | 7960.13 | 8.9822 | 5.2216 |  |  | 199 | 7960.94 | 8.9823 | 5.1706
200 | 8002.826 | 8.9876 | 5.2524 |  |  | 200 | 8000.131 | 8.9872 | 5.22 |  |  | 200 | 8000.944 | 8.9873 | 5.1689
201 | 8042.84 | 8.9925 | 5.2509 |  |  | 201 | 8040.131 | 8.9922 | 5.2185 |  |  | 201 | 8040.949 | 8.9923 | 5.1673
202 | 8082.854 | 8.9975 | 5.2495 |  |  | 202 | 8080.132 | 8.9972 | 5.2169 |  |  | 202 | 8080.954 | 8.9973 | 5.1657
203 | 8122.868 | 9.0024 | 5.2481 |  |  | 203 | 8120.133 | 9.0021 | 5.2154 |  |  | 203 | 8120.958 | 9.0022 | 5.1641
204 | 8162.882 | 9.0074 | 5.2467 |  |  | 204 | 8160.133 | 9.007 | 5.2138 |  |  | 204 | 8160.963 | 9.0071 | 5.1625
205 | 8202.896 | 9.0122 | 5.2453 |  |  | 205 | 8200.134 | 9.0119 | 5.2123 |  |  | 205 | 8200.968 | 9.012 | 5.1609
206 | 8242.91 | 9.0171 | 5.2439 |  |  | 206 | 8240.135 | 9.0168 | 5.2108 |  |  | 206 | 8240.973 | 9.0169 | 5.1593
207 | 8282.925 | 9.022 | 5.2425 |  |  | 207 | 8280.135 | 9.0216 | 5.2093 |  |  | 207 | 8280.977 | 9.0217 | 5.1578
208 | 8322.939 | 9.0268 | 5.2411 |  |  | 208 | 8320.136 | 9.0264 | 5.2079 |  |  | 208 | 8320.982 | 9.0265 | 5.1562
209 | 8362.953 | 9.0316 | 5.2398 |  |  | 209 | 8360.136 | 9.0312 | 5.2064 |  |  | 209 | 8360.987 | 9.0313 | 5.1547
210 | 8402.967 | 9.0363 | 5.2384 |  |  | 210 | 8400.137 | 9.036 | 5.205 |  |  | 210 | 8400.992 | 9.0361 | 5.1532
211 | 8442.981 | 9.0411 | 5.2371 |  |  | 211 | 8440.138 | 9.0408 | 5.2035 |  |  | 211 | 8440.996 | 9.0409 | 5.1517
212 | 8482.995 | 9.0458 | 5.2357 |  |  | 212 | 8480.138 | 9.0455 | 5.2021 |  |  | 212 | 8481.001 | 9.0456 | 5.1502
213 | 8523.009 | 9.0505 | 5.2344 |  |  | 213 | 8520.139 | 9.0502 | 5.2007 |  |  | 213 | 8521.006 | 9.0503 | 5.1487
214 | 8563.024 | 9.0552 | 5.2331 |  |  | 214 | 8560.14 | 9.0549 | 5.1993 |  |  | 214 | 8561.01 | 9.055 | 5.1472
215 | 8603.038 | 9.0599 | 5.2318 |  |  | 215 | 8600.14 | 9.0595 | 5.1979 |  |  | 215 | 8601.015 | 9.0596 | 5.1458
216 | 8643.052 | 9.0645 | 5.2305 |  |  | 216 | 8640.141 | 9.0642 | 5.1966 |  |  | 216 | 8641.02 | 9.0643 | 5.1444
217 | 8683.066 | 9.0691 | 5.2292 |  |  | 217 | 8680.142 | 9.0688 | 5.1952 |  |  | 217 | 8681.025 | 9.0689 | 5.1429
218 | 8723.08 | 9.0737 | 5.2279 |  |  | 218 | 8720.142 | 9.0734 | 5.1939 |  |  | 218 | 8721.029 | 9.0735 | 5.1415
219 | 8763.094 | 9.0783 | 5.2266 |  |  | 219 | 8760.143 | 9.078 | 5.1925 |  |  | 219 | 8761.034 | 9.0781 | 5.1401
220 | 8803.108 | 9.0829 | 5.2253 |  |  | 220 | 8800.144 | 9.0825 | 5.1912 |  |  | 220 | 8801.039 | 9.0826 | 5.1387
221 | 8843.122 | 9.0874 | 5.2241 |  |  | 221 | 8840.144 | 9.0871 | 5.1899 |  |  | 221 | 8841.043 | 9.0872 | 5.1374
222 | 8883.137 | 9.0919 | 5.2228 |  |  | 222 | 8880.145 | 9.0916 | 5.1886 |  |  | 222 | 8881.048 | 9.0917 | 5.136
223 | 8923.151 | 9.0964 | 5.2216 |  |  | 223 | 8920.146 | 9.0961 | 5.1873 |  |  | 223 | 8921.053 | 9.0962 | 5.1347
224 | 8963.165 | 9.1009 | 5.2203 |  |  | 224 | 8960.146 | 9.1005 | 5.1861 |  |  | 224 | 8961.058 | 9.1006 | 5.1333
225 | 9003.179 | 9.1053 | 5.2191 |  |  | 225 | 9000.147 | 9.105 | 5.1848 |  |  | 225 | 9001.062 | 9.1051 | 5.132
226 | 9043.193 | 9.1098 | 5.2179 |  |  | 226 | 9040.148 | 9.1094 | 5.1835 |  |  | 226 | 9041.067 | 9.1095 | 5.1307
227 | 9083.207 | 9.1142 | 5.2167 |  |  | 227 | 9080.148 | 9.1138 | 5.1823 |  |  | 227 | 9081.072 | 9.1139 | 5.1294
228 | 9123.221 | 9.1186 | 5.2154 |  |  | 228 | 9120.149 | 9.1182 | 5.1811 |  |  | 228 | 9121.077 | 9.1183 | 5.1281
229 | 9163.235 | 9.123 | 5.2142 |  |  | 229 | 9160.15 | 9.1226 | 5.1798 |  |  | 229 | 9161.081 | 9.1227 | 5.1268
230 | 9203.25 | 9.1273 | 5.213 |  |  | 230 | 9200.15 | 9.127 | 5.1786 |  |  | 230 | 9201.086 | 9.1271 | 5.1255
231 | 9243.264 | 9.1317 | 5.2119 |  |  | 231 | 9240.151 | 9.1313 | 5.1774 |  |  | 231 | 9241.091 | 9.1314 | 5.1242
232 | 9283.278 | 9.136 | 5.2107 |  |  | 232 | 9280.152 | 9.1356 | 5.1763 |  |  | 232 | 9281.095 | 9.1357 | 5.123
233 | 9323.292 | 9.1403 | 5.2095 |  |  | 233 | 9320.152 | 9.1399 | 5.1751 |  |  | 233 | 9321.1 | 9.14 | 5.1218
234 | 9363.306 | 9.1446 | 5.2083 |  |  | 234 | 9360.153 | 9.1442 | 5.1739 |  |  | 234 | 9361.105 | 9.1443 | 5.1205
235 | 9403.32 | 9.1488 | 5.2072 |  |  | 235 | 9400.153 | 9.1485 | 5.1728 |  |  | 235 | 9401.11 | 9.1486 | 5.1193
236 | 9443.334 | 9.1531 | 5.206 |  |  | 236 | 9440.154 | 9.1527 | 5.1716 |  |  | 236 | 9441.114 | 9.1528 | 5.1181
237 | 9483.348 | 9.1573 | 5.2049 |  |  | 237 | 9480.155 | 9.157 | 5.1705 |  |  | 237 | 9481.119 | 9.1571 | 5.1169
238 | 9523.363 | 9.1615 | 5.2037 |  |  | 238 | 9520.155 | 9.1612 | 5.1693 |  |  | 238 | 9521.124 | 9.1613 | 5.1157
239 | 9563.377 | 9.1657 | 5.2026 |  |  | 239 | 9560.156 | 9.1654 | 5.1682 |  |  | 239 | 9561.128 | 9.1655 | 5.1145
240 | 9603.391 | 9.1699 | 5.2014 |  |  | 240 | 9600.157 | 9.1695 | 5.1671 |  |  | 240 | 9601.133 | 9.1696 | 5.1134
241 | 9643.405 | 9.174 | 5.2003 |  |  | 241 | 9640.157 | 9.1737 | 5.166 |  |  | 241 | 9641.138 | 9.1738 | 5.1122
242 | 9683.419 | 9.1782 | 5.1992 |  |  | 242 | 9680.158 | 9.1778 | 5.1649 |  |  | 242 | 9681.143 | 9.1779 | 5.111
243 | 9723.433 | 9.1823 | 5.1981 |  |  | 243 | 9720.159 | 9.182 | 5.1638 |  |  | 243 | 9721.147 | 9.1821 | 5.1099
244 | 9763.447 | 9.1864 | 5.197 |  |  | 244 | 9760.159 | 9.1861 | 5.1628 |  |  | 244 | 9761.152 | 9.1862 | 5.1088
245 | 9803.461 | 9.1905 | 5.1959 |  |  | 245 | 9800.16 | 9.1902 | 5.1617 |  |  | 245 | 9801.157 | 9.1903 | 5.1076
246 | 9843.476 | 9.1946 | 5.1948 |  |  | 246 | 9840.161 | 9.1942 | 5.1607 |  |  | 246 | 9841.162 | 9.1943 | 5.1065
247 | 9883.49 | 9.1986 | 5.1937 |  |  | 247 | 9880.161 | 9.1983 | 5.1596 |  |  | 247 | 9881.166 | 9.1984 | 5.1054
248 | 9923.504 | 9.2027 | 5.1926 |  |  | 248 | 9920.162 | 9.2023 | 5.1586 |  |  | 248 | 9921.171 | 9.2024 | 5.1043
249 | 9963.518 | 9.2067 | 5.1915 |  |  | 249 | 9960.163 | 9.2063 | 5.1576 |  |  | 249 | 9961.176 | 9.2065 | 5.1032
250 | 10003.532 | 9.2107 | 5.1905 |  |  | 250 | 10000.163 | 9.2104 | 5.1565 |  |  | 250 | 10001.18 | 9.2105 | 5.1022
251 | 10043.546 | 9.2147 | 5.1894 |  |  | 251 | 10040.164 | 9.2143 | 5.1555 |  |  | 251 | 10041.185 | 9.2145 | 5.1011
252 | 10083.56 | 9.2187 | 5.1883 |  |  | 252 | 10080.165 | 9.2183 | 5.1545 |  |  | 252 | 10081.19 | 9.2184 | 5.1
253 | 10123.575 | 9.2226 | 5.1873 |  |  | 253 | 10120.165 | 9.2223 | 5.1535 |  |  | 253 | 10121.195 | 9.2224 | 5.099
254 | 10163.589 | 9.2266 | 5.1862 |  |  | 254 | 10160.166 | 9.2262 | 5.1525 |  |  | 254 | 10161.199 | 9.2263 | 5.098
255 | 10203.603 | 9.2305 | 5.1852 |  |  | 255 | 10200.167 | 9.2302 | 5.1516 |  |  | 255 | 10201.204 | 9.2303 | 5.0969
256 | 10243.617 | 9.2344 | 5.1842 |  |  | 256 | 10240.167 | 9.2341 | 5.1506 |  |  | 256 | 10241.209 | 9.2342 | 5.0959
257 | 10283.631 | 9.2383 | 5.1831 |  |  | 257 | 10280.168 | 9.238 | 5.1496 |  |  | 257 | 10281.213 | 9.2381 | 5.0949
258 | 10323.645 | 9.2422 | 5.1821 |  |  | 258 | 10320.168 | 9.2419 | 5.1487 |  |  | 258 | 10321.218 | 9.242 | 5.0939
259 | 10363.659 | 9.2461 | 5.1811 |  |  | 259 | 10360.169 | 9.2457 | 5.1478 |  |  | 259 | 10361.223 | 9.2458 | 5.0929
260 | 10403.673 | 9.2499 | 5.18 |  |  | 260 | 10400.17 | 9.2496 | 5.1468 |  |  | 260 | 10401.228 | 9.2497 | 5.0919
261 | 10443.688 | 9.2538 | 5.179 |  |  | 261 | 10440.17 | 9.2534 | 5.1459 |  |  | 261 | 10441.232 | 9.2535 | 5.0909
262 | 10483.702 | 9.2576 | 5.178 |  |  | 262 | 10480.171 | 9.2572 | 5.145 |  |  | 262 | 10481.237 | 9.2573 | 5.0899
263 | 10523.716 | 9.2614 | 5.177 |  |  | 263 | 10520.172 | 9.261 | 5.1441 |  |  | 263 | 10521.242 | 9.2612 | 5.0889
264 | 10563.73 | 9.2652 | 5.176 |  |  | 264 | 10560.172 | 9.2648 | 5.1432 |  |  | 264 | 10561.247 | 9.2649 | 5.088
265 | 10603.744 | 9.269 | 5.175 |  |  | 265 | 10600.173 | 9.2686 | 5.1423 |  |  | 265 | 10601.251 | 9.2687 | 5.087
266 | 10643.758 | 9.2727 | 5.174 |  |  | 266 | 10640.174 | 9.2724 | 5.1414 |  |  | 266 | 10641.256 | 9.2725 | 5.0861
267 | 10683.772 | 9.2765 | 5.173 |  |  | 267 | 10680.174 | 9.2761 | 5.1405 |  |  | 267 | 10681.261 | 9.2762 | 5.0851
268 | 10723.786 | 9.2802 | 5.1721 |  |  | 268 | 10720.175 | 9.2799 | 5.1396 |  |  | 268 | 10721.265 | 9.28 | 5.0842
269 | 10763.801 | 9.2839 | 5.1711 |  |  | 269 | 10760.176 | 9.2836 | 5.1387 |  |  | 269 | 10761.27 | 9.2837 | 5.0833
270 | 10803.815 | 9.2877 | 5.1701 |  |  | 270 | 10800.176 | 9.2873 | 5.1379 |  |  | 270 | 10801.275 | 9.2874 | 5.0824
271 | 10843.829 | 9.2914 | 5.1691 |  |  | 271 | 10840.177 | 9.291 | 5.137 |  |  | 271 | 10841.28 | 9.2911 | 5.0815
272 | 10883.843 | 9.295 | 5.1682 |  |  | 272 | 10880.178 | 9.2947 | 5.1362 |  |  | 272 | 10881.284 | 9.2948 | 5.0806
273 | 10923.857 | 9.2987 | 5.1672 |  |  | 273 | 10920.178 | 9.2984 | 5.1353 |  |  | 273 | 10921.289 | 9.2985 | 5.0797
274 | 10963.871 | 9.3024 | 5.1663 |  |  | 274 | 10960.179 | 9.302 | 5.1345 |  |  | 274 | 10961.294 | 9.3021 | 5.0788
275 | 11003.885 | 9.306 | 5.1653 |  |  | 275 | 11000.18 | 9.3057 | 5.1337 |  |  | 275 | 11001.298 | 9.3058 | 5.0779
276 | 11043.899 | 9.3096 | 5.1644 |  |  | 276 | 11040.18 | 9.3093 | 5.1329 |  |  | 276 | 11041.303 | 9.3094 | 5.077
277 | 11083.914 | 9.3133 | 5.1634 |  |  | 277 | 11080.181 | 9.3129 | 5.132 |  |  | 277 | 11081.308 | 9.313 | 5.0762
278 | 11123.928 | 9.3169 | 5.1625 |  |  | 278 | 11120.182 | 9.3165 | 5.1312 |  |  | 278 | 11121.313 | 9.3166 | 5.0753
279 | 11163.942 | 9.3204 | 5.1616 |  |  | 279 | 11160.182 | 9.3201 | 5.1304 |  |  | 279 | 11161.317 | 9.3202 | 5.0745
280 | 11203.956 | 9.324 | 5.1606 |  |  | 280 | 11200.183 | 9.3237 | 5.1297 |  |  | 280 | 11201.322 | 9.3238 | 5.0736
281 | 11243.97 | 9.3276 | 5.1597 |  |  | 281 | 11240.184 | 9.3273 | 5.1289 |  |  | 281 | 11241.327 | 9.3274 | 5.0728
282 | 11283.984 | 9.3311 | 5.1588 |  |  | 282 | 11280.184 | 9.3308 | 5.1281 |  |  | 282 | 11281.331 | 9.3309 | 5.0719
283 | 11323.998 | 9.3347 | 5.1579 |  |  | 283 | 11320.185 | 9.3343 | 5.1273 |  |  | 283 | 11321.336 | 9.3344 | 5.0711
284 | 11364.013 | 9.3382 | 5.157 |  |  | 284 | 11360.185 | 9.3379 | 5.1265 |  |  | 284 | 11361.341 | 9.338 | 5.0703
285 | 11404.027 | 9.3417 | 5.156 |  |  | 285 | 11400.186 | 9.3414 | 5.1258 |  |  | 285 | 11401.346 | 9.3415 | 5.0695
286 | 11444.041 | 9.3452 | 5.1551 |  |  | 286 | 11440.187 | 9.3449 | 5.125 |  |  | 286 | 11441.35 | 9.345 | 5.0687
287 | 11484.055 | 9.3487 | 5.1542 |  |  | 287 | 11480.187 | 9.3484 | 5.1243 |  |  | 287 | 11481.355 | 9.3485 | 5.0679
288 | 11524.069 | 9.3522 | 5.1533 |  |  | 288 | 11520.188 | 9.3519 | 5.1235 |  |  | 288 | 11521.36 | 9.352 | 5.0671
289 | 11564.083 | 9.3557 | 5.1524 |  |  | 289 | 11560.189 | 9.3553 | 5.1228 |  |  | 289 | 11561.365 | 9.3554 | 5.0663
290 | 11604.097 | 9.3591 | 5.1516 |  |  | 290 | 11600.189 | 9.3588 | 5.1221 |  |  | 290 | 11601.369 | 9.3589 | 5.0655
291 | 11644.111 | 9.3626 | 5.1507 |  |  | 291 | 11640.19 | 9.3622 | 5.1213 |  |  | 291 | 11641.374 | 9.3623 | 5.0647
292 | 11684.126 | 9.366 | 5.1498 |  |  | 292 | 11680.191 | 9.3656 | 5.1206 |  |  | 292 | 11681.379 | 9.3658 | 5.064
293 | 11724.14 | 9.3694 | 5.1489 |  |  | 293 | 11720.191 | 9.3691 | 5.1199 |  |  | 293 | 11721.383 | 9.3692 | 5.0632
294 | 11764.154 | 9.3728 | 5.148 |  |  | 294 | 11760.192 | 9.3725 | 5.1192 |  |  | 294 | 11761.388 | 9.3726 | 5.0624
295 | 11804.168 | 9.3762 | 5.1472 |  |  | 295 | 11800.193 | 9.3759 | 5.1185 |  |  | 295 | 11801.393 | 9.376 | 5.0617
296 | 11844.182 | 9.3796 | 5.1463 |  |  | 296 | 11840.193 | 9.3793 | 5.1178 |  |  | 296 | 11841.398 | 9.3794 | 5.0609
297 | 11884.196 | 9.383 | 5.1454 |  |  | 297 | 11880.194 | 9.3826 | 5.1171 |  |  | 297 | 11881.402 | 9.3827 | 5.0602
298 | 11924.21 | 9.3863 | 5.1446 |  |  | 298 | 11920.195 | 9.386 | 5.1164 |  |  | 298 | 11921.407 | 9.3861 | 5.0595
299 | 11964.224 | 9.3897 | 5.1437 |  |  | 299 | 11960.195 | 9.3893 | 5.1157 |  |  | 299 | 11961.412 | 9.3894 | 5.0587
300 | 12004.239 | 9.393 | 5.1429 |  |  | 300 | 12000.196 | 9.3927 | 5.1151 |  |  | 300 | 12001.416 | 9.3928 | 5.058
301 | 12044.253 | 9.3963 | 5.142 |  |  | 301 | 12040.197 | 9.396 | 5.1144 |  |  | 301 | 12041.421 | 9.3961 | 5.0573
302 | 12084.267 | 9.3997 | 5.1412 |  |  | 302 | 12080.197 | 9.3993 | 5.1137 |  |  | 302 | 12081.426 | 9.3994 | 5.0566
303 | 12124.281 | 9.403 | 5.1403 |  |  | 303 | 12120.198 | 9.4026 | 5.1131 |  |  | 303 | 12121.431 | 9.4027 | 5.0559
304 | 12164.295 | 9.4063 | 5.1395 |  |  | 304 | 12160.199 | 9.4059 | 5.1124 |  |  | 304 | 12161.435 | 9.406 | 5.0552
305 | 12204.309 | 9.4095 | 5.1386 |  |  | 305 | 12200.199 | 9.4092 | 5.1118 |  |  | 305 | 12201.44 | 9.4093 | 5.0545
306 | 12244.323 | 9.4128 | 5.1378 |  |  | 306 | 12240.2 | 9.4125 | 5.1111 |  |  | 306 | 12241.445 | 9.4126 | 5.0538
307 | 12284.337 | 9.4161 | 5.137 |  |  | 307 | 12280.2 | 9.4157 | 5.1105 |  |  | 307 | 12281.45 | 9.4158 | 5.0531
308 | 12324.352 | 9.4193 | 5.1361 |  |  | 308 | 12320.201 | 9.419 | 5.1099 |  |  | 308 | 12321.454 | 9.4191 | 5.0524
309 | 12364.366 | 9.4226 | 5.1353 |  |  | 309 | 12360.202 | 9.4222 | 5.1092 |  |  | 309 | 12361.459 | 9.4223 | 5.0517
310 | 12404.38 | 9.4258 | 5.1345 |  |  | 310 | 12400.202 | 9.4255 | 5.1086 |  |  | 310 | 12401.464 | 9.4256 | 5.0511
311 | 12444.394 | 9.429 | 5.1337 |  |  | 311 | 12440.203 | 9.4287 | 5.108 |  |  | 311 | 12441.468 | 9.4288 | 5.0504
312 | 12484.408 | 9.4322 | 5.1329 |  |  | 312 | 12480.204 | 9.4319 | 5.1074 |  |  | 312 | 12481.473 | 9.432 | 5.0497
313 | 12524.422 | 9.4354 | 5.132 |  |  | 313 | 12520.204 | 9.4351 | 5.1068 |  |  | 313 | 12521.478 | 9.4352 | 5.0491
314 | 12564.436 | 9.4386 | 5.1312 |  |  | 314 | 12560.205 | 9.4383 | 5.1062 |  |  | 314 | 12561.483 | 9.4384 | 5.0484
315 | 12604.45 | 9.4418 | 5.1304 |  |  | 315 | 12600.206 | 9.4415 | 5.1056 |  |  | 315 | 12601.487 | 9.4416 | 5.0478
316 | 12644.465 | 9.445 | 5.1296 |  |  | 316 | 12640.206 | 9.4446 | 5.105 |  |  | 316 | 12641.492 | 9.4447 | 5.0471
317 | 12684.479 | 9.4481 | 5.1288 |  |  | 317 | 12680.207 | 9.4478 | 5.1044 |  |  | 317 | 12681.497 | 9.4479 | 5.0465
318 | 12724.493 | 9.4513 | 5.128 |  |  | 318 | 12720.208 | 9.4509 | 5.1038 |  |  | 318 | 12721.501 | 9.451 | 5.0459
319 | 12764.507 | 9.4544 | 5.1272 |  |  | 319 | 12760.208 | 9.4541 | 5.1032 |  |  | 319 | 12761.506 | 9.4542 | 5.0452
320 | 12804.521 | 9.4576 | 5.1264 |  |  | 320 | 12800.209 | 9.4572 | 5.1026 |  |  | 320 | 12801.511 | 9.4573 | 5.0446
321 | 12844.535 | 9.4607 | 5.1257 |  |  | 321 | 12840.21 | 9.4603 | 5.102 |  |  | 321 | 12841.516 | 9.4604 | 5.044
322 | 12884.549 | 9.4638 | 5.1249 |  |  | 322 | 12880.21 | 9.4634 | 5.1015 |  |  | 322 | 12881.52 | 9.4635 | 5.0434
323 | 12924.564 | 9.4669 | 5.1241 |  |  | 323 | 12920.211 | 9.4665 | 5.1009 |  |  | 323 | 12921.525 | 9.4666 | 5.0428
324 | 12964.578 | 9.47 | 5.1233 |  |  | 324 | 12960.212 | 9.4696 | 5.1004 |  |  | 324 | 12961.53 | 9.4697 | 5.0422
325 | 13004.592 | 9.4731 | 5.1225 |  |  | 325 | 13000.212 | 9.4727 | 5.0998 |  |  | 325 | 13001.535 | 9.4728 | 5.0416
326 | 13044.606 | 9.4761 | 5.1218 |  |  | 326 | 13040.213 | 9.4758 | 5.0992 |  |  | 326 | 13041.539 | 9.4759 | 5.041
327 | 13084.62 | 9.4792 | 5.121 |  |  | 327 | 13080.214 | 9.4789 | 5.0987 |  |  | 327 | 13081.544 | 9.479 | 5.0404
328 | 13124.634 | 9.4822 | 5.1202 |  |  | 328 | 13120.214 | 9.4819 | 5.0982 |  |  | 328 | 13121.549 | 9.482 | 5.0398
329 | 13164.648 | 9.4853 | 5.1195 |  |  | 329 | 13160.215 | 9.485 | 5.0976 |  |  | 329 | 13161.553 | 9.4851 | 5.0392
330 | 13204.662 | 9.4883 | 5.1187 |  |  | 330 | 13200.216 | 9.488 | 5.0971 |  |  | 330 | 13201.558 | 9.4881 | 5.0386
331 | 13244.677 | 9.4914 | 5.1179 |  |  | 331 | 13240.216 | 9.491 | 5.0966 |  |  | 331 | 13241.563 | 9.4911 | 5.038
332 | 13284.691 | 9.4944 | 5.1172 |  |  | 332 | 13280.217 | 9.494 | 5.096 |  |  | 332 | 13281.568 | 9.4941 | 5.0375
333 | 13324.705 | 9.4974 | 5.1164 |  |  | 333 | 13320.217 | 9.497 | 5.0955 |  |  | 333 | 13321.572 | 9.4971 | 5.0369
334 | 13364.719 | 9.5004 | 5.1157 |  |  | 334 | 13360.218 | 9.5 | 5.095 |  |  | 334 | 13361.577 | 9.5001 | 5.0363
335 | 13404.733 | 9.5034 | 5.1149 |  |  | 335 | 13400.219 | 9.503 | 5.0945 |  |  | 335 | 13401.582 | 9.5031 | 5.0358
336 | 13444.747 | 9.5063 | 5.1142 |  |  | 336 | 13440.219 | 9.506 | 5.094 |  |  | 336 | 13441.586 | 9.5061 | 5.0352
337 | 13484.761 | 9.5093 | 5.1134 |  |  | 337 | 13480.22 | 9.509 | 5.0934 |  |  | 337 | 13481.591 | 9.5091 | 5.0347
338 | 13524.775 | 9.5123 | 5.1127 |  |  | 338 | 13520.221 | 9.5119 | 5.0929 |  |  | 338 | 13521.596 | 9.512 | 5.0341
339 | 13564.79 | 9.5152 | 5.1119 |  |  | 339 | 13560.221 | 9.5149 | 5.0924 |  |  | 339 | 13561.601 | 9.515 | 5.0336
340 | 13604.804 | 9.5182 | 5.1112 |  |  | 340 | 13600.222 | 9.5178 | 5.092 |  |  | 340 | 13601.605 | 9.5179 | 5.033
341 | 13644.818 | 9.5211 | 5.1105 |  |  | 341 | 13640.223 | 9.5208 | 5.0915 |  |  | 341 | 13641.61 | 9.5209 | 5.0325
342 | 13684.832 | 9.524 | 5.1097 |  |  | 342 | 13680.223 | 9.5237 | 5.091 |  |  | 342 | 13681.615 | 9.5238 | 5.032
343 | 13724.846 | 9.527 | 5.109 |  |  | 343 | 13720.224 | 9.5266 | 5.0905 |  |  | 343 | 13721.62 | 9.5267 | 5.0314
344 | 13764.86 | 9.5299 | 5.1083 |  |  | 344 | 13760.225 | 9.5295 | 5.09 |  |  | 344 | 13761.624 | 9.5296 | 5.0309
345 | 13804.874 | 9.5328 | 5.1076 |  |  | 345 | 13800.225 | 9.5324 | 5.0895 |  |  | 345 | 13801.629 | 9.5325 | 5.0304
346 | 13844.888 | 9.5357 | 5.1068 |  |  | 346 | 13840.226 | 9.5353 | 5.0891 |  |  | 346 | 13841.634 | 9.5354 | 5.0299
347 | 13884.903 | 9.5386 | 5.1061 |  |  | 347 | 13880.227 | 9.5382 | 5.0886 |  |  | 347 | 13881.638 | 9.5383 | 5.0294
348 | 13924.917 | 9.5414 | 5.1054 |  |  | 348 | 13920.227 | 9.5411 | 5.0881 |  |  | 348 | 13921.643 | 9.5412 | 5.0289
349 | 13964.931 | 9.5443 | 5.1047 |  |  | 349 | 13960.228 | 9.544 | 5.0876 |  |  | 349 | 13961.648 | 9.5441 | 5.0284
350 | 14004.945 | 9.5472 | 5.104 |  |  | 350 | 14000.229 | 9.5468 | 5.0872 |  |  | 350 | 14001.653 | 9.5469 | 5.0279
351 | 14044.959 | 9.55 | 5.1033 |  |  | 351 | 14040.229 | 9.5497 | 5.0867 |  |  | 351 | 14041.657 | 9.5498 | 5.0274
352 | 14084.973 | 9.5529 | 5.1026 |  |  | 352 | 14080.23 | 9.5525 | 5.0863 |  |  | 352 | 14081.662 | 9.5526 | 5.0269
353 | 14124.987 | 9.5557 | 5.1019 |  |  | 353 | 14120.231 | 9.5554 | 5.0858 |  |  | 353 | 14121.667 | 9.5555 | 5.0264
354 | 14165.001 | 9.5585 | 5.1012 |  |  | 354 | 14160.231 | 9.5582 | 5.0854 |  |  | 354 | 14161.671 | 9.5583 | 5.0259
355 | 14205.016 | 9.5614 | 5.1005 |  |  | 355 | 14200.232 | 9.561 | 5.0849 |  |  | 355 | 14201.676 | 9.5611 | 5.0254
356 | 14245.03 | 9.5642 | 5.0998 |  |  | 356 | 14240.232 | 9.5638 | 5.0845 |  |  | 356 | 14241.681 | 9.5639 | 5.0249
357 | 14285.044 | 9.567 | 5.0991 |  |  | 357 | 14280.233 | 9.5666 | 5.0841 |  |  | 357 | 14281.686 | 9.5667 | 5.0244
358 | 14325.058 | 9.5698 | 5.0984 |  |  | 358 | 14320.234 | 9.5694 | 5.0836 |  |  | 358 | 14321.69 | 9.5695 | 5.024
359 | 14365.072 | 9.5726 | 5.0977 |  |  | 359 | 14360.234 | 9.5722 | 5.0832 |  |  | 359 | 14361.695 | 9.5723 | 5.0235
360 | 14405.086 | 9.5753 | 5.097 |  |  | 360 | 14400.235 | 9.575 | 5.0828 |  |  | 360 | 14401.7 | 9.5751 | 5.023
361 | 14445.1 | 9.5781 | 5.0963 |  |  | 361 | 14440.236 | 9.5778 | 5.0823 |  |  | 361 | 14441.705 | 9.5779 | 5.0226
362 | 14485.115 | 9.5809 | 5.0956 |  |  | 362 | 14480.236 | 9.5805 | 5.0819 |  |  | 362 | 14481.709 | 9.5806 | 5.0221
363 | 14525.129 | 9.5836 | 5.0949 |  |  | 363 | 14520.237 | 9.5833 | 5.0815 |  |  | 363 | 14521.714 | 9.5834 | 5.0216
364 | 14565.143 | 9.5864 | 5.0943 |  |  | 364 | 14560.238 | 9.586 | 5.0811 |  |  | 364 | 14561.719 | 9.5862 | 5.0212
365 | 14605.157 | 9.5891 | 5.0936 |  |  | 365 | 14600.238 | 9.5888 | 5.0807 |  |  | 365 | 14601.723 | 9.5889 | 5.0207
366 | 14645.171 | 9.5919 | 5.0929 |  |  | 366 | 14640.239 | 9.5915 | 5.0803 |  |  | 366 | 14641.728 | 9.5916 | 5.0203
367 | 14685.185 | 9.5946 | 5.0922 |  |  | 367 | 14680.24 | 9.5943 | 5.0799 |  |  | 367 | 14681.733 | 9.5944 | 5.0198
368 | 14725.199 | 9.5973 | 5.0916 |  |  | 368 | 14720.24 | 9.597 | 5.0795 |  |  | 368 | 14721.738 | 9.5971 | 5.0194
369 | 14765.213 | 9.6 | 5.0909 |  |  | 369 | 14760.241 | 9.5997 | 5.0791 |  |  | 369 | 14761.742 | 9.5998 | 5.0189
370 | 14805.228 | 9.6027 | 5.0902 |  |  | 370 | 14800.242 | 9.6024 | 5.0787 |  |  | 370 | 14801.747 | 9.6025 | 5.0185
371 | 14845.242 | 9.6054 | 5.0896 |  |  | 371 | 14840.242 | 9.6051 | 5.0783 |  |  | 371 | 14841.752 | 9.6052 | 5.0181
372 | 14885.256 | 9.6081 | 5.0889 |  |  | 372 | 14880.243 | 9.6078 | 5.0779 |  |  | 372 | 14881.756 | 9.6079 | 5.0176
373 | 14925.27 | 9.6108 | 5.0882 |  |  | 373 | 14920.244 | 9.6105 | 5.0775 |  |  | 373 | 14921.761 | 9.6106 | 5.0172
374 | 14965.284 | 9.6135 | 5.0876 |  |  | 374 | 14960.244 | 9.6132 | 5.0771 |  |  | 374 | 14961.766 | 9.6133 | 5.0168
375 | 15005.298 | 9.6162 | 5.0869 |  |  | 375 | 15000.245 | 9.6158 | 5.0767 |  |  | 375 | 15001.771 | 9.6159 | 5.0164
376 | 15045.312 | 9.6188 | 5.0863 |  |  | 376 | 15040.246 | 9.6185 | 5.0763 |  |  | 376 | 15041.775 | 9.6186 | 5.0159
377 | 15085.326 | 9.6215 | 5.0856 |  |  | 377 | 15080.246 | 9.6211 | 5.0759 |  |  | 377 | 15081.78 | 9.6212 | 5.0155
378 | 15125.341 | 9.6241 | 5.085 |  |  | 378 | 15120.247 | 9.6238 | 5.0756 |  |  | 378 | 15121.785 | 9.6239 | 5.0151
379 | 15165.355 | 9.6268 | 5.0843 |  |  | 379 | 15160.248 | 9.6264 | 5.0752 |  |  | 379 | 15161.789 | 9.6265 | 5.0147
380 | 15205.369 | 9.6294 | 5.0837 |  |  | 380 | 15200.248 | 9.6291 | 5.0748 |  |  | 380 | 15201.794 | 9.6292 | 5.0143
381 | 15245.383 | 9.632 | 5.083 |  |  | 381 | 15240.249 | 9.6317 | 5.0745 |  |  | 381 | 15241.799 | 9.6318 | 5.0139
382 | 15285.397 | 9.6347 | 5.0824 |  |  | 382 | 15280.249 | 9.6343 | 5.0741 |  |  | 382 | 15281.804 | 9.6344 | 5.0135
383 | 15325.411 | 9.6373 | 5.0817 |  |  | 383 | 15320.25 | 9.6369 | 5.0737 |  |  | 383 | 15321.808 | 9.637 | 5.0131
384 | 15365.425 | 9.6399 | 5.0811 |  |  | 384 | 15360.251 | 9.6395 | 5.0734 |  |  | 384 | 15361.813 | 9.6396 | 5.0127
385 | 15405.439 | 9.6425 | 5.0804 |  |  | 385 | 15400.251 | 9.6421 | 5.073 |  |  | 385 | 15401.818 | 9.6422 | 5.0123
386 | 15445.454 | 9.6451 | 5.0798 |  |  | 386 | 15440.252 | 9.6447 | 5.0727 |  |  | 386 | 15441.823 | 9.6448 | 5.0119
387 | 15485.468 | 9.6477 | 5.0792 |  |  | 387 | 15480.253 | 9.6473 | 5.0723 |  |  | 387 | 15481.827 | 9.6474 | 5.0115
388 | 15525.482 | 9.6502 | 5.0786 |  |  | 388 | 15520.253 | 9.6499 | 5.072 |  |  | 388 | 15521.832 | 9.65 | 5.0111
389 | 15565.496 | 9.6528 | 5.0779 |  |  | 389 | 15560.254 | 9.6525 | 5.0716 |  |  | 389 | 15561.837 | 9.6526 | 5.0107
390 | 15605.51 | 9.6554 | 5.0773 |  |  | 390 | 15600.255 | 9.655 | 5.0713 |  |  | 390 | 15601.841 | 9.6551 | 5.0103
391 | 15645.524 | 9.6579 | 5.0767 |  |  | 391 | 15640.255 | 9.6576 | 5.0709 |  |  | 391 | 15641.846 | 9.6577 | 5.0099
392 | 15685.538 | 9.6605 | 5.076 |  |  | 392 | 15680.256 | 9.6602 | 5.0706 |  |  | 392 | 15681.851 | 9.6603 | 5.0096
393 | 15725.553 | 9.663 | 5.0754 |  |  | 393 | 15720.257 | 9.6627 | 5.0702 |  |  | 393 | 15721.856 | 9.6628 | 5.0092
394 | 15765.567 | 9.6656 | 5.0748 |  |  | 394 | 15760.257 | 9.6652 | 5.0699 |  |  | 394 | 15761.86 | 9.6653 | 5.0088
395 | 15805.581 | 9.6681 | 5.0742 |  |  | 395 | 15800.258 | 9.6678 | 5.0696 |  |  | 395 | 15801.865 | 9.6679 | 5.0084
396 | 15845.595 | 9.6706 | 5.0736 |  |  | 396 | 15840.259 | 9.6703 | 5.0692 |  |  | 396 | 15841.87 | 9.6704 | 5.0081
397 | 15885.609 | 9.6732 | 5.073 |  |  | 397 | 15880.259 | 9.6728 | 5.0689 |  |  | 397 | 15881.874 | 9.6729 | 5.0077
398 | 15925.623 | 9.6757 | 5.0723 |  |  | 398 | 15920.26 | 9.6753 | 5.0686 |  |  | 398 | 15921.879 | 9.6754 | 5.0073
399 | 15965.637 | 9.6782 | 5.0717 |  |  | 399 | 15960.261 | 9.6779 | 5.0682 |  |  | 399 | 15961.884 | 9.678 | 5.007
400 | 16005.651 | 9.6807 | 5.0711 |  |  | 400 | 16000.261 | 9.6804 | 5.0679 |  |  | 400 | 16001.889 | 9.6805 | 5.0066
401 | 16045.666 | 9.6832 | 5.0705 |  |  | 401 | 16040.262 | 9.6829 | 5.0676 |  |  | 401 | 16041.893 | 9.683 | 5.0063
402 | 16085.68 | 9.6857 | 5.0699 |  |  | 402 | 16080.263 | 9.6853 | 5.0673 |  |  | 402 | 16081.898 | 9.6854 | 5.0059
403 | 16125.694 | 9.6882 | 5.0693 |  |  | 403 | 16120.263 | 9.6878 | 5.067 |  |  | 403 | 16121.903 | 9.6879 | 5.0056
404 | 16165.708 | 9.6906 | 5.0687 |  |  | 404 | 16160.264 | 9.6903 | 5.0667 |  |  | 404 | 16161.908 | 9.6904 | 5.0052
405 | 16205.722 | 9.6931 | 5.0681 |  |  | 405 | 16200.265 | 9.6928 | 5.0663 |  |  | 405 | 16201.912 | 9.6929 | 5.0049
406 | 16245.736 | 9.6956 | 5.0675 |  |  | 406 | 16240.265 | 9.6952 | 5.066 |  |  | 406 | 16241.917 | 9.6954 | 5.0045
407 | 16285.75 | 9.698 | 5.0669 |  |  | 407 | 16280.266 | 9.6977 | 5.0657 |  |  | 407 | 16281.922 | 9.6978 | 5.0042
408 | 16325.764 | 9.7005 | 5.0663 |  |  | 408 | 16320.266 | 9.7002 | 5.0654 |  |  | 408 | 16321.926 | 9.7003 | 5.0038
409 | 16365.779 | 9.7029 | 5.0657 |  |  | 409 | 16360.267 | 9.7026 | 5.0651 |  |  | 409 | 16361.931 | 9.7027 | 5.0035
410 | 16405.793 | 9.7054 | 5.0651 |  |  | 410 | 16400.268 | 9.7051 | 5.0648 |  |  | 410 | 16401.936 | 9.7052 | 5.0031
411 | 16445.807 | 9.7078 | 5.0645 |  |  | 411 | 16440.268 | 9.7075 | 5.0645 |  |  | 411 | 16441.941 | 9.7076 | 5.0028
412 | 16485.821 | 9.7103 | 5.0639 |  |  | 412 | 16480.269 | 9.7099 | 5.0642 |  |  | 412 | 16481.945 | 9.71 | 5.0025
413 | 16525.835 | 9.7127 | 5.0634 |  |  | 413 | 16520.27 | 9.7123 | 5.0639 |  |  | 413 | 16521.95 | 9.7124 | 5.0021
414 | 16565.849 | 9.7151 | 5.0628 |  |  | 414 | 16560.27 | 9.7148 | 5.0636 |  |  | 414 | 16561.955 | 9.7149 | 5.0018
415 | 16605.863 | 9.7175 | 5.0622 |  |  | 415 | 16600.271 | 9.7172 | 5.0633 |  |  | 415 | 16601.959 | 9.7173 | 5.0015
416 | 16645.877 | 9.7199 | 5.0616 |  |  | 416 | 16640.272 | 9.7196 | 5.063 |  |  | 416 | 16641.964 | 9.7197 | 5.0012
417 | 16685.892 | 9.7223 | 5.061 |  |  | 417 | 16680.272 | 9.722 | 5.0628 |  |  | 417 | 16681.969 | 9.7221 | 5.0008
418 | 16725.906 | 9.7247 | 5.0605 |  |  | 418 | 16720.273 | 9.7244 | 5.0625 |  |  | 418 | 16721.974 | 9.7245 | 5.0005
419 | 16765.92 | 9.7271 | 5.0599 |  |  | 419 | 16760.274 | 9.7268 | 5.0622 |  |  | 419 | 16761.978 | 9.7269 | 5.0002
420 | 16805.934 | 9.7295 | 5.0593 |  |  | 420 | 16800.274 | 9.7292 | 5.0619 |  |  | 420 | 16801.983 | 9.7293 | 4.9999
421 | 16845.948 | 9.7319 | 5.0587 |  |  | 421 | 16840.275 | 9.7315 | 5.0616 |  |  | 421 | 16841.988 | 9.7316 | 4.9996
422 | 16885.962 | 9.7342 | 5.0582 |  |  | 422 | 16880.276 | 9.7339 | 5.0613 |  |  | 422 | 16881.993 | 9.734 | 4.9993
423 | 16925.976 | 9.7366 | 5.0576 |  |  | 423 | 16920.276 | 9.7363 | 5.0611 |  |  | 423 | 16921.997 | 9.7364 | 4.9989
424 | 16965.99 | 9.739 | 5.057 |  |  | 424 | 16960.277 | 9.7386 | 5.0608 |  |  | 424 | 16962.002 | 9.7387 | 4.9986
425 | 17006.005 | 9.7413 | 5.0565 |  |  | 425 | 17000.278 | 9.741 | 5.0605 |  |  | 425 | 17002.007 | 9.7411 | 4.9983
426 | 17046.019 | 9.7437 | 5.0559 |  |  | 426 | 17040.278 | 9.7433 | 5.0602 |  |  | 426 | 17042.011 | 9.7434 | 4.998
427 | 17086.033 | 9.746 | 5.0553 |  |  | 427 | 17080.279 | 9.7457 | 5.06 |  |  | 427 | 17082.016 | 9.7458 | 4.9977
428 | 17126.047 | 9.7484 | 5.0548 |  |  | 428 | 17120.28 | 9.748 | 5.0597 |  |  | 428 | 17122.021 | 9.7481 | 4.9974
429 | 17166.061 | 9.7507 | 5.0542 |  |  | 429 | 17160.28 | 9.7504 | 5.0594 |  |  | 429 | 17162.026 | 9.7505 | 4.9971
430 | 17206.075 | 9.753 | 5.0536 |  |  | 430 | 17200.281 | 9.7527 | 5.0592 |  |  | 430 | 17202.03 | 9.7528 | 4.9968
431 | 17246.089 | 9.7553 | 5.0531 |  |  | 431 | 17240.281 | 9.755 | 5.0589 |  |  | 431 | 17242.035 | 9.7551 | 4.9965
432 | 17286.104 | 9.7577 | 5.0525 |  |  | 432 | 17280.282 | 9.7573 | 5.0586 |  |  | 432 | 17282.04 | 9.7574 | 4.9962
433 | 17326.118 | 9.76 | 5.052 |  |  | 433 | 17320.283 | 9.7596 | 5.0584 |  |  | 433 | 17322.044 | 9.7597 | 4.9959
434 | 17366.132 | 9.7623 | 5.0514 |  |  | 434 | 17360.283 | 9.7619 | 5.0581 |  |  | 434 | 17362.049 | 9.762 | 4.9956
435 | 17406.146 | 9.7646 | 5.0509 |  |  | 435 | 17400.284 | 9.7642 | 5.0579 |  |  | 435 | 17402.054 | 9.7643 | 4.9953
436 | 17446.16 | 9.7669 | 5.0503 |  |  | 436 | 17440.285 | 9.7665 | 5.0576 |  |  | 436 | 17442.059 | 9.7666 | 4.995
437 | 17486.174 | 9.7692 | 5.0498 |  |  | 437 | 17480.285 | 9.7688 | 5.0574 |  |  | 437 | 17482.063 | 9.7689 | 4.9948
438 | 17526.188 | 9.7715 | 5.0492 |  |  | 438 | 17520.286 | 9.7711 | 5.0571 |  |  | 438 | 17522.068 | 9.7712 | 4.9945
439 | 17566.202 | 9.7737 | 5.0487 |  |  | 439 | 17560.287 | 9.7734 | 5.0569 |  |  | 439 | 17562.073 | 9.7735 | 4.9942
440 | 17606.217 | 9.776 | 5.0481 |  |  | 440 | 17600.287 | 9.7757 | 5.0566 |  |  | 440 | 17602.078 | 9.7758 | 4.9939
441 | 17646.231 | 9.7783 | 5.0476 |  |  | 441 | 17640.288 | 9.7779 | 5.0564 |  |  | 441 | 17642.082 | 9.778 | 4.9936
442 | 17686.245 | 9.7805 | 5.0471 |  |  | 442 | 17680.289 | 9.7802 | 5.0561 |  |  | 442 | 17682.087 | 9.7803 | 4.9933
443 | 17726.259 | 9.7828 | 5.0465 |  |  | 443 | 17720.289 | 9.7825 | 5.0559 |  |  | 443 | 17722.092 | 9.7826 | 4.9931
444 | 17766.273 | 9.7851 | 5.046 |  |  | 444 | 17760.29 | 9.7847 | 5.0556 |  |  | 444 | 17762.096 | 9.7848 | 4.9928
445 | 17806.287 | 9.7873 | 5.0454 |  |  | 445 | 17800.291 | 9.787 | 5.0554 |  |  | 445 | 17802.101 | 9.7871 | 4.9925
446 | 17846.301 | 9.7896 | 5.0449 |  |  | 446 | 17840.291 | 9.7892 | 5.0551 |  |  | 446 | 17842.106 | 9.7893 | 4.9922
447 | 17886.315 | 9.7918 | 5.0444 |  |  | 447 | 17880.292 | 9.7915 | 5.0549 |  |  | 447 | 17882.111 | 9.7916 | 4.992
448 | 17926.33 | 9.794 | 5.0438 |  |  | 448 | 17920.293 | 9.7937 | 5.0547 |  |  | 448 | 17922.115 | 9.7938 | 4.9917
449 | 17966.344 | 9.7963 | 5.0433 |  |  | 449 | 17960.293 | 9.7959 | 5.0544 |  |  | 449 | 17962.12 | 9.796 | 4.9914
450 | 18006.358 | 9.7985 | 5.0428 |  |  | 450 | 18000.294 | 9.7981 | 5.0542 |  |  | 450 | 18002.125 | 9.7982 | 4.9912
451 | 18046.372 | 9.8007 | 5.0423 |  |  | 451 | 18040.295 | 9.8004 | 5.054 |  |  | 451 | 18042.129 | 9.8005 | 4.9909
452 | 18086.386 | 9.8029 | 5.0417 |  |  | 452 | 18080.295 | 9.8026 | 5.0537 |  |  | 452 | 18082.134 | 9.8027 | 4.9906
453 | 18126.4 | 9.8051 | 5.0412 |  |  | 453 | 18120.296 | 9.8048 | 5.0535 |  |  | 453 | 18122.139 | 9.8049 | 4.9904
454 | 18166.414 | 9.8073 | 5.0407 |  |  | 454 | 18160.297 | 9.807 | 5.0533 |  |  | 454 | 18162.144 | 9.8071 | 4.9901
455 | 18206.428 | 9.8095 | 5.0402 |  |  | 455 | 18200.297 | 9.8092 | 5.053 |  |  | 455 | 18202.148 | 9.8093 | 4.9898
456 | 18246.443 | 9.8117 | 5.0396 |  |  | 456 | 18240.298 | 9.8114 | 5.0528 |  |  | 456 | 18242.153 | 9.8115 | 4.9896
457 | 18286.457 | 9.8139 | 5.0391 |  |  | 457 | 18280.298 | 9.8136 | 5.0526 |  |  | 457 | 18282.158 | 9.8137 | 4.9893
458 | 18326.471 | 9.8161 | 5.0386 |  |  | 458 | 18320.299 | 9.8158 | 5.0524 |  |  | 458 | 18322.163 | 9.8159 | 4.9891
459 | 18366.485 | 9.8183 | 5.0381 |  |  | 459 | 18360.3 | 9.8179 | 5.0521 |  |  | 459 | 18362.167 | 9.818 | 4.9888
460 | 18406.499 | 9.8205 | 5.0376 |  |  | 460 | 18400.3 | 9.8201 | 5.0519 |  |  | 460 | 18402.172 | 9.8202 | 4.9886
461 | 18446.513 | 9.8226 | 5.0371 |  |  | 461 | 18440.301 | 9.8223 | 5.0517 |  |  | 461 | 18442.177 | 9.8224 | 4.9883
462 | 18486.527 | 9.8248 | 5.0365 |  |  | 462 | 18480.302 | 9.8245 | 5.0515 |  |  | 462 | 18482.181 | 9.8246 | 4.9881
463 | 18526.542 | 9.827 | 5.036 |  |  | 463 | 18520.302 | 9.8266 | 5.0513 |  |  | 463 | 18522.186 | 9.8267 | 4.9878
464 | 18566.556 | 9.8291 | 5.0355 |  |  | 464 | 18560.303 | 9.8288 | 5.0511 |  |  | 464 | 18562.191 | 9.8289 | 4.9876
465 | 18606.57 | 9.8313 | 5.035 |  |  | 465 | 18600.304 | 9.8309 | 5.0508 |  |  | 465 | 18602.196 | 9.831 | 4.9873
466 | 18646.584 | 9.8334 | 5.0345 |  |  | 466 | 18640.304 | 9.8331 | 5.0506 |  |  | 466 | 18642.2 | 9.8332 | 4.9871
467 | 18686.598 | 9.8356 | 5.034 |  |  | 467 | 18680.305 | 9.8352 | 5.0504 |  |  | 467 | 18682.205 | 9.8353 | 4.9868
468 | 18726.612 | 9.8377 | 5.0335 |  |  | 468 | 18720.306 | 9.8374 | 5.0502 |  |  | 468 | 18722.21 | 9.8375 | 4.9866
469 | 18766.626 | 9.8398 | 5.033 |  |  | 469 | 18760.306 | 9.8395 | 5.05 |  |  | 469 | 18762.214 | 9.8396 | 4.9863
470 | 18806.64 | 9.842 | 5.0325 |  |  | 470 | 18800.307 | 9.8416 | 5.0498 |  |  | 470 | 18802.219 | 9.8417 | 4.9861
471 | 18846.655 | 9.8441 | 5.032 |  |  | 471 | 18840.308 | 9.8438 | 5.0496 |  |  | 471 | 18842.224 | 9.8439 | 4.9859
472 | 18886.669 | 9.8462 | 5.0315 |  |  | 472 | 18880.308 | 9.8459 | 5.0494 |  |  | 472 | 18882.229 | 9.846 | 4.9856
473 | 18926.683 | 9.8483 | 5.031 |  |  | 473 | 18920.309 | 9.848 | 5.0492 |  |  | 473 | 18922.233 | 9.8481 | 4.9854
474 | 18966.697 | 9.8504 | 5.0305 |  |  | 474 | 18960.31 | 9.8501 | 5.049 |  |  | 474 | 18962.238 | 9.8502 | 4.9852
475 | 19006.711 | 9.8525 | 5.03 |  |  | 475 | 19000.31 | 9.8522 | 5.0488 |  |  | 475 | 19002.243 | 9.8523 | 4.9849
476 | 19046.725 | 9.8547 | 5.0295 |  |  | 476 | 19040.311 | 9.8543 | 5.0486 |  |  | 476 | 19042.247 | 9.8544 | 4.9847
477 | 19086.739 | 9.8567 | 5.029 |  |  | 477 | 19080.312 | 9.8564 | 5.0484 |  |  | 477 | 19082.252 | 9.8565 | 4.9845
478 | 19126.753 | 9.8588 | 5.0285 |  |  | 478 | 19120.312 | 9.8585 | 5.0482 |  |  | 478 | 19122.257 | 9.8586 | 4.9842
479 | 19166.768 | 9.8609 | 5.0281 |  |  | 479 | 19160.313 | 9.8606 | 5.048 |  |  | 479 | 19162.262 | 9.8607 | 4.984
480 | 19206.782 | 9.863 | 5.0276 |  |  | 480 | 19200.313 | 9.8627 | 5.0478 |  |  | 480 | 19202.266 | 9.8628 | 4.9838
481 | 19246.796 | 9.8651 | 5.0271 |  |  | 481 | 19240.314 | 9.8648 | 5.0476 |  |  | 481 | 19242.271 | 9.8649 | 4.9835
482 | 19286.81 | 9.8672 | 5.0266 |  |  | 482 | 19280.315 | 9.8668 | 5.0474 |  |  | 482 | 19282.276 | 9.8669 | 4.9833
483 | 19326.824 | 9.8692 | 5.0261 |  |  | 483 | 19320.315 | 9.8689 | 5.0472 |  |  | 483 | 19322.281 | 9.869 | 4.9831
484 | 19366.838 | 9.8713 | 5.0256 |  |  | 484 | 19360.316 | 9.871 | 5.047 |  |  | 484 | 19362.285 | 9.8711 | 4.9829
485 | 19406.852 | 9.8734 | 5.0251 |  |  | 485 | 19400.317 | 9.873 | 5.0468 |  |  | 485 | 19402.29 | 9.8731 | 4.9826
486 | 19446.866 | 9.8754 | 5.0247 |  |  | 486 | 19440.317 | 9.8751 | 5.0466 |  |  | 486 | 19442.295 | 9.8752 | 4.9824
487 | 19486.881 | 9.8775 | 5.0242 |  |  | 487 | 19480.318 | 9.8772 | 5.0464 |  |  | 487 | 19482.299 | 9.8773 | 4.9822
488 | 19526.895 | 9.8795 | 5.0237 |  |  | 488 | 19520.319 | 9.8792 | 5.0462 |  |  | 488 | 19522.304 | 9.8793 | 4.982
489 | 19566.909 | 9.8816 | 5.0232 |  |  | 489 | 19560.319 | 9.8813 | 5.046 |  |  | 489 | 19562.309 | 9.8814 | 4.9818
490 | 19606.923 | 9.8836 | 5.0228 |  |  | 490 | 19600.32 | 9.8833 | 5.0459 |  |  | 490 | 19602.314 | 9.8834 | 4.9816
491 | 19646.937 | 9.8857 | 5.0223 |  |  | 491 | 19640.321 | 9.8853 | 5.0457 |  |  | 491 | 19642.318 | 9.8854 | 4.9813
492 | 19686.951 | 9.8877 | 5.0218 |  |  | 492 | 19680.321 | 9.8874 | 5.0455 |  |  | 492 | 19682.323 | 9.8875 | 4.9811
493 | 19726.965 | 9.8897 | 5.0213 |  |  | 493 | 19720.322 | 9.8894 | 5.0453 |  |  | 493 | 19722.328 | 9.8895 | 4.9809
494 | 19766.979 | 9.8918 | 5.0209 |  |  | 494 | 19760.323 | 9.8914 | 5.0451 |  |  | 494 | 19762.332 | 9.8915 | 4.9807
495 | 19806.994 | 9.8938 | 5.0204 |  |  | 495 | 19800.323 | 9.8935 | 5.0449 |  |  | 495 | 19802.337 | 9.8936 | 4.9805
496 | 19847.008 | 9.8958 | 5.0199 |  |  | 496 | 19840.324 | 9.8955 | 5.0448 |  |  | 496 | 19842.342 | 9.8956 | 4.9803
497 | 19887.022 | 9.8978 | 5.0195 |  |  | 497 | 19880.325 | 9.8975 | 5.0446 |  |  | 497 | 19882.347 | 9.8976 | 4.9801
498 | 19927.036 | 9.8998 | 5.019 |  |  | 498 | 19920.325 | 9.8995 | 5.0444 |  |  | 498 | 19922.351 | 9.8996 | 4.9799
499 | 19967.05 | 9.9018 | 5.0185 |  |  | 499 | 19960.326 | 9.9015 | 5.0442 |  |  | 499 | 19962.356 | 9.9016 | 4.9796
500 | 20007.064 | 9.9038 | 5.0181 |  |  | 500 | 20000.327 | 9.9035 | 5.044 |  |  | 500 | 20002.361 | 9.9036 | 4.9794
501 | 20047.078 | 9.9058 | 5.0176 |  |  | 501 | 20040.327 | 9.9055 | 5.0439 |  |  | 501 | 20042.366 | 9.9056 | 4.9792
502 | 20087.093 | 9.9078 | 5.0172 |  |  | 502 | 20080.328 | 9.9075 | 5.0437 |  |  | 502 | 20082.37 | 9.9076 | 4.979
503 | 20127.107 | 9.9098 | 5.0167 |  |  | 503 | 20120.329 | 9.9095 | 5.0435 |  |  | 503 | 20122.375 | 9.9096 | 4.9788
504 | 20167.121 | 9.9118 | 5.0162 |  |  | 504 | 20160.329 | 9.9115 | 5.0433 |  |  | 504 | 20162.38 | 9.9116 | 4.9786
505 | 20207.135 | 9.9138 | 5.0158 |  |  | 505 | 20200.33 | 9.9135 | 5.0432 |  |  | 505 | 20202.384 | 9.9136 | 4.9784
506 | 20247.149 | 9.9158 | 5.0153 |  |  | 506 | 20240.33 | 9.9154 | 5.043 |  |  | 506 | 20242.389 | 9.9155 | 4.9782
507 | 20287.163 | 9.9177 | 5.0149 |  |  | 507 | 20280.331 | 9.9174 | 5.0428 |  |  | 507 | 20282.394 | 9.9175 | 4.978
508 | 20327.177 | 9.9197 | 5.0144 |  |  | 508 | 20320.332 | 9.9194 | 5.0427 |  |  | 508 | 20322.399 | 9.9195 | 4.9778
509 | 20367.191 | 9.9217 | 5.014 |  |  | 509 | 20360.332 | 9.9213 | 5.0425 |  |  | 509 | 20362.403 | 9.9214 | 4.9776
510 | 20407.206 | 9.9236 | 5.0135 |  |  | 510 | 20400.333 | 9.9233 | 5.0423 |  |  | 510 | 20402.408 | 9.9234 | 4.9774
511 | 20447.22 | 9.9256 | 5.0131 |  |  | 511 | 20440.334 | 9.9253 | 5.0422 |  |  | 511 | 20442.413 | 9.9254 | 4.9772
512 | 20487.234 | 9.9276 | 5.0126 |  |  | 512 | 20480.334 | 9.9272 | 5.042 |  |  | 512 | 20482.417 | 9.9273 | 4.977
513 | 20527.248 | 9.9295 | 5.0122 |  |  | 513 | 20520.335 | 9.9292 | 5.0418 |  |  | 513 | 20522.422 | 9.9293 | 4.9768
514 | 20567.262 | 9.9315 | 5.0117 |  |  | 514 | 20560.336 | 9.9311 | 5.0417 |  |  | 514 | 20562.427 | 9.9312 | 4.9766
515 | 20607.276 | 9.9334 | 5.0113 |  |  | 515 | 20600.336 | 9.9331 | 5.0415 |  |  | 515 | 20602.432 | 9.9332 | 4.9764
516 | 20647.29 | 9.9353 | 5.0108 |  |  | 516 | 20640.337 | 9.935 | 5.0413 |  |  | 516 | 20642.436 | 9.9351 | 4.9762
517 | 20687.304 | 9.9373 | 5.0104 |  |  | 517 | 20680.338 | 9.9369 | 5.0412 |  |  | 517 | 20682.441 | 9.937 | 4.9761
518 | 20727.319 | 9.9392 | 5.01 |  |  | 518 | 20720.338 | 9.9389 | 5.041 |  |  | 518 | 20722.446 | 9.939 | 4.9759
519 | 20767.333 | 9.9411 | 5.0095 |  |  | 519 | 20760.339 | 9.9408 | 5.0408 |  |  | 519 | 20762.451 | 9.9409 | 4.9757
520 | 20807.347 | 9.9431 | 5.0091 |  |  | 520 | 20800.34 | 9.9427 | 5.0407 |  |  | 520 | 20802.455 | 9.9428 | 4.9755
521 | 20847.361 | 9.945 | 5.0086 |  |  | 521 | 20840.34 | 9.9446 | 5.0405 |  |  | 521 | 20842.46 | 9.9447 | 4.9753
522 | 20887.375 | 9.9469 | 5.0082 |  |  | 522 | 20880.341 | 9.9466 | 5.0404 |  |  | 522 | 20882.465 | 9.9467 | 4.9751
523 | 20927.389 | 9.9488 | 5.0078 |  |  | 523 | 20920.342 | 9.9485 | 5.0402 |  |  | 523 | 20922.469 | 9.9486 | 4.9749
524 | 20967.403 | 9.9507 | 5.0073 |  |  | 524 | 20960.342 | 9.9504 | 5.0401 |  |  | 524 | 20962.474 | 9.9505 | 4.9747
525 | 21007.417 | 9.9526 | 5.0069 |  |  | 525 | 21000.343 | 9.9523 | 5.0399 |  |  | 525 | 21002.479 | 9.9524 | 4.9745
526 | 21047.432 | 9.9545 | 5.0065 |  |  | 526 | 21040.344 | 9.9542 | 5.0397 |  |  | 526 | 21042.484 | 9.9543 | 4.9744
527 | 21087.446 | 9.9564 | 5.006 |  |  | 527 | 21080.344 | 9.9561 | 5.0396 |  |  | 527 | 21082.488 | 9.9562 | 4.9742
528 | 21127.46 | 9.9583 | 5.0056 |  |  | 528 | 21120.345 | 9.958 | 5.0394 |  |  | 528 | 21122.493 | 9.9581 | 4.974
529 | 21167.474 | 9.9602 | 5.0052 |  |  | 529 | 21160.345 | 9.9599 | 5.0393 |  |  | 529 | 21162.498 | 9.96 | 4.9738
530 | 21207.488 | 9.9621 | 5.0048 |  |  | 530 | 21200.346 | 9.9618 | 5.0391 |  |  | 530 | 21202.502 | 9.9619 | 4.9736
531 | 21247.502 | 9.964 | 5.0043 |  |  | 531 | 21240.347 | 9.9637 | 5.039 |  |  | 531 | 21242.507 | 9.9638 | 4.9734
532 | 21287.516 | 9.9659 | 5.0039 |  |  | 532 | 21280.347 | 9.9655 | 5.0388 |  |  | 532 | 21282.512 | 9.9656 | 4.9733
533 | 21327.531 | 9.9678 | 5.0035 |  |  | 533 | 21320.348 | 9.9674 | 5.0387 |  |  | 533 | 21322.517 | 9.9675 | 4.9731
534 | 21367.545 | 9.9696 | 5.0031 |  |  | 534 | 21360.349 | 9.9693 | 5.0385 |  |  | 534 | 21362.521 | 9.9694 | 4.9729
535 | 21407.559 | 9.9715 | 5.0026 |  |  | 535 | 21400.349 | 9.9712 | 5.0384 |  |  | 535 | 21402.526 | 9.9713 | 4.9727
536 | 21447.573 | 9.9734 | 5.0022 |  |  | 536 | 21440.35 | 9.973 | 5.0382 |  |  | 536 | 21442.531 | 9.9731 | 4.9726
537 | 21487.587 | 9.9752 | 5.0018 |  |  | 537 | 21480.351 | 9.9749 | 5.0381 |  |  | 537 | 21482.536 | 9.975 | 4.9724
538 | 21527.601 | 9.9771 | 5.0014 |  |  | 538 | 21520.351 | 9.9768 | 5.0379 |  |  | 538 | 21522.54 | 9.9769 | 4.9722
539 | 21567.615 | 9.9789 | 5.001 |  |  | 539 | 21560.352 | 9.9786 | 5.0378 |  |  | 539 | 21562.545 | 9.9787 | 4.972
540 | 21607.629 | 9.9808 | 5.0005 |  |  | 540 | 21600.353 | 9.9805 | 5.0376 |  |  | 540 | 21602.55 | 9.9806 | 4.9719
541 | 21647.644 | 9.9827 | 5.0001 |  |  | 541 | 21640.353 | 9.9823 | 5.0375 |  |  | 541 | 21642.554 | 9.9824 | 4.9717
542 | 21687.658 | 9.9845 | 4.9997 |  |  | 542 | 21680.354 | 9.9842 | 5.0373 |  |  | 542 | 21682.559 | 9.9843 | 4.9715
543 | 21727.672 | 9.9863 | 4.9993 |  |  | 543 | 21720.355 | 9.986 | 5.0372 |  |  | 543 | 21722.564 | 9.9861 | 4.9713
544 | 21767.686 | 9.9882 | 4.9989 |  |  | 544 | 21760.355 | 9.9878 | 5.0371 |  |  | 544 | 21762.569 | 9.9879 | 4.9712
545 | 21807.7 | 9.99 | 4.9985 |  |  | 545 | 21800.356 | 9.9897 | 5.0369 |  |  | 545 | 21802.573 | 9.9898 | 4.971
546 | 21847.714 | 9.9919 | 4.9981 |  |  | 546 | 21840.357 | 9.9915 | 5.0368 |  |  | 546 | 21842.578 | 9.9916 | 4.9708
547 | 21887.728 | 9.9937 | 4.9977 |  |  | 547 | 21880.357 | 9.9933 | 5.0366 |  |  | 547 | 21882.583 | 9.9934 | 4.9706
548 | 21927.742 | 9.9955 | 4.9972 |  |  | 548 | 21920.358 | 9.9952 | 5.0365 |  |  | 548 | 21922.587 | 9.9953 | 4.9705
549 | 21967.757 | 9.9973 | 4.9968 |  |  | 549 | 21960.359 | 9.997 | 5.0363 |  |  | 549 | 21962.592 | 9.9971 | 4.9703
550 | 22007.771 | 9.9992 | 4.9964 |  |  | 550 | 22000.359 | 9.9988 | 5.0362 |  |  | 550 | 22002.597 | 9.9989 | 4.9701
551 | 22047.785 | 10.001 | 4.996 |  |  | 551 | 22040.36 | 10.0006 | 5.0361 |  |  | 551 | 22042.602 | 10.0007 | 4.97
552 | 22087.799 | 10.0028 | 4.9956 |  |  | 552 | 22080.361 | 10.0024 | 5.0359 |  |  | 552 | 22082.606 | 10.0025 | 4.9698
553 | 22127.813 | 10.0046 | 4.9952 |  |  | 553 | 22120.361 | 10.0043 | 5.0358 |  |  | 553 | 22122.611 | 10.0044 | 4.9696
554 | 22167.827 | 10.0064 | 4.9948 |  |  | 554 | 22160.362 | 10.0061 | 5.0356 |  |  | 554 | 22162.616 | 10.0062 | 4.9695
555 | 22207.841 | 10.0082 | 4.9944 |  |  | 555 | 22200.362 | 10.0079 | 5.0355 |  |  | 555 | 22202.621 | 10.008 | 4.9693
556 | 22247.855 | 10.01 | 4.994 |  |  | 556 | 22240.363 | 10.0097 | 5.0354 |  |  | 556 | 22242.625 | 10.0098 | 4.9691
557 | 22287.87 | 10.0118 | 4.9936 |  |  | 557 | 22280.364 | 10.0115 | 5.0352 |  |  | 557 | 22282.63 | 10.0116 | 4.969
558 | 22327.884 | 10.0136 | 4.9932 |  |  | 558 | 22320.364 | 10.0133 | 5.0351 |  |  | 558 | 22322.635 | 10.0134 | 4.9688
559 | 22367.898 | 10.0154 | 4.9928 |  |  | 559 | 22360.365 | 10.015 | 5.035 |  |  | 559 | 22362.639 | 10.0151 | 4.9686
560 | 22407.912 | 10.0172 | 4.9924 |  |  | 560 | 22400.366 | 10.0168 | 5.0348 |  |  | 560 | 22402.644 | 10.0169 | 4.9685
561 | 22447.926 | 10.019 | 4.992 |  |  | 561 | 22440.366 | 10.0186 | 5.0347 |  |  | 561 | 22442.649 | 10.0187 | 4.9683
562 | 22487.94 | 10.0207 | 4.9916 |  |  | 562 | 22480.367 | 10.0204 | 5.0346 |  |  | 562 | 22482.654 | 10.0205 | 4.9682
563 | 22527.954 | 10.0225 | 4.9913 |  |  | 563 | 22520.368 | 10.0222 | 5.0344 |  |  | 563 | 22522.658 | 10.0223 | 4.968
564 | 22567.968 | 10.0243 | 4.9909 |  |  | 564 | 22560.368 | 10.0239 | 5.0343 |  |  | 564 | 22562.663 | 10.0241 | 4.9678
565 | 22607.983 | 10.0261 | 4.9905 |  |  | 565 | 22600.369 | 10.0257 | 5.0342 |  |  | 565 | 22602.668 | 10.0258 | 4.9677
566 | 22647.997 | 10.0278 | 4.9901 |  |  | 566 | 22640.37 | 10.0275 | 5.034 |  |  | 566 | 22642.672 | 10.0276 | 4.9675
567 | 22688.011 | 10.0296 | 4.9897 |  |  | 567 | 22680.37 | 10.0293 | 5.0339 |  |  | 567 | 22682.677 | 10.0294 | 4.9674
568 | 22728.025 | 10.0314 | 4.9893 |  |  | 568 | 22720.371 | 10.031 | 5.0338 |  |  | 568 | 22722.682 | 10.0311 | 4.9672
569 | 22768.039 | 10.0331 | 4.9889 |  |  | 569 | 22760.372 | 10.0328 | 5.0336 |  |  | 569 | 22762.687 | 10.0329 | 4.967
570 | 22808.053 | 10.0349 | 4.9885 |  |  | 570 | 22800.372 | 10.0345 | 5.0335 |  |  | 570 | 22802.691 | 10.0346 | 4.9669
571 | 22848.067 | 10.0366 | 4.9882 |  |  | 571 | 22840.373 | 10.0363 | 5.0334 |  |  | 571 | 22842.696 | 10.0364 | 4.9667
572 | 22888.082 | 10.0384 | 4.9878 |  |  | 572 | 22880.374 | 10.038 | 5.0333 |  |  | 572 | 22882.701 | 10.0381 | 4.9666
573 | 22928.096 | 10.0401 | 4.9874 |  |  | 573 | 22920.374 | 10.0398 | 5.0331 |  |  | 573 | 22922.705 | 10.0399 | 4.9664
574 | 22968.11 | 10.0419 | 4.987 |  |  | 574 | 22960.375 | 10.0415 | 5.033 |  |  | 574 | 22962.71 | 10.0416 | 4.9663
575 | 23008.124 | 10.0436 | 4.9866 |  |  | 575 | 23000.376 | 10.0433 | 5.0329 |  |  | 575 | 23002.715 | 10.0434 | 4.9661
576 | 23048.138 | 10.0453 | 4.9863 |  |  | 576 | 23040.376 | 10.045 | 5.0327 |  |  | 576 | 23042.72 | 10.0451 | 4.966
577 | 23088.152 | 10.0471 | 4.9859 |  |  | 577 | 23080.377 | 10.0467 | 5.0326 |  |  | 577 | 23082.724 | 10.0468 | 4.9658
578 | 23128.166 | 10.0488 | 4.9855 |  |  | 578 | 23120.377 | 10.0485 | 5.0325 |  |  | 578 | 23122.729 | 10.0486 | 4.9657
579 | 23168.18 | 10.0505 | 4.9851 |  |  | 579 | 23160.378 | 10.0502 | 5.0324 |  |  | 579 | 23162.734 | 10.0503 | 4.9655
580 | 23208.195 | 10.0523 | 4.9847 |  |  | 580 | 23200.379 | 10.0519 | 5.0322 |  |  | 580 | 23202.739 | 10.052 | 4.9653
581 | 23248.209 | 10.054 | 4.9844 |  |  | 581 | 23240.379 | 10.0536 | 5.0321 |  |  | 581 | 23242.743 | 10.0537 | 4.9652
582 | 23288.223 | 10.0557 | 4.984 |  |  | 582 | 23280.38 | 10.0554 | 5.032 |  |  | 582 | 23282.748 | 10.0555 | 4.965
583 | 23328.237 | 10.0574 | 4.9836 |  |  | 583 | 23320.381 | 10.0571 | 5.0319 |  |  | 583 | 23322.753 | 10.0572 | 4.9649
584 | 23368.251 | 10.0591 | 4.9833 |  |  | 584 | 23360.381 | 10.0588 | 5.0317 |  |  | 584 | 23362.757 | 10.0589 | 4.9647
585 | 23408.265 | 10.0608 | 4.9829 |  |  | 585 | 23400.382 | 10.0605 | 5.0316 |  |  | 585 | 23402.762 | 10.0606 | 4.9646
586 | 23448.279 | 10.0626 | 4.9825 |  |  | 586 | 23440.383 | 10.0622 | 5.0315 |  |  | 586 | 23442.767 | 10.0623 | 4.9644
587 | 23488.293 | 10.0643 | 4.9821 |  |  | 587 | 23480.383 | 10.0639 | 5.0314 |  |  | 587 | 23482.772 | 10.064 | 4.9643
588 | 23528.308 | 10.066 | 4.9818 |  |  | 588 | 23520.384 | 10.0656 | 5.0313 |  |  | 588 | 23522.776 | 10.0657 | 4.9641
589 | 23568.322 | 10.0677 | 4.9814 |  |  | 589 | 23560.385 | 10.0673 | 5.0311 |  |  | 589 | 23562.781 | 10.0674 | 4.964
590 | 23608.336 | 10.0694 | 4.981 |  |  | 590 | 23600.385 | 10.069 | 5.031 |  |  | 590 | 23602.786 | 10.0691 | 4.9639
591 | 23648.35 | 10.071 | 4.9807 |  |  | 591 | 23640.386 | 10.0707 | 5.0309 |  |  | 591 | 23642.79 | 10.0708 | 4.9637
592 | 23688.364 | 10.0727 | 4.9803 |  |  | 592 | 23680.387 | 10.0724 | 5.0308 |  |  | 592 | 23682.795 | 10.0725 | 4.9636
593 | 23728.378 | 10.0744 | 4.98 |  |  | 593 | 23720.387 | 10.0741 | 5.0307 |  |  | 593 | 23722.8 | 10.0742 | 4.9634
594 | 23768.392 | 10.0761 | 4.9796 |  |  | 594 | 23760.388 | 10.0758 | 5.0305 |  |  | 594 | 23762.805 | 10.0759 | 4.9633
595 | 23808.406 | 10.0778 | 4.9792 |  |  | 595 | 23800.389 | 10.0775 | 5.0304 |  |  | 595 | 23802.809 | 10.0776 | 4.9631
596 | 23848.421 | 10.0795 | 4.9789 |  |  | 596 | 23840.389 | 10.0791 | 5.0303 |  |  | 596 | 23842.814 | 10.0792 | 4.963
597 | 23888.435 | 10.0811 | 4.9785 |  |  | 597 | 23880.39 | 10.0808 | 5.0302 |  |  | 597 | 23882.819 | 10.0809 | 4.9628
598 | 23928.449 | 10.0828 | 4.9782 |  |  | 598 | 23920.391 | 10.0825 | 5.0301 |  |  | 598 | 23922.824 | 10.0826 | 4.9627
599 | 23968.463 | 10.0845 | 4.9778 |  |  | 599 | 23960.391 | 10.0842 | 5.0299 |  |  | 599 | 23962.828 | 10.0843 | 4.9625
600 | 24008.477 | 10.0862 | 4.9774 |  |  | 600 | 24000.392 | 10.0858 | 5.0298 |  |  | 600 | 24002.833 | 10.0859 | 4.9624
601 | 24048.491 | 10.0878 | 4.9771 |  |  | 601 | 24040.393 | 10.0875 | 5.0297 |  |  | 601 | 24042.838 | 10.0876 | 4.9623
602 | 24088.505 | 10.0895 | 4.9767 |  |  | 602 | 24080.393 | 10.0892 | 5.0296 |  |  | 602 | 24082.842 | 10.0893 | 4.9621
603 | 24128.519 | 10.0911 | 4.9764 |  |  | 603 | 24120.394 | 10.0908 | 5.0295 |  |  | 603 | 24122.847 | 10.0909 | 4.962
604 | 24168.534 | 10.0928 | 4.976 |  |  | 604 | 24160.394 | 10.0925 | 5.0294 |  |  | 604 | 24162.852 | 10.0926 | 4.9618
605 | 24208.548 | 10.0945 | 4.9757 |  |  | 605 | 24200.395 | 10.0941 | 5.0292 |  |  | 605 | 24202.857 | 10.0942 | 4.9617
606 | 24248.562 | 10.0961 | 4.9753 |  |  | 606 | 24240.396 | 10.0958 | 5.0291 |  |  | 606 | 24242.861 | 10.0959 | 4.9615
607 | 24288.576 | 10.0978 | 4.975 |  |  | 607 | 24280.396 | 10.0974 | 5.029 |  |  | 607 | 24282.866 | 10.0975 | 4.9614
608 | 24328.59 | 10.0994 | 4.9746 |  |  | 608 | 24320.397 | 10.0991 | 5.0289 |  |  | 608 | 24322.871 | 10.0992 | 4.9613
609 | 24368.604 | 10.1011 | 4.9743 |  |  | 609 | 24360.398 | 10.1007 | 5.0288 |  |  | 609 | 24362.875 | 10.1008 | 4.9611
610 | 24408.618 | 10.1027 | 4.9739 |  |  | 610 | 24400.398 | 10.1024 | 5.0287 |  |  | 610 | 24402.88 | 10.1025 | 4.961
611 | 24448.633 | 10.1043 | 4.9736 |  |  | 611 | 24440.399 | 10.104 | 5.0286 |  |  | 611 | 24442.885 | 10.1041 | 4.9608
612 | 24488.647 | 10.106 | 4.9732 |  |  | 612 | 24480.4 | 10.1056 | 5.0284 |  |  | 612 | 24482.89 | 10.1057 | 4.9607
613 | 24528.661 | 10.1076 | 4.9729 |  |  | 613 | 24520.4 | 10.1073 | 5.0283 |  |  | 613 | 24522.894 | 10.1074 | 4.9606
614 | 24568.675 | 10.1092 | 4.9726 |  |  | 614 | 24560.401 | 10.1089 | 5.0282 |  |  | 614 | 24562.899 | 10.109 | 4.9604
615 | 24608.689 | 10.1109 | 4.9722 |  |  | 615 | 24600.402 | 10.1105 | 5.0281 |  |  | 615 | 24602.904 | 10.1106 | 4.9603
616 | 24648.703 | 10.1125 | 4.9719 |  |  | 616 | 24640.402 | 10.1121 | 5.028 |  |  | 616 | 24642.909 | 10.1122 | 4.9601
617 | 24688.717 | 10.1141 | 4.9715 |  |  | 617 | 24680.403 | 10.1138 | 5.0279 |  |  | 617 | 24682.913 | 10.1139 | 4.96
618 | 24728.731 | 10.1157 | 4.9712 |  |  | 618 | 24720.404 | 10.1154 | 5.0278 |  |  | 618 | 24722.918 | 10.1155 | 4.9599
619 | 24768.746 | 10.1173 | 4.9709 |  |  | 619 | 24760.404 | 10.117 | 5.0277 |  |  | 619 | 24762.923 | 10.1171 | 4.9597
620 | 24808.76 | 10.119 | 4.9705 |  |  | 620 | 24800.405 | 10.1186 | 5.0275 |  |  | 620 | 24802.927 | 10.1187 | 4.9596
621 | 24848.774 | 10.1206 | 4.9702 |  |  | 621 | 24840.406 | 10.1202 | 5.0274 |  |  | 621 | 24842.932 | 10.1203 | 4.9595
622 | 24888.788 | 10.1222 | 4.9698 |  |  | 622 | 24880.406 | 10.1218 | 5.0273 |  |  | 622 | 24882.937 | 10.1219 | 4.9593
623 | 24928.802 | 10.1238 | 4.9695 |  |  | 623 | 24920.407 | 10.1234 | 5.0272 |  |  | 623 | 24922.942 | 10.1235 | 4.9592
624 | 24968.816 | 10.1254 | 4.9692 |  |  | 624 | 24960.408 | 10.125 | 5.0271 |  |  | 624 | 24962.946 | 10.1251 | 4.959
625 | 25008.83 | 10.127 | 4.9688 |  |  | 625 | 25000.408 | 10.1266 | 5.027 |  |  | 625 | 25002.951 | 10.1267 | 4.9589
626 | 25048.844 | 10.1286 | 4.9685 |  |  | 626 | 25040.409 | 10.1282 | 5.0269 |  |  | 626 | 25042.956 | 10.1283 | 4.9588
627 | 25088.859 | 10.1302 | 4.9682 |  |  | 627 | 25080.409 | 10.1298 | 5.0268 |  |  | 627 | 25082.96 | 10.1299 | 4.9586
628 | 25128.873 | 10.1318 | 4.9678 |  |  | 628 | 25120.41 | 10.1314 | 5.0267 |  |  | 628 | 25122.965 | 10.1315 | 4.9585
629 | 25168.887 | 10.1334 | 4.9675 |  |  | 629 | 25160.411 | 10.133 | 5.0266 |  |  | 629 | 25162.97 | 10.1331 | 4.9584
630 | 25208.901 | 10.135 | 4.9672 |  |  | 630 | 25200.411 | 10.1346 | 5.0265 |  |  | 630 | 25202.975 | 10.1347 | 4.9582
631 | 25248.915 | 10.1365 | 4.9669 |  |  | 631 | 25240.412 | 10.1362 | 5.0263 |  |  | 631 | 25242.979 | 10.1363 | 4.9581
632 | 25288.929 | 10.1381 | 4.9665 |  |  | 632 | 25280.413 | 10.1378 | 5.0262 |  |  | 632 | 25282.984 | 10.1379 | 4.958
633 | 25328.943 | 10.1397 | 4.9662 |  |  | 633 | 25320.413 | 10.1394 | 5.0261 |  |  | 633 | 25322.989 | 10.1395 | 4.9578
634 | 25368.957 | 10.1413 | 4.9659 |  |  | 634 | 25360.414 | 10.1409 | 5.026 |  |  | 634 | 25362.994 | 10.141 | 4.9577
635 | 25408.972 | 10.1429 | 4.9656 |  |  | 635 | 25400.415 | 10.1425 | 5.0259 |  |  | 635 | 25402.998 | 10.1426 | 4.9576
636 | 25448.986 | 10.1444 | 4.9652 |  |  | 636 | 25440.415 | 10.1441 | 5.0258 |  |  | 636 | 25443.003 | 10.1442 | 4.9574
637 | 25489 | 10.146 | 4.9649 |  |  | 637 | 25480.416 | 10.1457 | 5.0257 |  |  | 637 | 25483.008 | 10.1458 | 4.9573
638 | 25529.014 | 10.1476 | 4.9646 |  |  | 638 | 25520.417 | 10.1472 | 5.0256 |  |  | 638 | 25523.012 | 10.1473 | 4.9572
639 | 25569.028 | 10.1491 | 4.9643 |  |  | 639 | 25560.417 | 10.1488 | 5.0255 |  |  | 639 | 25563.017 | 10.1489 | 4.957
640 | 25609.042 | 10.1507 | 4.9639 |  |  | 640 | 25600.418 | 10.1504 | 5.0254 |  |  | 640 | 25603.022 | 10.1505 | 4.9569
641 | 25649.056 | 10.1523 | 4.9636 |  |  | 641 | 25640.419 | 10.1519 | 5.0253 |  |  | 641 | 25643.027 | 10.152 | 4.9568
642 | 25689.071 | 10.1538 | 4.9633 |  |  | 642 | 25680.419 | 10.1535 | 5.0252 |  |  | 642 | 25683.031 | 10.1536 | 4.9567
643 | 25729.085 | 10.1554 | 4.963 |  |  | 643 | 25720.42 | 10.155 | 5.0251 |  |  | 643 | 25723.036 | 10.1551 | 4.9565
644 | 25769.099 | 10.1569 | 4.9627 |  |  | 644 | 25760.421 | 10.1566 | 5.025 |  |  | 644 | 25763.041 | 10.1567 | 4.9564
645 | 25809.113 | 10.1585 | 4.9624 |  |  | 645 | 25800.421 | 10.1581 | 5.0249 |  |  | 645 | 25803.045 | 10.1582 | 4.9563
646 | 25849.127 | 10.16 | 4.962 |  |  | 646 | 25840.422 | 10.1597 | 5.0248 |  |  | 646 | 25843.05 | 10.1598 | 4.9561
647 | 25889.141 | 10.1616 | 4.9617 |  |  | 647 | 25880.423 | 10.1612 | 5.0247 |  |  | 647 | 25883.055 | 10.1613 | 4.956
648 | 25929.155 | 10.1631 | 4.9614 |  |  | 648 | 25920.423 | 10.1628 | 5.0246 |  |  | 648 | 25923.06 | 10.1629 | 4.9559
649 | 25969.169 | 10.1647 | 4.9611 |  |  | 649 | 25960.424 | 10.1643 | 5.0245 |  |  | 649 | 25963.064 | 10.1644 | 4.9557
650 | 26009.184 | 10.1662 | 4.9608 |  |  | 650 | 26000.425 | 10.1659 | 5.0243 |  |  | 650 | 26003.069 | 10.166 | 4.9556
651 | 26049.198 | 10.1677 | 4.9605 |  |  | 651 | 26040.425 | 10.1674 | 5.0242 |  |  | 651 | 26043.074 | 10.1675 | 4.9555
652 | 26089.212 | 10.1693 | 4.9602 |  |  | 652 | 26080.426 | 10.1689 | 5.0241 |  |  | 652 | 26083.079 | 10.169 | 4.9554
653 | 26129.226 | 10.1708 | 4.9599 |  |  | 653 | 26120.426 | 10.1705 | 5.024 |  |  | 653 | 26123.083 | 10.1706 | 4.9552
654 | 26169.24 | 10.1723 | 4.9595 |  |  | 654 | 26160.427 | 10.172 | 5.0239 |  |  | 654 | 26163.088 | 10.1721 | 4.9551
655 | 26209.254 | 10.1739 | 4.9592 |  |  | 655 | 26200.428 | 10.1735 | 5.0238 |  |  | 655 | 26203.093 | 10.1736 | 4.955
656 | 26249.268 | 10.1754 | 4.9589 |  |  | 656 | 26240.428 | 10.1751 | 5.0237 |  |  | 656 | 26243.097 | 10.1752 | 4.9549
657 | 26289.282 | 10.1769 | 4.9586 |  |  | 657 | 26280.429 | 10.1766 | 5.0236 |  |  | 657 | 26283.102 | 10.1767 | 4.9547
658 | 26329.297 | 10.1784 | 4.9583 |  |  | 658 | 26320.43 | 10.1781 | 5.0235 |  |  | 658 | 26323.107 | 10.1782 | 4.9546
659 | 26369.311 | 10.18 | 4.958 |  |  | 659 | 26360.43 | 10.1796 | 5.0234 |  |  | 659 | 26363.112 | 10.1797 | 4.9545
660 | 26409.325 | 10.1815 | 4.9577 |  |  | 660 | 26400.431 | 10.1811 | 5.0233 |  |  | 660 | 26403.116 | 10.1812 | 4.9543
661 | 26449.339 | 10.183 | 4.9574 |  |  | 661 | 26440.432 | 10.1826 | 5.0232 |  |  | 661 | 26443.121 | 10.1828 | 4.9542
662 | 26489.353 | 10.1845 | 4.9571 |  |  | 662 | 26480.432 | 10.1842 | 5.0231 |  |  | 662 | 26483.126 | 10.1843 | 4.9541
663 | 26529.367 | 10.186 | 4.9568 |  |  | 663 | 26520.433 | 10.1857 | 5.023 |  |  | 663 | 26523.13 | 10.1858 | 4.954
664 | 26569.381 | 10.1875 | 4.9565 |  |  | 664 | 26560.434 | 10.1872 | 5.0229 |  |  | 664 | 26563.135 | 10.1873 | 4.9538
665 | 26609.395 | 10.189 | 4.9562 |  |  | 665 | 26600.434 | 10.1887 | 5.0228 |  |  | 665 | 26603.14 | 10.1888 | 4.9537
666 | 26649.41 | 10.1905 | 4.9559 |  |  | 666 | 26640.435 | 10.1902 | 5.0227 |  |  | 666 | 26643.145 | 10.1903 | 4.9536
667 | 26689.424 | 10.192 | 4.9556 |  |  | 667 | 26680.436 | 10.1917 | 5.0226 |  |  | 667 | 26683.149 | 10.1918 | 4.9535
668 | 26729.438 | 10.1935 | 4.9553 |  |  | 668 | 26720.436 | 10.1932 | 5.0225 |  |  | 668 | 26723.154 | 10.1933 | 4.9533
669 | 26769.452 | 10.195 | 4.955 |  |  | 669 | 26760.437 | 10.1947 | 5.0224 |  |  | 669 | 26763.159 | 10.1948 | 4.9532
670 | 26809.466 | 10.1965 | 4.9547 |  |  | 670 | 26800.438 | 10.1962 | 5.0223 |  |  | 670 | 26803.163 | 10.1963 | 4.9531
671 | 26849.48 | 10.198 | 4.9544 |  |  | 671 | 26840.438 | 10.1977 | 5.0222 |  |  | 671 | 26843.168 | 10.1978 | 4.953
672 | 26889.494 | 10.1995 | 4.9541 |  |  | 672 | 26880.439 | 10.1992 | 5.0221 |  |  | 672 | 26883.173 | 10.1993 | 4.9528
673 | 26929.508 | 10.201 | 4.9538 |  |  | 673 | 26920.44 | 10.2006 | 5.022 |  |  | 673 | 26923.178 | 10.2007 | 4.9527
674 | 26969.523 | 10.2025 | 4.9535 |  |  | 674 | 26960.44 | 10.2021 | 5.0219 |  |  | 674 | 26963.182 | 10.2022 | 4.9526
675 | 27009.537 | 10.2039 | 4.9533 |  |  | 675 | 27000.441 | 10.2036 | 5.0218 |  |  | 675 | 27003.187 | 10.2037 | 4.9525
676 | 27049.551 | 10.2054 | 4.953 |  |  | 676 | 27040.441 | 10.2051 | 5.0217 |  |  | 676 | 27043.192 | 10.2052 | 4.9523
677 | 27089.565 | 10.2069 | 4.9527 |  |  | 677 | 27080.442 | 10.2066 | 5.0216 |  |  | 677 | 27083.197 | 10.2067 | 4.9522
678 | 27129.579 | 10.2084 | 4.9524 |  |  | 678 | 27120.443 | 10.208 | 5.0215 |  |  | 678 | 27123.201 | 10.2081 | 4.9521
679 | 27169.593 | 10.2099 | 4.9521 |  |  | 679 | 27160.443 | 10.2095 | 5.0214 |  |  | 679 | 27163.206 | 10.2096 | 4.952
680 | 27209.607 | 10.2113 | 4.9518 |  |  | 680 | 27200.444 | 10.211 | 5.0213 |  |  | 680 | 27203.211 | 10.2111 | 4.9519
681 | 27249.622 | 10.2128 | 4.9515 |  |  | 681 | 27240.445 | 10.2125 | 5.0212 |  |  | 681 | 27243.215 | 10.2126 | 4.9517
682 | 27289.636 | 10.2143 | 4.9512 |  |  | 682 | 27280.445 | 10.2139 | 5.0211 |  |  | 682 | 27283.22 | 10.214 | 4.9516
683 | 27329.65 | 10.2157 | 4.951 |  |  | 683 | 27320.446 | 10.2154 | 5.021 |  |  | 683 | 27323.225 | 10.2155 | 4.9515
684 | 27369.664 | 10.2172 | 4.9507 |  |  | 684 | 27360.447 | 10.2169 | 5.0209 |  |  | 684 | 27363.23 | 10.217 | 4.9514
685 | 27409.678 | 10.2187 | 4.9504 |  |  | 685 | 27400.447 | 10.2183 | 5.0209 |  |  | 685 | 27403.234 | 10.2184 | 4.9512
686 | 27449.692 | 10.2201 | 4.9501 |  |  | 686 | 27440.448 | 10.2198 | 5.0208 |  |  | 686 | 27443.239 | 10.2199 | 4.9511
687 | 27489.706 | 10.2216 | 4.9498 |  |  | 687 | 27480.449 | 10.2212 | 5.0207 |  |  | 687 | 27483.244 | 10.2213 | 4.951
688 | 27529.72 | 10.223 | 4.9495 |  |  | 688 | 27520.449 | 10.2227 | 5.0206 |  |  | 688 | 27523.248 | 10.2228 | 4.9509
689 | 27569.735 | 10.2245 | 4.9493 |  |  | 689 | 27560.45 | 10.2241 | 5.0205 |  |  | 689 | 27563.253 | 10.2242 | 4.9508
690 | 27609.749 | 10.2259 | 4.949 |  |  | 690 | 27600.451 | 10.2256 | 5.0204 |  |  | 690 | 27603.258 | 10.2257 | 4.9506
691 | 27649.763 | 10.2274 | 4.9487 |  |  | 691 | 27640.451 | 10.227 | 5.0203 |  |  | 691 | 27643.263 | 10.2271 | 4.9505
692 | 27689.777 | 10.2288 | 4.9484 |  |  | 692 | 27680.452 | 10.2285 | 5.0202 |  |  | 692 | 27683.267 | 10.2286 | 4.9504
693 | 27729.791 | 10.2303 | 4.9481 |  |  | 693 | 27720.453 | 10.2299 | 5.0201 |  |  | 693 | 27723.272 | 10.23 | 4.9503
694 | 27769.805 | 10.2317 | 4.9479 |  |  | 694 | 27760.453 | 10.2314 | 5.02 |  |  | 694 | 27763.277 | 10.2315 | 4.9501
695 | 27809.819 | 10.2331 | 4.9476 |  |  | 695 | 27800.454 | 10.2328 | 5.0199 |  |  | 695 | 27803.282 | 10.2329 | 4.95
696 | 27849.833 | 10.2346 | 4.9473 |  |  | 696 | 27840.455 | 10.2342 | 5.0198 |  |  | 696 | 27843.286 | 10.2343 | 4.9499
697 | 27889.848 | 10.236 | 4.947 |  |  | 697 | 27880.455 | 10.2357 | 5.0197 |  |  | 697 | 27883.291 | 10.2358 | 4.9498
698 | 27929.862 | 10.2375 | 4.9468 |  |  | 698 | 27920.456 | 10.2371 | 5.0196 |  |  | 698 | 27923.296 | 10.2372 | 4.9497
699 | 27969.876 | 10.2389 | 4.9465 |  |  | 699 | 27960.457 | 10.2385 | 5.0195 |  |  | 699 | 27963.3 | 10.2386 | 4.9495
700 | 28009.89 | 10.2403 | 4.9462 |  |  | 700 | 28000.457 | 10.24 | 5.0194 |  |  | 700 | 28003.305 | 10.2401 | 4.9494
701 | 28049.904 | 10.2417 | 4.9459 |  |  | 701 | 28040.458 | 10.2414 | 5.0193 |  |  | 701 | 28043.31 | 10.2415 | 4.9493
702 | 28089.918 | 10.2432 | 4.9457 |  |  | 702 | 28080.458 | 10.2428 | 5.0192 |  |  | 702 | 28083.315 | 10.2429 | 4.9492
703 | 28129.932 | 10.2446 | 4.9454 |  |  | 703 | 28120.459 | 10.2443 | 5.0191 |  |  | 703 | 28123.319 | 10.2444 | 4.9491
704 | 28169.946 | 10.246 | 4.9451 |  |  | 704 | 28160.46 | 10.2457 | 5.019 |  |  | 704 | 28163.324 | 10.2458 | 4.9489
705 | 28209.961 | 10.2474 | 4.9449 |  |  | 705 | 28200.46 | 10.2471 | 5.0189 |  |  | 705 | 28203.329 | 10.2472 | 4.9488
706 | 28249.975 | 10.2488 | 4.9446 |  |  | 706 | 28240.461 | 10.2485 | 5.0188 |  |  | 706 | 28243.333 | 10.2486 | 4.9487
707 | 28289.989 | 10.2503 | 4.9443 |  |  | 707 | 28280.462 | 10.2499 | 5.0187 |  |  | 707 | 28283.338 | 10.25 | 4.9486
708 | 28330.003 | 10.2517 | 4.9441 |  |  | 708 | 28320.462 | 10.2513 | 5.0186 |  |  | 708 | 28323.343 | 10.2514 | 4.9485
709 | 28370.017 | 10.2531 | 4.9438 |  |  | 709 | 28360.463 | 10.2528 | 5.0186 |  |  | 709 | 28363.348 | 10.2529 | 4.9483
710 | 28410.031 | 10.2545 | 4.9435 |  |  | 710 | 28400.464 | 10.2542 | 5.0185 |  |  | 710 | 28403.352 | 10.2543 | 4.9482
711 | 28450.045 | 10.2559 | 4.9433 |  |  | 711 | 28440.464 | 10.2556 | 5.0184 |  |  | 711 | 28443.357 | 10.2557 | 4.9481
712 | 28490.06 | 10.2573 | 4.943 |  |  | 712 | 28480.465 | 10.257 | 5.0183 |  |  | 712 | 28483.362 | 10.2571 | 4.948
713 | 28530.074 | 10.2587 | 4.9427 |  |  | 713 | 28520.466 | 10.2584 | 5.0182 |  |  | 713 | 28523.367 | 10.2585 | 4.9479
714 | 28570.088 | 10.2601 | 4.9425 |  |  | 714 | 28560.466 | 10.2598 | 5.0181 |  |  | 714 | 28563.371 | 10.2599 | 4.9478
715 | 28610.102 | 10.2615 | 4.9422 |  |  | 715 | 28600.467 | 10.2612 | 5.018 |  |  | 715 | 28603.376 | 10.2613 | 4.9476
716 | 28650.116 | 10.2629 | 4.942 |  |  | 716 | 28640.468 | 10.2626 | 5.0179 |  |  | 716 | 28643.381 | 10.2627 | 4.9475
717 | 28690.13 | 10.2643 | 4.9417 |  |  | 717 | 28680.468 | 10.264 | 5.0178 |  |  | 717 | 28683.385 | 10.2641 | 4.9474
718 | 28730.144 | 10.2657 | 4.9414 |  |  | 718 | 28720.469 | 10.2654 | 5.0177 |  |  | 718 | 28723.39 | 10.2655 | 4.9473
719 | 28770.158 | 10.2671 | 4.9412 |  |  | 719 | 28760.47 | 10.2668 | 5.0176 |  |  | 719 | 28763.395 | 10.2669 | 4.9472
720 | 28810.173 | 10.2685 | 4.9409 |  |  | 720 | 28800.47 | 10.2681 | 5.0175 |  |  | 720 | 28803.4 | 10.2682 | 4.947
721 | 28850.187 | 10.2699 | 4.9407 |  |  | 721 | 28840.471 | 10.2695 | 5.0174 |  |  | 721 | 28843.404 | 10.2696 | 4.9469
722 | 28890.201 | 10.2713 | 4.9404 |  |  | 722 | 28880.472 | 10.2709 | 5.0173 |  |  | 722 | 28883.409 | 10.271 | 4.9468
723 | 28930.215 | 10.2726 | 4.9402 |  |  | 723 | 28920.472 | 10.2723 | 5.0172 |  |  | 723 | 28923.414 | 10.2724 | 4.9467
724 | 28970.229 | 10.274 | 4.9399 |  |  | 724 | 28960.473 | 10.2737 | 5.0171 |  |  | 724 | 28963.418 | 10.2738 | 4.9466
725 | 29010.243 | 10.2754 | 4.9396 |  |  | 725 | 29000.473 | 10.2751 | 5.0171 |  |  | 725 | 29003.423 | 10.2752 | 4.9465
726 | 29050.257 | 10.2768 | 4.9394 |  |  | 726 | 29040.474 | 10.2764 | 5.017 |  |  | 726 | 29043.428 | 10.2765 | 4.9463
727 | 29090.271 | 10.2782 | 4.9391 |  |  | 727 | 29080.475 | 10.2778 | 5.0169 |  |  | 727 | 29083.433 | 10.2779 | 4.9462
728 | 29130.286 | 10.2795 | 4.9389 |  |  | 728 | 29120.475 | 10.2792 | 5.0168 |  |  | 728 | 29123.437 | 10.2793 | 4.9461
729 | 29170.3 | 10.2809 | 4.9386 |  |  | 729 | 29160.476 | 10.2806 | 5.0167 |  |  | 729 | 29163.442 | 10.2807 | 4.946
730 | 29210.314 | 10.2823 | 4.9384 |  |  | 730 | 29200.477 | 10.2819 | 5.0166 |  |  | 730 | 29203.447 | 10.282 | 4.9459
731 | 29250.328 | 10.2836 | 4.9381 |  |  | 731 | 29240.477 | 10.2833 | 5.0165 |  |  | 731 | 29243.452 | 10.2834 | 4.9458
732 | 29290.342 | 10.285 | 4.9379 |  |  | 732 | 29280.478 | 10.2847 | 5.0164 |  |  | 732 | 29283.456 | 10.2848 | 4.9456
733 | 29330.356 | 10.2864 | 4.9376 |  |  | 733 | 29320.479 | 10.286 | 5.0163 |  |  | 733 | 29323.461 | 10.2861 | 4.9455
734 | 29370.37 | 10.2877 | 4.9374 |  |  | 734 | 29360.479 | 10.2874 | 5.0162 |  |  | 734 | 29363.466 | 10.2875 | 4.9454
735 | 29410.384 | 10.2891 | 4.9371 |  |  | 735 | 29400.48 | 10.2888 | 5.0161 |  |  | 735 | 29403.47 | 10.2889 | 4.9453
736 | 29450.399 | 10.2905 | 4.9369 |  |  | 736 | 29440.481 | 10.2901 | 5.016 |  |  | 736 | 29443.475 | 10.2902 | 4.9452
737 | 29490.413 | 10.2918 | 4.9366 |  |  | 737 | 29480.481 | 10.2915 | 5.0159 |  |  | 737 | 29483.48 | 10.2916 | 4.9451
738 | 29530.427 | 10.2932 | 4.9364 |  |  | 738 | 29520.482 | 10.2928 | 5.0158 |  |  | 738 | 29523.485 | 10.2929 | 4.9449
739 | 29570.441 | 10.2945 | 4.9362 |  |  | 739 | 29560.483 | 10.2942 | 5.0158 |  |  | 739 | 29563.489 | 10.2943 | 4.9448
740 | 29610.455 | 10.2959 | 4.9359 |  |  | 740 | 29600.483 | 10.2955 | 5.0157 |  |  | 740 | 29603.494 | 10.2956 | 4.9447
741 | 29650.469 | 10.2972 | 4.9357 |  |  | 741 | 29640.484 | 10.2969 | 5.0156 |  |  | 741 | 29643.499 | 10.297 | 4.9446
742 | 29690.483 | 10.2986 | 4.9354 |  |  | 742 | 29680.485 | 10.2982 | 5.0155 |  |  | 742 | 29683.503 | 10.2983 | 4.9445
743 | 29730.497 | 10.2999 | 4.9352 |  |  | 743 | 29720.485 | 10.2996 | 5.0154 |  |  | 743 | 29723.508 | 10.2997 | 4.9444
744 | 29770.512 | 10.3013 | 4.9349 |  |  | 744 | 29760.486 | 10.3009 | 5.0153 |  |  | 744 | 29763.513 | 10.301 | 4.9442
745 | 29810.526 | 10.3026 | 4.9347 |  |  | 745 | 29800.487 | 10.3023 | 5.0152 |  |  | 745 | 29803.518 | 10.3024 | 4.9441
746 | 29850.54 | 10.304 | 4.9345 |  |  | 746 | 29840.487 | 10.3036 | 5.0151 |  |  | 746 | 29843.522 | 10.3037 | 4.944
747 | 29890.554 | 10.3053 | 4.9342 |  |  | 747 | 29880.488 | 10.305 | 5.015 |  |  | 747 | 29883.527 | 10.3051 | 4.9439
748 | 29930.568 | 10.3066 | 4.934 |  |  | 748 | 29920.489 | 10.3063 | 5.0149 |  |  | 748 | 29923.532 | 10.3064 | 4.9438
749 | 29970.582 | 10.308 | 4.9338 |  |  | 749 | 29960.489 | 10.3076 | 5.0148 |  |  | 749 | 29963.536 | 10.3077 | 4.9437
750 | 30010.596 | 10.3093 | 4.9335 |  |  | 750 | 30000.49 | 10.309 | 5.0147 |  |  | 750 | 30003.541 | 10.3091 | 4.9435
751 | 30050.611 | 10.3106 | 4.9333 |  |  | 751 | 30040.49 | 10.3103 | 5.0147 |  |  | 751 | 30043.546 | 10.3104 | 4.9434
752 | 30090.625 | 10.312 | 4.933 |  |  | 752 | 30080.491 | 10.3116 | 5.0146 |  |  | 752 | 30083.551 | 10.3117 | 4.9433
753 | 30130.639 | 10.3133 | 4.9328 |  |  | 753 | 30120.492 | 10.313 | 5.0145 |  |  | 753 | 30123.555 | 10.3131 | 4.9432
754 | 30170.653 | 10.3146 | 4.9326 |  |  | 754 | 30160.492 | 10.3143 | 5.0144 |  |  | 754 | 30163.56 | 10.3144 | 4.9431
755 | 30210.667 | 10.316 | 4.9323 |  |  | 755 | 30200.493 | 10.3156 | 5.0143 |  |  | 755 | 30203.565 | 10.3157 | 4.943
756 | 30250.681 | 10.3173 | 4.9321 |  |  | 756 | 30240.494 | 10.3169 | 5.0142 |  |  | 756 | 30243.57 | 10.317 | 4.9428
757 | 30290.695 | 10.3186 | 4.9319 |  |  | 757 | 30280.494 | 10.3183 | 5.0141 |  |  | 757 | 30283.574 | 10.3184 | 4.9427
758 | 30330.709 | 10.3199 | 4.9316 |  |  | 758 | 30320.495 | 10.3196 | 5.014 |  |  | 758 | 30323.579 | 10.3197 | 4.9426
759 | 30370.724 | 10.3212 | 4.9314 |  |  | 759 | 30360.496 | 10.3209 | 5.0139 |  |  | 759 | 30363.584 | 10.321 | 4.9425
760 | 30410.738 | 10.3226 | 4.9312 |  |  | 760 | 30400.496 | 10.3222 | 5.0138 |  |  | 760 | 30403.588 | 10.3223 | 4.9424
761 | 30450.752 | 10.3239 | 4.931 |  |  | 761 | 30440.497 | 10.3235 | 5.0137 |  |  | 761 | 30443.593 | 10.3236 | 4.9423
762 | 30490.766 | 10.3252 | 4.9307 |  |  | 762 | 30480.498 | 10.3248 | 5.0136 |  |  | 762 | 30483.598 | 10.3249 | 4.9422
763 | 30530.78 | 10.3265 | 4.9305 |  |  | 763 | 30520.498 | 10.3262 | 5.0136 |  |  | 763 | 30523.603 | 10.3263 | 4.942
764 | 30570.794 | 10.3278 | 4.9303 |  |  | 764 | 30560.499 | 10.3275 | 5.0135 |  |  | 764 | 30563.607 | 10.3276 | 4.9419
765 | 30610.808 | 10.3291 | 4.93 |  |  | 765 | 30600.5 | 10.3288 | 5.0134 |  |  | 765 | 30603.612 | 10.3289 | 4.9418
766 | 30650.822 | 10.3304 | 4.9298 |  |  | 766 | 30640.5 | 10.3301 | 5.0133 |  |  | 766 | 30643.617 | 10.3302 | 4.9417
767 | 30690.837 | 10.3317 | 4.9296 |  |  | 767 | 30680.501 | 10.3314 | 5.0132 |  |  | 767 | 30683.621 | 10.3315 | 4.9416
768 | 30730.851 | 10.333 | 4.9294 |  |  | 768 | 30720.502 | 10.3327 | 5.0131 |  |  | 768 | 30723.626 | 10.3328 | 4.9415
769 | 30770.865 | 10.3343 | 4.9291 |  |  | 769 | 30760.502 | 10.334 | 5.013 |  |  | 769 | 30763.631 | 10.3341 | 4.9414
770 | 30810.879 | 10.3356 | 4.9289 |  |  | 770 | 30800.503 | 10.3353 | 5.0129 |  |  | 770 | 30803.636 | 10.3354 | 4.9412
771 | 30850.893 | 10.3369 | 4.9287 |  |  | 771 | 30840.504 | 10.3366 | 5.0128 |  |  | 771 | 30843.64 | 10.3367 | 4.9411
772 | 30890.907 | 10.3382 | 4.9285 |  |  | 772 | 30880.504 | 10.3379 | 5.0127 |  |  | 772 | 30883.645 | 10.338 | 4.941
773 | 30930.921 | 10.3395 | 4.9282 |  |  | 773 | 30920.505 | 10.3392 | 5.0126 |  |  | 773 | 30923.65 | 10.3393 | 4.9409
774 | 30970.935 | 10.3408 | 4.928 |  |  | 774 | 30960.505 | 10.3405 | 5.0126 |  |  | 774 | 30963.655 | 10.3406 | 4.9408
775 | 31010.95 | 10.3421 | 4.9278 |  |  | 775 | 31000.506 | 10.3418 | 5.0125 |  |  | 775 | 31003.659 | 10.3419 | 4.9407
776 | 31050.964 | 10.3434 | 4.9276 |  |  | 776 | 31040.507 | 10.343 | 5.0124 |  |  | 776 | 31043.664 | 10.3431 | 4.9405
777 | 31090.978 | 10.3447 | 4.9274 |  |  | 777 | 31080.507 | 10.3443 | 5.0123 |  |  | 777 | 31083.669 | 10.3444 | 4.9404
778 | 31130.992 | 10.346 | 4.9271 |  |  | 778 | 31120.508 | 10.3456 | 5.0122 |  |  | 778 | 31123.673 | 10.3457 | 4.9403
779 | 31171.006 | 10.3472 | 4.9269 |  |  | 779 | 31160.509 | 10.3469 | 5.0121 |  |  | 779 | 31163.678 | 10.347 | 4.9402
780 | 31211.02 | 10.3485 | 4.9267 |  |  | 780 | 31200.509 | 10.3482 | 5.012 |  |  | 780 | 31203.683 | 10.3483 | 4.9401
781 | 31251.034 | 10.3498 | 4.9265 |  |  | 781 | 31240.51 | 10.3495 | 5.0119 |  |  | 781 | 31243.688 | 10.3496 | 4.94
782 | 31291.049 | 10.3511 | 4.9263 |  |  | 782 | 31280.511 | 10.3508 | 5.0118 |  |  | 782 | 31283.692 | 10.3509 | 4.9399
783 | 31331.063 | 10.3524 | 4.9261 |  |  | 783 | 31320.511 | 10.352 | 5.0117 |  |  | 783 | 31323.697 | 10.3521 | 4.9397
784 | 31371.077 | 10.3536 | 4.9258 |  |  | 784 | 31360.512 | 10.3533 | 5.0116 |  |  | 784 | 31363.702 | 10.3534 | 4.9396
785 | 31411.091 | 10.3549 | 4.9256 |  |  | 785 | 31400.513 | 10.3546 | 5.0116 |  |  | 785 | 31403.706 | 10.3547 | 4.9395
786 | 31451.105 | 10.3562 | 4.9254 |  |  | 786 | 31440.513 | 10.3559 | 5.0115 |  |  | 786 | 31443.711 | 10.356 | 4.9394
787 | 31491.119 | 10.3575 | 4.9252 |  |  | 787 | 31480.514 | 10.3571 | 5.0114 |  |  | 787 | 31483.716 | 10.3572 | 4.9393
788 | 31531.133 | 10.3587 | 4.925 |  |  | 788 | 31520.515 | 10.3584 | 5.0113 |  |  | 788 | 31523.721 | 10.3585 | 4.9392
789 | 31571.147 | 10.36 | 4.9248 |  |  | 789 | 31560.515 | 10.3597 | 5.0112 |  |  | 789 | 31563.725 | 10.3598 | 4.9391
790 | 31611.162 | 10.3613 | 4.9246 |  |  | 790 | 31600.516 | 10.3609 | 5.0111 |  |  | 790 | 31603.73 | 10.361 | 4.9389
791 | 31651.176 | 10.3625 | 4.9244 |  |  | 791 | 31640.517 | 10.3622 | 5.011 |  |  | 791 | 31643.735 | 10.3623 | 4.9388
792 | 31691.19 | 10.3638 | 4.9242 |  |  | 792 | 31680.517 | 10.3635 | 5.0109 |  |  | 792 | 31683.74 | 10.3636 | 4.9387
793 | 31731.204 | 10.3651 | 4.9239 |  |  | 793 | 31720.518 | 10.3647 | 5.0108 |  |  | 793 | 31723.744 | 10.3648 | 4.9386
794 | 31771.218 | 10.3663 | 4.9237 |  |  | 794 | 31760.519 | 10.366 | 5.0107 |  |  | 794 | 31763.749 | 10.3661 | 4.9385
795 | 31811.232 | 10.3676 | 4.9235 |  |  | 795 | 31800.519 | 10.3672 | 5.0107 |  |  | 795 | 31803.754 | 10.3673 | 4.9384
796 | 31851.246 | 10.3688 | 4.9233 |  |  | 796 | 31840.52 | 10.3685 | 5.0106 |  |  | 796 | 31843.758 | 10.3686 | 4.9383
797 | 31891.26 | 10.3701 | 4.9231 |  |  | 797 | 31880.521 | 10.3698 | 5.0105 |  |  | 797 | 31883.763 | 10.3699 | 4.9381
798 | 31931.275 | 10.3713 | 4.9229 |  |  | 798 | 31920.521 | 10.371 | 5.0104 |  |  | 798 | 31923.768 | 10.3711 | 4.938
799 | 31971.289 | 10.3726 | 4.9227 |  |  | 799 | 31960.522 | 10.3723 | 5.0103 |  |  | 799 | 31963.773 | 10.3724 | 4.9379
800 | 32011.303 | 10.3738 | 4.9225 |  |  | 800 | 32000.522 | 10.3735 | 5.0102 |  |  | 800 | 32003.777 | 10.3736 | 4.9378
801 | 32051.317 | 10.3751 | 4.9223 |  |  | 801 | 32040.523 | 10.3748 | 5.0101 |  |  | 801 | 32043.782 | 10.3749 | 4.9377
802 | 32091.331 | 10.3763 | 4.9221 |  |  | 802 | 32080.524 | 10.376 | 5.01 |  |  | 802 | 32083.787 | 10.3761 | 4.9376
803 | 32131.345 | 10.3776 | 4.9219 |  |  | 803 | 32120.524 | 10.3773 | 5.0099 |  |  | 803 | 32123.791 | 10.3774 | 4.9375
804 | 32171.359 | 10.3788 | 4.9217 |  |  | 804 | 32160.525 | 10.3785 | 5.0098 |  |  | 804 | 32163.796 | 10.3786 | 4.9373
805 | 32211.373 | 10.3801 | 4.9215 |  |  | 805 | 32200.526 | 10.3797 | 5.0097 |  |  | 805 | 32203.801 | 10.3798 | 4.9372
806 | 32251.388 | 10.3813 | 4.9213 |  |  | 806 | 32240.526 | 10.381 | 5.0097 |  |  | 806 | 32243.806 | 10.3811 | 4.9371
807 | 32291.402 | 10.3826 | 4.9211 |  |  | 807 | 32280.527 | 10.3822 | 5.0096 |  |  | 807 | 32283.81 | 10.3823 | 4.937
808 | 32331.416 | 10.3838 | 4.9209 |  |  | 808 | 32320.528 | 10.3835 | 5.0095 |  |  | 808 | 32323.815 | 10.3836 | 4.9369
809 | 32371.43 | 10.385 | 4.9207 |  |  | 809 | 32360.528 | 10.3847 | 5.0094 |  |  | 809 | 32363.82 | 10.3848 | 4.9368
810 | 32411.444 | 10.3863 | 4.9205 |  |  | 810 | 32400.529 | 10.3859 | 5.0093 |  |  | 810 | 32403.825 | 10.386 | 4.9367
811 | 32451.458 | 10.3875 | 4.9203 |  |  | 811 | 32440.53 | 10.3872 | 5.0092 |  |  | 811 | 32443.829 | 10.3873 | 4.9365
812 | 32491.472 | 10.3887 | 4.9201 |  |  | 812 | 32480.53 | 10.3884 | 5.0091 |  |  | 812 | 32483.834 | 10.3885 | 4.9364
813 | 32531.486 | 10.39 | 4.9199 |  |  | 813 | 32520.531 | 10.3896 | 5.009 |  |  | 813 | 32523.839 | 10.3897 | 4.9363
814 | 32571.501 | 10.3912 | 4.9197 |  |  | 814 | 32560.532 | 10.3909 | 5.0089 |  |  | 814 | 32563.843 | 10.391 | 4.9362
815 | 32611.515 | 10.3924 | 4.9195 |  |  | 815 | 32600.532 | 10.3921 | 5.0088 |  |  | 815 | 32603.848 | 10.3922 | 4.9361
816 | 32651.529 | 10.3936 | 4.9193 |  |  | 816 | 32640.533 | 10.3933 | 5.0088 |  |  | 816 | 32643.853 | 10.3934 | 4.936
817 | 32691.543 | 10.3949 | 4.9191 |  |  | 817 | 32680.534 | 10.3945 | 5.0087 |  |  | 817 | 32683.858 | 10.3946 | 4.9359
818 | 32731.557 | 10.3961 | 4.9189 |  |  | 818 | 32720.534 | 10.3958 | 5.0086 |  |  | 818 | 32723.862 | 10.3959 | 4.9357
819 | 32771.571 | 10.3973 | 4.9187 |  |  | 819 | 32760.535 | 10.397 | 5.0085 |  |  | 819 | 32763.867 | 10.3971 | 4.9356
820 | 32811.585 | 10.3985 | 4.9185 |  |  | 820 | 32800.536 | 10.3982 | 5.0084 |  |  | 820 | 32803.872 | 10.3983 | 4.9355
821 | 32851.6 | 10.3998 | 4.9183 |  |  | 821 | 32840.536 | 10.3994 | 5.0083 |  |  | 821 | 32843.876 | 10.3995 | 4.9354
822 | 32891.614 | 10.401 | 4.9181 |  |  | 822 | 32880.537 | 10.4006 | 5.0082 |  |  | 822 | 32883.881 | 10.4007 | 4.9353
823 | 32931.628 | 10.4022 | 4.9179 |  |  | 823 | 32920.537 | 10.4019 | 5.0081 |  |  | 823 | 32923.886 | 10.402 | 4.9352
824 | 32971.642 | 10.4034 | 4.9177 |  |  | 824 | 32960.538 | 10.4031 | 5.008 |  |  | 824 | 32963.891 | 10.4032 | 4.9351
825 | 33011.656 | 10.4046 | 4.9176 |  |  | 825 | 33000.539 | 10.4043 | 5.0079 |  |  | 825 | 33003.895 | 10.4044 | 4.9349
826 | 33051.67 | 10.4058 | 4.9174 |  |  | 826 | 33040.539 | 10.4055 | 5.0078 |  |  | 826 | 33043.9 | 10.4056 | 4.9348
827 | 33091.684 | 10.407 | 4.9172 |  |  | 827 | 33080.54 | 10.4067 | 5.0078 |  |  | 827 | 33083.905 | 10.4068 | 4.9347
828 | 33131.698 | 10.4082 | 4.917 |  |  | 828 | 33120.541 | 10.4079 | 5.0077 |  |  | 828 | 33123.91 | 10.408 | 4.9346
829 | 33171.713 | 10.4095 | 4.9168 |  |  | 829 | 33160.541 | 10.4091 | 5.0076 |  |  | 829 | 33163.914 | 10.4092 | 4.9345
830 | 33211.727 | 10.4107 | 4.9166 |  |  | 830 | 33200.542 | 10.4103 | 5.0075 |  |  | 830 | 33203.919 | 10.4104 | 4.9344
831 | 33251.741 | 10.4119 | 4.9164 |  |  | 831 | 33240.543 | 10.4115 | 5.0074 |  |  | 831 | 33243.924 | 10.4116 | 4.9343
832 | 33291.755 | 10.4131 | 4.9162 |  |  | 832 | 33280.543 | 10.4127 | 5.0073 |  |  | 832 | 33283.928 | 10.4128 | 4.9341
833 | 33331.769 | 10.4143 | 4.9161 |  |  | 833 | 33320.544 | 10.4139 | 5.0072 |  |  | 833 | 33323.933 | 10.414 | 4.934
834 | 33371.783 | 10.4155 | 4.9159 |  |  | 834 | 33360.545 | 10.4151 | 5.0071 |  |  | 834 | 33363.938 | 10.4152 | 4.9339
835 | 33411.797 | 10.4167 | 4.9157 |  |  | 835 | 33400.545 | 10.4163 | 5.007 |  |  | 835 | 33403.943 | 10.4164 | 4.9338
836 | 33451.811 | 10.4179 | 4.9155 |  |  | 836 | 33440.546 | 10.4175 | 5.0069 |  |  | 836 | 33443.947 | 10.4176 | 4.9337
837 | 33491.826 | 10.4191 | 4.9153 |  |  | 837 | 33480.547 | 10.4187 | 5.0068 |  |  | 837 | 33483.952 | 10.4188 | 4.9336
838 | 33531.84 | 10.4203 | 4.9151 |  |  | 838 | 33520.547 | 10.4199 | 5.0068 |  |  | 838 | 33523.957 | 10.42 | 4.9335
839 | 33571.854 | 10.4214 | 4.9149 |  |  | 839 | 33560.548 | 10.4211 | 5.0067 |  |  | 839 | 33563.961 | 10.4212 | 4.9333
840 | 33611.868 | 10.4226 | 4.9148 |  |  | 840 | 33600.549 | 10.4223 | 5.0066 |  |  | 840 | 33603.966 | 10.4224 | 4.9332
841 | 33651.882 | 10.4238 | 4.9146 |  |  | 841 | 33640.549 | 10.4235 | 5.0065 |  |  | 841 | 33643.971 | 10.4236 | 4.9331
842 | 33691.896 | 10.425 | 4.9144 |  |  | 842 | 33680.55 | 10.4247 | 5.0064 |  |  | 842 | 33683.976 | 10.4248 | 4.933
843 | 33731.91 | 10.4262 | 4.9142 |  |  | 843 | 33720.551 | 10.4259 | 5.0063 |  |  | 843 | 33723.98 | 10.426 | 4.9329
844 | 33771.924 | 10.4274 | 4.914 |  |  | 844 | 33760.551 | 10.427 | 5.0062 |  |  | 844 | 33763.985 | 10.4271 | 4.9328
845 | 33811.939 | 10.4286 | 4.9139 |  |  | 845 | 33800.552 | 10.4282 | 5.0061 |  |  | 845 | 33803.99 | 10.4283 | 4.9327
846 | 33851.953 | 10.4298 | 4.9137 |  |  | 846 | 33840.553 | 10.4294 | 5.006 |  |  | 846 | 33843.994 | 10.4295 | 4.9325
847 | 33891.967 | 10.4309 | 4.9135 |  |  | 847 | 33880.553 | 10.4306 | 5.0059 |  |  | 847 | 33883.999 | 10.4307 | 4.9324
848 | 33931.981 | 10.4321 | 4.9133 |  |  | 848 | 33920.554 | 10.4318 | 5.0058 |  |  | 848 | 33924.004 | 10.4319 | 4.9323
849 | 33971.995 | 10.4333 | 4.9132 |  |  | 849 | 33960.554 | 10.433 | 5.0058 |  |  | 849 | 33964.009 | 10.4331 | 4.9322
850 | 34012.009 | 10.4345 | 4.913 |  |  | 850 | 34000.555 | 10.4341 | 5.0057 |  |  | 850 | 34004.013 | 10.4342 | 4.9321
851 | 34052.023 | 10.4356 | 4.9128 |  |  | 851 | 34040.556 | 10.4353 | 5.0056 |  |  | 851 | 34044.018 | 10.4354 | 4.932
852 | 34092.038 | 10.4368 | 4.9126 |  |  | 852 | 34080.556 | 10.4365 | 5.0055 |  |  | 852 | 34084.023 | 10.4366 | 4.9319
853 | 34132.052 | 10.438 | 4.9124 |  |  | 853 | 34120.557 | 10.4377 | 5.0054 |  |  | 853 | 34124.028 | 10.4378 | 4.9317
854 | 34172.066 | 10.4392 | 4.9123 |  |  | 854 | 34160.558 | 10.4388 | 5.0053 |  |  | 854 | 34164.032 | 10.4389 | 4.9316
855 | 34212.08 | 10.4403 | 4.9121 |  |  | 855 | 34200.558 | 10.44 | 5.0052 |  |  | 855 | 34204.037 | 10.4401 | 4.9315
856 | 34252.094 | 10.4415 | 4.9119 |  |  | 856 | 34240.559 | 10.4412 | 5.0051 |  |  | 856 | 34244.042 | 10.4413 | 4.9314
857 | 34292.108 | 10.4427 | 4.9118 |  |  | 857 | 34280.56 | 10.4423 | 5.005 |  |  | 857 | 34284.046 | 10.4424 | 4.9313
858 | 34332.122 | 10.4438 | 4.9116 |  |  | 858 | 34320.56 | 10.4435 | 5.0049 |  |  | 858 | 34324.051 | 10.4436 | 4.9312
859 | 34372.136 | 10.445 | 4.9114 |  |  | 859 | 34360.561 | 10.4447 | 5.0048 |  |  | 859 | 34364.056 | 10.4448 | 4.9311
860 | 34412.151 | 10.4462 | 4.9112 |  |  | 860 | 34400.562 | 10.4458 | 5.0047 |  |  | 860 | 34404.061 | 10.4459 | 4.9309
861 | 34452.165 | 10.4473 | 4.9111 |  |  | 861 | 34440.562 | 10.447 | 5.0047 |  |  | 861 | 34444.065 | 10.4471 | 4.9308
862 | 34492.179 | 10.4485 | 4.9109 |  |  | 862 | 34480.563 | 10.4482 | 5.0046 |  |  | 862 | 34484.07 | 10.4483 | 4.9307
863 | 34532.193 | 10.4496 | 4.9107 |  |  | 863 | 34520.564 | 10.4493 | 5.0045 |  |  | 863 | 34524.075 | 10.4494 | 4.9306
864 | 34572.207 | 10.4508 | 4.9106 |  |  | 864 | 34560.564 | 10.4505 | 5.0044 |  |  | 864 | 34564.079 | 10.4506 | 4.9305
865 | 34612.221 | 10.452 | 4.9104 |  |  | 865 | 34600.565 | 10.4516 | 5.0043 |  |  | 865 | 34604.084 | 10.4517 | 4.9304
866 | 34652.235 | 10.4531 | 4.9102 |  |  | 866 | 34640.566 | 10.4528 | 5.0042 |  |  | 866 | 34644.089 | 10.4529 | 4.9303
867 | 34692.249 | 10.4543 | 4.9101 |  |  | 867 | 34680.566 | 10.4539 | 5.0041 |  |  | 867 | 34684.094 | 10.454 | 4.9301
868 | 34732.264 | 10.4554 | 4.9099 |  |  | 868 | 34720.567 | 10.4551 | 5.004 |  |  | 868 | 34724.098 | 10.4552 | 4.93
869 | 34772.278 | 10.4566 | 4.9097 |  |  | 869 | 34760.568 | 10.4562 | 5.0039 |  |  | 869 | 34764.103 | 10.4563 | 4.9299
870 | 34812.292 | 10.4577 | 4.9096 |  |  | 870 | 34800.568 | 10.4574 | 5.0038 |  |  | 870 | 34804.108 | 10.4575 | 4.9298
871 | 34852.306 | 10.4589 | 4.9094 |  |  | 871 | 34840.569 | 10.4585 | 5.0037 |  |  | 871 | 34844.113 | 10.4586 | 4.9297
872 | 34892.32 | 10.46 | 4.9092 |  |  | 872 | 34880.569 | 10.4597 | 5.0036 |  |  | 872 | 34884.117 | 10.4598 | 4.9296
873 | 34932.334 | 10.4612 | 4.9091 |  |  | 873 | 34920.57 | 10.4608 | 5.0036 |  |  | 873 | 34924.122 | 10.4609 | 4.9294
874 | 34972.348 | 10.4623 | 4.9089 |  |  | 874 | 34960.571 | 10.462 | 5.0035 |  |  | 874 | 34964.127 | 10.4621 | 4.9293
875 | 35012.362 | 10.4635 | 4.9087 |  |  | 875 | 35000.571 | 10.4631 | 5.0034 |  |  | 875 | 35004.131 | 10.4632 | 4.9292
876 | 35052.377 | 10.4646 | 4.9086 |  |  | 876 | 35040.572 | 10.4643 | 5.0033 |  |  | 876 | 35044.136 | 10.4644 | 4.9291
877 | 35092.391 | 10.4657 | 4.9084 |  |  | 877 | 35080.573 | 10.4654 | 5.0032 |  |  | 877 | 35084.141 | 10.4655 | 4.929
878 | 35132.405 | 10.4669 | 4.9082 |  |  | 878 | 35120.573 | 10.4665 | 5.0031 |  |  | 878 | 35124.146 | 10.4666 | 4.9289
879 | 35172.419 | 10.468 | 4.9081 |  |  | 879 | 35160.574 | 10.4677 | 5.003 |  |  | 879 | 35164.15 | 10.4678 | 4.9288
880 | 35212.433 | 10.4692 | 4.9079 |  |  | 880 | 35200.575 | 10.4688 | 5.0029 |  |  | 880 | 35204.155 | 10.4689 | 4.9286
881 | 35252.447 | 10.4703 | 4.9078 |  |  | 881 | 35240.575 | 10.47 | 5.0028 |  |  | 881 | 35244.16 | 10.4701 | 4.9285
882 | 35292.461 | 10.4714 | 4.9076 |  |  | 882 | 35280.576 | 10.4711 | 5.0027 |  |  | 882 | 35284.164 | 10.4712 | 4.9284
883 | 35332.475 | 10.4726 | 4.9074 |  |  | 883 | 35320.577 | 10.4722 | 5.0026 |  |  | 883 | 35324.169 | 10.4723 | 4.9283
884 | 35372.49 | 10.4737 | 4.9073 |  |  | 884 | 35360.577 | 10.4734 | 5.0025 |  |  | 884 | 35364.174 | 10.4735 | 4.9282
885 | 35412.504 | 10.4748 | 4.9071 |  |  | 885 | 35400.578 | 10.4745 | 5.0025 |  |  | 885 | 35404.179 | 10.4746 | 4.9281
886 | 35452.518 | 10.4759 | 4.907 |  |  | 886 | 35440.579 | 10.4756 | 5.0024 |  |  | 886 | 35444.183 | 10.4757 | 4.9279
887 | 35492.532 | 10.4771 | 4.9068 |  |  | 887 | 35480.579 | 10.4767 | 5.0023 |  |  | 887 | 35484.188 | 10.4768 | 4.9278
888 | 35532.546 | 10.4782 | 4.9066 |  |  | 888 | 35520.58 | 10.4779 | 5.0022 |  |  | 888 | 35524.193 | 10.478 | 4.9277
889 | 35572.56 | 10.4793 | 4.9065 |  |  | 889 | 35560.581 | 10.479 | 5.0021 |  |  | 889 | 35564.198 | 10.4791 | 4.9276
890 | 35612.574 | 10.4805 | 4.9063 |  |  | 890 | 35600.581 | 10.4801 | 5.002 |  |  | 890 | 35604.202 | 10.4802 | 4.9275
891 | 35652.589 | 10.4816 | 4.9062 |  |  | 891 | 35640.582 | 10.4812 | 5.0019 |  |  | 891 | 35644.207 | 10.4813 | 4.9274
892 | 35692.603 | 10.4827 | 4.906 |  |  | 892 | 35680.583 | 10.4824 | 5.0018 |  |  | 892 | 35684.212 | 10.4825 | 4.9273
893 | 35732.617 | 10.4838 | 4.9059 |  |  | 893 | 35720.583 | 10.4835 | 5.0017 |  |  | 893 | 35724.216 | 10.4836 | 4.9271
894 | 35772.631 | 10.4849 | 4.9057 |  |  | 894 | 35760.584 | 10.4846 | 5.0016 |  |  | 894 | 35764.221 | 10.4847 | 4.927
895 | 35812.645 | 10.4861 | 4.9056 |  |  | 895 | 35800.585 | 10.4857 | 5.0015 |  |  | 895 | 35804.226 | 10.4858 | 4.9269
896 | 35852.659 | 10.4872 | 4.9054 |  |  | 896 | 35840.585 | 10.4868 | 5.0014 |  |  | 896 | 35844.231 | 10.4869 | 4.9268
897 | 35892.673 | 10.4883 | 4.9053 |  |  | 897 | 35880.586 | 10.488 | 5.0013 |  |  | 897 | 35884.235 | 10.4881 | 4.9267
898 | 35932.687 | 10.4894 | 4.9051 |  |  | 898 | 35920.586 | 10.4891 | 5.0013 |  |  | 898 | 35924.24 | 10.4892 | 4.9266
899 | 35972.702 | 10.4905 | 4.905 |  |  | 899 | 35960.587 | 10.4902 | 5.0012 |  |  | 899 | 35964.245 | 10.4903 | 4.9264
900 | 36012.716 | 10.4916 | 4.9048 |  |  | 900 | 36000.588 | 10.4913 | 5.0011 |  |  | 900 | 36004.249 | 10.4914 | 4.9263
901 | 36052.73 | 10.4927 | 4.9047 |  |  | 901 | 36040.588 | 10.4924 | 5.001 |  |  | 901 | 36044.254 | 10.4925 | 4.9262
902 | 36092.744 | 10.4938 | 4.9045 |  |  | 902 | 36080.589 | 10.4935 | 5.0009 |  |  | 902 | 36084.259 | 10.4936 | 4.9261
903 | 36132.758 | 10.495 | 4.9043 |  |  | 903 | 36120.59 | 10.4946 | 5.0008 |  |  | 903 | 36124.264 | 10.4947 | 4.926
904 | 36172.772 | 10.4961 | 4.9042 |  |  | 904 | 36160.59 | 10.4957 | 5.0007 |  |  | 904 | 36164.268 | 10.4958 | 4.9259
905 | 36212.786 | 10.4972 | 4.9041 |  |  | 905 | 36200.591 | 10.4968 | 5.0006 |  |  | 905 | 36204.273 | 10.4969 | 4.9257
906 | 36252.8 | 10.4983 | 4.9039 |  |  | 906 | 36240.592 | 10.4979 | 5.0005 |  |  | 906 | 36244.278 | 10.498 | 4.9256
907 | 36292.815 | 10.4994 | 4.9038 |  |  | 907 | 36280.592 | 10.499 | 5.0004 |  |  | 907 | 36284.283 | 10.4991 | 4.9255
908 | 36332.829 | 10.5005 | 4.9036 |  |  | 908 | 36320.593 | 10.5001 | 5.0003 |  |  | 908 | 36324.287 | 10.5002 | 4.9254
909 | 36372.843 | 10.5016 | 4.9035 |  |  | 909 | 36360.594 | 10.5012 | 5.0002 |  |  | 909 | 36364.292 | 10.5013 | 4.9253
910 | 36412.857 | 10.5027 | 4.9033 |  |  | 910 | 36400.594 | 10.5023 | 5.0001 |  |  | 910 | 36404.297 | 10.5024 | 4.9252
911 | 36452.871 | 10.5038 | 4.9032 |  |  | 911 | 36440.595 | 10.5034 | 5 |  |  | 911 | 36444.301 | 10.5035 | 4.9251
912 | 36492.885 | 10.5049 | 4.903 |  |  | 912 | 36480.596 | 10.5045 | 5 |  |  | 912 | 36484.306 | 10.5046 | 4.9249
913 | 36532.899 | 10.506 | 4.9029 |  |  | 913 | 36520.596 | 10.5056 | 4.9999 |  |  | 913 | 36524.311 | 10.5057 | 4.9248
914 | 36572.913 | 10.5071 | 4.9027 |  |  | 914 | 36560.597 | 10.5067 | 4.9998 |  |  | 914 | 36564.316 | 10.5068 | 4.9247
915 | 36612.928 | 10.5082 | 4.9026 |  |  | 915 | 36600.598 | 10.5078 | 4.9997 |  |  | 915 | 36604.32 | 10.5079 | 4.9246
916 | 36652.942 | 10.5092 | 4.9024 |  |  | 916 | 36640.598 | 10.5089 | 4.9996 |  |  | 916 | 36644.325 | 10.509 | 4.9245
917 | 36692.956 | 10.5103 | 4.9023 |  |  | 917 | 36680.599 | 10.51 | 4.9995 |  |  | 917 | 36684.33 | 10.5101 | 4.9244
918 | 36732.97 | 10.5114 | 4.9022 |  |  | 918 | 36720.6 | 10.5111 | 4.9994 |  |  | 918 | 36724.334 | 10.5112 | 4.9242
919 | 36772.984 | 10.5125 | 4.902 |  |  | 919 | 36760.6 | 10.5122 | 4.9993 |  |  | 919 | 36764.339 | 10.5123 | 4.9241
920 | 36812.998 | 10.5136 | 4.9019 |  |  | 920 | 36800.601 | 10.5133 | 4.9992 |  |  | 920 | 36804.344 | 10.5134 | 4.924
921 | 36853.012 | 10.5147 | 4.9017 |  |  | 921 | 36840.601 | 10.5144 | 4.9991 |  |  | 921 | 36844.349 | 10.5145 | 4.9239
922 | 36893.026 | 10.5158 | 4.9016 |  |  | 922 | 36880.602 | 10.5154 | 4.999 |  |  | 922 | 36884.353 | 10.5155 | 4.9238
923 | 36933.041 | 10.5169 | 4.9015 |  |  | 923 | 36920.603 | 10.5165 | 4.9989 |  |  | 923 | 36924.358 | 10.5166 | 4.9237
924 | 36973.055 | 10.5179 | 4.9013 |  |  | 924 | 36960.603 | 10.5176 | 4.9988 |  |  | 924 | 36964.363 | 10.5177 | 4.9235
925 | 37013.069 | 10.519 | 4.9012 |  |  | 925 | 37000.604 | 10.5187 | 4.9987 |  |  | 925 | 37004.368 | 10.5188 | 4.9234
926 | 37053.083 | 10.5201 | 4.901 |  |  | 926 | 37040.605 | 10.5198 | 4.9986 |  |  | 926 | 37044.372 | 10.5199 | 4.9233
927 | 37093.097 | 10.5212 | 4.9009 |  |  | 927 | 37080.605 | 10.5208 | 4.9985 |  |  | 927 | 37084.377 | 10.521 | 4.9232
928 | 37133.111 | 10.5223 | 4.9008 |  |  | 928 | 37120.606 | 10.5219 | 4.9985 |  |  | 928 | 37124.382 | 10.522 | 4.9231
929 | 37173.125 | 10.5233 | 4.9006 |  |  | 929 | 37160.607 | 10.523 | 4.9984 |  |  | 929 | 37164.386 | 10.5231 | 4.923
930 | 37213.14 | 10.5244 | 4.9005 |  |  | 930 | 37200.607 | 10.5241 | 4.9983 |  |  | 930 | 37204.391 | 10.5242 | 4.9228
931 | 37253.154 | 10.5255 | 4.9003 |  |  | 931 | 37240.608 | 10.5252 | 4.9982 |  |  | 931 | 37244.396 | 10.5253 | 4.9227
932 | 37293.168 | 10.5266 | 4.9002 |  |  | 932 | 37280.609 | 10.5262 | 4.9981 |  |  | 932 | 37284.401 | 10.5263 | 4.9226
933 | 37333.182 | 10.5276 | 4.9001 |  |  | 933 | 37320.609 | 10.5273 | 4.998 |  |  | 933 | 37324.405 | 10.5274 | 4.9225
934 | 37373.196 | 10.5287 | 4.8999 |  |  | 934 | 37360.61 | 10.5284 | 4.9979 |  |  | 934 | 37364.41 | 10.5285 | 4.9224
935 | 37413.21 | 10.5298 | 4.8998 |  |  | 935 | 37400.611 | 10.5294 | 4.9978 |  |  | 935 | 37404.415 | 10.5295 | 4.9222
936 | 37453.224 | 10.5308 | 4.8997 |  |  | 936 | 37440.611 | 10.5305 | 4.9977 |  |  | 936 | 37444.419 | 10.5306 | 4.9221
937 | 37493.238 | 10.5319 | 4.8995 |  |  | 937 | 37480.612 | 10.5316 | 4.9976 |  |  | 937 | 37484.424 | 10.5317 | 4.922
938 | 37533.253 | 10.533 | 4.8994 |  |  | 938 | 37520.613 | 10.5326 | 4.9975 |  |  | 938 | 37524.429 | 10.5327 | 4.9219
939 | 37573.267 | 10.534 | 4.8993 |  |  | 939 | 37560.613 | 10.5337 | 4.9974 |  |  | 939 | 37564.434 | 10.5338 | 4.9218
940 | 37613.281 | 10.5351 | 4.8991 |  |  | 940 | 37600.614 | 10.5348 | 4.9973 |  |  | 940 | 37604.438 | 10.5349 | 4.9217
941 | 37653.295 | 10.5362 | 4.899 |  |  | 941 | 37640.615 | 10.5358 | 4.9972 |  |  | 941 | 37644.443 | 10.5359 | 4.9215
942 | 37693.309 | 10.5372 | 4.8989 |  |  | 942 | 37680.615 | 10.5369 | 4.9971 |  |  | 942 | 37684.448 | 10.537 | 4.9214
943 | 37733.323 | 10.5383 | 4.8987 |  |  | 943 | 37720.616 | 10.538 | 4.997 |  |  | 943 | 37724.452 | 10.5381 | 4.9213
944 | 37773.337 | 10.5394 | 4.8986 |  |  | 944 | 37760.617 | 10.539 | 4.9969 |  |  | 944 | 37764.457 | 10.5391 | 4.9212
945 | 37813.351 | 10.5404 | 4.8985 |  |  | 945 | 37800.617 | 10.5401 | 4.9968 |  |  | 945 | 37804.462 | 10.5402 | 4.9211
946 | 37853.366 | 10.5415 | 4.8983 |  |  | 946 | 37840.618 | 10.5411 | 4.9968 |  |  | 946 | 37844.467 | 10.5412 | 4.921
947 | 37893.38 | 10.5425 | 4.8982 |  |  | 947 | 37880.618 | 10.5422 | 4.9967 |  |  | 947 | 37884.471 | 10.5423 | 4.9208
948 | 37933.394 | 10.5436 | 4.8981 |  |  | 948 | 37920.619 | 10.5433 | 4.9966 |  |  | 948 | 37924.476 | 10.5434 | 4.9207
949 | 37973.408 | 10.5446 | 4.898 |  |  | 949 | 37960.62 | 10.5443 | 4.9965 |  |  | 949 | 37964.481 | 10.5444 | 4.9206
950 | 38013.422 | 10.5457 | 4.8978 |  |  | 950 | 38000.62 | 10.5454 | 4.9964 |  |  | 950 | 38004.486 | 10.5455 | 4.9205
951 | 38053.436 | 10.5467 | 4.8977 |  |  | 951 | 38040.621 | 10.5464 | 4.9963 |  |  | 951 | 38044.49 | 10.5465 | 4.9204
952 | 38093.45 | 10.5478 | 4.8976 |  |  | 952 | 38080.622 | 10.5475 | 4.9962 |  |  | 952 | 38084.495 | 10.5476 | 4.9202
953 | 38133.464 | 10.5488 | 4.8974 |  |  | 953 | 38120.622 | 10.5485 | 4.9961 |  |  | 953 | 38124.5 | 10.5486 | 4.9201
954 | 38173.479 | 10.5499 | 4.8973 |  |  | 954 | 38160.623 | 10.5496 | 4.996 |  |  | 954 | 38164.504 | 10.5497 | 4.92
955 | 38213.493 | 10.5509 | 4.8972 |  |  | 955 | 38200.624 | 10.5506 | 4.9959 |  |  | 955 | 38204.509 | 10.5507 | 4.9199
956 | 38253.507 | 10.552 | 4.8971 |  |  | 956 | 38240.624 | 10.5517 | 4.9958 |  |  | 956 | 38244.514 | 10.5518 | 4.9198
957 | 38293.521 | 10.553 | 4.8969 |  |  | 957 | 38280.625 | 10.5527 | 4.9957 |  |  | 957 | 38284.519 | 10.5528 | 4.9197
958 | 38333.535 | 10.5541 | 4.8968 |  |  | 958 | 38320.626 | 10.5537 | 4.9956 |  |  | 958 | 38324.523 | 10.5538 | 4.9195
959 | 38373.549 | 10.5551 | 4.8967 |  |  | 959 | 38360.626 | 10.5548 | 4.9955 |  |  | 959 | 38364.528 | 10.5549 | 4.9194
960 | 38413.563 | 10.5562 | 4.8966 |  |  | 960 | 38400.627 | 10.5558 | 4.9954 |  |  | 960 | 38404.533 | 10.5559 | 4.9193
961 | 38453.578 | 10.5572 | 4.8964 |  |  | 961 | 38440.628 | 10.5569 | 4.9953 |  |  | 961 | 38444.537 | 10.557 | 4.9192
962 | 38493.592 | 10.5582 | 4.8963 |  |  | 962 | 38480.628 | 10.5579 | 4.9952 |  |  | 962 | 38484.542 | 10.558 | 4.9191
963 | 38533.606 | 10.5593 | 4.8962 |  |  | 963 | 38520.629 | 10.5589 | 4.9951 |  |  | 963 | 38524.547 | 10.5591 | 4.9189
964 | 38573.62 | 10.5603 | 4.8961 |  |  | 964 | 38560.63 | 10.56 | 4.995 |  |  | 964 | 38564.552 | 10.5601 | 4.9188
965 | 38613.634 | 10.5614 | 4.8959 |  |  | 965 | 38600.63 | 10.561 | 4.9949 |  |  | 965 | 38604.556 | 10.5611 | 4.9187
966 | 38653.648 | 10.5624 | 4.8958 |  |  | 966 | 38640.631 | 10.5621 | 4.9949 |  |  | 966 | 38644.561 | 10.5622 | 4.9186
967 | 38693.662 | 10.5634 | 4.8957 |  |  | 967 | 38680.632 | 10.5631 | 4.9948 |  |  | 967 | 38684.566 | 10.5632 | 4.9185
968 | 38733.676 | 10.5645 | 4.8956 |  |  | 968 | 38720.632 | 10.5641 | 4.9947 |  |  | 968 | 38724.571 | 10.5642 | 4.9184
969 | 38773.691 | 10.5655 | 4.8955 |  |  | 969 | 38760.633 | 10.5652 | 4.9946 |  |  | 969 | 38764.575 | 10.5653 | 4.9182
970 | 38813.705 | 10.5665 | 4.8953 |  |  | 970 | 38800.633 | 10.5662 | 4.9945 |  |  | 970 | 38804.58 | 10.5663 | 4.9181
971 | 38853.719 | 10.5676 | 4.8952 |  |  | 971 | 38840.634 | 10.5672 | 4.9944 |  |  | 971 | 38844.585 | 10.5673 | 4.918
972 | 38893.733 | 10.5686 | 4.8951 |  |  | 972 | 38880.635 | 10.5683 | 4.9943 |  |  | 972 | 38884.589 | 10.5684 | 4.9179
973 | 38933.747 | 10.5696 | 4.895 |  |  | 973 | 38920.635 | 10.5693 | 4.9942 |  |  | 973 | 38924.594 | 10.5694 | 4.9178
974 | 38973.761 | 10.5706 | 4.8949 |  |  | 974 | 38960.636 | 10.5703 | 4.9941 |  |  | 974 | 38964.599 | 10.5704 | 4.9176
975 | 39013.775 | 10.5717 | 4.8947 |  |  | 975 | 39000.637 | 10.5713 | 4.994 |  |  | 975 | 39004.604 | 10.5714 | 4.9175
976 | 39053.789 | 10.5727 | 4.8946 |  |  | 976 | 39040.637 | 10.5724 | 4.9939 |  |  | 976 | 39044.608 | 10.5725 | 4.9174
977 | 39093.804 | 10.5737 | 4.8945 |  |  | 977 | 39080.638 | 10.5734 | 4.9938 |  |  | 977 | 39084.613 | 10.5735 | 4.9173
978 | 39133.818 | 10.5747 | 4.8944 |  |  | 978 | 39120.639 | 10.5744 | 4.9937 |  |  | 978 | 39124.618 | 10.5745 | 4.9172
979 | 39173.832 | 10.5758 | 4.8943 |  |  | 979 | 39160.639 | 10.5754 | 4.9936 |  |  | 979 | 39164.622 | 10.5755 | 4.917
980 | 39213.846 | 10.5768 | 4.8942 |  |  | 980 | 39200.64 | 10.5764 | 4.9935 |  |  | 980 | 39204.627 | 10.5766 | 4.9169
981 | 39253.86 | 10.5778 | 4.894 |  |  | 981 | 39240.641 | 10.5775 | 4.9934 |  |  | 981 | 39244.632 | 10.5776 | 4.9168
982 | 39293.874 | 10.5788 | 4.8939 |  |  | 982 | 39280.641 | 10.5785 | 4.9933 |  |  | 982 | 39284.637 | 10.5786 | 4.9167
983 | 39333.888 | 10.5798 | 4.8938 |  |  | 983 | 39320.642 | 10.5795 | 4.9932 |  |  | 983 | 39324.641 | 10.5796 | 4.9166
984 | 39373.902 | 10.5809 | 4.8937 |  |  | 984 | 39360.643 | 10.5805 | 4.9931 |  |  | 984 | 39364.646 | 10.5806 | 4.9164
985 | 39413.917 | 10.5819 | 4.8936 |  |  | 985 | 39400.643 | 10.5815 | 4.993 |  |  | 985 | 39404.651 | 10.5816 | 4.9163
986 | 39453.931 | 10.5829 | 4.8935 |  |  | 986 | 39440.644 | 10.5826 | 4.9929 |  |  | 986 | 39444.656 | 10.5827 | 4.9162
987 | 39493.945 | 10.5839 | 4.8933 |  |  | 987 | 39480.645 | 10.5836 | 4.9928 |  |  | 987 | 39484.66 | 10.5837 | 4.9161
988 | 39533.959 | 10.5849 | 4.8932 |  |  | 988 | 39520.645 | 10.5846 | 4.9927 |  |  | 988 | 39524.665 | 10.5847 | 4.916
989 | 39573.973 | 10.5859 | 4.8931 |  |  | 989 | 39560.646 | 10.5856 | 4.9926 |  |  | 989 | 39564.67 | 10.5857 | 4.9158
990 | 39613.987 | 10.5869 | 4.893 |  |  | 990 | 39600.647 | 10.5866 | 4.9925 |  |  | 990 | 39604.674 | 10.5867 | 4.9157

## Run2
MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR
File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run2\1 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run2\2 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run2\3 01 001.RAW
Cursor Location: |  |  | (179, 193) |  |  | Cursor Location: |  |  | (178, 194) |  |  | Cursor Location: |  |  | (178, 193)
Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15)
Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev
1 | 40.005 | 3.689 | 7.7396 | 5.1222 | 0.32499 | 1 | 40.005 | 3.689 | 7.7553 | 5.1409 | 0.32574 | 1 | 40.005 | 3.689 | 7.7604 | 5.1416 | 0.32679
2 | 80.009 | 4.3821 | 7.4042 |  |  | 2 | 80.009 | 4.3821 | 7.4217 |  |  | 2 | 80.009 | 4.3821 | 7.4274
3 | 120.014 | 4.7876 | 7.2091 |  |  | 3 | 120.014 | 4.7876 | 7.2279 |  |  | 3 | 120.014 | 4.7876 | 7.2332
4 | 160.019 | 5.0753 | 7.0689 |  |  | 4 | 160.019 | 5.0753 | 7.0883 |  |  | 4 | 160.019 | 5.0753 | 7.0933
5 | 200.024 | 5.2984 | 6.9587 |  |  | 5 | 200.024 | 5.2984 | 6.9785 |  |  | 5 | 200.024 | 5.2984 | 6.9833
6 | 240.028 | 5.4808 | 6.8679 |  |  | 6 | 240.028 | 5.4808 | 6.8879 |  |  | 6 | 240.028 | 5.4808 | 6.8925
7 | 280.033 | 5.6349 | 6.7905 |  |  | 7 | 280.033 | 5.6349 | 6.8106 |  |  | 7 | 280.033 | 5.6349 | 6.8151
8 | 320.038 | 5.7684 | 6.7232 |  |  | 8 | 320.038 | 5.7684 | 6.7432 |  |  | 8 | 320.038 | 5.7684 | 6.7477
9 | 360.042 | 5.8862 | 6.6635 |  |  | 9 | 360.042 | 5.8862 | 6.6835 |  |  | 9 | 360.042 | 5.8862 | 6.688
10 | 400.047 | 5.9916 | 6.6099 |  |  | 10 | 400.047 | 5.9916 | 6.6299 |  |  | 10 | 400.047 | 5.9916 | 6.6344
11 | 440.052 | 6.0869 | 6.5614 |  |  | 11 | 440.052 | 6.0869 | 6.5813 |  |  | 11 | 440.052 | 6.0869 | 6.5858
12 | 480.057 | 6.1739 | 6.5169 |  |  | 12 | 480.057 | 6.1739 | 6.5369 |  |  | 12 | 480.057 | 6.1739 | 6.5413
13 | 520.061 | 6.2539 | 6.476 |  |  | 13 | 520.061 | 6.2539 | 6.496 |  |  | 13 | 520.061 | 6.2539 | 6.5003
14 | 560.066 | 6.3281 | 6.4381 |  |  | 14 | 560.066 | 6.3281 | 6.458 |  |  | 14 | 560.066 | 6.3281 | 6.4624
15 | 600.071 | 6.397 | 6.4027 |  |  | 15 | 600.071 | 6.397 | 6.4227 |  |  | 15 | 600.071 | 6.397 | 6.427
16 | 640.076 | 6.4616 | 6.3696 |  |  | 16 | 640.076 | 6.4616 | 6.3896 |  |  | 16 | 640.076 | 6.4616 | 6.3939
17 | 680.08 | 6.5222 | 6.3385 |  |  | 17 | 680.08 | 6.5222 | 6.3585 |  |  | 17 | 680.08 | 6.5222 | 6.3628
18 | 720.085 | 6.5794 | 6.3092 |  |  | 18 | 720.085 | 6.5794 | 6.3292 |  |  | 18 | 720.085 | 6.5794 | 6.3335
19 | 760.09 | 6.6334 | 6.2814 |  |  | 19 | 760.09 | 6.6334 | 6.3014 |  |  | 19 | 760.09 | 6.6334 | 6.3057
20 | 800.094 | 6.6847 | 6.2551 |  |  | 20 | 800.094 | 6.6847 | 6.2751 |  |  | 20 | 800.094 | 6.6847 | 6.2794
21 | 840.099 | 6.7335 | 6.23 |  |  | 21 | 840.099 | 6.7335 | 6.2501 |  |  | 21 | 840.099 | 6.7335 | 6.2543
22 | 880.104 | 6.78 | 6.2061 |  |  | 22 | 880.104 | 6.78 | 6.2263 |  |  | 22 | 880.104 | 6.78 | 6.2304
23 | 920.109 | 6.8245 | 6.1833 |  |  | 23 | 920.109 | 6.8245 | 6.2035 |  |  | 23 | 920.109 | 6.8245 | 6.2076
24 | 960.113 | 6.8671 | 6.1615 |  |  | 24 | 960.113 | 6.8671 | 6.1817 |  |  | 24 | 960.113 | 6.8671 | 6.1858
25 | 1000.118 | 6.9079 | 6.1405 |  |  | 25 | 1000.118 | 6.9079 | 6.1607 |  |  | 25 | 1000.118 | 6.9079 | 6.1648
26 | 1040.123 | 6.9471 | 6.1204 |  |  | 26 | 1040.123 | 6.9471 | 6.1407 |  |  | 26 | 1040.123 | 6.9471 | 6.1447
27 | 1080.127 | 6.9848 | 6.101 |  |  | 27 | 1080.127 | 6.9848 | 6.1213 |  |  | 27 | 1080.127 | 6.9848 | 6.1254
28 | 1120.132 | 7.0212 | 6.0823 |  |  | 28 | 1120.132 | 7.0212 | 6.1027 |  |  | 28 | 1120.132 | 7.0212 | 6.1067
29 | 1160.137 | 7.0563 | 6.0643 |  |  | 29 | 1160.137 | 7.0563 | 6.0848 |  |  | 29 | 1160.137 | 7.0563 | 6.0887
30 | 1200.142 | 7.0902 | 6.047 |  |  | 30 | 1200.142 | 7.0902 | 6.0675 |  |  | 30 | 1200.142 | 7.0902 | 6.0714
31 | 1240.146 | 7.123 | 6.0302 |  |  | 31 | 1240.146 | 7.123 | 6.0507 |  |  | 31 | 1240.146 | 7.123 | 6.0546
32 | 1280.151 | 7.1547 | 6.0139 |  |  | 32 | 1280.151 | 7.1547 | 6.0345 |  |  | 32 | 1280.151 | 7.1547 | 6.0383
33 | 1320.156 | 7.1855 | 5.9981 |  |  | 33 | 1320.156 | 7.1855 | 6.0188 |  |  | 33 | 1320.156 | 7.1855 | 6.0226
34 | 1360.161 | 7.2154 | 5.9829 |  |  | 34 | 1360.161 | 7.2154 | 6.0036 |  |  | 34 | 1360.161 | 7.2154 | 6.0073
35 | 1400.165 | 7.2443 | 5.968 |  |  | 35 | 1400.165 | 7.2443 | 5.9888 |  |  | 35 | 1400.165 | 7.2443 | 5.9925
36 | 1440.17 | 7.2725 | 5.9536 |  |  | 36 | 1440.17 | 7.2725 | 5.9744 |  |  | 36 | 1440.17 | 7.2725 | 5.9781
37 | 1480.175 | 7.2999 | 5.9396 |  |  | 37 | 1480.175 | 7.2999 | 5.9605 |  |  | 37 | 1480.175 | 7.2999 | 5.9642
38 | 1520.179 | 7.3266 | 5.926 |  |  | 38 | 1520.179 | 7.3266 | 5.9469 |  |  | 38 | 1520.179 | 7.3266 | 5.9506
39 | 1560.184 | 7.3526 | 5.9128 |  |  | 39 | 1560.184 | 7.3526 | 5.9337 |  |  | 39 | 1560.184 | 7.3526 | 5.9373
40 | 1600.189 | 7.3779 | 5.8999 |  |  | 40 | 1600.189 | 7.3779 | 5.9209 |  |  | 40 | 1600.189 | 7.3779 | 5.9244
41 | 1640.194 | 7.4026 | 5.8873 |  |  | 41 | 1640.194 | 7.4026 | 5.9083 |  |  | 41 | 1640.194 | 7.4026 | 5.9119
42 | 1680.198 | 7.4267 | 5.875 |  |  | 42 | 1680.198 | 7.4267 | 5.8961 |  |  | 42 | 1680.198 | 7.4267 | 5.8996
43 | 1720.203 | 7.4502 | 5.8631 |  |  | 43 | 1720.203 | 7.4502 | 5.8842 |  |  | 43 | 1720.203 | 7.4502 | 5.8877
44 | 1760.208 | 7.4732 | 5.8514 |  |  | 44 | 1760.208 | 7.4732 | 5.8726 |  |  | 44 | 1760.208 | 7.4732 | 5.876
45 | 1800.212 | 7.4957 | 5.84 |  |  | 45 | 1800.212 | 7.4957 | 5.8612 |  |  | 45 | 1800.212 | 7.4957 | 5.8646
46 | 1840.217 | 7.5176 | 5.8288 |  |  | 46 | 1840.217 | 7.5176 | 5.8501 |  |  | 46 | 1840.217 | 7.5176 | 5.8535
47 | 1880.222 | 7.5391 | 5.818 |  |  | 47 | 1880.222 | 7.5391 | 5.8393 |  |  | 47 | 1880.222 | 7.5391 | 5.8426
48 | 1920.227 | 7.5602 | 5.8073 |  |  | 48 | 1920.227 | 7.5602 | 5.8287 |  |  | 48 | 1920.227 | 7.5602 | 5.8319
49 | 1960.231 | 7.5808 | 5.7969 |  |  | 49 | 1960.231 | 7.5808 | 5.8183 |  |  | 49 | 1960.231 | 7.5808 | 5.8215
50 | 2000.236 | 7.601 | 5.7867 |  |  | 50 | 2000.236 | 7.601 | 5.8081 |  |  | 50 | 2000.236 | 7.601 | 5.8113
51 | 2040.241 | 7.6208 | 5.7767 |  |  | 51 | 2040.241 | 7.6208 | 5.7982 |  |  | 51 | 2040.241 | 7.6208 | 5.8014
52 | 2080.246 | 7.6402 | 5.7669 |  |  | 52 | 2080.246 | 7.6402 | 5.7884 |  |  | 52 | 2080.246 | 7.6402 | 5.7916
53 | 2120.25 | 7.6593 | 5.7573 |  |  | 53 | 2120.25 | 7.6593 | 5.7789 |  |  | 53 | 2120.25 | 7.6593 | 5.782
54 | 2160.255 | 7.678 | 5.748 |  |  | 54 | 2160.255 | 7.678 | 5.7695 |  |  | 54 | 2160.255 | 7.678 | 5.7726
55 | 2200.26 | 7.6963 | 5.7388 |  |  | 55 | 2200.26 | 7.6963 | 5.7604 |  |  | 55 | 2200.26 | 7.6963 | 5.7634
56 | 2240.264 | 7.7143 | 5.7297 |  |  | 56 | 2240.264 | 7.7143 | 5.7514 |  |  | 56 | 2240.264 | 7.7143 | 5.7544
57 | 2280.269 | 7.732 | 5.7209 |  |  | 57 | 2280.269 | 7.732 | 5.7426 |  |  | 57 | 2280.269 | 7.732 | 5.7456
58 | 2320.274 | 7.7494 | 5.7122 |  |  | 58 | 2320.274 | 7.7494 | 5.7339 |  |  | 58 | 2320.274 | 7.7494 | 5.7369
59 | 2360.279 | 7.7665 | 5.7037 |  |  | 59 | 2360.279 | 7.7665 | 5.7254 |  |  | 59 | 2360.279 | 7.7665 | 5.7283
60 | 2400.283 | 7.7833 | 5.6953 |  |  | 60 | 2400.283 | 7.7833 | 5.7171 |  |  | 60 | 2400.283 | 7.7833 | 5.72
61 | 2440.288 | 7.7999 | 5.6871 |  |  | 61 | 2440.288 | 7.7999 | 5.7089 |  |  | 61 | 2440.288 | 7.7999 | 5.7117
62 | 2480.293 | 7.8161 | 5.679 |  |  | 62 | 2480.293 | 7.8161 | 5.7008 |  |  | 62 | 2480.293 | 7.8161 | 5.7037
63 | 2520.297 | 7.8321 | 5.6711 |  |  | 63 | 2520.297 | 7.8321 | 5.6929 |  |  | 63 | 2520.297 | 7.8321 | 5.6957
64 | 2560.302 | 7.8479 | 5.6633 |  |  | 64 | 2560.302 | 7.8479 | 5.6851 |  |  | 64 | 2560.302 | 7.8479 | 5.6879
65 | 2600.307 | 7.8634 | 5.6556 |  |  | 65 | 2600.307 | 7.8634 | 5.6775 |  |  | 65 | 2600.307 | 7.8634 | 5.6803
66 | 2640.312 | 7.8787 | 5.6481 |  |  | 66 | 2640.312 | 7.8787 | 5.67 |  |  | 66 | 2640.312 | 7.8787 | 5.6727
67 | 2680.316 | 7.8937 | 5.6407 |  |  | 67 | 2680.316 | 7.8937 | 5.6626 |  |  | 67 | 2680.316 | 7.8937 | 5.6653
68 | 2720.321 | 7.9085 | 5.6334 |  |  | 68 | 2720.321 | 7.9085 | 5.6554 |  |  | 68 | 2720.321 | 7.9085 | 5.6581
69 | 2760.326 | 7.9231 | 5.6263 |  |  | 69 | 2760.326 | 7.9231 | 5.6482 |  |  | 69 | 2760.326 | 7.9231 | 5.6509
70 | 2800.331 | 7.9375 | 5.6192 |  |  | 70 | 2800.331 | 7.9375 | 5.6412 |  |  | 70 | 2800.331 | 7.9375 | 5.6438
71 | 2840.335 | 7.9517 | 5.6123 |  |  | 71 | 2840.335 | 7.9517 | 5.6343 |  |  | 71 | 2840.335 | 7.9517 | 5.6369
72 | 2880.34 | 7.9657 | 5.6055 |  |  | 72 | 2880.34 | 7.9657 | 5.6275 |  |  | 72 | 2880.34 | 7.9657 | 5.6301
73 | 2920.345 | 7.9795 | 5.5988 |  |  | 73 | 2920.345 | 7.9795 | 5.6208 |  |  | 73 | 2920.345 | 7.9795 | 5.6233
74 | 2960.349 | 7.9931 | 5.5922 |  |  | 74 | 2960.349 | 7.9931 | 5.6142 |  |  | 74 | 2960.349 | 7.9931 | 5.6167
75 | 3000.354 | 8.0065 | 5.5857 |  |  | 75 | 3000.354 | 8.0065 | 5.6077 |  |  | 75 | 3000.354 | 8.0065 | 5.6102
76 | 3040.359 | 8.0197 | 5.5793 |  |  | 76 | 3040.359 | 8.0197 | 5.6013 |  |  | 76 | 3040.359 | 8.0197 | 5.6038
77 | 3080.364 | 8.0328 | 5.573 |  |  | 77 | 3080.364 | 8.0328 | 5.595 |  |  | 77 | 3080.364 | 8.0328 | 5.5975
78 | 3120.368 | 8.0457 | 5.5668 |  |  | 78 | 3120.368 | 8.0457 | 5.5888 |  |  | 78 | 3120.368 | 8.0457 | 5.5912
79 | 3160.373 | 8.0584 | 5.5606 |  |  | 79 | 3160.373 | 8.0584 | 5.5827 |  |  | 79 | 3160.373 | 8.0584 | 5.5851
80 | 3200.378 | 8.071 | 5.5546 |  |  | 80 | 3200.378 | 8.071 | 5.5766 |  |  | 80 | 3200.378 | 8.071 | 5.5791
81 | 3240.382 | 8.0834 | 5.5487 |  |  | 81 | 3240.382 | 8.0834 | 5.5707 |  |  | 81 | 3240.382 | 8.0834 | 5.5731
82 | 3280.387 | 8.0957 | 5.5428 |  |  | 82 | 3280.387 | 8.0957 | 5.5649 |  |  | 82 | 3280.387 | 8.0957 | 5.5672
83 | 3320.392 | 8.1078 | 5.537 |  |  | 83 | 3320.392 | 8.1078 | 5.5591 |  |  | 83 | 3320.392 | 8.1078 | 5.5614
84 | 3360.397 | 8.1198 | 5.5313 |  |  | 84 | 3360.397 | 8.1198 | 5.5534 |  |  | 84 | 3360.397 | 8.1198 | 5.5557
85 | 3400.401 | 8.1316 | 5.5257 |  |  | 85 | 3400.401 | 8.1316 | 5.5478 |  |  | 85 | 3400.401 | 8.1316 | 5.5501
86 | 3440.406 | 8.1433 | 5.5202 |  |  | 86 | 3440.406 | 8.1433 | 5.5423 |  |  | 86 | 3440.406 | 8.1433 | 5.5446
87 | 3480.411 | 8.1549 | 5.5148 |  |  | 87 | 3480.411 | 8.1549 | 5.5368 |  |  | 87 | 3480.411 | 8.1549 | 5.5391
88 | 3520.416 | 8.1663 | 5.5094 |  |  | 88 | 3520.416 | 8.1663 | 5.5314 |  |  | 88 | 3520.416 | 8.1663 | 5.5337
89 | 3560.42 | 8.1776 | 5.5041 |  |  | 89 | 3560.42 | 8.1776 | 5.5261 |  |  | 89 | 3560.42 | 8.1776 | 5.5284
90 | 3600.425 | 8.1888 | 5.4988 |  |  | 90 | 3600.425 | 8.1888 | 5.5209 |  |  | 90 | 3600.425 | 8.1888 | 5.5231
91 | 3640.43 | 8.1999 | 5.4937 |  |  | 91 | 3640.43 | 8.1999 | 5.5157 |  |  | 91 | 3640.43 | 8.1999 | 5.5179
92 | 3680.434 | 8.2108 | 5.4886 |  |  | 92 | 3680.434 | 8.2108 | 5.5106 |  |  | 92 | 3680.434 | 8.2108 | 5.5128
93 | 3720.439 | 8.2216 | 5.4836 |  |  | 93 | 3720.439 | 8.2216 | 5.5056 |  |  | 93 | 3720.439 | 8.2216 | 5.5078
94 | 3760.444 | 8.2323 | 5.4786 |  |  | 94 | 3760.444 | 8.2323 | 5.5006 |  |  | 94 | 3760.444 | 8.2323 | 5.5028
95 | 3800.449 | 8.2429 | 5.4737 |  |  | 95 | 3800.449 | 8.2429 | 5.4957 |  |  | 95 | 3800.449 | 8.2429 | 5.4979
96 | 3840.453 | 8.2533 | 5.4689 |  |  | 96 | 3840.453 | 8.2533 | 5.4909 |  |  | 96 | 3840.453 | 8.2533 | 5.493
97 | 3880.458 | 8.2637 | 5.4641 |  |  | 97 | 3880.458 | 8.2637 | 5.4861 |  |  | 97 | 3880.458 | 8.2637 | 5.4882
98 | 3920.463 | 8.274 | 5.4594 |  |  | 98 | 3920.463 | 8.274 | 5.4814 |  |  | 98 | 3920.463 | 8.274 | 5.4835
99 | 3960.467 | 8.2841 | 5.4548 |  |  | 99 | 3960.467 | 8.2841 | 5.4768 |  |  | 99 | 3960.467 | 8.2841 | 5.4788
100 | 4000.472 | 8.2942 | 5.4502 |  |  | 100 | 4000.472 | 8.2942 | 5.4722 |  |  | 100 | 4000.472 | 8.2942 | 5.4742
101 | 4040.477 | 8.3041 | 5.4457 |  |  | 101 | 4040.477 | 8.3041 | 5.4676 |  |  | 101 | 4040.477 | 8.3041 | 5.4697
102 | 4080.482 | 8.314 | 5.4412 |  |  | 102 | 4080.482 | 8.314 | 5.4632 |  |  | 102 | 4080.482 | 8.314 | 5.4652
103 | 4120.486 | 8.3237 | 5.4368 |  |  | 103 | 4120.486 | 8.3237 | 5.4587 |  |  | 103 | 4120.486 | 8.3237 | 5.4607
104 | 4160.491 | 8.3334 | 5.4324 |  |  | 104 | 4160.491 | 8.3334 | 5.4544 |  |  | 104 | 4160.491 | 8.3334 | 5.4564
105 | 4200.496 | 8.343 | 5.4281 |  |  | 105 | 4200.496 | 8.343 | 5.45 |  |  | 105 | 4200.496 | 8.343 | 5.452
106 | 4240.5 | 8.3524 | 5.4239 |  |  | 106 | 4240.5 | 8.3524 | 5.4458 |  |  | 106 | 4240.5 | 8.3524 | 5.4478
107 | 4280.505 | 8.3618 | 5.4197 |  |  | 107 | 4280.505 | 8.3618 | 5.4416 |  |  | 107 | 4280.505 | 8.3618 | 5.4435
108 | 4320.51 | 8.3711 | 5.4155 |  |  | 108 | 4320.51 | 8.3711 | 5.4374 |  |  | 108 | 4320.51 | 8.3711 | 5.4394
109 | 4360.515 | 8.3803 | 5.4114 |  |  | 109 | 4360.515 | 8.3803 | 5.4333 |  |  | 109 | 4360.515 | 8.3803 | 5.4352
110 | 4400.519 | 8.3895 | 5.4074 |  |  | 110 | 4400.519 | 8.3895 | 5.4292 |  |  | 110 | 4400.519 | 8.3895 | 5.4312
111 | 4440.524 | 8.3985 | 5.4034 |  |  | 111 | 4440.524 | 8.3985 | 5.4252 |  |  | 111 | 4440.524 | 8.3985 | 5.4271
112 | 4480.529 | 8.4075 | 5.3994 |  |  | 112 | 4480.529 | 8.4075 | 5.4213 |  |  | 112 | 4480.529 | 8.4075 | 5.4231
113 | 4520.534 | 8.4164 | 5.3955 |  |  | 113 | 4520.534 | 8.4164 | 5.4173 |  |  | 113 | 4520.534 | 8.4164 | 5.4192
114 | 4560.538 | 8.4252 | 5.3917 |  |  | 114 | 4560.538 | 8.4252 | 5.4135 |  |  | 114 | 4560.538 | 8.4252 | 5.4153
115 | 4600.543 | 8.4339 | 5.3879 |  |  | 115 | 4600.543 | 8.4339 | 5.4096 |  |  | 115 | 4600.543 | 8.4339 | 5.4115
116 | 4640.548 | 8.4426 | 5.3841 |  |  | 116 | 4640.548 | 8.4426 | 5.4058 |  |  | 116 | 4640.548 | 8.4426 | 5.4077
117 | 4680.552 | 8.4512 | 5.3804 |  |  | 117 | 4680.552 | 8.4512 | 5.4021 |  |  | 117 | 4680.552 | 8.4512 | 5.4039
118 | 4720.557 | 8.4597 | 5.3767 |  |  | 118 | 4720.557 | 8.4597 | 5.3984 |  |  | 118 | 4720.557 | 8.4597 | 5.4002
119 | 4760.562 | 8.4681 | 5.3731 |  |  | 119 | 4760.562 | 8.4681 | 5.3947 |  |  | 119 | 4760.562 | 8.4681 | 5.3966
120 | 4800.567 | 8.4765 | 5.3695 |  |  | 120 | 4800.567 | 8.4765 | 5.3911 |  |  | 120 | 4800.567 | 8.4765 | 5.3929
121 | 4840.571 | 8.4848 | 5.3659 |  |  | 121 | 4840.571 | 8.4848 | 5.3876 |  |  | 121 | 4840.571 | 8.4848 | 5.3894
122 | 4880.576 | 8.493 | 5.3624 |  |  | 122 | 4880.576 | 8.493 | 5.384 |  |  | 122 | 4880.576 | 8.493 | 5.3858
123 | 4920.581 | 8.5012 | 5.3589 |  |  | 123 | 4920.581 | 8.5012 | 5.3805 |  |  | 123 | 4920.581 | 8.5012 | 5.3823
124 | 4960.585 | 8.5093 | 5.3555 |  |  | 124 | 4960.585 | 8.5093 | 5.3771 |  |  | 124 | 4960.585 | 8.5093 | 5.3788
125 | 5000.59 | 8.5173 | 5.3521 |  |  | 125 | 5000.59 | 8.5173 | 5.3737 |  |  | 125 | 5000.59 | 8.5173 | 5.3754
126 | 5040.595 | 8.5253 | 5.3487 |  |  | 126 | 5040.595 | 8.5253 | 5.3703 |  |  | 126 | 5040.595 | 8.5253 | 5.372
127 | 5080.6 | 8.5332 | 5.3454 |  |  | 127 | 5080.6 | 8.5332 | 5.3669 |  |  | 127 | 5080.6 | 8.5332 | 5.3687
128 | 5120.604 | 8.541 | 5.3421 |  |  | 128 | 5120.604 | 8.541 | 5.3636 |  |  | 128 | 5120.604 | 8.541 | 5.3654
129 | 5160.609 | 8.5488 | 5.3389 |  |  | 129 | 5160.609 | 8.5488 | 5.3604 |  |  | 129 | 5160.609 | 8.5488 | 5.3621
130 | 5200.614 | 8.5565 | 5.3357 |  |  | 130 | 5200.614 | 8.5565 | 5.3571 |  |  | 130 | 5200.614 | 8.5565 | 5.3588
131 | 5240.619 | 8.5642 | 5.3325 |  |  | 131 | 5240.619 | 8.5642 | 5.3539 |  |  | 131 | 5240.619 | 8.5642 | 5.3556
132 | 5280.623 | 8.5718 | 5.3293 |  |  | 132 | 5280.623 | 8.5718 | 5.3508 |  |  | 132 | 5280.623 | 8.5718 | 5.3525
133 | 5320.628 | 8.5793 | 5.3262 |  |  | 133 | 5320.628 | 8.5793 | 5.3476 |  |  | 133 | 5320.628 | 8.5793 | 5.3493
134 | 5360.633 | 8.5868 | 5.3232 |  |  | 134 | 5360.633 | 8.5868 | 5.3445 |  |  | 134 | 5360.633 | 8.5868 | 5.3462
135 | 5400.637 | 8.5943 | 5.3201 |  |  | 135 | 5400.637 | 8.5943 | 5.3415 |  |  | 135 | 5400.637 | 8.5943 | 5.3432
136 | 5440.642 | 8.6017 | 5.3171 |  |  | 136 | 5440.642 | 8.6017 | 5.3385 |  |  | 136 | 5440.642 | 8.6017 | 5.3401
137 | 5480.647 | 8.609 | 5.3141 |  |  | 137 | 5480.647 | 8.609 | 5.3355 |  |  | 137 | 5480.647 | 8.609 | 5.3371
138 | 5520.652 | 8.6163 | 5.3112 |  |  | 138 | 5520.652 | 8.6163 | 5.3325 |  |  | 138 | 5520.652 | 8.6163 | 5.3341
139 | 5560.656 | 8.6235 | 5.3083 |  |  | 139 | 5560.656 | 8.6235 | 5.3296 |  |  | 139 | 5560.656 | 8.6235 | 5.3312
140 | 5600.661 | 8.6306 | 5.3054 |  |  | 140 | 5600.661 | 8.6306 | 5.3266 |  |  | 140 | 5600.661 | 8.6306 | 5.3283
141 | 5640.666 | 8.6378 | 5.3026 |  |  | 141 | 5640.666 | 8.6378 | 5.3238 |  |  | 141 | 5640.666 | 8.6378 | 5.3254
142 | 5680.67 | 8.6448 | 5.2997 |  |  | 142 | 5680.67 | 8.6448 | 5.3209 |  |  | 142 | 5680.67 | 8.6448 | 5.3225
143 | 5720.675 | 8.6518 | 5.2969 |  |  | 143 | 5720.675 | 8.6518 | 5.3181 |  |  | 143 | 5720.675 | 8.6518 | 5.3197
144 | 5760.68 | 8.6588 | 5.2942 |  |  | 144 | 5760.68 | 8.6588 | 5.3153 |  |  | 144 | 5760.68 | 8.6588 | 5.3169
145 | 5800.685 | 8.6657 | 5.2915 |  |  | 145 | 5800.685 | 8.6657 | 5.3126 |  |  | 145 | 5800.685 | 8.6657 | 5.3142
146 | 5840.689 | 8.6726 | 5.2888 |  |  | 146 | 5840.689 | 8.6726 | 5.3098 |  |  | 146 | 5840.689 | 8.6726 | 5.3114
147 | 5880.694 | 8.6794 | 5.2861 |  |  | 147 | 5880.694 | 8.6794 | 5.3072 |  |  | 147 | 5880.694 | 8.6794 | 5.3087
148 | 5920.699 | 8.6862 | 5.2834 |  |  | 148 | 5920.699 | 8.6862 | 5.3045 |  |  | 148 | 5920.699 | 8.6862 | 5.3061
149 | 5960.704 | 8.6929 | 5.2808 |  |  | 149 | 5960.704 | 8.6929 | 5.3018 |  |  | 149 | 5960.704 | 8.6929 | 5.3034
150 | 6000.708 | 8.6996 | 5.2782 |  |  | 150 | 6000.708 | 8.6996 | 5.2992 |  |  | 150 | 6000.708 | 8.6996 | 5.3008
151 | 6040.713 | 8.7063 | 5.2757 |  |  | 151 | 6040.713 | 8.7063 | 5.2966 |  |  | 151 | 6040.713 | 8.7063 | 5.2982
152 | 6080.718 | 8.7129 | 5.2731 |  |  | 152 | 6080.718 | 8.7129 | 5.2941 |  |  | 152 | 6080.718 | 8.7129 | 5.2956
153 | 6120.722 | 8.7194 | 5.2706 |  |  | 153 | 6120.722 | 8.7194 | 5.2915 |  |  | 153 | 6120.722 | 8.7194 | 5.2931
154 | 6160.727 | 8.726 | 5.2681 |  |  | 154 | 6160.727 | 8.726 | 5.289 |  |  | 154 | 6160.727 | 8.726 | 5.2905
155 | 6200.732 | 8.7324 | 5.2657 |  |  | 155 | 6200.732 | 8.7324 | 5.2865 |  |  | 155 | 6200.732 | 8.7324 | 5.2881
156 | 6240.737 | 8.7389 | 5.2632 |  |  | 156 | 6240.737 | 8.7389 | 5.2841 |  |  | 156 | 6240.737 | 8.7389 | 5.2856
157 | 6280.741 | 8.7452 | 5.2608 |  |  | 157 | 6280.741 | 8.7452 | 5.2816 |  |  | 157 | 6280.741 | 8.7452 | 5.2831
158 | 6320.746 | 8.7516 | 5.2584 |  |  | 158 | 6320.746 | 8.7516 | 5.2792 |  |  | 158 | 6320.746 | 8.7516 | 5.2807
159 | 6360.751 | 8.7579 | 5.256 |  |  | 159 | 6360.751 | 8.7579 | 5.2768 |  |  | 159 | 6360.751 | 8.7579 | 5.2783
160 | 6400.755 | 8.7642 | 5.2537 |  |  | 160 | 6400.755 | 8.7642 | 5.2744 |  |  | 160 | 6400.755 | 8.7642 | 5.276
161 | 6440.76 | 8.7704 | 5.2514 |  |  | 161 | 6440.76 | 8.7704 | 5.2721 |  |  | 161 | 6440.76 | 8.7704 | 5.2736
162 | 6480.765 | 8.7766 | 5.2491 |  |  | 162 | 6480.765 | 8.7766 | 5.2698 |  |  | 162 | 6480.765 | 8.7766 | 5.2713
163 | 6520.77 | 8.7827 | 5.2468 |  |  | 163 | 6520.77 | 8.7827 | 5.2675 |  |  | 163 | 6520.77 | 8.7827 | 5.269
164 | 6560.774 | 8.7889 | 5.2446 |  |  | 164 | 6560.774 | 8.7889 | 5.2652 |  |  | 164 | 6560.774 | 8.7889 | 5.2667
165 | 6600.779 | 8.7949 | 5.2423 |  |  | 165 | 6600.779 | 8.7949 | 5.263 |  |  | 165 | 6600.779 | 8.7949 | 5.2644
166 | 6640.784 | 8.801 | 5.2401 |  |  | 166 | 6640.784 | 8.801 | 5.2607 |  |  | 166 | 6640.784 | 8.801 | 5.2622
167 | 6680.789 | 8.807 | 5.238 |  |  | 167 | 6680.789 | 8.807 | 5.2585 |  |  | 167 | 6680.789 | 8.807 | 5.26
168 | 6720.793 | 8.813 | 5.2358 |  |  | 168 | 6720.793 | 8.813 | 5.2563 |  |  | 168 | 6720.793 | 8.813 | 5.2578
169 | 6760.798 | 8.8189 | 5.2337 |  |  | 169 | 6760.798 | 8.8189 | 5.2542 |  |  | 169 | 6760.798 | 8.8189 | 5.2556
170 | 6800.803 | 8.8248 | 5.2315 |  |  | 170 | 6800.803 | 8.8248 | 5.252 |  |  | 170 | 6800.803 | 8.8248 | 5.2535
171 | 6840.807 | 8.8307 | 5.2294 |  |  | 171 | 6840.807 | 8.8307 | 5.2499 |  |  | 171 | 6840.807 | 8.8307 | 5.2513
172 | 6880.812 | 8.8365 | 5.2274 |  |  | 172 | 6880.812 | 8.8365 | 5.2478 |  |  | 172 | 6880.812 | 8.8365 | 5.2492
173 | 6920.817 | 8.8423 | 5.2253 |  |  | 173 | 6920.817 | 8.8423 | 5.2457 |  |  | 173 | 6920.817 | 8.8423 | 5.2471
174 | 6960.822 | 8.8481 | 5.2233 |  |  | 174 | 6960.822 | 8.8481 | 5.2436 |  |  | 174 | 6960.822 | 8.8481 | 5.2451
175 | 7000.826 | 8.8538 | 5.2212 |  |  | 175 | 7000.826 | 8.8538 | 5.2416 |  |  | 175 | 7000.826 | 8.8538 | 5.243
176 | 7040.831 | 8.8595 | 5.2192 |  |  | 176 | 7040.831 | 8.8595 | 5.2396 |  |  | 176 | 7040.831 | 8.8595 | 5.241
177 | 7080.836 | 8.8651 | 5.2173 |  |  | 177 | 7080.836 | 8.8651 | 5.2376 |  |  | 177 | 7080.836 | 8.8651 | 5.239
178 | 7120.84 | 8.8708 | 5.2153 |  |  | 178 | 7120.84 | 8.8708 | 5.2356 |  |  | 178 | 7120.84 | 8.8708 | 5.237
179 | 7160.845 | 8.8764 | 5.2134 |  |  | 179 | 7160.845 | 8.8764 | 5.2336 |  |  | 179 | 7160.845 | 8.8764 | 5.235
180 | 7200.85 | 8.882 | 5.2114 |  |  | 180 | 7200.85 | 8.882 | 5.2317 |  |  | 180 | 7200.85 | 8.882 | 5.2331
181 | 7240.855 | 8.8875 | 5.2095 |  |  | 181 | 7240.855 | 8.8875 | 5.2297 |  |  | 181 | 7240.855 | 8.8875 | 5.2311
182 | 7280.859 | 8.893 | 5.2076 |  |  | 182 | 7280.859 | 8.893 | 5.2278 |  |  | 182 | 7280.859 | 8.893 | 5.2292
183 | 7320.864 | 8.8985 | 5.2058 |  |  | 183 | 7320.864 | 8.8985 | 5.2259 |  |  | 183 | 7320.864 | 8.8985 | 5.2273
184 | 7360.869 | 8.9039 | 5.2039 |  |  | 184 | 7360.869 | 8.9039 | 5.224 |  |  | 184 | 7360.869 | 8.9039 | 5.2254
185 | 7400.874 | 8.9094 | 5.2021 |  |  | 185 | 7400.874 | 8.9094 | 5.2222 |  |  | 185 | 7400.874 | 8.9094 | 5.2236
186 | 7440.878 | 8.9147 | 5.2003 |  |  | 186 | 7440.878 | 8.9147 | 5.2203 |  |  | 186 | 7440.878 | 8.9147 | 5.2217
187 | 7480.883 | 8.9201 | 5.1985 |  |  | 187 | 7480.883 | 8.9201 | 5.2185 |  |  | 187 | 7480.883 | 8.9201 | 5.2199
188 | 7520.888 | 8.9254 | 5.1967 |  |  | 188 | 7520.888 | 8.9254 | 5.2167 |  |  | 188 | 7520.888 | 8.9254 | 5.2181
189 | 7560.892 | 8.9307 | 5.1949 |  |  | 189 | 7560.892 | 8.9307 | 5.2149 |  |  | 189 | 7560.892 | 8.9307 | 5.2163
190 | 7600.897 | 8.936 | 5.1932 |  |  | 190 | 7600.897 | 8.936 | 5.2131 |  |  | 190 | 7600.897 | 8.936 | 5.2145
191 | 7640.902 | 8.9413 | 5.1914 |  |  | 191 | 7640.902 | 8.9413 | 5.2114 |  |  | 191 | 7640.902 | 8.9413 | 5.2127
192 | 7680.907 | 8.9465 | 5.1897 |  |  | 192 | 7680.907 | 8.9465 | 5.2096 |  |  | 192 | 7680.907 | 8.9465 | 5.211
193 | 7720.911 | 8.9517 | 5.188 |  |  | 193 | 7720.911 | 8.9517 | 5.2079 |  |  | 193 | 7720.911 | 8.9517 | 5.2093
194 | 7760.916 | 8.9569 | 5.1863 |  |  | 194 | 7760.916 | 8.9569 | 5.2062 |  |  | 194 | 7760.916 | 8.9569 | 5.2076
195 | 7800.921 | 8.962 | 5.1847 |  |  | 195 | 7800.921 | 8.962 | 5.2045 |  |  | 195 | 7800.921 | 8.962 | 5.2059
196 | 7840.925 | 8.9671 | 5.183 |  |  | 196 | 7840.925 | 8.9671 | 5.2028 |  |  | 196 | 7840.925 | 8.9671 | 5.2042
197 | 7880.93 | 8.9722 | 5.1814 |  |  | 197 | 7880.93 | 8.9722 | 5.2011 |  |  | 197 | 7880.93 | 8.9722 | 5.2025
198 | 7920.935 | 8.9773 | 5.1797 |  |  | 198 | 7920.935 | 8.9773 | 5.1995 |  |  | 198 | 7920.935 | 8.9773 | 5.2009
199 | 7960.94 | 8.9823 | 5.1781 |  |  | 199 | 7960.94 | 8.9823 | 5.1979 |  |  | 199 | 7960.94 | 8.9823 | 5.1992
200 | 8000.944 | 8.9873 | 5.1765 |  |  | 200 | 8000.944 | 8.9873 | 5.1962 |  |  | 200 | 8000.944 | 8.9873 | 5.1976
201 | 8040.949 | 8.9923 | 5.175 |  |  | 201 | 8040.949 | 8.9923 | 5.1946 |  |  | 201 | 8040.949 | 8.9923 | 5.196
202 | 8080.954 | 8.9973 | 5.1734 |  |  | 202 | 8080.954 | 8.9973 | 5.193 |  |  | 202 | 8080.954 | 8.9973 | 5.1944
203 | 8120.958 | 9.0022 | 5.1718 |  |  | 203 | 8120.958 | 9.0022 | 5.1915 |  |  | 203 | 8120.958 | 9.0022 | 5.1928
204 | 8160.963 | 9.0071 | 5.1703 |  |  | 204 | 8160.963 | 9.0071 | 5.1899 |  |  | 204 | 8160.963 | 9.0071 | 5.1912
205 | 8200.968 | 9.012 | 5.1688 |  |  | 205 | 8200.968 | 9.012 | 5.1884 |  |  | 205 | 8200.968 | 9.012 | 5.1897
206 | 8240.973 | 9.0169 | 5.1673 |  |  | 206 | 8240.973 | 9.0169 | 5.1868 |  |  | 206 | 8240.973 | 9.0169 | 5.1882
207 | 8280.977 | 9.0217 | 5.1658 |  |  | 207 | 8280.977 | 9.0217 | 5.1853 |  |  | 207 | 8280.977 | 9.0217 | 5.1866
208 | 8320.982 | 9.0265 | 5.1643 |  |  | 208 | 8320.982 | 9.0265 | 5.1838 |  |  | 208 | 8320.982 | 9.0265 | 5.1851
209 | 8360.987 | 9.0313 | 5.1628 |  |  | 209 | 8360.987 | 9.0313 | 5.1823 |  |  | 209 | 8360.987 | 9.0313 | 5.1836
210 | 8400.992 | 9.0361 | 5.1614 |  |  | 210 | 8400.992 | 9.0361 | 5.1808 |  |  | 210 | 8400.992 | 9.0361 | 5.1821
211 | 8440.996 | 9.0409 | 5.1599 |  |  | 211 | 8440.996 | 9.0409 | 5.1793 |  |  | 211 | 8440.996 | 9.0409 | 5.1807
212 | 8481.001 | 9.0456 | 5.1585 |  |  | 212 | 8481.001 | 9.0456 | 5.1779 |  |  | 212 | 8481.001 | 9.0456 | 5.1792
213 | 8521.006 | 9.0503 | 5.1571 |  |  | 213 | 8521.006 | 9.0503 | 5.1765 |  |  | 213 | 8521.006 | 9.0503 | 5.1778
214 | 8561.01 | 9.055 | 5.1557 |  |  | 214 | 8561.01 | 9.055 | 5.175 |  |  | 214 | 8561.01 | 9.055 | 5.1763
215 | 8601.015 | 9.0596 | 5.1543 |  |  | 215 | 8601.015 | 9.0596 | 5.1736 |  |  | 215 | 8601.015 | 9.0596 | 5.1749
216 | 8641.02 | 9.0643 | 5.1529 |  |  | 216 | 8641.02 | 9.0643 | 5.1722 |  |  | 216 | 8641.02 | 9.0643 | 5.1735
217 | 8681.025 | 9.0689 | 5.1515 |  |  | 217 | 8681.025 | 9.0689 | 5.1708 |  |  | 217 | 8681.025 | 9.0689 | 5.1721
218 | 8721.029 | 9.0735 | 5.1502 |  |  | 218 | 8721.029 | 9.0735 | 5.1694 |  |  | 218 | 8721.029 | 9.0735 | 5.1707
219 | 8761.034 | 9.0781 | 5.1488 |  |  | 219 | 8761.034 | 9.0781 | 5.1681 |  |  | 219 | 8761.034 | 9.0781 | 5.1694
220 | 8801.039 | 9.0826 | 5.1475 |  |  | 220 | 8801.039 | 9.0826 | 5.1667 |  |  | 220 | 8801.039 | 9.0826 | 5.168
221 | 8841.043 | 9.0872 | 5.1462 |  |  | 221 | 8841.043 | 9.0872 | 5.1654 |  |  | 221 | 8841.043 | 9.0872 | 5.1667
222 | 8881.048 | 9.0917 | 5.1449 |  |  | 222 | 8881.048 | 9.0917 | 5.164 |  |  | 222 | 8881.048 | 9.0917 | 5.1653
223 | 8921.053 | 9.0962 | 5.1436 |  |  | 223 | 8921.053 | 9.0962 | 5.1627 |  |  | 223 | 8921.053 | 9.0962 | 5.164
224 | 8961.058 | 9.1006 | 5.1423 |  |  | 224 | 8961.058 | 9.1006 | 5.1614 |  |  | 224 | 8961.058 | 9.1006 | 5.1627
225 | 9001.062 | 9.1051 | 5.141 |  |  | 225 | 9001.062 | 9.1051 | 5.1601 |  |  | 225 | 9001.062 | 9.1051 | 5.1614
226 | 9041.067 | 9.1095 | 5.1397 |  |  | 226 | 9041.067 | 9.1095 | 5.1588 |  |  | 226 | 9041.067 | 9.1095 | 5.1601
227 | 9081.072 | 9.1139 | 5.1385 |  |  | 227 | 9081.072 | 9.1139 | 5.1575 |  |  | 227 | 9081.072 | 9.1139 | 5.1588
228 | 9121.077 | 9.1183 | 5.1372 |  |  | 228 | 9121.077 | 9.1183 | 5.1563 |  |  | 228 | 9121.077 | 9.1183 | 5.1576
229 | 9161.081 | 9.1227 | 5.136 |  |  | 229 | 9161.081 | 9.1227 | 5.155 |  |  | 229 | 9161.081 | 9.1227 | 5.1563
230 | 9201.086 | 9.1271 | 5.1348 |  |  | 230 | 9201.086 | 9.1271 | 5.1538 |  |  | 230 | 9201.086 | 9.1271 | 5.1551
231 | 9241.091 | 9.1314 | 5.1336 |  |  | 231 | 9241.091 | 9.1314 | 5.1525 |  |  | 231 | 9241.091 | 9.1314 | 5.1538
232 | 9281.095 | 9.1357 | 5.1324 |  |  | 232 | 9281.095 | 9.1357 | 5.1513 |  |  | 232 | 9281.095 | 9.1357 | 5.1526
233 | 9321.1 | 9.14 | 5.1312 |  |  | 233 | 9321.1 | 9.14 | 5.1501 |  |  | 233 | 9321.1 | 9.14 | 5.1514
234 | 9361.105 | 9.1443 | 5.13 |  |  | 234 | 9361.105 | 9.1443 | 5.1489 |  |  | 234 | 9361.105 | 9.1443 | 5.1502
235 | 9401.11 | 9.1486 | 5.1288 |  |  | 235 | 9401.11 | 9.1486 | 5.1477 |  |  | 235 | 9401.11 | 9.1486 | 5.149
236 | 9441.114 | 9.1528 | 5.1277 |  |  | 236 | 9441.114 | 9.1528 | 5.1465 |  |  | 236 | 9441.114 | 9.1528 | 5.1478
237 | 9481.119 | 9.1571 | 5.1265 |  |  | 237 | 9481.119 | 9.1571 | 5.1454 |  |  | 237 | 9481.119 | 9.1571 | 5.1466
238 | 9521.124 | 9.1613 | 5.1254 |  |  | 238 | 9521.124 | 9.1613 | 5.1442 |  |  | 238 | 9521.124 | 9.1613 | 5.1455
239 | 9561.128 | 9.1655 | 5.1243 |  |  | 239 | 9561.128 | 9.1655 | 5.1431 |  |  | 239 | 9561.128 | 9.1655 | 5.1443
240 | 9601.133 | 9.1696 | 5.1231 |  |  | 240 | 9601.133 | 9.1696 | 5.1419 |  |  | 240 | 9601.133 | 9.1696 | 5.1432
241 | 9641.138 | 9.1738 | 5.122 |  |  | 241 | 9641.138 | 9.1738 | 5.1408 |  |  | 241 | 9641.138 | 9.1738 | 5.142
242 | 9681.143 | 9.1779 | 5.1209 |  |  | 242 | 9681.143 | 9.1779 | 5.1397 |  |  | 242 | 9681.143 | 9.1779 | 5.1409
243 | 9721.147 | 9.1821 | 5.1198 |  |  | 243 | 9721.147 | 9.1821 | 5.1386 |  |  | 243 | 9721.147 | 9.1821 | 5.1398
244 | 9761.152 | 9.1862 | 5.1188 |  |  | 244 | 9761.152 | 9.1862 | 5.1374 |  |  | 244 | 9761.152 | 9.1862 | 5.1387
245 | 9801.157 | 9.1903 | 5.1177 |  |  | 245 | 9801.157 | 9.1903 | 5.1364 |  |  | 245 | 9801.157 | 9.1903 | 5.1376
246 | 9841.162 | 9.1943 | 5.1166 |  |  | 246 | 9841.162 | 9.1943 | 5.1353 |  |  | 246 | 9841.162 | 9.1943 | 5.1365
247 | 9881.166 | 9.1984 | 5.1156 |  |  | 247 | 9881.166 | 9.1984 | 5.1342 |  |  | 247 | 9881.166 | 9.1984 | 5.1354
248 | 9921.171 | 9.2024 | 5.1145 |  |  | 248 | 9921.171 | 9.2024 | 5.1331 |  |  | 248 | 9921.171 | 9.2024 | 5.1344
249 | 9961.176 | 9.2065 | 5.1135 |  |  | 249 | 9961.176 | 9.2065 | 5.1321 |  |  | 249 | 9961.176 | 9.2065 | 5.1333
250 | 10001.18 | 9.2105 | 5.1125 |  |  | 250 | 10001.18 | 9.2105 | 5.131 |  |  | 250 | 10001.18 | 9.2105 | 5.1323
251 | 10041.185 | 9.2145 | 5.1114 |  |  | 251 | 10041.185 | 9.2145 | 5.13 |  |  | 251 | 10041.185 | 9.2145 | 5.1312
252 | 10081.19 | 9.2184 | 5.1104 |  |  | 252 | 10081.19 | 9.2184 | 5.129 |  |  | 252 | 10081.19 | 9.2184 | 5.1302
253 | 10121.195 | 9.2224 | 5.1094 |  |  | 253 | 10121.195 | 9.2224 | 5.1279 |  |  | 253 | 10121.195 | 9.2224 | 5.1292
254 | 10161.199 | 9.2263 | 5.1084 |  |  | 254 | 10161.199 | 9.2263 | 5.1269 |  |  | 254 | 10161.199 | 9.2263 | 5.1281
255 | 10201.204 | 9.2303 | 5.1074 |  |  | 255 | 10201.204 | 9.2303 | 5.1259 |  |  | 255 | 10201.204 | 9.2303 | 5.1271
256 | 10241.209 | 9.2342 | 5.1065 |  |  | 256 | 10241.209 | 9.2342 | 5.1249 |  |  | 256 | 10241.209 | 9.2342 | 5.1261
257 | 10281.213 | 9.2381 | 5.1055 |  |  | 257 | 10281.213 | 9.2381 | 5.1239 |  |  | 257 | 10281.213 | 9.2381 | 5.1251
258 | 10321.218 | 9.242 | 5.1045 |  |  | 258 | 10321.218 | 9.242 | 5.1229 |  |  | 258 | 10321.218 | 9.242 | 5.1242
259 | 10361.223 | 9.2458 | 5.1036 |  |  | 259 | 10361.223 | 9.2458 | 5.122 |  |  | 259 | 10361.223 | 9.2458 | 5.1232
260 | 10401.228 | 9.2497 | 5.1026 |  |  | 260 | 10401.228 | 9.2497 | 5.121 |  |  | 260 | 10401.228 | 9.2497 | 5.1222
261 | 10441.232 | 9.2535 | 5.1017 |  |  | 261 | 10441.232 | 9.2535 | 5.12 |  |  | 261 | 10441.232 | 9.2535 | 5.1213
262 | 10481.237 | 9.2573 | 5.1008 |  |  | 262 | 10481.237 | 9.2573 | 5.1191 |  |  | 262 | 10481.237 | 9.2573 | 5.1203
263 | 10521.242 | 9.2612 | 5.0998 |  |  | 263 | 10521.242 | 9.2612 | 5.1181 |  |  | 263 | 10521.242 | 9.2612 | 5.1194
264 | 10561.247 | 9.2649 | 5.0989 |  |  | 264 | 10561.247 | 9.2649 | 5.1172 |  |  | 264 | 10561.247 | 9.2649 | 5.1184
265 | 10601.251 | 9.2687 | 5.098 |  |  | 265 | 10601.251 | 9.2687 | 5.1163 |  |  | 265 | 10601.251 | 9.2687 | 5.1175
266 | 10641.256 | 9.2725 | 5.0971 |  |  | 266 | 10641.256 | 9.2725 | 5.1154 |  |  | 266 | 10641.256 | 9.2725 | 5.1166
267 | 10681.261 | 9.2762 | 5.0962 |  |  | 267 | 10681.261 | 9.2762 | 5.1145 |  |  | 267 | 10681.261 | 9.2762 | 5.1157
268 | 10721.265 | 9.28 | 5.0953 |  |  | 268 | 10721.265 | 9.28 | 5.1136 |  |  | 268 | 10721.265 | 9.28 | 5.1147
269 | 10761.27 | 9.2837 | 5.0945 |  |  | 269 | 10761.27 | 9.2837 | 5.1127 |  |  | 269 | 10761.27 | 9.2837 | 5.1138
270 | 10801.275 | 9.2874 | 5.0936 |  |  | 270 | 10801.275 | 9.2874 | 5.1118 |  |  | 270 | 10801.275 | 9.2874 | 5.113
271 | 10841.28 | 9.2911 | 5.0927 |  |  | 271 | 10841.28 | 9.2911 | 5.1109 |  |  | 271 | 10841.28 | 9.2911 | 5.1121
272 | 10881.284 | 9.2948 | 5.0919 |  |  | 272 | 10881.284 | 9.2948 | 5.11 |  |  | 272 | 10881.284 | 9.2948 | 5.1112
273 | 10921.289 | 9.2985 | 5.091 |  |  | 273 | 10921.289 | 9.2985 | 5.1091 |  |  | 273 | 10921.289 | 9.2985 | 5.1103
274 | 10961.294 | 9.3021 | 5.0902 |  |  | 274 | 10961.294 | 9.3021 | 5.1083 |  |  | 274 | 10961.294 | 9.3021 | 5.1095
275 | 11001.298 | 9.3058 | 5.0893 |  |  | 275 | 11001.298 | 9.3058 | 5.1074 |  |  | 275 | 11001.298 | 9.3058 | 5.1086
276 | 11041.303 | 9.3094 | 5.0885 |  |  | 276 | 11041.303 | 9.3094 | 5.1066 |  |  | 276 | 11041.303 | 9.3094 | 5.1078
277 | 11081.308 | 9.313 | 5.0877 |  |  | 277 | 11081.308 | 9.313 | 5.1057 |  |  | 277 | 11081.308 | 9.313 | 5.1069
278 | 11121.313 | 9.3166 | 5.0869 |  |  | 278 | 11121.313 | 9.3166 | 5.1049 |  |  | 278 | 11121.313 | 9.3166 | 5.1061
279 | 11161.317 | 9.3202 | 5.086 |  |  | 279 | 11161.317 | 9.3202 | 5.1041 |  |  | 279 | 11161.317 | 9.3202 | 5.1052
280 | 11201.322 | 9.3238 | 5.0852 |  |  | 280 | 11201.322 | 9.3238 | 5.1033 |  |  | 280 | 11201.322 | 9.3238 | 5.1044
281 | 11241.327 | 9.3274 | 5.0844 |  |  | 281 | 11241.327 | 9.3274 | 5.1024 |  |  | 281 | 11241.327 | 9.3274 | 5.1036
282 | 11281.331 | 9.3309 | 5.0836 |  |  | 282 | 11281.331 | 9.3309 | 5.1016 |  |  | 282 | 11281.331 | 9.3309 | 5.1028
283 | 11321.336 | 9.3344 | 5.0829 |  |  | 283 | 11321.336 | 9.3344 | 5.1008 |  |  | 283 | 11321.336 | 9.3344 | 5.102
284 | 11361.341 | 9.338 | 5.0821 |  |  | 284 | 11361.341 | 9.338 | 5.1 |  |  | 284 | 11361.341 | 9.338 | 5.1012
285 | 11401.346 | 9.3415 | 5.0813 |  |  | 285 | 11401.346 | 9.3415 | 5.0992 |  |  | 285 | 11401.346 | 9.3415 | 5.1004
286 | 11441.35 | 9.345 | 5.0805 |  |  | 286 | 11441.35 | 9.345 | 5.0985 |  |  | 286 | 11441.35 | 9.345 | 5.0996
287 | 11481.355 | 9.3485 | 5.0798 |  |  | 287 | 11481.355 | 9.3485 | 5.0977 |  |  | 287 | 11481.355 | 9.3485 | 5.0988
288 | 11521.36 | 9.352 | 5.079 |  |  | 288 | 11521.36 | 9.352 | 5.0969 |  |  | 288 | 11521.36 | 9.352 | 5.0981
289 | 11561.365 | 9.3554 | 5.0783 |  |  | 289 | 11561.365 | 9.3554 | 5.0961 |  |  | 289 | 11561.365 | 9.3554 | 5.0973
290 | 11601.369 | 9.3589 | 5.0775 |  |  | 290 | 11601.369 | 9.3589 | 5.0954 |  |  | 290 | 11601.369 | 9.3589 | 5.0965
291 | 11641.374 | 9.3623 | 5.0768 |  |  | 291 | 11641.374 | 9.3623 | 5.0946 |  |  | 291 | 11641.374 | 9.3623 | 5.0958
292 | 11681.379 | 9.3658 | 5.0761 |  |  | 292 | 11681.379 | 9.3658 | 5.0939 |  |  | 292 | 11681.379 | 9.3658 | 5.095
293 | 11721.383 | 9.3692 | 5.0753 |  |  | 293 | 11721.383 | 9.3692 | 5.0932 |  |  | 293 | 11721.383 | 9.3692 | 5.0943
294 | 11761.388 | 9.3726 | 5.0746 |  |  | 294 | 11761.388 | 9.3726 | 5.0924 |  |  | 294 | 11761.388 | 9.3726 | 5.0936
295 | 11801.393 | 9.376 | 5.0739 |  |  | 295 | 11801.393 | 9.376 | 5.0917 |  |  | 295 | 11801.393 | 9.376 | 5.0928
296 | 11841.398 | 9.3794 | 5.0732 |  |  | 296 | 11841.398 | 9.3794 | 5.091 |  |  | 296 | 11841.398 | 9.3794 | 5.0921
297 | 11881.402 | 9.3827 | 5.0725 |  |  | 297 | 11881.402 | 9.3827 | 5.0902 |  |  | 297 | 11881.402 | 9.3827 | 5.0914
298 | 11921.407 | 9.3861 | 5.0718 |  |  | 298 | 11921.407 | 9.3861 | 5.0895 |  |  | 298 | 11921.407 | 9.3861 | 5.0907
299 | 11961.412 | 9.3894 | 5.0711 |  |  | 299 | 11961.412 | 9.3894 | 5.0888 |  |  | 299 | 11961.412 | 9.3894 | 5.09
300 | 12001.416 | 9.3928 | 5.0704 |  |  | 300 | 12001.416 | 9.3928 | 5.0881 |  |  | 300 | 12001.416 | 9.3928 | 5.0893
301 | 12041.421 | 9.3961 | 5.0697 |  |  | 301 | 12041.421 | 9.3961 | 5.0874 |  |  | 301 | 12041.421 | 9.3961 | 5.0886
302 | 12081.426 | 9.3994 | 5.0691 |  |  | 302 | 12081.426 | 9.3994 | 5.0867 |  |  | 302 | 12081.426 | 9.3994 | 5.0879
303 | 12121.431 | 9.4027 | 5.0684 |  |  | 303 | 12121.431 | 9.4027 | 5.0861 |  |  | 303 | 12121.431 | 9.4027 | 5.0872
304 | 12161.435 | 9.406 | 5.0677 |  |  | 304 | 12161.435 | 9.406 | 5.0854 |  |  | 304 | 12161.435 | 9.406 | 5.0865
305 | 12201.44 | 9.4093 | 5.0671 |  |  | 305 | 12201.44 | 9.4093 | 5.0847 |  |  | 305 | 12201.44 | 9.4093 | 5.0858
306 | 12241.445 | 9.4126 | 5.0664 |  |  | 306 | 12241.445 | 9.4126 | 5.084 |  |  | 306 | 12241.445 | 9.4126 | 5.0852
307 | 12281.45 | 9.4158 | 5.0658 |  |  | 307 | 12281.45 | 9.4158 | 5.0834 |  |  | 307 | 12281.45 | 9.4158 | 5.0845
308 | 12321.454 | 9.4191 | 5.0651 |  |  | 308 | 12321.454 | 9.4191 | 5.0827 |  |  | 308 | 12321.454 | 9.4191 | 5.0838
309 | 12361.459 | 9.4223 | 5.0645 |  |  | 309 | 12361.459 | 9.4223 | 5.0821 |  |  | 309 | 12361.459 | 9.4223 | 5.0832
310 | 12401.464 | 9.4256 | 5.0638 |  |  | 310 | 12401.464 | 9.4256 | 5.0814 |  |  | 310 | 12401.464 | 9.4256 | 5.0825
311 | 12441.468 | 9.4288 | 5.0632 |  |  | 311 | 12441.468 | 9.4288 | 5.0808 |  |  | 311 | 12441.468 | 9.4288 | 5.0819
312 | 12481.473 | 9.432 | 5.0626 |  |  | 312 | 12481.473 | 9.432 | 5.0801 |  |  | 312 | 12481.473 | 9.432 | 5.0812
313 | 12521.478 | 9.4352 | 5.0619 |  |  | 313 | 12521.478 | 9.4352 | 5.0795 |  |  | 313 | 12521.478 | 9.4352 | 5.0806
314 | 12561.483 | 9.4384 | 5.0613 |  |  | 314 | 12561.483 | 9.4384 | 5.0789 |  |  | 314 | 12561.483 | 9.4384 | 5.08
315 | 12601.487 | 9.4416 | 5.0607 |  |  | 315 | 12601.487 | 9.4416 | 5.0782 |  |  | 315 | 12601.487 | 9.4416 | 5.0793
316 | 12641.492 | 9.4447 | 5.0601 |  |  | 316 | 12641.492 | 9.4447 | 5.0776 |  |  | 316 | 12641.492 | 9.4447 | 5.0787
317 | 12681.497 | 9.4479 | 5.0595 |  |  | 317 | 12681.497 | 9.4479 | 5.077 |  |  | 317 | 12681.497 | 9.4479 | 5.0781
318 | 12721.501 | 9.451 | 5.0589 |  |  | 318 | 12721.501 | 9.451 | 5.0764 |  |  | 318 | 12721.501 | 9.451 | 5.0775
319 | 12761.506 | 9.4542 | 5.0583 |  |  | 319 | 12761.506 | 9.4542 | 5.0758 |  |  | 319 | 12761.506 | 9.4542 | 5.0769
320 | 12801.511 | 9.4573 | 5.0577 |  |  | 320 | 12801.511 | 9.4573 | 5.0752 |  |  | 320 | 12801.511 | 9.4573 | 5.0763
321 | 12841.516 | 9.4604 | 5.0571 |  |  | 321 | 12841.516 | 9.4604 | 5.0746 |  |  | 321 | 12841.516 | 9.4604 | 5.0757
322 | 12881.52 | 9.4635 | 5.0566 |  |  | 322 | 12881.52 | 9.4635 | 5.074 |  |  | 322 | 12881.52 | 9.4635 | 5.0751
323 | 12921.525 | 9.4666 | 5.056 |  |  | 323 | 12921.525 | 9.4666 | 5.0734 |  |  | 323 | 12921.525 | 9.4666 | 5.0745
324 | 12961.53 | 9.4697 | 5.0554 |  |  | 324 | 12961.53 | 9.4697 | 5.0728 |  |  | 324 | 12961.53 | 9.4697 | 5.0739
325 | 13001.535 | 9.4728 | 5.0548 |  |  | 325 | 13001.535 | 9.4728 | 5.0722 |  |  | 325 | 13001.535 | 9.4728 | 5.0733
326 | 13041.539 | 9.4759 | 5.0543 |  |  | 326 | 13041.539 | 9.4759 | 5.0717 |  |  | 326 | 13041.539 | 9.4759 | 5.0727
327 | 13081.544 | 9.479 | 5.0537 |  |  | 327 | 13081.544 | 9.479 | 5.0711 |  |  | 327 | 13081.544 | 9.479 | 5.0722
328 | 13121.549 | 9.482 | 5.0532 |  |  | 328 | 13121.549 | 9.482 | 5.0705 |  |  | 328 | 13121.549 | 9.482 | 5.0716
329 | 13161.553 | 9.4851 | 5.0526 |  |  | 329 | 13161.553 | 9.4851 | 5.07 |  |  | 329 | 13161.553 | 9.4851 | 5.071
330 | 13201.558 | 9.4881 | 5.0521 |  |  | 330 | 13201.558 | 9.4881 | 5.0694 |  |  | 330 | 13201.558 | 9.4881 | 5.0705
331 | 13241.563 | 9.4911 | 5.0515 |  |  | 331 | 13241.563 | 9.4911 | 5.0688 |  |  | 331 | 13241.563 | 9.4911 | 5.0699
332 | 13281.568 | 9.4941 | 5.051 |  |  | 332 | 13281.568 | 9.4941 | 5.0683 |  |  | 332 | 13281.568 | 9.4941 | 5.0694
333 | 13321.572 | 9.4971 | 5.0504 |  |  | 333 | 13321.572 | 9.4971 | 5.0678 |  |  | 333 | 13321.572 | 9.4971 | 5.0688
334 | 13361.577 | 9.5001 | 5.0499 |  |  | 334 | 13361.577 | 9.5001 | 5.0672 |  |  | 334 | 13361.577 | 9.5001 | 5.0683
335 | 13401.582 | 9.5031 | 5.0494 |  |  | 335 | 13401.582 | 9.5031 | 5.0667 |  |  | 335 | 13401.582 | 9.5031 | 5.0677
336 | 13441.586 | 9.5061 | 5.0489 |  |  | 336 | 13441.586 | 9.5061 | 5.0661 |  |  | 336 | 13441.586 | 9.5061 | 5.0672
337 | 13481.591 | 9.5091 | 5.0483 |  |  | 337 | 13481.591 | 9.5091 | 5.0656 |  |  | 337 | 13481.591 | 9.5091 | 5.0667
338 | 13521.596 | 9.512 | 5.0478 |  |  | 338 | 13521.596 | 9.512 | 5.0651 |  |  | 338 | 13521.596 | 9.512 | 5.0661
339 | 13561.601 | 9.515 | 5.0473 |  |  | 339 | 13561.601 | 9.515 | 5.0646 |  |  | 339 | 13561.601 | 9.515 | 5.0656
340 | 13601.605 | 9.5179 | 5.0468 |  |  | 340 | 13601.605 | 9.5179 | 5.064 |  |  | 340 | 13601.605 | 9.5179 | 5.0651
341 | 13641.61 | 9.5209 | 5.0463 |  |  | 341 | 13641.61 | 9.5209 | 5.0635 |  |  | 341 | 13641.61 | 9.5209 | 5.0646
342 | 13681.615 | 9.5238 | 5.0458 |  |  | 342 | 13681.615 | 9.5238 | 5.063 |  |  | 342 | 13681.615 | 9.5238 | 5.0641
343 | 13721.62 | 9.5267 | 5.0453 |  |  | 343 | 13721.62 | 9.5267 | 5.0625 |  |  | 343 | 13721.62 | 9.5267 | 5.0635
344 | 13761.624 | 9.5296 | 5.0448 |  |  | 344 | 13761.624 | 9.5296 | 5.062 |  |  | 344 | 13761.624 | 9.5296 | 5.063
345 | 13801.629 | 9.5325 | 5.0443 |  |  | 345 | 13801.629 | 9.5325 | 5.0615 |  |  | 345 | 13801.629 | 9.5325 | 5.0625
346 | 13841.634 | 9.5354 | 5.0438 |  |  | 346 | 13841.634 | 9.5354 | 5.061 |  |  | 346 | 13841.634 | 9.5354 | 5.062
347 | 13881.638 | 9.5383 | 5.0433 |  |  | 347 | 13881.638 | 9.5383 | 5.0605 |  |  | 347 | 13881.638 | 9.5383 | 5.0615
348 | 13921.643 | 9.5412 | 5.0429 |  |  | 348 | 13921.643 | 9.5412 | 5.06 |  |  | 348 | 13921.643 | 9.5412 | 5.0611
349 | 13961.648 | 9.5441 | 5.0424 |  |  | 349 | 13961.648 | 9.5441 | 5.0595 |  |  | 349 | 13961.648 | 9.5441 | 5.0606
350 | 14001.653 | 9.5469 | 5.0419 |  |  | 350 | 14001.653 | 9.5469 | 5.0591 |  |  | 350 | 14001.653 | 9.5469 | 5.0601
351 | 14041.657 | 9.5498 | 5.0414 |  |  | 351 | 14041.657 | 9.5498 | 5.0586 |  |  | 351 | 14041.657 | 9.5498 | 5.0596
352 | 14081.662 | 9.5526 | 5.041 |  |  | 352 | 14081.662 | 9.5526 | 5.0581 |  |  | 352 | 14081.662 | 9.5526 | 5.0591
353 | 14121.667 | 9.5555 | 5.0405 |  |  | 353 | 14121.667 | 9.5555 | 5.0576 |  |  | 353 | 14121.667 | 9.5555 | 5.0587
354 | 14161.671 | 9.5583 | 5.04 |  |  | 354 | 14161.671 | 9.5583 | 5.0572 |  |  | 354 | 14161.671 | 9.5583 | 5.0582
355 | 14201.676 | 9.5611 | 5.0396 |  |  | 355 | 14201.676 | 9.5611 | 5.0567 |  |  | 355 | 14201.676 | 9.5611 | 5.0577
356 | 14241.681 | 9.5639 | 5.0391 |  |  | 356 | 14241.681 | 9.5639 | 5.0562 |  |  | 356 | 14241.681 | 9.5639 | 5.0572
357 | 14281.686 | 9.5667 | 5.0387 |  |  | 357 | 14281.686 | 9.5667 | 5.0558 |  |  | 357 | 14281.686 | 9.5667 | 5.0568
358 | 14321.69 | 9.5695 | 5.0382 |  |  | 358 | 14321.69 | 9.5695 | 5.0553 |  |  | 358 | 14321.69 | 9.5695 | 5.0563
359 | 14361.695 | 9.5723 | 5.0378 |  |  | 359 | 14361.695 | 9.5723 | 5.0549 |  |  | 359 | 14361.695 | 9.5723 | 5.0559
360 | 14401.7 | 9.5751 | 5.0373 |  |  | 360 | 14401.7 | 9.5751 | 5.0544 |  |  | 360 | 14401.7 | 9.5751 | 5.0554
361 | 14441.705 | 9.5779 | 5.0369 |  |  | 361 | 14441.705 | 9.5779 | 5.054 |  |  | 361 | 14441.705 | 9.5779 | 5.055
362 | 14481.709 | 9.5806 | 5.0365 |  |  | 362 | 14481.709 | 9.5806 | 5.0535 |  |  | 362 | 14481.709 | 9.5806 | 5.0545
363 | 14521.714 | 9.5834 | 5.036 |  |  | 363 | 14521.714 | 9.5834 | 5.0531 |  |  | 363 | 14521.714 | 9.5834 | 5.0541
364 | 14561.719 | 9.5862 | 5.0356 |  |  | 364 | 14561.719 | 9.5862 | 5.0527 |  |  | 364 | 14561.719 | 9.5862 | 5.0536
365 | 14601.723 | 9.5889 | 5.0352 |  |  | 365 | 14601.723 | 9.5889 | 5.0522 |  |  | 365 | 14601.723 | 9.5889 | 5.0532
366 | 14641.728 | 9.5916 | 5.0348 |  |  | 366 | 14641.728 | 9.5916 | 5.0518 |  |  | 366 | 14641.728 | 9.5916 | 5.0528
367 | 14681.733 | 9.5944 | 5.0343 |  |  | 367 | 14681.733 | 9.5944 | 5.0514 |  |  | 367 | 14681.733 | 9.5944 | 5.0524
368 | 14721.738 | 9.5971 | 5.0339 |  |  | 368 | 14721.738 | 9.5971 | 5.0509 |  |  | 368 | 14721.738 | 9.5971 | 5.0519
369 | 14761.742 | 9.5998 | 5.0335 |  |  | 369 | 14761.742 | 9.5998 | 5.0505 |  |  | 369 | 14761.742 | 9.5998 | 5.0515
370 | 14801.747 | 9.6025 | 5.0331 |  |  | 370 | 14801.747 | 9.6025 | 5.0501 |  |  | 370 | 14801.747 | 9.6025 | 5.0511
371 | 14841.752 | 9.6052 | 5.0327 |  |  | 371 | 14841.752 | 9.6052 | 5.0497 |  |  | 371 | 14841.752 | 9.6052 | 5.0507
372 | 14881.756 | 9.6079 | 5.0323 |  |  | 372 | 14881.756 | 9.6079 | 5.0493 |  |  | 372 | 14881.756 | 9.6079 | 5.0502
373 | 14921.761 | 9.6106 | 5.0319 |  |  | 373 | 14921.761 | 9.6106 | 5.0489 |  |  | 373 | 14921.761 | 9.6106 | 5.0498
374 | 14961.766 | 9.6133 | 5.0315 |  |  | 374 | 14961.766 | 9.6133 | 5.0485 |  |  | 374 | 14961.766 | 9.6133 | 5.0494
375 | 15001.771 | 9.6159 | 5.0311 |  |  | 375 | 15001.771 | 9.6159 | 5.048 |  |  | 375 | 15001.771 | 9.6159 | 5.049
376 | 15041.775 | 9.6186 | 5.0307 |  |  | 376 | 15041.775 | 9.6186 | 5.0476 |  |  | 376 | 15041.775 | 9.6186 | 5.0486
377 | 15081.78 | 9.6212 | 5.0303 |  |  | 377 | 15081.78 | 9.6212 | 5.0472 |  |  | 377 | 15081.78 | 9.6212 | 5.0482
378 | 15121.785 | 9.6239 | 5.0299 |  |  | 378 | 15121.785 | 9.6239 | 5.0469 |  |  | 378 | 15121.785 | 9.6239 | 5.0478
379 | 15161.789 | 9.6265 | 5.0295 |  |  | 379 | 15161.789 | 9.6265 | 5.0465 |  |  | 379 | 15161.789 | 9.6265 | 5.0474
380 | 15201.794 | 9.6292 | 5.0291 |  |  | 380 | 15201.794 | 9.6292 | 5.0461 |  |  | 380 | 15201.794 | 9.6292 | 5.047
381 | 15241.799 | 9.6318 | 5.0287 |  |  | 381 | 15241.799 | 9.6318 | 5.0457 |  |  | 381 | 15241.799 | 9.6318 | 5.0466
382 | 15281.804 | 9.6344 | 5.0284 |  |  | 382 | 15281.804 | 9.6344 | 5.0453 |  |  | 382 | 15281.804 | 9.6344 | 5.0462
383 | 15321.808 | 9.637 | 5.028 |  |  | 383 | 15321.808 | 9.637 | 5.0449 |  |  | 383 | 15321.808 | 9.637 | 5.0459
384 | 15361.813 | 9.6396 | 5.0276 |  |  | 384 | 15361.813 | 9.6396 | 5.0445 |  |  | 384 | 15361.813 | 9.6396 | 5.0455
385 | 15401.818 | 9.6422 | 5.0272 |  |  | 385 | 15401.818 | 9.6422 | 5.0442 |  |  | 385 | 15401.818 | 9.6422 | 5.0451
386 | 15441.823 | 9.6448 | 5.0269 |  |  | 386 | 15441.823 | 9.6448 | 5.0438 |  |  | 386 | 15441.823 | 9.6448 | 5.0447
387 | 15481.827 | 9.6474 | 5.0265 |  |  | 387 | 15481.827 | 9.6474 | 5.0434 |  |  | 387 | 15481.827 | 9.6474 | 5.0443
388 | 15521.832 | 9.65 | 5.0261 |  |  | 388 | 15521.832 | 9.65 | 5.043 |  |  | 388 | 15521.832 | 9.65 | 5.044
389 | 15561.837 | 9.6526 | 5.0258 |  |  | 389 | 15561.837 | 9.6526 | 5.0427 |  |  | 389 | 15561.837 | 9.6526 | 5.0436
390 | 15601.841 | 9.6551 | 5.0254 |  |  | 390 | 15601.841 | 9.6551 | 5.0423 |  |  | 390 | 15601.841 | 9.6551 | 5.0432
391 | 15641.846 | 9.6577 | 5.0251 |  |  | 391 | 15641.846 | 9.6577 | 5.0419 |  |  | 391 | 15641.846 | 9.6577 | 5.0429
392 | 15681.851 | 9.6603 | 5.0247 |  |  | 392 | 15681.851 | 9.6603 | 5.0416 |  |  | 392 | 15681.851 | 9.6603 | 5.0425
393 | 15721.856 | 9.6628 | 5.0243 |  |  | 393 | 15721.856 | 9.6628 | 5.0412 |  |  | 393 | 15721.856 | 9.6628 | 5.0421
394 | 15761.86 | 9.6653 | 5.024 |  |  | 394 | 15761.86 | 9.6653 | 5.0409 |  |  | 394 | 15761.86 | 9.6653 | 5.0418
395 | 15801.865 | 9.6679 | 5.0236 |  |  | 395 | 15801.865 | 9.6679 | 5.0405 |  |  | 395 | 15801.865 | 9.6679 | 5.0414
396 | 15841.87 | 9.6704 | 5.0233 |  |  | 396 | 15841.87 | 9.6704 | 5.0402 |  |  | 396 | 15841.87 | 9.6704 | 5.0411
397 | 15881.874 | 9.6729 | 5.023 |  |  | 397 | 15881.874 | 9.6729 | 5.0398 |  |  | 397 | 15881.874 | 9.6729 | 5.0407
398 | 15921.879 | 9.6754 | 5.0226 |  |  | 398 | 15921.879 | 9.6754 | 5.0395 |  |  | 398 | 15921.879 | 9.6754 | 5.0404
399 | 15961.884 | 9.678 | 5.0223 |  |  | 399 | 15961.884 | 9.678 | 5.0391 |  |  | 399 | 15961.884 | 9.678 | 5.04
400 | 16001.889 | 9.6805 | 5.0219 |  |  | 400 | 16001.889 | 9.6805 | 5.0388 |  |  | 400 | 16001.889 | 9.6805 | 5.0397
401 | 16041.893 | 9.683 | 5.0216 |  |  | 401 | 16041.893 | 9.683 | 5.0384 |  |  | 401 | 16041.893 | 9.683 | 5.0393
402 | 16081.898 | 9.6854 | 5.0213 |  |  | 402 | 16081.898 | 9.6854 | 5.0381 |  |  | 402 | 16081.898 | 9.6854 | 5.039
403 | 16121.903 | 9.6879 | 5.0209 |  |  | 403 | 16121.903 | 9.6879 | 5.0378 |  |  | 403 | 16121.903 | 9.6879 | 5.0387
404 | 16161.908 | 9.6904 | 5.0206 |  |  | 404 | 16161.908 | 9.6904 | 5.0374 |  |  | 404 | 16161.908 | 9.6904 | 5.0383
405 | 16201.912 | 9.6929 | 5.0203 |  |  | 405 | 16201.912 | 9.6929 | 5.0371 |  |  | 405 | 16201.912 | 9.6929 | 5.038
406 | 16241.917 | 9.6954 | 5.02 |  |  | 406 | 16241.917 | 9.6954 | 5.0368 |  |  | 406 | 16241.917 | 9.6954 | 5.0377
407 | 16281.922 | 9.6978 | 5.0196 |  |  | 407 | 16281.922 | 9.6978 | 5.0365 |  |  | 407 | 16281.922 | 9.6978 | 5.0373
408 | 16321.926 | 9.7003 | 5.0193 |  |  | 408 | 16321.926 | 9.7003 | 5.0361 |  |  | 408 | 16321.926 | 9.7003 | 5.037
409 | 16361.931 | 9.7027 | 5.019 |  |  | 409 | 16361.931 | 9.7027 | 5.0358 |  |  | 409 | 16361.931 | 9.7027 | 5.0367
410 | 16401.936 | 9.7052 | 5.0187 |  |  | 410 | 16401.936 | 9.7052 | 5.0355 |  |  | 410 | 16401.936 | 9.7052 | 5.0364
411 | 16441.941 | 9.7076 | 5.0184 |  |  | 411 | 16441.941 | 9.7076 | 5.0352 |  |  | 411 | 16441.941 | 9.7076 | 5.036
412 | 16481.945 | 9.71 | 5.018 |  |  | 412 | 16481.945 | 9.71 | 5.0349 |  |  | 412 | 16481.945 | 9.71 | 5.0357
413 | 16521.95 | 9.7124 | 5.0177 |  |  | 413 | 16521.95 | 9.7124 | 5.0345 |  |  | 413 | 16521.95 | 9.7124 | 5.0354
414 | 16561.955 | 9.7149 | 5.0174 |  |  | 414 | 16561.955 | 9.7149 | 5.0342 |  |  | 414 | 16561.955 | 9.7149 | 5.0351
415 | 16601.959 | 9.7173 | 5.0171 |  |  | 415 | 16601.959 | 9.7173 | 5.0339 |  |  | 415 | 16601.959 | 9.7173 | 5.0348
416 | 16641.964 | 9.7197 | 5.0168 |  |  | 416 | 16641.964 | 9.7197 | 5.0336 |  |  | 416 | 16641.964 | 9.7197 | 5.0345
417 | 16681.969 | 9.7221 | 5.0165 |  |  | 417 | 16681.969 | 9.7221 | 5.0333 |  |  | 417 | 16681.969 | 9.7221 | 5.0342
418 | 16721.974 | 9.7245 | 5.0162 |  |  | 418 | 16721.974 | 9.7245 | 5.033 |  |  | 418 | 16721.974 | 9.7245 | 5.0339
419 | 16761.978 | 9.7269 | 5.0159 |  |  | 419 | 16761.978 | 9.7269 | 5.0327 |  |  | 419 | 16761.978 | 9.7269 | 5.0336
420 | 16801.983 | 9.7293 | 5.0156 |  |  | 420 | 16801.983 | 9.7293 | 5.0324 |  |  | 420 | 16801.983 | 9.7293 | 5.0332
421 | 16841.988 | 9.7316 | 5.0153 |  |  | 421 | 16841.988 | 9.7316 | 5.0321 |  |  | 421 | 16841.988 | 9.7316 | 5.0329
422 | 16881.993 | 9.734 | 5.015 |  |  | 422 | 16881.993 | 9.734 | 5.0318 |  |  | 422 | 16881.993 | 9.734 | 5.0326
423 | 16921.997 | 9.7364 | 5.0147 |  |  | 423 | 16921.997 | 9.7364 | 5.0315 |  |  | 423 | 16921.997 | 9.7364 | 5.0324
424 | 16962.002 | 9.7387 | 5.0144 |  |  | 424 | 16962.002 | 9.7387 | 5.0312 |  |  | 424 | 16962.002 | 9.7387 | 5.0321
425 | 17002.007 | 9.7411 | 5.0141 |  |  | 425 | 17002.007 | 9.7411 | 5.0309 |  |  | 425 | 17002.007 | 9.7411 | 5.0318
426 | 17042.011 | 9.7434 | 5.0139 |  |  | 426 | 17042.011 | 9.7434 | 5.0306 |  |  | 426 | 17042.011 | 9.7434 | 5.0315
427 | 17082.016 | 9.7458 | 5.0136 |  |  | 427 | 17082.016 | 9.7458 | 5.0303 |  |  | 427 | 17082.016 | 9.7458 | 5.0312
428 | 17122.021 | 9.7481 | 5.0133 |  |  | 428 | 17122.021 | 9.7481 | 5.0301 |  |  | 428 | 17122.021 | 9.7481 | 5.0309
429 | 17162.026 | 9.7505 | 5.013 |  |  | 429 | 17162.026 | 9.7505 | 5.0298 |  |  | 429 | 17162.026 | 9.7505 | 5.0306
430 | 17202.03 | 9.7528 | 5.0127 |  |  | 430 | 17202.03 | 9.7528 | 5.0295 |  |  | 430 | 17202.03 | 9.7528 | 5.0303
431 | 17242.035 | 9.7551 | 5.0124 |  |  | 431 | 17242.035 | 9.7551 | 5.0292 |  |  | 431 | 17242.035 | 9.7551 | 5.03
432 | 17282.04 | 9.7574 | 5.0122 |  |  | 432 | 17282.04 | 9.7574 | 5.0289 |  |  | 432 | 17282.04 | 9.7574 | 5.0298
433 | 17322.044 | 9.7597 | 5.0119 |  |  | 433 | 17322.044 | 9.7597 | 5.0287 |  |  | 433 | 17322.044 | 9.7597 | 5.0295
434 | 17362.049 | 9.762 | 5.0116 |  |  | 434 | 17362.049 | 9.762 | 5.0284 |  |  | 434 | 17362.049 | 9.762 | 5.0292
435 | 17402.054 | 9.7643 | 5.0113 |  |  | 435 | 17402.054 | 9.7643 | 5.0281 |  |  | 435 | 17402.054 | 9.7643 | 5.0289
436 | 17442.059 | 9.7666 | 5.0111 |  |  | 436 | 17442.059 | 9.7666 | 5.0278 |  |  | 436 | 17442.059 | 9.7666 | 5.0286
437 | 17482.063 | 9.7689 | 5.0108 |  |  | 437 | 17482.063 | 9.7689 | 5.0276 |  |  | 437 | 17482.063 | 9.7689 | 5.0284
438 | 17522.068 | 9.7712 | 5.0105 |  |  | 438 | 17522.068 | 9.7712 | 5.0273 |  |  | 438 | 17522.068 | 9.7712 | 5.0281
439 | 17562.073 | 9.7735 | 5.0103 |  |  | 439 | 17562.073 | 9.7735 | 5.027 |  |  | 439 | 17562.073 | 9.7735 | 5.0278
440 | 17602.078 | 9.7758 | 5.01 |  |  | 440 | 17602.078 | 9.7758 | 5.0268 |  |  | 440 | 17602.078 | 9.7758 | 5.0276
441 | 17642.082 | 9.778 | 5.0097 |  |  | 441 | 17642.082 | 9.778 | 5.0265 |  |  | 441 | 17642.082 | 9.778 | 5.0273
442 | 17682.087 | 9.7803 | 5.0095 |  |  | 442 | 17682.087 | 9.7803 | 5.0262 |  |  | 442 | 17682.087 | 9.7803 | 5.027
443 | 17722.092 | 9.7826 | 5.0092 |  |  | 443 | 17722.092 | 9.7826 | 5.026 |  |  | 443 | 17722.092 | 9.7826 | 5.0268
444 | 17762.096 | 9.7848 | 5.009 |  |  | 444 | 17762.096 | 9.7848 | 5.0257 |  |  | 444 | 17762.096 | 9.7848 | 5.0265
445 | 17802.101 | 9.7871 | 5.0087 |  |  | 445 | 17802.101 | 9.7871 | 5.0255 |  |  | 445 | 17802.101 | 9.7871 | 5.0262
446 | 17842.106 | 9.7893 | 5.0085 |  |  | 446 | 17842.106 | 9.7893 | 5.0252 |  |  | 446 | 17842.106 | 9.7893 | 5.026
447 | 17882.111 | 9.7916 | 5.0082 |  |  | 447 | 17882.111 | 9.7916 | 5.025 |  |  | 447 | 17882.111 | 9.7916 | 5.0257
448 | 17922.115 | 9.7938 | 5.0079 |  |  | 448 | 17922.115 | 9.7938 | 5.0247 |  |  | 448 | 17922.115 | 9.7938 | 5.0255
449 | 17962.12 | 9.796 | 5.0077 |  |  | 449 | 17962.12 | 9.796 | 5.0244 |  |  | 449 | 17962.12 | 9.796 | 5.0252
450 | 18002.125 | 9.7982 | 5.0074 |  |  | 450 | 18002.125 | 9.7982 | 5.0242 |  |  | 450 | 18002.125 | 9.7982 | 5.025
451 | 18042.129 | 9.8005 | 5.0072 |  |  | 451 | 18042.129 | 9.8005 | 5.0239 |  |  | 451 | 18042.129 | 9.8005 | 5.0247
452 | 18082.134 | 9.8027 | 5.007 |  |  | 452 | 18082.134 | 9.8027 | 5.0237 |  |  | 452 | 18082.134 | 9.8027 | 5.0245
453 | 18122.139 | 9.8049 | 5.0067 |  |  | 453 | 18122.139 | 9.8049 | 5.0235 |  |  | 453 | 18122.139 | 9.8049 | 5.0242
454 | 18162.144 | 9.8071 | 5.0065 |  |  | 454 | 18162.144 | 9.8071 | 5.0232 |  |  | 454 | 18162.144 | 9.8071 | 5.024
455 | 18202.148 | 9.8093 | 5.0062 |  |  | 455 | 18202.148 | 9.8093 | 5.023 |  |  | 455 | 18202.148 | 9.8093 | 5.0237
456 | 18242.153 | 9.8115 | 5.006 |  |  | 456 | 18242.153 | 9.8115 | 5.0227 |  |  | 456 | 18242.153 | 9.8115 | 5.0235
457 | 18282.158 | 9.8137 | 5.0057 |  |  | 457 | 18282.158 | 9.8137 | 5.0225 |  |  | 457 | 18282.158 | 9.8137 | 5.0232
458 | 18322.163 | 9.8159 | 5.0055 |  |  | 458 | 18322.163 | 9.8159 | 5.0222 |  |  | 458 | 18322.163 | 9.8159 | 5.023
459 | 18362.167 | 9.818 | 5.0053 |  |  | 459 | 18362.167 | 9.818 | 5.022 |  |  | 459 | 18362.167 | 9.818 | 5.0227
460 | 18402.172 | 9.8202 | 5.005 |  |  | 460 | 18402.172 | 9.8202 | 5.0218 |  |  | 460 | 18402.172 | 9.8202 | 5.0225
461 | 18442.177 | 9.8224 | 5.0048 |  |  | 461 | 18442.177 | 9.8224 | 5.0215 |  |  | 461 | 18442.177 | 9.8224 | 5.0223
462 | 18482.181 | 9.8246 | 5.0046 |  |  | 462 | 18482.181 | 9.8246 | 5.0213 |  |  | 462 | 18482.181 | 9.8246 | 5.022
463 | 18522.186 | 9.8267 | 5.0043 |  |  | 463 | 18522.186 | 9.8267 | 5.0211 |  |  | 463 | 18522.186 | 9.8267 | 5.0218
464 | 18562.191 | 9.8289 | 5.0041 |  |  | 464 | 18562.191 | 9.8289 | 5.0208 |  |  | 464 | 18562.191 | 9.8289 | 5.0216
465 | 18602.196 | 9.831 | 5.0039 |  |  | 465 | 18602.196 | 9.831 | 5.0206 |  |  | 465 | 18602.196 | 9.831 | 5.0213
466 | 18642.2 | 9.8332 | 5.0036 |  |  | 466 | 18642.2 | 9.8332 | 5.0204 |  |  | 466 | 18642.2 | 9.8332 | 5.0211
467 | 18682.205 | 9.8353 | 5.0034 |  |  | 467 | 18682.205 | 9.8353 | 5.0202 |  |  | 467 | 18682.205 | 9.8353 | 5.0209
468 | 18722.21 | 9.8375 | 5.0032 |  |  | 468 | 18722.21 | 9.8375 | 5.0199 |  |  | 468 | 18722.21 | 9.8375 | 5.0206
469 | 18762.214 | 9.8396 | 5.0029 |  |  | 469 | 18762.214 | 9.8396 | 5.0197 |  |  | 469 | 18762.214 | 9.8396 | 5.0204
470 | 18802.219 | 9.8417 | 5.0027 |  |  | 470 | 18802.219 | 9.8417 | 5.0195 |  |  | 470 | 18802.219 | 9.8417 | 5.0202
471 | 18842.224 | 9.8439 | 5.0025 |  |  | 471 | 18842.224 | 9.8439 | 5.0193 |  |  | 471 | 18842.224 | 9.8439 | 5.02
472 | 18882.229 | 9.846 | 5.0023 |  |  | 472 | 18882.229 | 9.846 | 5.019 |  |  | 472 | 18882.229 | 9.846 | 5.0197
473 | 18922.233 | 9.8481 | 5.0021 |  |  | 473 | 18922.233 | 9.8481 | 5.0188 |  |  | 473 | 18922.233 | 9.8481 | 5.0195
474 | 18962.238 | 9.8502 | 5.0018 |  |  | 474 | 18962.238 | 9.8502 | 5.0186 |  |  | 474 | 18962.238 | 9.8502 | 5.0193
475 | 19002.243 | 9.8523 | 5.0016 |  |  | 475 | 19002.243 | 9.8523 | 5.0184 |  |  | 475 | 19002.243 | 9.8523 | 5.0191
476 | 19042.247 | 9.8544 | 5.0014 |  |  | 476 | 19042.247 | 9.8544 | 5.0182 |  |  | 476 | 19042.247 | 9.8544 | 5.0189
477 | 19082.252 | 9.8565 | 5.0012 |  |  | 477 | 19082.252 | 9.8565 | 5.018 |  |  | 477 | 19082.252 | 9.8565 | 5.0186
478 | 19122.257 | 9.8586 | 5.001 |  |  | 478 | 19122.257 | 9.8586 | 5.0177 |  |  | 478 | 19122.257 | 9.8586 | 5.0184
479 | 19162.262 | 9.8607 | 5.0008 |  |  | 479 | 19162.262 | 9.8607 | 5.0175 |  |  | 479 | 19162.262 | 9.8607 | 5.0182
480 | 19202.266 | 9.8628 | 5.0005 |  |  | 480 | 19202.266 | 9.8628 | 5.0173 |  |  | 480 | 19202.266 | 9.8628 | 5.018
481 | 19242.271 | 9.8649 | 5.0003 |  |  | 481 | 19242.271 | 9.8649 | 5.0171 |  |  | 481 | 19242.271 | 9.8649 | 5.0178
482 | 19282.276 | 9.8669 | 5.0001 |  |  | 482 | 19282.276 | 9.8669 | 5.0169 |  |  | 482 | 19282.276 | 9.8669 | 5.0176
483 | 19322.281 | 9.869 | 4.9999 |  |  | 483 | 19322.281 | 9.869 | 5.0167 |  |  | 483 | 19322.281 | 9.869 | 5.0174
484 | 19362.285 | 9.8711 | 4.9997 |  |  | 484 | 19362.285 | 9.8711 | 5.0165 |  |  | 484 | 19362.285 | 9.8711 | 5.0171
485 | 19402.29 | 9.8731 | 4.9995 |  |  | 485 | 19402.29 | 9.8731 | 5.0163 |  |  | 485 | 19402.29 | 9.8731 | 5.0169
486 | 19442.295 | 9.8752 | 4.9993 |  |  | 486 | 19442.295 | 9.8752 | 5.0161 |  |  | 486 | 19442.295 | 9.8752 | 5.0167
487 | 19482.299 | 9.8773 | 4.9991 |  |  | 487 | 19482.299 | 9.8773 | 5.0159 |  |  | 487 | 19482.299 | 9.8773 | 5.0165
488 | 19522.304 | 9.8793 | 4.9989 |  |  | 488 | 19522.304 | 9.8793 | 5.0157 |  |  | 488 | 19522.304 | 9.8793 | 5.0163
489 | 19562.309 | 9.8814 | 4.9987 |  |  | 489 | 19562.309 | 9.8814 | 5.0155 |  |  | 489 | 19562.309 | 9.8814 | 5.0161
490 | 19602.314 | 9.8834 | 4.9985 |  |  | 490 | 19602.314 | 9.8834 | 5.0153 |  |  | 490 | 19602.314 | 9.8834 | 5.0159
491 | 19642.318 | 9.8854 | 4.9983 |  |  | 491 | 19642.318 | 9.8854 | 5.0151 |  |  | 491 | 19642.318 | 9.8854 | 5.0157
492 | 19682.323 | 9.8875 | 4.9981 |  |  | 492 | 19682.323 | 9.8875 | 5.0149 |  |  | 492 | 19682.323 | 9.8875 | 5.0155
493 | 19722.328 | 9.8895 | 4.9979 |  |  | 493 | 19722.328 | 9.8895 | 5.0147 |  |  | 493 | 19722.328 | 9.8895 | 5.0153
494 | 19762.332 | 9.8915 | 4.9977 |  |  | 494 | 19762.332 | 9.8915 | 5.0145 |  |  | 494 | 19762.332 | 9.8915 | 5.0151
495 | 19802.337 | 9.8936 | 4.9975 |  |  | 495 | 19802.337 | 9.8936 | 5.0143 |  |  | 495 | 19802.337 | 9.8936 | 5.0149
496 | 19842.342 | 9.8956 | 4.9973 |  |  | 496 | 19842.342 | 9.8956 | 5.0141 |  |  | 496 | 19842.342 | 9.8956 | 5.0147
497 | 19882.347 | 9.8976 | 4.9971 |  |  | 497 | 19882.347 | 9.8976 | 5.0139 |  |  | 497 | 19882.347 | 9.8976 | 5.0145
498 | 19922.351 | 9.8996 | 4.9969 |  |  | 498 | 19922.351 | 9.8996 | 5.0137 |  |  | 498 | 19922.351 | 9.8996 | 5.0143
499 | 19962.356 | 9.9016 | 4.9967 |  |  | 499 | 19962.356 | 9.9016 | 5.0135 |  |  | 499 | 19962.356 | 9.9016 | 5.0141
500 | 20002.361 | 9.9036 | 4.9965 |  |  | 500 | 20002.361 | 9.9036 | 5.0133 |  |  | 500 | 20002.361 | 9.9036 | 5.0139
501 | 20042.366 | 9.9056 | 4.9963 |  |  | 501 | 20042.366 | 9.9056 | 5.0131 |  |  | 501 | 20042.366 | 9.9056 | 5.0137
502 | 20082.37 | 9.9076 | 4.9961 |  |  | 502 | 20082.37 | 9.9076 | 5.0129 |  |  | 502 | 20082.37 | 9.9076 | 5.0135
503 | 20122.375 | 9.9096 | 4.9959 |  |  | 503 | 20122.375 | 9.9096 | 5.0127 |  |  | 503 | 20122.375 | 9.9096 | 5.0133
504 | 20162.38 | 9.9116 | 4.9957 |  |  | 504 | 20162.38 | 9.9116 | 5.0126 |  |  | 504 | 20162.38 | 9.9116 | 5.0132
505 | 20202.384 | 9.9136 | 4.9955 |  |  | 505 | 20202.384 | 9.9136 | 5.0124 |  |  | 505 | 20202.384 | 9.9136 | 5.013
506 | 20242.389 | 9.9155 | 4.9954 |  |  | 506 | 20242.389 | 9.9155 | 5.0122 |  |  | 506 | 20242.389 | 9.9155 | 5.0128
507 | 20282.394 | 9.9175 | 4.9952 |  |  | 507 | 20282.394 | 9.9175 | 5.012 |  |  | 507 | 20282.394 | 9.9175 | 5.0126
508 | 20322.399 | 9.9195 | 4.995 |  |  | 508 | 20322.399 | 9.9195 | 5.0118 |  |  | 508 | 20322.399 | 9.9195 | 5.0124
509 | 20362.403 | 9.9214 | 4.9948 |  |  | 509 | 20362.403 | 9.9214 | 5.0116 |  |  | 509 | 20362.403 | 9.9214 | 5.0122
510 | 20402.408 | 9.9234 | 4.9946 |  |  | 510 | 20402.408 | 9.9234 | 5.0115 |  |  | 510 | 20402.408 | 9.9234 | 5.012
511 | 20442.413 | 9.9254 | 4.9944 |  |  | 511 | 20442.413 | 9.9254 | 5.0113 |  |  | 511 | 20442.413 | 9.9254 | 5.0118
512 | 20482.417 | 9.9273 | 4.9942 |  |  | 512 | 20482.417 | 9.9273 | 5.0111 |  |  | 512 | 20482.417 | 9.9273 | 5.0117
513 | 20522.422 | 9.9293 | 4.9941 |  |  | 513 | 20522.422 | 9.9293 | 5.0109 |  |  | 513 | 20522.422 | 9.9293 | 5.0115
514 | 20562.427 | 9.9312 | 4.9939 |  |  | 514 | 20562.427 | 9.9312 | 5.0107 |  |  | 514 | 20562.427 | 9.9312 | 5.0113
515 | 20602.432 | 9.9332 | 4.9937 |  |  | 515 | 20602.432 | 9.9332 | 5.0106 |  |  | 515 | 20602.432 | 9.9332 | 5.0111
516 | 20642.436 | 9.9351 | 4.9935 |  |  | 516 | 20642.436 | 9.9351 | 5.0104 |  |  | 516 | 20642.436 | 9.9351 | 5.0109
517 | 20682.441 | 9.937 | 4.9933 |  |  | 517 | 20682.441 | 9.937 | 5.0102 |  |  | 517 | 20682.441 | 9.937 | 5.0108
518 | 20722.446 | 9.939 | 4.9932 |  |  | 518 | 20722.446 | 9.939 | 5.01 |  |  | 518 | 20722.446 | 9.939 | 5.0106
519 | 20762.451 | 9.9409 | 4.993 |  |  | 519 | 20762.451 | 9.9409 | 5.0099 |  |  | 519 | 20762.451 | 9.9409 | 5.0104
520 | 20802.455 | 9.9428 | 4.9928 |  |  | 520 | 20802.455 | 9.9428 | 5.0097 |  |  | 520 | 20802.455 | 9.9428 | 5.0102
521 | 20842.46 | 9.9447 | 4.9926 |  |  | 521 | 20842.46 | 9.9447 | 5.0095 |  |  | 521 | 20842.46 | 9.9447 | 5.0101
522 | 20882.465 | 9.9467 | 4.9925 |  |  | 522 | 20882.465 | 9.9467 | 5.0093 |  |  | 522 | 20882.465 | 9.9467 | 5.0099
523 | 20922.469 | 9.9486 | 4.9923 |  |  | 523 | 20922.469 | 9.9486 | 5.0092 |  |  | 523 | 20922.469 | 9.9486 | 5.0097
524 | 20962.474 | 9.9505 | 4.9921 |  |  | 524 | 20962.474 | 9.9505 | 5.009 |  |  | 524 | 20962.474 | 9.9505 | 5.0095
525 | 21002.479 | 9.9524 | 4.9919 |  |  | 525 | 21002.479 | 9.9524 | 5.0088 |  |  | 525 | 21002.479 | 9.9524 | 5.0094
526 | 21042.484 | 9.9543 | 4.9918 |  |  | 526 | 21042.484 | 9.9543 | 5.0087 |  |  | 526 | 21042.484 | 9.9543 | 5.0092
527 | 21082.488 | 9.9562 | 4.9916 |  |  | 527 | 21082.488 | 9.9562 | 5.0085 |  |  | 527 | 21082.488 | 9.9562 | 5.009
528 | 21122.493 | 9.9581 | 4.9914 |  |  | 528 | 21122.493 | 9.9581 | 5.0083 |  |  | 528 | 21122.493 | 9.9581 | 5.0088
529 | 21162.498 | 9.96 | 4.9913 |  |  | 529 | 21162.498 | 9.96 | 5.0082 |  |  | 529 | 21162.498 | 9.96 | 5.0087
530 | 21202.502 | 9.9619 | 4.9911 |  |  | 530 | 21202.502 | 9.9619 | 5.008 |  |  | 530 | 21202.502 | 9.9619 | 5.0085
531 | 21242.507 | 9.9638 | 4.9909 |  |  | 531 | 21242.507 | 9.9638 | 5.0078 |  |  | 531 | 21242.507 | 9.9638 | 5.0083
532 | 21282.512 | 9.9656 | 4.9907 |  |  | 532 | 21282.512 | 9.9656 | 5.0077 |  |  | 532 | 21282.512 | 9.9656 | 5.0082
533 | 21322.517 | 9.9675 | 4.9906 |  |  | 533 | 21322.517 | 9.9675 | 5.0075 |  |  | 533 | 21322.517 | 9.9675 | 5.008
534 | 21362.521 | 9.9694 | 4.9904 |  |  | 534 | 21362.521 | 9.9694 | 5.0073 |  |  | 534 | 21362.521 | 9.9694 | 5.0078
535 | 21402.526 | 9.9713 | 4.9902 |  |  | 535 | 21402.526 | 9.9713 | 5.0072 |  |  | 535 | 21402.526 | 9.9713 | 5.0077
536 | 21442.531 | 9.9731 | 4.9901 |  |  | 536 | 21442.531 | 9.9731 | 5.007 |  |  | 536 | 21442.531 | 9.9731 | 5.0075
537 | 21482.536 | 9.975 | 4.9899 |  |  | 537 | 21482.536 | 9.975 | 5.0068 |  |  | 537 | 21482.536 | 9.975 | 5.0073
538 | 21522.54 | 9.9769 | 4.9898 |  |  | 538 | 21522.54 | 9.9769 | 5.0067 |  |  | 538 | 21522.54 | 9.9769 | 5.0072
539 | 21562.545 | 9.9787 | 4.9896 |  |  | 539 | 21562.545 | 9.9787 | 5.0065 |  |  | 539 | 21562.545 | 9.9787 | 5.007
540 | 21602.55 | 9.9806 | 4.9894 |  |  | 540 | 21602.55 | 9.9806 | 5.0064 |  |  | 540 | 21602.55 | 9.9806 | 5.0069
541 | 21642.554 | 9.9824 | 4.9893 |  |  | 541 | 21642.554 | 9.9824 | 5.0062 |  |  | 541 | 21642.554 | 9.9824 | 5.0067
542 | 21682.559 | 9.9843 | 4.9891 |  |  | 542 | 21682.559 | 9.9843 | 5.0061 |  |  | 542 | 21682.559 | 9.9843 | 5.0065
543 | 21722.564 | 9.9861 | 4.9889 |  |  | 543 | 21722.564 | 9.9861 | 5.0059 |  |  | 543 | 21722.564 | 9.9861 | 5.0064
544 | 21762.569 | 9.9879 | 4.9888 |  |  | 544 | 21762.569 | 9.9879 | 5.0057 |  |  | 544 | 21762.569 | 9.9879 | 5.0062
545 | 21802.573 | 9.9898 | 4.9886 |  |  | 545 | 21802.573 | 9.9898 | 5.0056 |  |  | 545 | 21802.573 | 9.9898 | 5.0061
546 | 21842.578 | 9.9916 | 4.9885 |  |  | 546 | 21842.578 | 9.9916 | 5.0054 |  |  | 546 | 21842.578 | 9.9916 | 5.0059
547 | 21882.583 | 9.9934 | 4.9883 |  |  | 547 | 21882.583 | 9.9934 | 5.0053 |  |  | 547 | 21882.583 | 9.9934 | 5.0057
548 | 21922.587 | 9.9953 | 4.9881 |  |  | 548 | 21922.587 | 9.9953 | 5.0051 |  |  | 548 | 21922.587 | 9.9953 | 5.0056
549 | 21962.592 | 9.9971 | 4.988 |  |  | 549 | 21962.592 | 9.9971 | 5.005 |  |  | 549 | 21962.592 | 9.9971 | 5.0054
550 | 22002.597 | 9.9989 | 4.9878 |  |  | 550 | 22002.597 | 9.9989 | 5.0048 |  |  | 550 | 22002.597 | 9.9989 | 5.0053
551 | 22042.602 | 10.0007 | 4.9877 |  |  | 551 | 22042.602 | 10.0007 | 5.0047 |  |  | 551 | 22042.602 | 10.0007 | 5.0051
552 | 22082.606 | 10.0025 | 4.9875 |  |  | 552 | 22082.606 | 10.0025 | 5.0045 |  |  | 552 | 22082.606 | 10.0025 | 5.005
553 | 22122.611 | 10.0044 | 4.9874 |  |  | 553 | 22122.611 | 10.0044 | 5.0044 |  |  | 553 | 22122.611 | 10.0044 | 5.0048
554 | 22162.616 | 10.0062 | 4.9872 |  |  | 554 | 22162.616 | 10.0062 | 5.0042 |  |  | 554 | 22162.616 | 10.0062 | 5.0047
555 | 22202.621 | 10.008 | 4.9871 |  |  | 555 | 22202.621 | 10.008 | 5.0041 |  |  | 555 | 22202.621 | 10.008 | 5.0045
556 | 22242.625 | 10.0098 | 4.9869 |  |  | 556 | 22242.625 | 10.0098 | 5.0039 |  |  | 556 | 22242.625 | 10.0098 | 5.0044
557 | 22282.63 | 10.0116 | 4.9867 |  |  | 557 | 22282.63 | 10.0116 | 5.0038 |  |  | 557 | 22282.63 | 10.0116 | 5.0042
558 | 22322.635 | 10.0134 | 4.9866 |  |  | 558 | 22322.635 | 10.0134 | 5.0036 |  |  | 558 | 22322.635 | 10.0134 | 5.004
559 | 22362.639 | 10.0151 | 4.9864 |  |  | 559 | 22362.639 | 10.0151 | 5.0035 |  |  | 559 | 22362.639 | 10.0151 | 5.0039
560 | 22402.644 | 10.0169 | 4.9863 |  |  | 560 | 22402.644 | 10.0169 | 5.0033 |  |  | 560 | 22402.644 | 10.0169 | 5.0037
561 | 22442.649 | 10.0187 | 4.9861 |  |  | 561 | 22442.649 | 10.0187 | 5.0032 |  |  | 561 | 22442.649 | 10.0187 | 5.0036
562 | 22482.654 | 10.0205 | 4.986 |  |  | 562 | 22482.654 | 10.0205 | 5.003 |  |  | 562 | 22482.654 | 10.0205 | 5.0034
563 | 22522.658 | 10.0223 | 4.9858 |  |  | 563 | 22522.658 | 10.0223 | 5.0029 |  |  | 563 | 22522.658 | 10.0223 | 5.0033
564 | 22562.663 | 10.0241 | 4.9857 |  |  | 564 | 22562.663 | 10.0241 | 5.0027 |  |  | 564 | 22562.663 | 10.0241 | 5.0032
565 | 22602.668 | 10.0258 | 4.9855 |  |  | 565 | 22602.668 | 10.0258 | 5.0026 |  |  | 565 | 22602.668 | 10.0258 | 5.003
566 | 22642.672 | 10.0276 | 4.9854 |  |  | 566 | 22642.672 | 10.0276 | 5.0025 |  |  | 566 | 22642.672 | 10.0276 | 5.0029
567 | 22682.677 | 10.0294 | 4.9852 |  |  | 567 | 22682.677 | 10.0294 | 5.0023 |  |  | 567 | 22682.677 | 10.0294 | 5.0027
568 | 22722.682 | 10.0311 | 4.9851 |  |  | 568 | 22722.682 | 10.0311 | 5.0022 |  |  | 568 | 22722.682 | 10.0311 | 5.0026
569 | 22762.687 | 10.0329 | 4.985 |  |  | 569 | 22762.687 | 10.0329 | 5.002 |  |  | 569 | 22762.687 | 10.0329 | 5.0024
570 | 22802.691 | 10.0346 | 4.9848 |  |  | 570 | 22802.691 | 10.0346 | 5.0019 |  |  | 570 | 22802.691 | 10.0346 | 5.0023
571 | 22842.696 | 10.0364 | 4.9847 |  |  | 571 | 22842.696 | 10.0364 | 5.0018 |  |  | 571 | 22842.696 | 10.0364 | 5.0021
572 | 22882.701 | 10.0381 | 4.9845 |  |  | 572 | 22882.701 | 10.0381 | 5.0016 |  |  | 572 | 22882.701 | 10.0381 | 5.002
573 | 22922.705 | 10.0399 | 4.9844 |  |  | 573 | 22922.705 | 10.0399 | 5.0015 |  |  | 573 | 22922.705 | 10.0399 | 5.0018
574 | 22962.71 | 10.0416 | 4.9842 |  |  | 574 | 22962.71 | 10.0416 | 5.0013 |  |  | 574 | 22962.71 | 10.0416 | 5.0017
575 | 23002.715 | 10.0434 | 4.9841 |  |  | 575 | 23002.715 | 10.0434 | 5.0012 |  |  | 575 | 23002.715 | 10.0434 | 5.0016
576 | 23042.72 | 10.0451 | 4.9839 |  |  | 576 | 23042.72 | 10.0451 | 5.0011 |  |  | 576 | 23042.72 | 10.0451 | 5.0014
577 | 23082.724 | 10.0468 | 4.9838 |  |  | 577 | 23082.724 | 10.0468 | 5.0009 |  |  | 577 | 23082.724 | 10.0468 | 5.0013
578 | 23122.729 | 10.0486 | 4.9837 |  |  | 578 | 23122.729 | 10.0486 | 5.0008 |  |  | 578 | 23122.729 | 10.0486 | 5.0011
579 | 23162.734 | 10.0503 | 4.9835 |  |  | 579 | 23162.734 | 10.0503 | 5.0006 |  |  | 579 | 23162.734 | 10.0503 | 5.001
580 | 23202.739 | 10.052 | 4.9834 |  |  | 580 | 23202.739 | 10.052 | 5.0005 |  |  | 580 | 23202.739 | 10.052 | 5.0009
581 | 23242.743 | 10.0537 | 4.9832 |  |  | 581 | 23242.743 | 10.0537 | 5.0004 |  |  | 581 | 23242.743 | 10.0537 | 5.0007
582 | 23282.748 | 10.0555 | 4.9831 |  |  | 582 | 23282.748 | 10.0555 | 5.0002 |  |  | 582 | 23282.748 | 10.0555 | 5.0006
583 | 23322.753 | 10.0572 | 4.9829 |  |  | 583 | 23322.753 | 10.0572 | 5.0001 |  |  | 583 | 23322.753 | 10.0572 | 5.0004
584 | 23362.757 | 10.0589 | 4.9828 |  |  | 584 | 23362.757 | 10.0589 | 5 |  |  | 584 | 23362.757 | 10.0589 | 5.0003
585 | 23402.762 | 10.0606 | 4.9827 |  |  | 585 | 23402.762 | 10.0606 | 4.9998 |  |  | 585 | 23402.762 | 10.0606 | 5.0002
586 | 23442.767 | 10.0623 | 4.9825 |  |  | 586 | 23442.767 | 10.0623 | 4.9997 |  |  | 586 | 23442.767 | 10.0623 | 5
587 | 23482.772 | 10.064 | 4.9824 |  |  | 587 | 23482.772 | 10.064 | 4.9996 |  |  | 587 | 23482.772 | 10.064 | 4.9999
588 | 23522.776 | 10.0657 | 4.9823 |  |  | 588 | 23522.776 | 10.0657 | 4.9994 |  |  | 588 | 23522.776 | 10.0657 | 4.9998
589 | 23562.781 | 10.0674 | 4.9821 |  |  | 589 | 23562.781 | 10.0674 | 4.9993 |  |  | 589 | 23562.781 | 10.0674 | 4.9996
590 | 23602.786 | 10.0691 | 4.982 |  |  | 590 | 23602.786 | 10.0691 | 4.9992 |  |  | 590 | 23602.786 | 10.0691 | 4.9995
591 | 23642.79 | 10.0708 | 4.9818 |  |  | 591 | 23642.79 | 10.0708 | 4.999 |  |  | 591 | 23642.79 | 10.0708 | 4.9994
592 | 23682.795 | 10.0725 | 4.9817 |  |  | 592 | 23682.795 | 10.0725 | 4.9989 |  |  | 592 | 23682.795 | 10.0725 | 4.9992
593 | 23722.8 | 10.0742 | 4.9816 |  |  | 593 | 23722.8 | 10.0742 | 4.9988 |  |  | 593 | 23722.8 | 10.0742 | 4.9991
594 | 23762.805 | 10.0759 | 4.9814 |  |  | 594 | 23762.805 | 10.0759 | 4.9986 |  |  | 594 | 23762.805 | 10.0759 | 4.999
595 | 23802.809 | 10.0776 | 4.9813 |  |  | 595 | 23802.809 | 10.0776 | 4.9985 |  |  | 595 | 23802.809 | 10.0776 | 4.9988
596 | 23842.814 | 10.0792 | 4.9812 |  |  | 596 | 23842.814 | 10.0792 | 4.9984 |  |  | 596 | 23842.814 | 10.0792 | 4.9987
597 | 23882.819 | 10.0809 | 4.981 |  |  | 597 | 23882.819 | 10.0809 | 4.9983 |  |  | 597 | 23882.819 | 10.0809 | 4.9986
598 | 23922.824 | 10.0826 | 4.9809 |  |  | 598 | 23922.824 | 10.0826 | 4.9981 |  |  | 598 | 23922.824 | 10.0826 | 4.9984
599 | 23962.828 | 10.0843 | 4.9808 |  |  | 599 | 23962.828 | 10.0843 | 4.998 |  |  | 599 | 23962.828 | 10.0843 | 4.9983
600 | 24002.833 | 10.0859 | 4.9806 |  |  | 600 | 24002.833 | 10.0859 | 4.9979 |  |  | 600 | 24002.833 | 10.0859 | 4.9982
601 | 24042.838 | 10.0876 | 4.9805 |  |  | 601 | 24042.838 | 10.0876 | 4.9977 |  |  | 601 | 24042.838 | 10.0876 | 4.998
602 | 24082.842 | 10.0893 | 4.9804 |  |  | 602 | 24082.842 | 10.0893 | 4.9976 |  |  | 602 | 24082.842 | 10.0893 | 4.9979
603 | 24122.847 | 10.0909 | 4.9802 |  |  | 603 | 24122.847 | 10.0909 | 4.9975 |  |  | 603 | 24122.847 | 10.0909 | 4.9978
604 | 24162.852 | 10.0926 | 4.9801 |  |  | 604 | 24162.852 | 10.0926 | 4.9974 |  |  | 604 | 24162.852 | 10.0926 | 4.9976
605 | 24202.857 | 10.0942 | 4.98 |  |  | 605 | 24202.857 | 10.0942 | 4.9972 |  |  | 605 | 24202.857 | 10.0942 | 4.9975
606 | 24242.861 | 10.0959 | 4.9798 |  |  | 606 | 24242.861 | 10.0959 | 4.9971 |  |  | 606 | 24242.861 | 10.0959 | 4.9974
607 | 24282.866 | 10.0975 | 4.9797 |  |  | 607 | 24282.866 | 10.0975 | 4.997 |  |  | 607 | 24282.866 | 10.0975 | 4.9973
608 | 24322.871 | 10.0992 | 4.9796 |  |  | 608 | 24322.871 | 10.0992 | 4.9969 |  |  | 608 | 24322.871 | 10.0992 | 4.9971
609 | 24362.875 | 10.1008 | 4.9794 |  |  | 609 | 24362.875 | 10.1008 | 4.9967 |  |  | 609 | 24362.875 | 10.1008 | 4.997
610 | 24402.88 | 10.1025 | 4.9793 |  |  | 610 | 24402.88 | 10.1025 | 4.9966 |  |  | 610 | 24402.88 | 10.1025 | 4.9969
611 | 24442.885 | 10.1041 | 4.9792 |  |  | 611 | 24442.885 | 10.1041 | 4.9965 |  |  | 611 | 24442.885 | 10.1041 | 4.9967
612 | 24482.89 | 10.1057 | 4.979 |  |  | 612 | 24482.89 | 10.1057 | 4.9964 |  |  | 612 | 24482.89 | 10.1057 | 4.9966
613 | 24522.894 | 10.1074 | 4.9789 |  |  | 613 | 24522.894 | 10.1074 | 4.9962 |  |  | 613 | 24522.894 | 10.1074 | 4.9965
614 | 24562.899 | 10.109 | 4.9788 |  |  | 614 | 24562.899 | 10.109 | 4.9961 |  |  | 614 | 24562.899 | 10.109 | 4.9964
615 | 24602.904 | 10.1106 | 4.9787 |  |  | 615 | 24602.904 | 10.1106 | 4.996 |  |  | 615 | 24602.904 | 10.1106 | 4.9962
616 | 24642.909 | 10.1122 | 4.9785 |  |  | 616 | 24642.909 | 10.1122 | 4.9959 |  |  | 616 | 24642.909 | 10.1122 | 4.9961
617 | 24682.913 | 10.1139 | 4.9784 |  |  | 617 | 24682.913 | 10.1139 | 4.9958 |  |  | 617 | 24682.913 | 10.1139 | 4.996
618 | 24722.918 | 10.1155 | 4.9783 |  |  | 618 | 24722.918 | 10.1155 | 4.9956 |  |  | 618 | 24722.918 | 10.1155 | 4.9959
619 | 24762.923 | 10.1171 | 4.9781 |  |  | 619 | 24762.923 | 10.1171 | 4.9955 |  |  | 619 | 24762.923 | 10.1171 | 4.9957
620 | 24802.927 | 10.1187 | 4.978 |  |  | 620 | 24802.927 | 10.1187 | 4.9954 |  |  | 620 | 24802.927 | 10.1187 | 4.9956
621 | 24842.932 | 10.1203 | 4.9779 |  |  | 621 | 24842.932 | 10.1203 | 4.9953 |  |  | 621 | 24842.932 | 10.1203 | 4.9955
622 | 24882.937 | 10.1219 | 4.9778 |  |  | 622 | 24882.937 | 10.1219 | 4.9952 |  |  | 622 | 24882.937 | 10.1219 | 4.9954
623 | 24922.942 | 10.1235 | 4.9776 |  |  | 623 | 24922.942 | 10.1235 | 4.995 |  |  | 623 | 24922.942 | 10.1235 | 4.9952
624 | 24962.946 | 10.1251 | 4.9775 |  |  | 624 | 24962.946 | 10.1251 | 4.9949 |  |  | 624 | 24962.946 | 10.1251 | 4.9951
625 | 25002.951 | 10.1267 | 4.9774 |  |  | 625 | 25002.951 | 10.1267 | 4.9948 |  |  | 625 | 25002.951 | 10.1267 | 4.995
626 | 25042.956 | 10.1283 | 4.9773 |  |  | 626 | 25042.956 | 10.1283 | 4.9947 |  |  | 626 | 25042.956 | 10.1283 | 4.9949
627 | 25082.96 | 10.1299 | 4.9771 |  |  | 627 | 25082.96 | 10.1299 | 4.9946 |  |  | 627 | 25082.96 | 10.1299 | 4.9947
628 | 25122.965 | 10.1315 | 4.977 |  |  | 628 | 25122.965 | 10.1315 | 4.9944 |  |  | 628 | 25122.965 | 10.1315 | 4.9946
629 | 25162.97 | 10.1331 | 4.9769 |  |  | 629 | 25162.97 | 10.1331 | 4.9943 |  |  | 629 | 25162.97 | 10.1331 | 4.9945
630 | 25202.975 | 10.1347 | 4.9768 |  |  | 630 | 25202.975 | 10.1347 | 4.9942 |  |  | 630 | 25202.975 | 10.1347 | 4.9944
631 | 25242.979 | 10.1363 | 4.9766 |  |  | 631 | 25242.979 | 10.1363 | 4.9941 |  |  | 631 | 25242.979 | 10.1363 | 4.9943
632 | 25282.984 | 10.1379 | 4.9765 |  |  | 632 | 25282.984 | 10.1379 | 4.994 |  |  | 632 | 25282.984 | 10.1379 | 4.9941
633 | 25322.989 | 10.1395 | 4.9764 |  |  | 633 | 25322.989 | 10.1395 | 4.9938 |  |  | 633 | 25322.989 | 10.1395 | 4.994
634 | 25362.994 | 10.141 | 4.9763 |  |  | 634 | 25362.994 | 10.141 | 4.9937 |  |  | 634 | 25362.994 | 10.141 | 4.9939
635 | 25402.998 | 10.1426 | 4.9761 |  |  | 635 | 25402.998 | 10.1426 | 4.9936 |  |  | 635 | 25402.998 | 10.1426 | 4.9938
636 | 25443.003 | 10.1442 | 4.976 |  |  | 636 | 25443.003 | 10.1442 | 4.9935 |  |  | 636 | 25443.003 | 10.1442 | 4.9937
637 | 25483.008 | 10.1458 | 4.9759 |  |  | 637 | 25483.008 | 10.1458 | 4.9934 |  |  | 637 | 25483.008 | 10.1458 | 4.9935
638 | 25523.012 | 10.1473 | 4.9758 |  |  | 638 | 25523.012 | 10.1473 | 4.9933 |  |  | 638 | 25523.012 | 10.1473 | 4.9934
639 | 25563.017 | 10.1489 | 4.9756 |  |  | 639 | 25563.017 | 10.1489 | 4.9932 |  |  | 639 | 25563.017 | 10.1489 | 4.9933
640 | 25603.022 | 10.1505 | 4.9755 |  |  | 640 | 25603.022 | 10.1505 | 4.993 |  |  | 640 | 25603.022 | 10.1505 | 4.9932
641 | 25643.027 | 10.152 | 4.9754 |  |  | 641 | 25643.027 | 10.152 | 4.9929 |  |  | 641 | 25643.027 | 10.152 | 4.9931
642 | 25683.031 | 10.1536 | 4.9753 |  |  | 642 | 25683.031 | 10.1536 | 4.9928 |  |  | 642 | 25683.031 | 10.1536 | 4.9929
643 | 25723.036 | 10.1551 | 4.9752 |  |  | 643 | 25723.036 | 10.1551 | 4.9927 |  |  | 643 | 25723.036 | 10.1551 | 4.9928
644 | 25763.041 | 10.1567 | 4.975 |  |  | 644 | 25763.041 | 10.1567 | 4.9926 |  |  | 644 | 25763.041 | 10.1567 | 4.9927
645 | 25803.045 | 10.1582 | 4.9749 |  |  | 645 | 25803.045 | 10.1582 | 4.9925 |  |  | 645 | 25803.045 | 10.1582 | 4.9926
646 | 25843.05 | 10.1598 | 4.9748 |  |  | 646 | 25843.05 | 10.1598 | 4.9924 |  |  | 646 | 25843.05 | 10.1598 | 4.9925
647 | 25883.055 | 10.1613 | 4.9747 |  |  | 647 | 25883.055 | 10.1613 | 4.9922 |  |  | 647 | 25883.055 | 10.1613 | 4.9924
648 | 25923.06 | 10.1629 | 4.9746 |  |  | 648 | 25923.06 | 10.1629 | 4.9921 |  |  | 648 | 25923.06 | 10.1629 | 4.9922
649 | 25963.064 | 10.1644 | 4.9744 |  |  | 649 | 25963.064 | 10.1644 | 4.992 |  |  | 649 | 25963.064 | 10.1644 | 4.9921
650 | 26003.069 | 10.166 | 4.9743 |  |  | 650 | 26003.069 | 10.166 | 4.9919 |  |  | 650 | 26003.069 | 10.166 | 4.992
651 | 26043.074 | 10.1675 | 4.9742 |  |  | 651 | 26043.074 | 10.1675 | 4.9918 |  |  | 651 | 26043.074 | 10.1675 | 4.9919
652 | 26083.079 | 10.169 | 4.9741 |  |  | 652 | 26083.079 | 10.169 | 4.9917 |  |  | 652 | 26083.079 | 10.169 | 4.9918
653 | 26123.083 | 10.1706 | 4.974 |  |  | 653 | 26123.083 | 10.1706 | 4.9916 |  |  | 653 | 26123.083 | 10.1706 | 4.9917
654 | 26163.088 | 10.1721 | 4.9738 |  |  | 654 | 26163.088 | 10.1721 | 4.9915 |  |  | 654 | 26163.088 | 10.1721 | 4.9915
655 | 26203.093 | 10.1736 | 4.9737 |  |  | 655 | 26203.093 | 10.1736 | 4.9913 |  |  | 655 | 26203.093 | 10.1736 | 4.9914
656 | 26243.097 | 10.1752 | 4.9736 |  |  | 656 | 26243.097 | 10.1752 | 4.9912 |  |  | 656 | 26243.097 | 10.1752 | 4.9913
657 | 26283.102 | 10.1767 | 4.9735 |  |  | 657 | 26283.102 | 10.1767 | 4.9911 |  |  | 657 | 26283.102 | 10.1767 | 4.9912
658 | 26323.107 | 10.1782 | 4.9734 |  |  | 658 | 26323.107 | 10.1782 | 4.991 |  |  | 658 | 26323.107 | 10.1782 | 4.9911
659 | 26363.112 | 10.1797 | 4.9733 |  |  | 659 | 26363.112 | 10.1797 | 4.9909 |  |  | 659 | 26363.112 | 10.1797 | 4.991
660 | 26403.116 | 10.1812 | 4.9731 |  |  | 660 | 26403.116 | 10.1812 | 4.9908 |  |  | 660 | 26403.116 | 10.1812 | 4.9909
661 | 26443.121 | 10.1828 | 4.973 |  |  | 661 | 26443.121 | 10.1828 | 4.9907 |  |  | 661 | 26443.121 | 10.1828 | 4.9907
662 | 26483.126 | 10.1843 | 4.9729 |  |  | 662 | 26483.126 | 10.1843 | 4.9906 |  |  | 662 | 26483.126 | 10.1843 | 4.9906
663 | 26523.13 | 10.1858 | 4.9728 |  |  | 663 | 26523.13 | 10.1858 | 4.9905 |  |  | 663 | 26523.13 | 10.1858 | 4.9905
664 | 26563.135 | 10.1873 | 4.9727 |  |  | 664 | 26563.135 | 10.1873 | 4.9903 |  |  | 664 | 26563.135 | 10.1873 | 4.9904
665 | 26603.14 | 10.1888 | 4.9725 |  |  | 665 | 26603.14 | 10.1888 | 4.9902 |  |  | 665 | 26603.14 | 10.1888 | 4.9903
666 | 26643.145 | 10.1903 | 4.9724 |  |  | 666 | 26643.145 | 10.1903 | 4.9901 |  |  | 666 | 26643.145 | 10.1903 | 4.9902
667 | 26683.149 | 10.1918 | 4.9723 |  |  | 667 | 26683.149 | 10.1918 | 4.99 |  |  | 667 | 26683.149 | 10.1918 | 4.9901
668 | 26723.154 | 10.1933 | 4.9722 |  |  | 668 | 26723.154 | 10.1933 | 4.9899 |  |  | 668 | 26723.154 | 10.1933 | 4.99
669 | 26763.159 | 10.1948 | 4.9721 |  |  | 669 | 26763.159 | 10.1948 | 4.9898 |  |  | 669 | 26763.159 | 10.1948 | 4.9898
670 | 26803.163 | 10.1963 | 4.972 |  |  | 670 | 26803.163 | 10.1963 | 4.9897 |  |  | 670 | 26803.163 | 10.1963 | 4.9897
671 | 26843.168 | 10.1978 | 4.9719 |  |  | 671 | 26843.168 | 10.1978 | 4.9896 |  |  | 671 | 26843.168 | 10.1978 | 4.9896
672 | 26883.173 | 10.1993 | 4.9717 |  |  | 672 | 26883.173 | 10.1993 | 4.9895 |  |  | 672 | 26883.173 | 10.1993 | 4.9895
673 | 26923.178 | 10.2007 | 4.9716 |  |  | 673 | 26923.178 | 10.2007 | 4.9894 |  |  | 673 | 26923.178 | 10.2007 | 4.9894
674 | 26963.182 | 10.2022 | 4.9715 |  |  | 674 | 26963.182 | 10.2022 | 4.9893 |  |  | 674 | 26963.182 | 10.2022 | 4.9893
675 | 27003.187 | 10.2037 | 4.9714 |  |  | 675 | 27003.187 | 10.2037 | 4.9891 |  |  | 675 | 27003.187 | 10.2037 | 4.9892
676 | 27043.192 | 10.2052 | 4.9713 |  |  | 676 | 27043.192 | 10.2052 | 4.989 |  |  | 676 | 27043.192 | 10.2052 | 4.9891
677 | 27083.197 | 10.2067 | 4.9712 |  |  | 677 | 27083.197 | 10.2067 | 4.9889 |  |  | 677 | 27083.197 | 10.2067 | 4.989
678 | 27123.201 | 10.2081 | 4.971 |  |  | 678 | 27123.201 | 10.2081 | 4.9888 |  |  | 678 | 27123.201 | 10.2081 | 4.9888
679 | 27163.206 | 10.2096 | 4.9709 |  |  | 679 | 27163.206 | 10.2096 | 4.9887 |  |  | 679 | 27163.206 | 10.2096 | 4.9887
680 | 27203.211 | 10.2111 | 4.9708 |  |  | 680 | 27203.211 | 10.2111 | 4.9886 |  |  | 680 | 27203.211 | 10.2111 | 4.9886
681 | 27243.215 | 10.2126 | 4.9707 |  |  | 681 | 27243.215 | 10.2126 | 4.9885 |  |  | 681 | 27243.215 | 10.2126 | 4.9885
682 | 27283.22 | 10.214 | 4.9706 |  |  | 682 | 27283.22 | 10.214 | 4.9884 |  |  | 682 | 27283.22 | 10.214 | 4.9884
683 | 27323.225 | 10.2155 | 4.9705 |  |  | 683 | 27323.225 | 10.2155 | 4.9883 |  |  | 683 | 27323.225 | 10.2155 | 4.9883
684 | 27363.23 | 10.217 | 4.9704 |  |  | 684 | 27363.23 | 10.217 | 4.9882 |  |  | 684 | 27363.23 | 10.217 | 4.9882
685 | 27403.234 | 10.2184 | 4.9702 |  |  | 685 | 27403.234 | 10.2184 | 4.9881 |  |  | 685 | 27403.234 | 10.2184 | 4.9881
686 | 27443.239 | 10.2199 | 4.9701 |  |  | 686 | 27443.239 | 10.2199 | 4.988 |  |  | 686 | 27443.239 | 10.2199 | 4.988
687 | 27483.244 | 10.2213 | 4.97 |  |  | 687 | 27483.244 | 10.2213 | 4.9879 |  |  | 687 | 27483.244 | 10.2213 | 4.9879
688 | 27523.248 | 10.2228 | 4.9699 |  |  | 688 | 27523.248 | 10.2228 | 4.9878 |  |  | 688 | 27523.248 | 10.2228 | 4.9877
689 | 27563.253 | 10.2242 | 4.9698 |  |  | 689 | 27563.253 | 10.2242 | 4.9877 |  |  | 689 | 27563.253 | 10.2242 | 4.9876
690 | 27603.258 | 10.2257 | 4.9697 |  |  | 690 | 27603.258 | 10.2257 | 4.9876 |  |  | 690 | 27603.258 | 10.2257 | 4.9875
691 | 27643.263 | 10.2271 | 4.9696 |  |  | 691 | 27643.263 | 10.2271 | 4.9874 |  |  | 691 | 27643.263 | 10.2271 | 4.9874
692 | 27683.267 | 10.2286 | 4.9695 |  |  | 692 | 27683.267 | 10.2286 | 4.9873 |  |  | 692 | 27683.267 | 10.2286 | 4.9873
693 | 27723.272 | 10.23 | 4.9693 |  |  | 693 | 27723.272 | 10.23 | 4.9872 |  |  | 693 | 27723.272 | 10.23 | 4.9872
694 | 27763.277 | 10.2315 | 4.9692 |  |  | 694 | 27763.277 | 10.2315 | 4.9871 |  |  | 694 | 27763.277 | 10.2315 | 4.9871
695 | 27803.282 | 10.2329 | 4.9691 |  |  | 695 | 27803.282 | 10.2329 | 4.987 |  |  | 695 | 27803.282 | 10.2329 | 4.987
696 | 27843.286 | 10.2343 | 4.969 |  |  | 696 | 27843.286 | 10.2343 | 4.9869 |  |  | 696 | 27843.286 | 10.2343 | 4.9869
697 | 27883.291 | 10.2358 | 4.9689 |  |  | 697 | 27883.291 | 10.2358 | 4.9868 |  |  | 697 | 27883.291 | 10.2358 | 4.9868
698 | 27923.296 | 10.2372 | 4.9688 |  |  | 698 | 27923.296 | 10.2372 | 4.9867 |  |  | 698 | 27923.296 | 10.2372 | 4.9867
699 | 27963.3 | 10.2386 | 4.9687 |  |  | 699 | 27963.3 | 10.2386 | 4.9866 |  |  | 699 | 27963.3 | 10.2386 | 4.9866
700 | 28003.305 | 10.2401 | 4.9686 |  |  | 700 | 28003.305 | 10.2401 | 4.9865 |  |  | 700 | 28003.305 | 10.2401 | 4.9864
701 | 28043.31 | 10.2415 | 4.9684 |  |  | 701 | 28043.31 | 10.2415 | 4.9864 |  |  | 701 | 28043.31 | 10.2415 | 4.9863
702 | 28083.315 | 10.2429 | 4.9683 |  |  | 702 | 28083.315 | 10.2429 | 4.9863 |  |  | 702 | 28083.315 | 10.2429 | 4.9862
703 | 28123.319 | 10.2444 | 4.9682 |  |  | 703 | 28123.319 | 10.2444 | 4.9862 |  |  | 703 | 28123.319 | 10.2444 | 4.9861
704 | 28163.324 | 10.2458 | 4.9681 |  |  | 704 | 28163.324 | 10.2458 | 4.9861 |  |  | 704 | 28163.324 | 10.2458 | 4.986
705 | 28203.329 | 10.2472 | 4.968 |  |  | 705 | 28203.329 | 10.2472 | 4.986 |  |  | 705 | 28203.329 | 10.2472 | 4.9859
706 | 28243.333 | 10.2486 | 4.9679 |  |  | 706 | 28243.333 | 10.2486 | 4.9859 |  |  | 706 | 28243.333 | 10.2486 | 4.9858
707 | 28283.338 | 10.25 | 4.9678 |  |  | 707 | 28283.338 | 10.25 | 4.9858 |  |  | 707 | 28283.338 | 10.25 | 4.9857
708 | 28323.343 | 10.2514 | 4.9677 |  |  | 708 | 28323.343 | 10.2514 | 4.9857 |  |  | 708 | 28323.343 | 10.2514 | 4.9856
709 | 28363.348 | 10.2529 | 4.9676 |  |  | 709 | 28363.348 | 10.2529 | 4.9856 |  |  | 709 | 28363.348 | 10.2529 | 4.9855
710 | 28403.352 | 10.2543 | 4.9674 |  |  | 710 | 28403.352 | 10.2543 | 4.9855 |  |  | 710 | 28403.352 | 10.2543 | 4.9854
711 | 28443.357 | 10.2557 | 4.9673 |  |  | 711 | 28443.357 | 10.2557 | 4.9854 |  |  | 711 | 28443.357 | 10.2557 | 4.9853
712 | 28483.362 | 10.2571 | 4.9672 |  |  | 712 | 28483.362 | 10.2571 | 4.9853 |  |  | 712 | 28483.362 | 10.2571 | 4.9852
713 | 28523.367 | 10.2585 | 4.9671 |  |  | 713 | 28523.367 | 10.2585 | 4.9852 |  |  | 713 | 28523.367 | 10.2585 | 4.9851
714 | 28563.371 | 10.2599 | 4.967 |  |  | 714 | 28563.371 | 10.2599 | 4.9851 |  |  | 714 | 28563.371 | 10.2599 | 4.985
715 | 28603.376 | 10.2613 | 4.9669 |  |  | 715 | 28603.376 | 10.2613 | 4.985 |  |  | 715 | 28603.376 | 10.2613 | 4.9848
716 | 28643.381 | 10.2627 | 4.9668 |  |  | 716 | 28643.381 | 10.2627 | 4.9849 |  |  | 716 | 28643.381 | 10.2627 | 4.9847
717 | 28683.385 | 10.2641 | 4.9667 |  |  | 717 | 28683.385 | 10.2641 | 4.9848 |  |  | 717 | 28683.385 | 10.2641 | 4.9846
718 | 28723.39 | 10.2655 | 4.9666 |  |  | 718 | 28723.39 | 10.2655 | 4.9847 |  |  | 718 | 28723.39 | 10.2655 | 4.9845
719 | 28763.395 | 10.2669 | 4.9665 |  |  | 719 | 28763.395 | 10.2669 | 4.9846 |  |  | 719 | 28763.395 | 10.2669 | 4.9844
720 | 28803.4 | 10.2682 | 4.9663 |  |  | 720 | 28803.4 | 10.2682 | 4.9845 |  |  | 720 | 28803.4 | 10.2682 | 4.9843
721 | 28843.404 | 10.2696 | 4.9662 |  |  | 721 | 28843.404 | 10.2696 | 4.9844 |  |  | 721 | 28843.404 | 10.2696 | 4.9842
722 | 28883.409 | 10.271 | 4.9661 |  |  | 722 | 28883.409 | 10.271 | 4.9843 |  |  | 722 | 28883.409 | 10.271 | 4.9841
723 | 28923.414 | 10.2724 | 4.966 |  |  | 723 | 28923.414 | 10.2724 | 4.9842 |  |  | 723 | 28923.414 | 10.2724 | 4.984
724 | 28963.418 | 10.2738 | 4.9659 |  |  | 724 | 28963.418 | 10.2738 | 4.9841 |  |  | 724 | 28963.418 | 10.2738 | 4.9839
725 | 29003.423 | 10.2752 | 4.9658 |  |  | 725 | 29003.423 | 10.2752 | 4.9839 |  |  | 725 | 29003.423 | 10.2752 | 4.9838
726 | 29043.428 | 10.2765 | 4.9657 |  |  | 726 | 29043.428 | 10.2765 | 4.9838 |  |  | 726 | 29043.428 | 10.2765 | 4.9837
727 | 29083.433 | 10.2779 | 4.9656 |  |  | 727 | 29083.433 | 10.2779 | 4.9837 |  |  | 727 | 29083.433 | 10.2779 | 4.9836
728 | 29123.437 | 10.2793 | 4.9655 |  |  | 728 | 29123.437 | 10.2793 | 4.9836 |  |  | 728 | 29123.437 | 10.2793 | 4.9835
729 | 29163.442 | 10.2807 | 4.9654 |  |  | 729 | 29163.442 | 10.2807 | 4.9835 |  |  | 729 | 29163.442 | 10.2807 | 4.9834
730 | 29203.447 | 10.282 | 4.9653 |  |  | 730 | 29203.447 | 10.282 | 4.9834 |  |  | 730 | 29203.447 | 10.282 | 4.9833
731 | 29243.452 | 10.2834 | 4.9651 |  |  | 731 | 29243.452 | 10.2834 | 4.9833 |  |  | 731 | 29243.452 | 10.2834 | 4.9832
732 | 29283.456 | 10.2848 | 4.965 |  |  | 732 | 29283.456 | 10.2848 | 4.9832 |  |  | 732 | 29283.456 | 10.2848 | 4.9831
733 | 29323.461 | 10.2861 | 4.9649 |  |  | 733 | 29323.461 | 10.2861 | 4.9831 |  |  | 733 | 29323.461 | 10.2861 | 4.983
734 | 29363.466 | 10.2875 | 4.9648 |  |  | 734 | 29363.466 | 10.2875 | 4.983 |  |  | 734 | 29363.466 | 10.2875 | 4.9829
735 | 29403.47 | 10.2889 | 4.9647 |  |  | 735 | 29403.47 | 10.2889 | 4.9829 |  |  | 735 | 29403.47 | 10.2889 | 4.9828
736 | 29443.475 | 10.2902 | 4.9646 |  |  | 736 | 29443.475 | 10.2902 | 4.9828 |  |  | 736 | 29443.475 | 10.2902 | 4.9827
737 | 29483.48 | 10.2916 | 4.9645 |  |  | 737 | 29483.48 | 10.2916 | 4.9827 |  |  | 737 | 29483.48 | 10.2916 | 4.9825
738 | 29523.485 | 10.2929 | 4.9644 |  |  | 738 | 29523.485 | 10.2929 | 4.9826 |  |  | 738 | 29523.485 | 10.2929 | 4.9824
739 | 29563.489 | 10.2943 | 4.9643 |  |  | 739 | 29563.489 | 10.2943 | 4.9825 |  |  | 739 | 29563.489 | 10.2943 | 4.9823
740 | 29603.494 | 10.2956 | 4.9642 |  |  | 740 | 29603.494 | 10.2956 | 4.9824 |  |  | 740 | 29603.494 | 10.2956 | 4.9822
741 | 29643.499 | 10.297 | 4.9641 |  |  | 741 | 29643.499 | 10.297 | 4.9823 |  |  | 741 | 29643.499 | 10.297 | 4.9821
742 | 29683.503 | 10.2983 | 4.964 |  |  | 742 | 29683.503 | 10.2983 | 4.9822 |  |  | 742 | 29683.503 | 10.2983 | 4.982
743 | 29723.508 | 10.2997 | 4.9638 |  |  | 743 | 29723.508 | 10.2997 | 4.9821 |  |  | 743 | 29723.508 | 10.2997 | 4.9819
744 | 29763.513 | 10.301 | 4.9637 |  |  | 744 | 29763.513 | 10.301 | 4.982 |  |  | 744 | 29763.513 | 10.301 | 4.9818
745 | 29803.518 | 10.3024 | 4.9636 |  |  | 745 | 29803.518 | 10.3024 | 4.9819 |  |  | 745 | 29803.518 | 10.3024 | 4.9817
746 | 29843.522 | 10.3037 | 4.9635 |  |  | 746 | 29843.522 | 10.3037 | 4.9818 |  |  | 746 | 29843.522 | 10.3037 | 4.9816
747 | 29883.527 | 10.3051 | 4.9634 |  |  | 747 | 29883.527 | 10.3051 | 4.9817 |  |  | 747 | 29883.527 | 10.3051 | 4.9815
748 | 29923.532 | 10.3064 | 4.9633 |  |  | 748 | 29923.532 | 10.3064 | 4.9816 |  |  | 748 | 29923.532 | 10.3064 | 4.9814
749 | 29963.536 | 10.3077 | 4.9632 |  |  | 749 | 29963.536 | 10.3077 | 4.9815 |  |  | 749 | 29963.536 | 10.3077 | 4.9813
750 | 30003.541 | 10.3091 | 4.9631 |  |  | 750 | 30003.541 | 10.3091 | 4.9814 |  |  | 750 | 30003.541 | 10.3091 | 4.9812
751 | 30043.546 | 10.3104 | 4.963 |  |  | 751 | 30043.546 | 10.3104 | 4.9813 |  |  | 751 | 30043.546 | 10.3104 | 4.9811
752 | 30083.551 | 10.3117 | 4.9629 |  |  | 752 | 30083.551 | 10.3117 | 4.9812 |  |  | 752 | 30083.551 | 10.3117 | 4.981
753 | 30123.555 | 10.3131 | 4.9628 |  |  | 753 | 30123.555 | 10.3131 | 4.9811 |  |  | 753 | 30123.555 | 10.3131 | 4.9809
754 | 30163.56 | 10.3144 | 4.9627 |  |  | 754 | 30163.56 | 10.3144 | 4.981 |  |  | 754 | 30163.56 | 10.3144 | 4.9808
755 | 30203.565 | 10.3157 | 4.9626 |  |  | 755 | 30203.565 | 10.3157 | 4.981 |  |  | 755 | 30203.565 | 10.3157 | 4.9807
756 | 30243.57 | 10.317 | 4.9625 |  |  | 756 | 30243.57 | 10.317 | 4.9809 |  |  | 756 | 30243.57 | 10.317 | 4.9806
757 | 30283.574 | 10.3184 | 4.9623 |  |  | 757 | 30283.574 | 10.3184 | 4.9808 |  |  | 757 | 30283.574 | 10.3184 | 4.9805
758 | 30323.579 | 10.3197 | 4.9622 |  |  | 758 | 30323.579 | 10.3197 | 4.9807 |  |  | 758 | 30323.579 | 10.3197 | 4.9804
759 | 30363.584 | 10.321 | 4.9621 |  |  | 759 | 30363.584 | 10.321 | 4.9806 |  |  | 759 | 30363.584 | 10.321 | 4.9803
760 | 30403.588 | 10.3223 | 4.962 |  |  | 760 | 30403.588 | 10.3223 | 4.9805 |  |  | 760 | 30403.588 | 10.3223 | 4.9802
761 | 30443.593 | 10.3236 | 4.9619 |  |  | 761 | 30443.593 | 10.3236 | 4.9804 |  |  | 761 | 30443.593 | 10.3236 | 4.9801
762 | 30483.598 | 10.3249 | 4.9618 |  |  | 762 | 30483.598 | 10.3249 | 4.9803 |  |  | 762 | 30483.598 | 10.3249 | 4.98
763 | 30523.603 | 10.3263 | 4.9617 |  |  | 763 | 30523.603 | 10.3263 | 4.9802 |  |  | 763 | 30523.603 | 10.3263 | 4.9799
764 | 30563.607 | 10.3276 | 4.9616 |  |  | 764 | 30563.607 | 10.3276 | 4.9801 |  |  | 764 | 30563.607 | 10.3276 | 4.9798
765 | 30603.612 | 10.3289 | 4.9615 |  |  | 765 | 30603.612 | 10.3289 | 4.98 |  |  | 765 | 30603.612 | 10.3289 | 4.9797
766 | 30643.617 | 10.3302 | 4.9614 |  |  | 766 | 30643.617 | 10.3302 | 4.9799 |  |  | 766 | 30643.617 | 10.3302 | 4.9796
767 | 30683.621 | 10.3315 | 4.9613 |  |  | 767 | 30683.621 | 10.3315 | 4.9798 |  |  | 767 | 30683.621 | 10.3315 | 4.9795
768 | 30723.626 | 10.3328 | 4.9612 |  |  | 768 | 30723.626 | 10.3328 | 4.9797 |  |  | 768 | 30723.626 | 10.3328 | 4.9794
769 | 30763.631 | 10.3341 | 4.9611 |  |  | 769 | 30763.631 | 10.3341 | 4.9796 |  |  | 769 | 30763.631 | 10.3341 | 4.9793
770 | 30803.636 | 10.3354 | 4.961 |  |  | 770 | 30803.636 | 10.3354 | 4.9795 |  |  | 770 | 30803.636 | 10.3354 | 4.9792
771 | 30843.64 | 10.3367 | 4.9608 |  |  | 771 | 30843.64 | 10.3367 | 4.9794 |  |  | 771 | 30843.64 | 10.3367 | 4.9791
772 | 30883.645 | 10.338 | 4.9607 |  |  | 772 | 30883.645 | 10.338 | 4.9793 |  |  | 772 | 30883.645 | 10.338 | 4.979
773 | 30923.65 | 10.3393 | 4.9606 |  |  | 773 | 30923.65 | 10.3393 | 4.9792 |  |  | 773 | 30923.65 | 10.3393 | 4.9789
774 | 30963.655 | 10.3406 | 4.9605 |  |  | 774 | 30963.655 | 10.3406 | 4.9791 |  |  | 774 | 30963.655 | 10.3406 | 4.9788
775 | 31003.659 | 10.3419 | 4.9604 |  |  | 775 | 31003.659 | 10.3419 | 4.979 |  |  | 775 | 31003.659 | 10.3419 | 4.9787
776 | 31043.664 | 10.3431 | 4.9603 |  |  | 776 | 31043.664 | 10.3431 | 4.9789 |  |  | 776 | 31043.664 | 10.3431 | 4.9786
777 | 31083.669 | 10.3444 | 4.9602 |  |  | 777 | 31083.669 | 10.3444 | 4.9788 |  |  | 777 | 31083.669 | 10.3444 | 4.9785
778 | 31123.673 | 10.3457 | 4.9601 |  |  | 778 | 31123.673 | 10.3457 | 4.9787 |  |  | 778 | 31123.673 | 10.3457 | 4.9783
779 | 31163.678 | 10.347 | 4.96 |  |  | 779 | 31163.678 | 10.347 | 4.9786 |  |  | 779 | 31163.678 | 10.347 | 4.9782
780 | 31203.683 | 10.3483 | 4.9599 |  |  | 780 | 31203.683 | 10.3483 | 4.9785 |  |  | 780 | 31203.683 | 10.3483 | 4.9781
781 | 31243.688 | 10.3496 | 4.9598 |  |  | 781 | 31243.688 | 10.3496 | 4.9784 |  |  | 781 | 31243.688 | 10.3496 | 4.978
782 | 31283.692 | 10.3509 | 4.9597 |  |  | 782 | 31283.692 | 10.3509 | 4.9783 |  |  | 782 | 31283.692 | 10.3509 | 4.9779
783 | 31323.697 | 10.3521 | 4.9596 |  |  | 783 | 31323.697 | 10.3521 | 4.9782 |  |  | 783 | 31323.697 | 10.3521 | 4.9778
784 | 31363.702 | 10.3534 | 4.9595 |  |  | 784 | 31363.702 | 10.3534 | 4.9781 |  |  | 784 | 31363.702 | 10.3534 | 4.9777
785 | 31403.706 | 10.3547 | 4.9594 |  |  | 785 | 31403.706 | 10.3547 | 4.978 |  |  | 785 | 31403.706 | 10.3547 | 4.9776
786 | 31443.711 | 10.356 | 4.9593 |  |  | 786 | 31443.711 | 10.356 | 4.9779 |  |  | 786 | 31443.711 | 10.356 | 4.9775
787 | 31483.716 | 10.3572 | 4.9591 |  |  | 787 | 31483.716 | 10.3572 | 4.9778 |  |  | 787 | 31483.716 | 10.3572 | 4.9774
788 | 31523.721 | 10.3585 | 4.959 |  |  | 788 | 31523.721 | 10.3585 | 4.9777 |  |  | 788 | 31523.721 | 10.3585 | 4.9773
789 | 31563.725 | 10.3598 | 4.9589 |  |  | 789 | 31563.725 | 10.3598 | 4.9776 |  |  | 789 | 31563.725 | 10.3598 | 4.9772
790 | 31603.73 | 10.361 | 4.9588 |  |  | 790 | 31603.73 | 10.361 | 4.9775 |  |  | 790 | 31603.73 | 10.361 | 4.9771
791 | 31643.735 | 10.3623 | 4.9587 |  |  | 791 | 31643.735 | 10.3623 | 4.9774 |  |  | 791 | 31643.735 | 10.3623 | 4.977
792 | 31683.74 | 10.3636 | 4.9586 |  |  | 792 | 31683.74 | 10.3636 | 4.9773 |  |  | 792 | 31683.74 | 10.3636 | 4.9769
793 | 31723.744 | 10.3648 | 4.9585 |  |  | 793 | 31723.744 | 10.3648 | 4.9772 |  |  | 793 | 31723.744 | 10.3648 | 4.9768
794 | 31763.749 | 10.3661 | 4.9584 |  |  | 794 | 31763.749 | 10.3661 | 4.9771 |  |  | 794 | 31763.749 | 10.3661 | 4.9767
795 | 31803.754 | 10.3673 | 4.9583 |  |  | 795 | 31803.754 | 10.3673 | 4.977 |  |  | 795 | 31803.754 | 10.3673 | 4.9766
796 | 31843.758 | 10.3686 | 4.9582 |  |  | 796 | 31843.758 | 10.3686 | 4.9769 |  |  | 796 | 31843.758 | 10.3686 | 4.9765
797 | 31883.763 | 10.3699 | 4.9581 |  |  | 797 | 31883.763 | 10.3699 | 4.9768 |  |  | 797 | 31883.763 | 10.3699 | 4.9764
798 | 31923.768 | 10.3711 | 4.958 |  |  | 798 | 31923.768 | 10.3711 | 4.9767 |  |  | 798 | 31923.768 | 10.3711 | 4.9763
799 | 31963.773 | 10.3724 | 4.9579 |  |  | 799 | 31963.773 | 10.3724 | 4.9766 |  |  | 799 | 31963.773 | 10.3724 | 4.9762
800 | 32003.777 | 10.3736 | 4.9578 |  |  | 800 | 32003.777 | 10.3736 | 4.9765 |  |  | 800 | 32003.777 | 10.3736 | 4.9761
801 | 32043.782 | 10.3749 | 4.9577 |  |  | 801 | 32043.782 | 10.3749 | 4.9764 |  |  | 801 | 32043.782 | 10.3749 | 4.976
802 | 32083.787 | 10.3761 | 4.9576 |  |  | 802 | 32083.787 | 10.3761 | 4.9763 |  |  | 802 | 32083.787 | 10.3761 | 4.9759
803 | 32123.791 | 10.3774 | 4.9575 |  |  | 803 | 32123.791 | 10.3774 | 4.9762 |  |  | 803 | 32123.791 | 10.3774 | 4.9758
804 | 32163.796 | 10.3786 | 4.9573 |  |  | 804 | 32163.796 | 10.3786 | 4.9761 |  |  | 804 | 32163.796 | 10.3786 | 4.9757
805 | 32203.801 | 10.3798 | 4.9572 |  |  | 805 | 32203.801 | 10.3798 | 4.9761 |  |  | 805 | 32203.801 | 10.3798 | 4.9756
806 | 32243.806 | 10.3811 | 4.9571 |  |  | 806 | 32243.806 | 10.3811 | 4.976 |  |  | 806 | 32243.806 | 10.3811 | 4.9755
807 | 32283.81 | 10.3823 | 4.957 |  |  | 807 | 32283.81 | 10.3823 | 4.9759 |  |  | 807 | 32283.81 | 10.3823 | 4.9754
808 | 32323.815 | 10.3836 | 4.9569 |  |  | 808 | 32323.815 | 10.3836 | 4.9758 |  |  | 808 | 32323.815 | 10.3836 | 4.9753
809 | 32363.82 | 10.3848 | 4.9568 |  |  | 809 | 32363.82 | 10.3848 | 4.9757 |  |  | 809 | 32363.82 | 10.3848 | 4.9752
810 | 32403.825 | 10.386 | 4.9567 |  |  | 810 | 32403.825 | 10.386 | 4.9756 |  |  | 810 | 32403.825 | 10.386 | 4.9751
811 | 32443.829 | 10.3873 | 4.9566 |  |  | 811 | 32443.829 | 10.3873 | 4.9755 |  |  | 811 | 32443.829 | 10.3873 | 4.975
812 | 32483.834 | 10.3885 | 4.9565 |  |  | 812 | 32483.834 | 10.3885 | 4.9754 |  |  | 812 | 32483.834 | 10.3885 | 4.9749
813 | 32523.839 | 10.3897 | 4.9564 |  |  | 813 | 32523.839 | 10.3897 | 4.9753 |  |  | 813 | 32523.839 | 10.3897 | 4.9748
814 | 32563.843 | 10.391 | 4.9563 |  |  | 814 | 32563.843 | 10.391 | 4.9752 |  |  | 814 | 32563.843 | 10.391 | 4.9747
815 | 32603.848 | 10.3922 | 4.9562 |  |  | 815 | 32603.848 | 10.3922 | 4.9751 |  |  | 815 | 32603.848 | 10.3922 | 4.9746
816 | 32643.853 | 10.3934 | 4.9561 |  |  | 816 | 32643.853 | 10.3934 | 4.975 |  |  | 816 | 32643.853 | 10.3934 | 4.9745
817 | 32683.858 | 10.3946 | 4.956 |  |  | 817 | 32683.858 | 10.3946 | 4.9749 |  |  | 817 | 32683.858 | 10.3946 | 4.9744
818 | 32723.862 | 10.3959 | 4.9559 |  |  | 818 | 32723.862 | 10.3959 | 4.9748 |  |  | 818 | 32723.862 | 10.3959 | 4.9743
819 | 32763.867 | 10.3971 | 4.9558 |  |  | 819 | 32763.867 | 10.3971 | 4.9747 |  |  | 819 | 32763.867 | 10.3971 | 4.9742
820 | 32803.872 | 10.3983 | 4.9557 |  |  | 820 | 32803.872 | 10.3983 | 4.9746 |  |  | 820 | 32803.872 | 10.3983 | 4.9741
821 | 32843.876 | 10.3995 | 4.9555 |  |  | 821 | 32843.876 | 10.3995 | 4.9745 |  |  | 821 | 32843.876 | 10.3995 | 4.974
822 | 32883.881 | 10.4007 | 4.9554 |  |  | 822 | 32883.881 | 10.4007 | 4.9744 |  |  | 822 | 32883.881 | 10.4007 | 4.9739
823 | 32923.886 | 10.402 | 4.9553 |  |  | 823 | 32923.886 | 10.402 | 4.9743 |  |  | 823 | 32923.886 | 10.402 | 4.9738
824 | 32963.891 | 10.4032 | 4.9552 |  |  | 824 | 32963.891 | 10.4032 | 4.9742 |  |  | 824 | 32963.891 | 10.4032 | 4.9737
825 | 33003.895 | 10.4044 | 4.9551 |  |  | 825 | 33003.895 | 10.4044 | 4.9741 |  |  | 825 | 33003.895 | 10.4044 | 4.9736
826 | 33043.9 | 10.4056 | 4.955 |  |  | 826 | 33043.9 | 10.4056 | 4.974 |  |  | 826 | 33043.9 | 10.4056 | 4.9735
827 | 33083.905 | 10.4068 | 4.9549 |  |  | 827 | 33083.905 | 10.4068 | 4.9739 |  |  | 827 | 33083.905 | 10.4068 | 4.9734
828 | 33123.91 | 10.408 | 4.9548 |  |  | 828 | 33123.91 | 10.408 | 4.9738 |  |  | 828 | 33123.91 | 10.408 | 4.9733
829 | 33163.914 | 10.4092 | 4.9547 |  |  | 829 | 33163.914 | 10.4092 | 4.9737 |  |  | 829 | 33163.914 | 10.4092 | 4.9732
830 | 33203.919 | 10.4104 | 4.9546 |  |  | 830 | 33203.919 | 10.4104 | 4.9736 |  |  | 830 | 33203.919 | 10.4104 | 4.9731
831 | 33243.924 | 10.4116 | 4.9545 |  |  | 831 | 33243.924 | 10.4116 | 4.9735 |  |  | 831 | 33243.924 | 10.4116 | 4.973
832 | 33283.928 | 10.4128 | 4.9544 |  |  | 832 | 33283.928 | 10.4128 | 4.9734 |  |  | 832 | 33283.928 | 10.4128 | 4.9729
833 | 33323.933 | 10.414 | 4.9543 |  |  | 833 | 33323.933 | 10.414 | 4.9733 |  |  | 833 | 33323.933 | 10.414 | 4.9728
834 | 33363.938 | 10.4152 | 4.9542 |  |  | 834 | 33363.938 | 10.4152 | 4.9732 |  |  | 834 | 33363.938 | 10.4152 | 4.9727
835 | 33403.943 | 10.4164 | 4.9541 |  |  | 835 | 33403.943 | 10.4164 | 4.9731 |  |  | 835 | 33403.943 | 10.4164 | 4.9726
836 | 33443.947 | 10.4176 | 4.954 |  |  | 836 | 33443.947 | 10.4176 | 4.973 |  |  | 836 | 33443.947 | 10.4176 | 4.9725
837 | 33483.952 | 10.4188 | 4.9539 |  |  | 837 | 33483.952 | 10.4188 | 4.9729 |  |  | 837 | 33483.952 | 10.4188 | 4.9724
838 | 33523.957 | 10.42 | 4.9537 |  |  | 838 | 33523.957 | 10.42 | 4.9728 |  |  | 838 | 33523.957 | 10.42 | 4.9723
839 | 33563.961 | 10.4212 | 4.9536 |  |  | 839 | 33563.961 | 10.4212 | 4.9727 |  |  | 839 | 33563.961 | 10.4212 | 4.9722
840 | 33603.966 | 10.4224 | 4.9535 |  |  | 840 | 33603.966 | 10.4224 | 4.9726 |  |  | 840 | 33603.966 | 10.4224 | 4.9721
841 | 33643.971 | 10.4236 | 4.9534 |  |  | 841 | 33643.971 | 10.4236 | 4.9725 |  |  | 841 | 33643.971 | 10.4236 | 4.972
842 | 33683.976 | 10.4248 | 4.9533 |  |  | 842 | 33683.976 | 10.4248 | 4.9725 |  |  | 842 | 33683.976 | 10.4248 | 4.9719
843 | 33723.98 | 10.426 | 4.9532 |  |  | 843 | 33723.98 | 10.426 | 4.9724 |  |  | 843 | 33723.98 | 10.426 | 4.9718
844 | 33763.985 | 10.4271 | 4.9531 |  |  | 844 | 33763.985 | 10.4271 | 4.9723 |  |  | 844 | 33763.985 | 10.4271 | 4.9717
845 | 33803.99 | 10.4283 | 4.953 |  |  | 845 | 33803.99 | 10.4283 | 4.9722 |  |  | 845 | 33803.99 | 10.4283 | 4.9716
846 | 33843.994 | 10.4295 | 4.9529 |  |  | 846 | 33843.994 | 10.4295 | 4.9721 |  |  | 846 | 33843.994 | 10.4295 | 4.9715
847 | 33883.999 | 10.4307 | 4.9528 |  |  | 847 | 33883.999 | 10.4307 | 4.972 |  |  | 847 | 33883.999 | 10.4307 | 4.9714
848 | 33924.004 | 10.4319 | 4.9527 |  |  | 848 | 33924.004 | 10.4319 | 4.9719 |  |  | 848 | 33924.004 | 10.4319 | 4.9713
849 | 33964.009 | 10.4331 | 4.9526 |  |  | 849 | 33964.009 | 10.4331 | 4.9718 |  |  | 849 | 33964.009 | 10.4331 | 4.9712
850 | 34004.013 | 10.4342 | 4.9525 |  |  | 850 | 34004.013 | 10.4342 | 4.9717 |  |  | 850 | 34004.013 | 10.4342 | 4.9711
851 | 34044.018 | 10.4354 | 4.9524 |  |  | 851 | 34044.018 | 10.4354 | 4.9716 |  |  | 851 | 34044.018 | 10.4354 | 4.971
852 | 34084.023 | 10.4366 | 4.9523 |  |  | 852 | 34084.023 | 10.4366 | 4.9715 |  |  | 852 | 34084.023 | 10.4366 | 4.9709
853 | 34124.028 | 10.4378 | 4.9522 |  |  | 853 | 34124.028 | 10.4378 | 4.9714 |  |  | 853 | 34124.028 | 10.4378 | 4.9708
854 | 34164.032 | 10.4389 | 4.9521 |  |  | 854 | 34164.032 | 10.4389 | 4.9713 |  |  | 854 | 34164.032 | 10.4389 | 4.9707
855 | 34204.037 | 10.4401 | 4.9519 |  |  | 855 | 34204.037 | 10.4401 | 4.9712 |  |  | 855 | 34204.037 | 10.4401 | 4.9706
856 | 34244.042 | 10.4413 | 4.9518 |  |  | 856 | 34244.042 | 10.4413 | 4.9711 |  |  | 856 | 34244.042 | 10.4413 | 4.9705
857 | 34284.046 | 10.4424 | 4.9517 |  |  | 857 | 34284.046 | 10.4424 | 4.971 |  |  | 857 | 34284.046 | 10.4424 | 4.9704
858 | 34324.051 | 10.4436 | 4.9516 |  |  | 858 | 34324.051 | 10.4436 | 4.9709 |  |  | 858 | 34324.051 | 10.4436 | 4.9703
859 | 34364.056 | 10.4448 | 4.9515 |  |  | 859 | 34364.056 | 10.4448 | 4.9708 |  |  | 859 | 34364.056 | 10.4448 | 4.9702
860 | 34404.061 | 10.4459 | 4.9514 |  |  | 860 | 34404.061 | 10.4459 | 4.9707 |  |  | 860 | 34404.061 | 10.4459 | 4.9701
861 | 34444.065 | 10.4471 | 4.9513 |  |  | 861 | 34444.065 | 10.4471 | 4.9706 |  |  | 861 | 34444.065 | 10.4471 | 4.97
862 | 34484.07 | 10.4483 | 4.9512 |  |  | 862 | 34484.07 | 10.4483 | 4.9705 |  |  | 862 | 34484.07 | 10.4483 | 4.9699
863 | 34524.075 | 10.4494 | 4.9511 |  |  | 863 | 34524.075 | 10.4494 | 4.9704 |  |  | 863 | 34524.075 | 10.4494 | 4.9698
864 | 34564.079 | 10.4506 | 4.951 |  |  | 864 | 34564.079 | 10.4506 | 4.9703 |  |  | 864 | 34564.079 | 10.4506 | 4.9697
865 | 34604.084 | 10.4517 | 4.9509 |  |  | 865 | 34604.084 | 10.4517 | 4.9702 |  |  | 865 | 34604.084 | 10.4517 | 4.9696
866 | 34644.089 | 10.4529 | 4.9508 |  |  | 866 | 34644.089 | 10.4529 | 4.9701 |  |  | 866 | 34644.089 | 10.4529 | 4.9695
867 | 34684.094 | 10.454 | 4.9507 |  |  | 867 | 34684.094 | 10.454 | 4.97 |  |  | 867 | 34684.094 | 10.454 | 4.9694
868 | 34724.098 | 10.4552 | 4.9506 |  |  | 868 | 34724.098 | 10.4552 | 4.9699 |  |  | 868 | 34724.098 | 10.4552 | 4.9693
869 | 34764.103 | 10.4563 | 4.9505 |  |  | 869 | 34764.103 | 10.4563 | 4.9698 |  |  | 869 | 34764.103 | 10.4563 | 4.9692
870 | 34804.108 | 10.4575 | 4.9504 |  |  | 870 | 34804.108 | 10.4575 | 4.9697 |  |  | 870 | 34804.108 | 10.4575 | 4.9691
871 | 34844.113 | 10.4586 | 4.9502 |  |  | 871 | 34844.113 | 10.4586 | 4.9696 |  |  | 871 | 34844.113 | 10.4586 | 4.969
872 | 34884.117 | 10.4598 | 4.9501 |  |  | 872 | 34884.117 | 10.4598 | 4.9695 |  |  | 872 | 34884.117 | 10.4598 | 4.9689
873 | 34924.122 | 10.4609 | 4.95 |  |  | 873 | 34924.122 | 10.4609 | 4.9694 |  |  | 873 | 34924.122 | 10.4609 | 4.9688
874 | 34964.127 | 10.4621 | 4.9499 |  |  | 874 | 34964.127 | 10.4621 | 4.9693 |  |  | 874 | 34964.127 | 10.4621 | 4.9686
875 | 35004.131 | 10.4632 | 4.9498 |  |  | 875 | 35004.131 | 10.4632 | 4.9692 |  |  | 875 | 35004.131 | 10.4632 | 4.9685
876 | 35044.136 | 10.4644 | 4.9497 |  |  | 876 | 35044.136 | 10.4644 | 4.9691 |  |  | 876 | 35044.136 | 10.4644 | 4.9684
877 | 35084.141 | 10.4655 | 4.9496 |  |  | 877 | 35084.141 | 10.4655 | 4.969 |  |  | 877 | 35084.141 | 10.4655 | 4.9683
878 | 35124.146 | 10.4666 | 4.9495 |  |  | 878 | 35124.146 | 10.4666 | 4.9689 |  |  | 878 | 35124.146 | 10.4666 | 4.9682
879 | 35164.15 | 10.4678 | 4.9494 |  |  | 879 | 35164.15 | 10.4678 | 4.9688 |  |  | 879 | 35164.15 | 10.4678 | 4.9681
880 | 35204.155 | 10.4689 | 4.9493 |  |  | 880 | 35204.155 | 10.4689 | 4.9687 |  |  | 880 | 35204.155 | 10.4689 | 4.968
881 | 35244.16 | 10.4701 | 4.9492 |  |  | 881 | 35244.16 | 10.4701 | 4.9686 |  |  | 881 | 35244.16 | 10.4701 | 4.9679
882 | 35284.164 | 10.4712 | 4.9491 |  |  | 882 | 35284.164 | 10.4712 | 4.9685 |  |  | 882 | 35284.164 | 10.4712 | 4.9678
883 | 35324.169 | 10.4723 | 4.949 |  |  | 883 | 35324.169 | 10.4723 | 4.9684 |  |  | 883 | 35324.169 | 10.4723 | 4.9677
884 | 35364.174 | 10.4735 | 4.9489 |  |  | 884 | 35364.174 | 10.4735 | 4.9684 |  |  | 884 | 35364.174 | 10.4735 | 4.9676
885 | 35404.179 | 10.4746 | 4.9488 |  |  | 885 | 35404.179 | 10.4746 | 4.9683 |  |  | 885 | 35404.179 | 10.4746 | 4.9675
886 | 35444.183 | 10.4757 | 4.9487 |  |  | 886 | 35444.183 | 10.4757 | 4.9682 |  |  | 886 | 35444.183 | 10.4757 | 4.9674
887 | 35484.188 | 10.4768 | 4.9485 |  |  | 887 | 35484.188 | 10.4768 | 4.9681 |  |  | 887 | 35484.188 | 10.4768 | 4.9673
888 | 35524.193 | 10.478 | 4.9484 |  |  | 888 | 35524.193 | 10.478 | 4.968 |  |  | 888 | 35524.193 | 10.478 | 4.9672
889 | 35564.198 | 10.4791 | 4.9483 |  |  | 889 | 35564.198 | 10.4791 | 4.9679 |  |  | 889 | 35564.198 | 10.4791 | 4.9671
890 | 35604.202 | 10.4802 | 4.9482 |  |  | 890 | 35604.202 | 10.4802 | 4.9678 |  |  | 890 | 35604.202 | 10.4802 | 4.967
891 | 35644.207 | 10.4813 | 4.9481 |  |  | 891 | 35644.207 | 10.4813 | 4.9677 |  |  | 891 | 35644.207 | 10.4813 | 4.9669
892 | 35684.212 | 10.4825 | 4.948 |  |  | 892 | 35684.212 | 10.4825 | 4.9676 |  |  | 892 | 35684.212 | 10.4825 | 4.9668
893 | 35724.216 | 10.4836 | 4.9479 |  |  | 893 | 35724.216 | 10.4836 | 4.9675 |  |  | 893 | 35724.216 | 10.4836 | 4.9667
894 | 35764.221 | 10.4847 | 4.9478 |  |  | 894 | 35764.221 | 10.4847 | 4.9674 |  |  | 894 | 35764.221 | 10.4847 | 4.9666
895 | 35804.226 | 10.4858 | 4.9477 |  |  | 895 | 35804.226 | 10.4858 | 4.9673 |  |  | 895 | 35804.226 | 10.4858 | 4.9665
896 | 35844.231 | 10.4869 | 4.9476 |  |  | 896 | 35844.231 | 10.4869 | 4.9672 |  |  | 896 | 35844.231 | 10.4869 | 4.9664
897 | 35884.235 | 10.4881 | 4.9475 |  |  | 897 | 35884.235 | 10.4881 | 4.9671 |  |  | 897 | 35884.235 | 10.4881 | 4.9663
898 | 35924.24 | 10.4892 | 4.9474 |  |  | 898 | 35924.24 | 10.4892 | 4.967 |  |  | 898 | 35924.24 | 10.4892 | 4.9662
899 | 35964.245 | 10.4903 | 4.9473 |  |  | 899 | 35964.245 | 10.4903 | 4.9669 |  |  | 899 | 35964.245 | 10.4903 | 4.9661
900 | 36004.249 | 10.4914 | 4.9472 |  |  | 900 | 36004.249 | 10.4914 | 4.9668 |  |  | 900 | 36004.249 | 10.4914 | 4.966
901 | 36044.254 | 10.4925 | 4.947 |  |  | 901 | 36044.254 | 10.4925 | 4.9667 |  |  | 901 | 36044.254 | 10.4925 | 4.9659
902 | 36084.259 | 10.4936 | 4.9469 |  |  | 902 | 36084.259 | 10.4936 | 4.9666 |  |  | 902 | 36084.259 | 10.4936 | 4.9658
903 | 36124.264 | 10.4947 | 4.9468 |  |  | 903 | 36124.264 | 10.4947 | 4.9665 |  |  | 903 | 36124.264 | 10.4947 | 4.9657
904 | 36164.268 | 10.4958 | 4.9467 |  |  | 904 | 36164.268 | 10.4958 | 4.9664 |  |  | 904 | 36164.268 | 10.4958 | 4.9656
905 | 36204.273 | 10.4969 | 4.9466 |  |  | 905 | 36204.273 | 10.4969 | 4.9663 |  |  | 905 | 36204.273 | 10.4969 | 4.9655
906 | 36244.278 | 10.498 | 4.9465 |  |  | 906 | 36244.278 | 10.498 | 4.9662 |  |  | 906 | 36244.278 | 10.498 | 4.9654
907 | 36284.283 | 10.4991 | 4.9464 |  |  | 907 | 36284.283 | 10.4991 | 4.9661 |  |  | 907 | 36284.283 | 10.4991 | 4.9653
908 | 36324.287 | 10.5002 | 4.9463 |  |  | 908 | 36324.287 | 10.5002 | 4.966 |  |  | 908 | 36324.287 | 10.5002 | 4.9652
909 | 36364.292 | 10.5013 | 4.9462 |  |  | 909 | 36364.292 | 10.5013 | 4.9659 |  |  | 909 | 36364.292 | 10.5013 | 4.9651
910 | 36404.297 | 10.5024 | 4.9461 |  |  | 910 | 36404.297 | 10.5024 | 4.9658 |  |  | 910 | 36404.297 | 10.5024 | 4.965
911 | 36444.301 | 10.5035 | 4.946 |  |  | 911 | 36444.301 | 10.5035 | 4.9657 |  |  | 911 | 36444.301 | 10.5035 | 4.9649
912 | 36484.306 | 10.5046 | 4.9459 |  |  | 912 | 36484.306 | 10.5046 | 4.9656 |  |  | 912 | 36484.306 | 10.5046 | 4.9648
913 | 36524.311 | 10.5057 | 4.9458 |  |  | 913 | 36524.311 | 10.5057 | 4.9655 |  |  | 913 | 36524.311 | 10.5057 | 4.9647
914 | 36564.316 | 10.5068 | 4.9457 |  |  | 914 | 36564.316 | 10.5068 | 4.9654 |  |  | 914 | 36564.316 | 10.5068 | 4.9646
915 | 36604.32 | 10.5079 | 4.9455 |  |  | 915 | 36604.32 | 10.5079 | 4.9653 |  |  | 915 | 36604.32 | 10.5079 | 4.9645
916 | 36644.325 | 10.509 | 4.9454 |  |  | 916 | 36644.325 | 10.509 | 4.9652 |  |  | 916 | 36644.325 | 10.509 | 4.9644
917 | 36684.33 | 10.5101 | 4.9453 |  |  | 917 | 36684.33 | 10.5101 | 4.9651 |  |  | 917 | 36684.33 | 10.5101 | 4.9643
918 | 36724.334 | 10.5112 | 4.9452 |  |  | 918 | 36724.334 | 10.5112 | 4.965 |  |  | 918 | 36724.334 | 10.5112 | 4.9642
919 | 36764.339 | 10.5123 | 4.9451 |  |  | 919 | 36764.339 | 10.5123 | 4.9649 |  |  | 919 | 36764.339 | 10.5123 | 4.9641
920 | 36804.344 | 10.5134 | 4.945 |  |  | 920 | 36804.344 | 10.5134 | 4.9648 |  |  | 920 | 36804.344 | 10.5134 | 4.964
921 | 36844.349 | 10.5145 | 4.9449 |  |  | 921 | 36844.349 | 10.5145 | 4.9647 |  |  | 921 | 36844.349 | 10.5145 | 4.9639
922 | 36884.353 | 10.5155 | 4.9448 |  |  | 922 | 36884.353 | 10.5155 | 4.9646 |  |  | 922 | 36884.353 | 10.5155 | 4.9638
923 | 36924.358 | 10.5166 | 4.9447 |  |  | 923 | 36924.358 | 10.5166 | 4.9645 |  |  | 923 | 36924.358 | 10.5166 | 4.9637
924 | 36964.363 | 10.5177 | 4.9446 |  |  | 924 | 36964.363 | 10.5177 | 4.9644 |  |  | 924 | 36964.363 | 10.5177 | 4.9635
925 | 37004.368 | 10.5188 | 4.9445 |  |  | 925 | 37004.368 | 10.5188 | 4.9643 |  |  | 925 | 37004.368 | 10.5188 | 4.9634
926 | 37044.372 | 10.5199 | 4.9444 |  |  | 926 | 37044.372 | 10.5199 | 4.9642 |  |  | 926 | 37044.372 | 10.5199 | 4.9633
927 | 37084.377 | 10.521 | 4.9443 |  |  | 927 | 37084.377 | 10.521 | 4.9641 |  |  | 927 | 37084.377 | 10.521 | 4.9632
928 | 37124.382 | 10.522 | 4.9441 |  |  | 928 | 37124.382 | 10.522 | 4.964 |  |  | 928 | 37124.382 | 10.522 | 4.9631
929 | 37164.386 | 10.5231 | 4.944 |  |  | 929 | 37164.386 | 10.5231 | 4.9639 |  |  | 929 | 37164.386 | 10.5231 | 4.963
930 | 37204.391 | 10.5242 | 4.9439 |  |  | 930 | 37204.391 | 10.5242 | 4.9638 |  |  | 930 | 37204.391 | 10.5242 | 4.9629
931 | 37244.396 | 10.5253 | 4.9438 |  |  | 931 | 37244.396 | 10.5253 | 4.9637 |  |  | 931 | 37244.396 | 10.5253 | 4.9628
932 | 37284.401 | 10.5263 | 4.9437 |  |  | 932 | 37284.401 | 10.5263 | 4.9636 |  |  | 932 | 37284.401 | 10.5263 | 4.9627
933 | 37324.405 | 10.5274 | 4.9436 |  |  | 933 | 37324.405 | 10.5274 | 4.9635 |  |  | 933 | 37324.405 | 10.5274 | 4.9626
934 | 37364.41 | 10.5285 | 4.9435 |  |  | 934 | 37364.41 | 10.5285 | 4.9634 |  |  | 934 | 37364.41 | 10.5285 | 4.9625
935 | 37404.415 | 10.5295 | 4.9434 |  |  | 935 | 37404.415 | 10.5295 | 4.9633 |  |  | 935 | 37404.415 | 10.5295 | 4.9624
936 | 37444.419 | 10.5306 | 4.9433 |  |  | 936 | 37444.419 | 10.5306 | 4.9632 |  |  | 936 | 37444.419 | 10.5306 | 4.9623
937 | 37484.424 | 10.5317 | 4.9432 |  |  | 937 | 37484.424 | 10.5317 | 4.9631 |  |  | 937 | 37484.424 | 10.5317 | 4.9622
938 | 37524.429 | 10.5327 | 4.9431 |  |  | 938 | 37524.429 | 10.5327 | 4.963 |  |  | 938 | 37524.429 | 10.5327 | 4.9621
939 | 37564.434 | 10.5338 | 4.943 |  |  | 939 | 37564.434 | 10.5338 | 4.9629 |  |  | 939 | 37564.434 | 10.5338 | 4.962
940 | 37604.438 | 10.5349 | 4.9428 |  |  | 940 | 37604.438 | 10.5349 | 4.9628 |  |  | 940 | 37604.438 | 10.5349 | 4.9619
941 | 37644.443 | 10.5359 | 4.9427 |  |  | 941 | 37644.443 | 10.5359 | 4.9627 |  |  | 941 | 37644.443 | 10.5359 | 4.9618
942 | 37684.448 | 10.537 | 4.9426 |  |  | 942 | 37684.448 | 10.537 | 4.9626 |  |  | 942 | 37684.448 | 10.537 | 4.9617
943 | 37724.452 | 10.5381 | 4.9425 |  |  | 943 | 37724.452 | 10.5381 | 4.9625 |  |  | 943 | 37724.452 | 10.5381 | 4.9616
944 | 37764.457 | 10.5391 | 4.9424 |  |  | 944 | 37764.457 | 10.5391 | 4.9624 |  |  | 944 | 37764.457 | 10.5391 | 4.9615
945 | 37804.462 | 10.5402 | 4.9423 |  |  | 945 | 37804.462 | 10.5402 | 4.9623 |  |  | 945 | 37804.462 | 10.5402 | 4.9614
946 | 37844.467 | 10.5412 | 4.9422 |  |  | 946 | 37844.467 | 10.5412 | 4.9622 |  |  | 946 | 37844.467 | 10.5412 | 4.9613
947 | 37884.471 | 10.5423 | 4.9421 |  |  | 947 | 37884.471 | 10.5423 | 4.9621 |  |  | 947 | 37884.471 | 10.5423 | 4.9612
948 | 37924.476 | 10.5434 | 4.942 |  |  | 948 | 37924.476 | 10.5434 | 4.962 |  |  | 948 | 37924.476 | 10.5434 | 4.9611
949 | 37964.481 | 10.5444 | 4.9419 |  |  | 949 | 37964.481 | 10.5444 | 4.9619 |  |  | 949 | 37964.481 | 10.5444 | 4.961
950 | 38004.486 | 10.5455 | 4.9418 |  |  | 950 | 38004.486 | 10.5455 | 4.9618 |  |  | 950 | 38004.486 | 10.5455 | 4.9609
951 | 38044.49 | 10.5465 | 4.9417 |  |  | 951 | 38044.49 | 10.5465 | 4.9617 |  |  | 951 | 38044.49 | 10.5465 | 4.9608
952 | 38084.495 | 10.5476 | 4.9415 |  |  | 952 | 38084.495 | 10.5476 | 4.9616 |  |  | 952 | 38084.495 | 10.5476 | 4.9607
953 | 38124.5 | 10.5486 | 4.9414 |  |  | 953 | 38124.5 | 10.5486 | 4.9615 |  |  | 953 | 38124.5 | 10.5486 | 4.9606
954 | 38164.504 | 10.5497 | 4.9413 |  |  | 954 | 38164.504 | 10.5497 | 4.9614 |  |  | 954 | 38164.504 | 10.5497 | 4.9604
955 | 38204.509 | 10.5507 | 4.9412 |  |  | 955 | 38204.509 | 10.5507 | 4.9613 |  |  | 955 | 38204.509 | 10.5507 | 4.9603
956 | 38244.514 | 10.5518 | 4.9411 |  |  | 956 | 38244.514 | 10.5518 | 4.9612 |  |  | 956 | 38244.514 | 10.5518 | 4.9602
957 | 38284.519 | 10.5528 | 4.941 |  |  | 957 | 38284.519 | 10.5528 | 4.9611 |  |  | 957 | 38284.519 | 10.5528 | 4.9601
958 | 38324.523 | 10.5538 | 4.9409 |  |  | 958 | 38324.523 | 10.5538 | 4.961 |  |  | 958 | 38324.523 | 10.5538 | 4.96
959 | 38364.528 | 10.5549 | 4.9408 |  |  | 959 | 38364.528 | 10.5549 | 4.9609 |  |  | 959 | 38364.528 | 10.5549 | 4.9599
960 | 38404.533 | 10.5559 | 4.9407 |  |  | 960 | 38404.533 | 10.5559 | 4.9608 |  |  | 960 | 38404.533 | 10.5559 | 4.9598
961 | 38444.537 | 10.557 | 4.9406 |  |  | 961 | 38444.537 | 10.557 | 4.9607 |  |  | 961 | 38444.537 | 10.557 | 4.9597
962 | 38484.542 | 10.558 | 4.9405 |  |  | 962 | 38484.542 | 10.558 | 4.9606 |  |  | 962 | 38484.542 | 10.558 | 4.9596
963 | 38524.547 | 10.5591 | 4.9404 |  |  | 963 | 38524.547 | 10.5591 | 4.9605 |  |  | 963 | 38524.547 | 10.5591 | 4.9595
964 | 38564.552 | 10.5601 | 4.9402 |  |  | 964 | 38564.552 | 10.5601 | 4.9604 |  |  | 964 | 38564.552 | 10.5601 | 4.9594
965 | 38604.556 | 10.5611 | 4.9401 |  |  | 965 | 38604.556 | 10.5611 | 4.9603 |  |  | 965 | 38604.556 | 10.5611 | 4.9593
966 | 38644.561 | 10.5622 | 4.94 |  |  | 966 | 38644.561 | 10.5622 | 4.9602 |  |  | 966 | 38644.561 | 10.5622 | 4.9592
967 | 38684.566 | 10.5632 | 4.9399 |  |  | 967 | 38684.566 | 10.5632 | 4.9601 |  |  | 967 | 38684.566 | 10.5632 | 4.9591
968 | 38724.571 | 10.5642 | 4.9398 |  |  | 968 | 38724.571 | 10.5642 | 4.96 |  |  | 968 | 38724.571 | 10.5642 | 4.959
969 | 38764.575 | 10.5653 | 4.9397 |  |  | 969 | 38764.575 | 10.5653 | 4.9599 |  |  | 969 | 38764.575 | 10.5653 | 4.9589
970 | 38804.58 | 10.5663 | 4.9396 |  |  | 970 | 38804.58 | 10.5663 | 4.9598 |  |  | 970 | 38804.58 | 10.5663 | 4.9588
971 | 38844.585 | 10.5673 | 4.9395 |  |  | 971 | 38844.585 | 10.5673 | 4.9597 |  |  | 971 | 38844.585 | 10.5673 | 4.9587
972 | 38884.589 | 10.5684 | 4.9394 |  |  | 972 | 38884.589 | 10.5684 | 4.9596 |  |  | 972 | 38884.589 | 10.5684 | 4.9586
973 | 38924.594 | 10.5694 | 4.9393 |  |  | 973 | 38924.594 | 10.5694 | 4.9595 |  |  | 973 | 38924.594 | 10.5694 | 4.9585
974 | 38964.599 | 10.5704 | 4.9391 |  |  | 974 | 38964.599 | 10.5704 | 4.9594 |  |  | 974 | 38964.599 | 10.5704 | 4.9584
975 | 39004.604 | 10.5714 | 4.939 |  |  | 975 | 39004.604 | 10.5714 | 4.9593 |  |  | 975 | 39004.604 | 10.5714 | 4.9583
976 | 39044.608 | 10.5725 | 4.9389 |  |  | 976 | 39044.608 | 10.5725 | 4.9592 |  |  | 976 | 39044.608 | 10.5725 | 4.9582
977 | 39084.613 | 10.5735 | 4.9388 |  |  | 977 | 39084.613 | 10.5735 | 4.9591 |  |  | 977 | 39084.613 | 10.5735 | 4.958
978 | 39124.618 | 10.5745 | 4.9387 |  |  | 978 | 39124.618 | 10.5745 | 4.959 |  |  | 978 | 39124.618 | 10.5745 | 4.9579
979 | 39164.622 | 10.5755 | 4.9386 |  |  | 979 | 39164.622 | 10.5755 | 4.9589 |  |  | 979 | 39164.622 | 10.5755 | 4.9578
980 | 39204.627 | 10.5766 | 4.9385 |  |  | 980 | 39204.627 | 10.5766 | 4.9588 |  |  | 980 | 39204.627 | 10.5766 | 4.9577
981 | 39244.632 | 10.5776 | 4.9384 |  |  | 981 | 39244.632 | 10.5776 | 4.9587 |  |  | 981 | 39244.632 | 10.5776 | 4.9576
982 | 39284.637 | 10.5786 | 4.9383 |  |  | 982 | 39284.637 | 10.5786 | 4.9586 |  |  | 982 | 39284.637 | 10.5786 | 4.9575
983 | 39324.641 | 10.5796 | 4.9382 |  |  | 983 | 39324.641 | 10.5796 | 4.9585 |  |  | 983 | 39324.641 | 10.5796 | 4.9574
984 | 39364.646 | 10.5806 | 4.9381 |  |  | 984 | 39364.646 | 10.5806 | 4.9584 |  |  | 984 | 39364.646 | 10.5806 | 4.9573
985 | 39404.651 | 10.5816 | 4.9379 |  |  | 985 | 39404.651 | 10.5816 | 4.9583 |  |  | 985 | 39404.651 | 10.5816 | 4.9572
986 | 39444.656 | 10.5827 | 4.9378 |  |  | 986 | 39444.656 | 10.5827 | 4.9582 |  |  | 986 | 39444.656 | 10.5827 | 4.9571
987 | 39484.66 | 10.5837 | 4.9377 |  |  | 987 | 39484.66 | 10.5837 | 4.9581 |  |  | 987 | 39484.66 | 10.5837 | 4.957
988 | 39524.665 | 10.5847 | 4.9376 |  |  | 988 | 39524.665 | 10.5847 | 4.958 |  |  | 988 | 39524.665 | 10.5847 | 4.9569
989 | 39564.67 | 10.5857 | 4.9375 |  |  | 989 | 39564.67 | 10.5857 | 4.9579 |  |  | 989 | 39564.67 | 10.5857 | 4.9568
990 | 39604.674 | 10.5867 | 4.9374 |  |  | 990 | 39604.674 | 10.5867 | 4.9578 |  |  | 990 | 39604.674 | 10.5867 | 4.9567

## Run3
MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR
File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run3\1 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run3\2 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run3\3 01 001.RAW
Cursor Location: |  |  | (178, 194) |  |  | Cursor Location: |  |  | (178, 194) |  |  | Cursor Location: |  |  | (178, 193)
Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15)
Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev
1 | 40.005 | 3.689 | 7.7554 | 5.1462 | 0.32479 | 1 | 40.005 | 3.689 | 7.7655 | 5.1619 | 0.32401 | 1 | 40.005 | 3.689 | 7.77 | 5.1593 | 0.32464
2 | 80.009 | 4.3821 | 7.4239 |  |  | 2 | 80.009 | 4.3821 | 7.4352 |  |  | 2 | 80.009 | 4.3821 | 7.4387
3 | 120.014 | 4.7876 | 7.2305 |  |  | 3 | 120.014 | 4.7876 | 7.2421 |  |  | 3 | 120.014 | 4.7876 | 7.2446
4 | 160.019 | 5.0753 | 7.0911 |  |  | 4 | 160.019 | 5.0753 | 7.1028 |  |  | 4 | 160.019 | 5.0753 | 7.1044
5 | 200.024 | 5.2984 | 6.9814 |  |  | 5 | 200.024 | 5.2984 | 6.993 |  |  | 5 | 200.024 | 5.2984 | 6.994
6 | 240.028 | 5.4808 | 6.8907 |  |  | 6 | 240.028 | 5.4808 | 6.9023 |  |  | 6 | 240.028 | 5.4808 | 6.9029
7 | 280.033 | 5.6349 | 6.8134 |  |  | 7 | 280.033 | 5.6349 | 6.825 |  |  | 7 | 280.033 | 5.6349 | 6.8253
8 | 320.038 | 5.7684 | 6.7461 |  |  | 8 | 320.038 | 5.7684 | 6.7576 |  |  | 8 | 320.038 | 5.7684 | 6.7576
9 | 360.042 | 5.8862 | 6.6864 |  |  | 9 | 360.042 | 5.8862 | 6.6979 |  |  | 9 | 360.042 | 5.8862 | 6.6977
10 | 400.047 | 5.9916 | 6.6328 |  |  | 10 | 400.047 | 5.9916 | 6.6442 |  |  | 10 | 400.047 | 5.9916 | 6.644
11 | 440.052 | 6.0869 | 6.5842 |  |  | 11 | 440.052 | 6.0869 | 6.5956 |  |  | 11 | 440.052 | 6.0869 | 6.5953
12 | 480.057 | 6.1739 | 6.5397 |  |  | 12 | 480.057 | 6.1739 | 6.5512 |  |  | 12 | 480.057 | 6.1739 | 6.5508
13 | 520.061 | 6.2539 | 6.4988 |  |  | 13 | 520.061 | 6.2539 | 6.5102 |  |  | 13 | 520.061 | 6.2539 | 6.5098
14 | 560.066 | 6.3281 | 6.4608 |  |  | 14 | 560.066 | 6.3281 | 6.4723 |  |  | 14 | 560.066 | 6.3281 | 6.4718
15 | 600.071 | 6.397 | 6.4255 |  |  | 15 | 600.071 | 6.397 | 6.4369 |  |  | 15 | 600.071 | 6.397 | 6.4365
16 | 640.076 | 6.4616 | 6.3924 |  |  | 16 | 640.076 | 6.4616 | 6.4038 |  |  | 16 | 640.076 | 6.4616 | 6.4034
17 | 680.08 | 6.5222 | 6.3613 |  |  | 17 | 680.08 | 6.5222 | 6.3728 |  |  | 17 | 680.08 | 6.5222 | 6.3723
18 | 720.085 | 6.5794 | 6.3319 |  |  | 18 | 720.085 | 6.5794 | 6.3435 |  |  | 18 | 720.085 | 6.5794 | 6.343
19 | 760.09 | 6.6334 | 6.3042 |  |  | 19 | 760.09 | 6.6334 | 6.3157 |  |  | 19 | 760.09 | 6.6334 | 6.3153
20 | 800.094 | 6.6847 | 6.2779 |  |  | 20 | 800.094 | 6.6847 | 6.2895 |  |  | 20 | 800.094 | 6.6847 | 6.289
21 | 840.099 | 6.7335 | 6.2529 |  |  | 21 | 840.099 | 6.7335 | 6.2645 |  |  | 21 | 840.099 | 6.7335 | 6.2641
22 | 880.104 | 6.78 | 6.229 |  |  | 22 | 880.104 | 6.78 | 6.2406 |  |  | 22 | 880.104 | 6.78 | 6.2403
23 | 920.109 | 6.8245 | 6.2062 |  |  | 23 | 920.109 | 6.8245 | 6.2179 |  |  | 23 | 920.109 | 6.8245 | 6.2175
24 | 960.113 | 6.8671 | 6.1844 |  |  | 24 | 960.113 | 6.8671 | 6.1961 |  |  | 24 | 960.113 | 6.8671 | 6.1958
25 | 1000.118 | 6.9079 | 6.1635 |  |  | 25 | 1000.118 | 6.9079 | 6.1752 |  |  | 25 | 1000.118 | 6.9079 | 6.1749
26 | 1040.123 | 6.9471 | 6.1434 |  |  | 26 | 1040.123 | 6.9471 | 6.1552 |  |  | 26 | 1040.123 | 6.9471 | 6.1549
27 | 1080.127 | 6.9848 | 6.1241 |  |  | 27 | 1080.127 | 6.9848 | 6.1359 |  |  | 27 | 1080.127 | 6.9848 | 6.1356
28 | 1120.132 | 7.0212 | 6.1054 |  |  | 28 | 1120.132 | 7.0212 | 6.1173 |  |  | 28 | 1120.132 | 7.0212 | 6.1171
29 | 1160.137 | 7.0563 | 6.0875 |  |  | 29 | 1160.137 | 7.0563 | 6.0994 |  |  | 29 | 1160.137 | 7.0563 | 6.0992
30 | 1200.142 | 7.0902 | 6.0702 |  |  | 30 | 1200.142 | 7.0902 | 6.0821 |  |  | 30 | 1200.142 | 7.0902 | 6.0819
31 | 1240.146 | 7.123 | 6.0534 |  |  | 31 | 1240.146 | 7.123 | 6.0654 |  |  | 31 | 1240.146 | 7.123 | 6.0652
32 | 1280.151 | 7.1547 | 6.0372 |  |  | 32 | 1280.151 | 7.1547 | 6.0493 |  |  | 32 | 1280.151 | 7.1547 | 6.049
33 | 1320.156 | 7.1855 | 6.0215 |  |  | 33 | 1320.156 | 7.1855 | 6.0336 |  |  | 33 | 1320.156 | 7.1855 | 6.0334
34 | 1360.161 | 7.2154 | 6.0063 |  |  | 34 | 1360.161 | 7.2154 | 6.0184 |  |  | 34 | 1360.161 | 7.2154 | 6.0182
35 | 1400.165 | 7.2443 | 5.9915 |  |  | 35 | 1400.165 | 7.2443 | 6.0037 |  |  | 35 | 1400.165 | 7.2443 | 6.0035
36 | 1440.17 | 7.2725 | 5.9772 |  |  | 36 | 1440.17 | 7.2725 | 5.9894 |  |  | 36 | 1440.17 | 7.2725 | 5.9892
37 | 1480.175 | 7.2999 | 5.9632 |  |  | 37 | 1480.175 | 7.2999 | 5.9755 |  |  | 37 | 1480.175 | 7.2999 | 5.9753
38 | 1520.179 | 7.3266 | 5.9497 |  |  | 38 | 1520.179 | 7.3266 | 5.962 |  |  | 38 | 1520.179 | 7.3266 | 5.9618
39 | 1560.184 | 7.3526 | 5.9365 |  |  | 39 | 1560.184 | 7.3526 | 5.9489 |  |  | 39 | 1560.184 | 7.3526 | 5.9487
40 | 1600.189 | 7.3779 | 5.9236 |  |  | 40 | 1600.189 | 7.3779 | 5.936 |  |  | 40 | 1600.189 | 7.3779 | 5.9359
41 | 1640.194 | 7.4026 | 5.9111 |  |  | 41 | 1640.194 | 7.4026 | 5.9236 |  |  | 41 | 1640.194 | 7.4026 | 5.9234
42 | 1680.198 | 7.4267 | 5.8989 |  |  | 42 | 1680.198 | 7.4267 | 5.9114 |  |  | 42 | 1680.198 | 7.4267 | 5.9112
43 | 1720.203 | 7.4502 | 5.887 |  |  | 43 | 1720.203 | 7.4502 | 5.8996 |  |  | 43 | 1720.203 | 7.4502 | 5.8994
44 | 1760.208 | 7.4732 | 5.8754 |  |  | 44 | 1760.208 | 7.4732 | 5.888 |  |  | 44 | 1760.208 | 7.4732 | 5.8878
45 | 1800.212 | 7.4957 | 5.864 |  |  | 45 | 1800.212 | 7.4957 | 5.8767 |  |  | 45 | 1800.212 | 7.4957 | 5.8765
46 | 1840.217 | 7.5176 | 5.8529 |  |  | 46 | 1840.217 | 7.5176 | 5.8656 |  |  | 46 | 1840.217 | 7.5176 | 5.8654
47 | 1880.222 | 7.5391 | 5.8421 |  |  | 47 | 1880.222 | 7.5391 | 5.8548 |  |  | 47 | 1880.222 | 7.5391 | 5.8546
48 | 1920.227 | 7.5602 | 5.8315 |  |  | 48 | 1920.227 | 7.5602 | 5.8443 |  |  | 48 | 1920.227 | 7.5602 | 5.844
49 | 1960.231 | 7.5808 | 5.8211 |  |  | 49 | 1960.231 | 7.5808 | 5.834 |  |  | 49 | 1960.231 | 7.5808 | 5.8337
50 | 2000.236 | 7.601 | 5.811 |  |  | 50 | 2000.236 | 7.601 | 5.8239 |  |  | 50 | 2000.236 | 7.601 | 5.8236
51 | 2040.241 | 7.6208 | 5.8011 |  |  | 51 | 2040.241 | 7.6208 | 5.814 |  |  | 51 | 2040.241 | 7.6208 | 5.8137
52 | 2080.246 | 7.6402 | 5.7913 |  |  | 52 | 2080.246 | 7.6402 | 5.8043 |  |  | 52 | 2080.246 | 7.6402 | 5.804
53 | 2120.25 | 7.6593 | 5.7818 |  |  | 53 | 2120.25 | 7.6593 | 5.7948 |  |  | 53 | 2120.25 | 7.6593 | 5.7945
54 | 2160.255 | 7.678 | 5.7724 |  |  | 54 | 2160.255 | 7.678 | 5.7855 |  |  | 54 | 2160.255 | 7.678 | 5.7852
55 | 2200.26 | 7.6963 | 5.7633 |  |  | 55 | 2200.26 | 7.6963 | 5.7764 |  |  | 55 | 2200.26 | 7.6963 | 5.776
56 | 2240.264 | 7.7143 | 5.7543 |  |  | 56 | 2240.264 | 7.7143 | 5.7674 |  |  | 56 | 2240.264 | 7.7143 | 5.7671
57 | 2280.269 | 7.732 | 5.7455 |  |  | 57 | 2280.269 | 7.732 | 5.7587 |  |  | 57 | 2280.269 | 7.732 | 5.7583
58 | 2320.274 | 7.7494 | 5.7369 |  |  | 58 | 2320.274 | 7.7494 | 5.7501 |  |  | 58 | 2320.274 | 7.7494 | 5.7497
59 | 2360.279 | 7.7665 | 5.7284 |  |  | 59 | 2360.279 | 7.7665 | 5.7416 |  |  | 59 | 2360.279 | 7.7665 | 5.7412
60 | 2400.283 | 7.7833 | 5.72 |  |  | 60 | 2400.283 | 7.7833 | 5.7334 |  |  | 60 | 2400.283 | 7.7833 | 5.7329
61 | 2440.288 | 7.7999 | 5.7119 |  |  | 61 | 2440.288 | 7.7999 | 5.7252 |  |  | 61 | 2440.288 | 7.7999 | 5.7248
62 | 2480.293 | 7.8161 | 5.7038 |  |  | 62 | 2480.293 | 7.8161 | 5.7172 |  |  | 62 | 2480.293 | 7.8161 | 5.7167
63 | 2520.297 | 7.8321 | 5.6959 |  |  | 63 | 2520.297 | 7.8321 | 5.7094 |  |  | 63 | 2520.297 | 7.8321 | 5.7089
64 | 2560.302 | 7.8479 | 5.6882 |  |  | 64 | 2560.302 | 7.8479 | 5.7017 |  |  | 64 | 2560.302 | 7.8479 | 5.7011
65 | 2600.307 | 7.8634 | 5.6806 |  |  | 65 | 2600.307 | 7.8634 | 5.6941 |  |  | 65 | 2600.307 | 7.8634 | 5.6935
66 | 2640.312 | 7.8787 | 5.6731 |  |  | 66 | 2640.312 | 7.8787 | 5.6866 |  |  | 66 | 2640.312 | 7.8787 | 5.6861
67 | 2680.316 | 7.8937 | 5.6657 |  |  | 67 | 2680.316 | 7.8937 | 5.6793 |  |  | 67 | 2680.316 | 7.8937 | 5.6787
68 | 2720.321 | 7.9085 | 5.6585 |  |  | 68 | 2720.321 | 7.9085 | 5.6721 |  |  | 68 | 2720.321 | 7.9085 | 5.6715
69 | 2760.326 | 7.9231 | 5.6513 |  |  | 69 | 2760.326 | 7.9231 | 5.665 |  |  | 69 | 2760.326 | 7.9231 | 5.6644
70 | 2800.331 | 7.9375 | 5.6443 |  |  | 70 | 2800.331 | 7.9375 | 5.658 |  |  | 70 | 2800.331 | 7.9375 | 5.6574
71 | 2840.335 | 7.9517 | 5.6374 |  |  | 71 | 2840.335 | 7.9517 | 5.6512 |  |  | 71 | 2840.335 | 7.9517 | 5.6505
72 | 2880.34 | 7.9657 | 5.6306 |  |  | 72 | 2880.34 | 7.9657 | 5.6444 |  |  | 72 | 2880.34 | 7.9657 | 5.6437
73 | 2920.345 | 7.9795 | 5.6239 |  |  | 73 | 2920.345 | 7.9795 | 5.6378 |  |  | 73 | 2920.345 | 7.9795 | 5.637
74 | 2960.349 | 7.9931 | 5.6174 |  |  | 74 | 2960.349 | 7.9931 | 5.6312 |  |  | 74 | 2960.349 | 7.9931 | 5.6305
75 | 3000.354 | 8.0065 | 5.6109 |  |  | 75 | 3000.354 | 8.0065 | 5.6248 |  |  | 75 | 3000.354 | 8.0065 | 5.624
76 | 3040.359 | 8.0197 | 5.6045 |  |  | 76 | 3040.359 | 8.0197 | 5.6185 |  |  | 76 | 3040.359 | 8.0197 | 5.6176
77 | 3080.364 | 8.0328 | 5.5982 |  |  | 77 | 3080.364 | 8.0328 | 5.6122 |  |  | 77 | 3080.364 | 8.0328 | 5.6113
78 | 3120.368 | 8.0457 | 5.592 |  |  | 78 | 3120.368 | 8.0457 | 5.6061 |  |  | 78 | 3120.368 | 8.0457 | 5.6052
79 | 3160.373 | 8.0584 | 5.5859 |  |  | 79 | 3160.373 | 8.0584 | 5.6 |  |  | 79 | 3160.373 | 8.0584 | 5.5991
80 | 3200.378 | 8.071 | 5.5799 |  |  | 80 | 3200.378 | 8.071 | 5.594 |  |  | 80 | 3200.378 | 8.071 | 5.5931
81 | 3240.382 | 8.0834 | 5.574 |  |  | 81 | 3240.382 | 8.0834 | 5.5881 |  |  | 81 | 3240.382 | 8.0834 | 5.5872
82 | 3280.387 | 8.0957 | 5.5681 |  |  | 82 | 3280.387 | 8.0957 | 5.5823 |  |  | 82 | 3280.387 | 8.0957 | 5.5813
83 | 3320.392 | 8.1078 | 5.5624 |  |  | 83 | 3320.392 | 8.1078 | 5.5766 |  |  | 83 | 3320.392 | 8.1078 | 5.5756
84 | 3360.397 | 8.1198 | 5.5567 |  |  | 84 | 3360.397 | 8.1198 | 5.571 |  |  | 84 | 3360.397 | 8.1198 | 5.5699
85 | 3400.401 | 8.1316 | 5.5511 |  |  | 85 | 3400.401 | 8.1316 | 5.5654 |  |  | 85 | 3400.401 | 8.1316 | 5.5643
86 | 3440.406 | 8.1433 | 5.5456 |  |  | 86 | 3440.406 | 8.1433 | 5.5599 |  |  | 86 | 3440.406 | 8.1433 | 5.5588
87 | 3480.411 | 8.1549 | 5.5402 |  |  | 87 | 3480.411 | 8.1549 | 5.5545 |  |  | 87 | 3480.411 | 8.1549 | 5.5534
88 | 3520.416 | 8.1663 | 5.5348 |  |  | 88 | 3520.416 | 8.1663 | 5.5492 |  |  | 88 | 3520.416 | 8.1663 | 5.548
89 | 3560.42 | 8.1776 | 5.5295 |  |  | 89 | 3560.42 | 8.1776 | 5.5439 |  |  | 89 | 3560.42 | 8.1776 | 5.5427
90 | 3600.425 | 8.1888 | 5.5243 |  |  | 90 | 3600.425 | 8.1888 | 5.5388 |  |  | 90 | 3600.425 | 8.1888 | 5.5375
91 | 3640.43 | 8.1999 | 5.5191 |  |  | 91 | 3640.43 | 8.1999 | 5.5336 |  |  | 91 | 3640.43 | 8.1999 | 5.5324
92 | 3680.434 | 8.2108 | 5.5141 |  |  | 92 | 3680.434 | 8.2108 | 5.5286 |  |  | 92 | 3680.434 | 8.2108 | 5.5273
93 | 3720.439 | 8.2216 | 5.509 |  |  | 93 | 3720.439 | 8.2216 | 5.5236 |  |  | 93 | 3720.439 | 8.2216 | 5.5223
94 | 3760.444 | 8.2323 | 5.5041 |  |  | 94 | 3760.444 | 8.2323 | 5.5187 |  |  | 94 | 3760.444 | 8.2323 | 5.5173
95 | 3800.449 | 8.2429 | 5.4992 |  |  | 95 | 3800.449 | 8.2429 | 5.5138 |  |  | 95 | 3800.449 | 8.2429 | 5.5125
96 | 3840.453 | 8.2533 | 5.4944 |  |  | 96 | 3840.453 | 8.2533 | 5.5091 |  |  | 96 | 3840.453 | 8.2533 | 5.5076
97 | 3880.458 | 8.2637 | 5.4896 |  |  | 97 | 3880.458 | 8.2637 | 5.5043 |  |  | 97 | 3880.458 | 8.2637 | 5.5029
98 | 3920.463 | 8.274 | 5.4849 |  |  | 98 | 3920.463 | 8.274 | 5.4997 |  |  | 98 | 3920.463 | 8.274 | 5.4982
99 | 3960.467 | 8.2841 | 5.4803 |  |  | 99 | 3960.467 | 8.2841 | 5.4951 |  |  | 99 | 3960.467 | 8.2841 | 5.4936
100 | 4000.472 | 8.2942 | 5.4757 |  |  | 100 | 4000.472 | 8.2942 | 5.4905 |  |  | 100 | 4000.472 | 8.2942 | 5.489
101 | 4040.477 | 8.3041 | 5.4712 |  |  | 101 | 4040.477 | 8.3041 | 5.486 |  |  | 101 | 4040.477 | 8.3041 | 5.4845
102 | 4080.482 | 8.314 | 5.4667 |  |  | 102 | 4080.482 | 8.314 | 5.4816 |  |  | 102 | 4080.482 | 8.314 | 5.48
103 | 4120.486 | 8.3237 | 5.4623 |  |  | 103 | 4120.486 | 8.3237 | 5.4772 |  |  | 103 | 4120.486 | 8.3237 | 5.4756
104 | 4160.491 | 8.3334 | 5.4579 |  |  | 104 | 4160.491 | 8.3334 | 5.4729 |  |  | 104 | 4160.491 | 8.3334 | 5.4712
105 | 4200.496 | 8.343 | 5.4536 |  |  | 105 | 4200.496 | 8.343 | 5.4686 |  |  | 105 | 4200.496 | 8.343 | 5.4669
106 | 4240.5 | 8.3524 | 5.4494 |  |  | 106 | 4240.5 | 8.3524 | 5.4644 |  |  | 106 | 4240.5 | 8.3524 | 5.4627
107 | 4280.505 | 8.3618 | 5.4452 |  |  | 107 | 4280.505 | 8.3618 | 5.4602 |  |  | 107 | 4280.505 | 8.3618 | 5.4585
108 | 4320.51 | 8.3711 | 5.441 |  |  | 108 | 4320.51 | 8.3711 | 5.4561 |  |  | 108 | 4320.51 | 8.3711 | 5.4543
109 | 4360.515 | 8.3803 | 5.4369 |  |  | 109 | 4360.515 | 8.3803 | 5.452 |  |  | 109 | 4360.515 | 8.3803 | 5.4502
110 | 4400.519 | 8.3895 | 5.4329 |  |  | 110 | 4400.519 | 8.3895 | 5.448 |  |  | 110 | 4400.519 | 8.3895 | 5.4462
111 | 4440.524 | 8.3985 | 5.4289 |  |  | 111 | 4440.524 | 8.3985 | 5.444 |  |  | 111 | 4440.524 | 8.3985 | 5.4422
112 | 4480.529 | 8.4075 | 5.4249 |  |  | 112 | 4480.529 | 8.4075 | 5.4401 |  |  | 112 | 4480.529 | 8.4075 | 5.4382
113 | 4520.534 | 8.4164 | 5.421 |  |  | 113 | 4520.534 | 8.4164 | 5.4362 |  |  | 113 | 4520.534 | 8.4164 | 5.4343
114 | 4560.538 | 8.4252 | 5.4172 |  |  | 114 | 4560.538 | 8.4252 | 5.4324 |  |  | 114 | 4560.538 | 8.4252 | 5.4305
115 | 4600.543 | 8.4339 | 5.4134 |  |  | 115 | 4600.543 | 8.4339 | 5.4286 |  |  | 115 | 4600.543 | 8.4339 | 5.4267
116 | 4640.548 | 8.4426 | 5.4096 |  |  | 116 | 4640.548 | 8.4426 | 5.4249 |  |  | 116 | 4640.548 | 8.4426 | 5.4229
117 | 4680.552 | 8.4512 | 5.4059 |  |  | 117 | 4680.552 | 8.4512 | 5.4212 |  |  | 117 | 4680.552 | 8.4512 | 5.4192
118 | 4720.557 | 8.4597 | 5.4022 |  |  | 118 | 4720.557 | 8.4597 | 5.4175 |  |  | 118 | 4720.557 | 8.4597 | 5.4155
119 | 4760.562 | 8.4681 | 5.3985 |  |  | 119 | 4760.562 | 8.4681 | 5.4139 |  |  | 119 | 4760.562 | 8.4681 | 5.4118
120 | 4800.567 | 8.4765 | 5.3949 |  |  | 120 | 4800.567 | 8.4765 | 5.4103 |  |  | 120 | 4800.567 | 8.4765 | 5.4082
121 | 4840.571 | 8.4848 | 5.3914 |  |  | 121 | 4840.571 | 8.4848 | 5.4068 |  |  | 121 | 4840.571 | 8.4848 | 5.4047
122 | 4880.576 | 8.493 | 5.3878 |  |  | 122 | 4880.576 | 8.493 | 5.4033 |  |  | 122 | 4880.576 | 8.493 | 5.4011
123 | 4920.581 | 8.5012 | 5.3844 |  |  | 123 | 4920.581 | 8.5012 | 5.3998 |  |  | 123 | 4920.581 | 8.5012 | 5.3977
124 | 4960.585 | 8.5093 | 5.3809 |  |  | 124 | 4960.585 | 8.5093 | 5.3964 |  |  | 124 | 4960.585 | 8.5093 | 5.3942
125 | 5000.59 | 8.5173 | 5.3775 |  |  | 125 | 5000.59 | 8.5173 | 5.393 |  |  | 125 | 5000.59 | 8.5173 | 5.3908
126 | 5040.595 | 8.5253 | 5.3741 |  |  | 126 | 5040.595 | 8.5253 | 5.3897 |  |  | 126 | 5040.595 | 8.5253 | 5.3874
127 | 5080.6 | 8.5332 | 5.3708 |  |  | 127 | 5080.6 | 8.5332 | 5.3864 |  |  | 127 | 5080.6 | 8.5332 | 5.3841
128 | 5120.604 | 8.541 | 5.3675 |  |  | 128 | 5120.604 | 8.541 | 5.3831 |  |  | 128 | 5120.604 | 8.541 | 5.3808
129 | 5160.609 | 8.5488 | 5.3643 |  |  | 129 | 5160.609 | 8.5488 | 5.3799 |  |  | 129 | 5160.609 | 8.5488 | 5.3776
130 | 5200.614 | 8.5565 | 5.361 |  |  | 130 | 5200.614 | 8.5565 | 5.3767 |  |  | 130 | 5200.614 | 8.5565 | 5.3743
131 | 5240.619 | 8.5642 | 5.3578 |  |  | 131 | 5240.619 | 8.5642 | 5.3735 |  |  | 131 | 5240.619 | 8.5642 | 5.3711
132 | 5280.623 | 8.5718 | 5.3547 |  |  | 132 | 5280.623 | 8.5718 | 5.3704 |  |  | 132 | 5280.623 | 8.5718 | 5.368
133 | 5320.628 | 8.5793 | 5.3516 |  |  | 133 | 5320.628 | 8.5793 | 5.3673 |  |  | 133 | 5320.628 | 8.5793 | 5.3649
134 | 5360.633 | 8.5868 | 5.3485 |  |  | 134 | 5360.633 | 8.5868 | 5.3642 |  |  | 134 | 5360.633 | 8.5868 | 5.3618
135 | 5400.637 | 8.5943 | 5.3454 |  |  | 135 | 5400.637 | 8.5943 | 5.3612 |  |  | 135 | 5400.637 | 8.5943 | 5.3587
136 | 5440.642 | 8.6017 | 5.3424 |  |  | 136 | 5440.642 | 8.6017 | 5.3582 |  |  | 136 | 5440.642 | 8.6017 | 5.3557
137 | 5480.647 | 8.609 | 5.3394 |  |  | 137 | 5480.647 | 8.609 | 5.3552 |  |  | 137 | 5480.647 | 8.609 | 5.3527
138 | 5520.652 | 8.6163 | 5.3365 |  |  | 138 | 5520.652 | 8.6163 | 5.3523 |  |  | 138 | 5520.652 | 8.6163 | 5.3498
139 | 5560.656 | 8.6235 | 5.3336 |  |  | 139 | 5560.656 | 8.6235 | 5.3494 |  |  | 139 | 5560.656 | 8.6235 | 5.3468
140 | 5600.661 | 8.6306 | 5.3307 |  |  | 140 | 5600.661 | 8.6306 | 5.3465 |  |  | 140 | 5600.661 | 8.6306 | 5.344
141 | 5640.666 | 8.6378 | 5.3278 |  |  | 141 | 5640.666 | 8.6378 | 5.3437 |  |  | 141 | 5640.666 | 8.6378 | 5.3411
142 | 5680.67 | 8.6448 | 5.325 |  |  | 142 | 5680.67 | 8.6448 | 5.3409 |  |  | 142 | 5680.67 | 8.6448 | 5.3383
143 | 5720.675 | 8.6518 | 5.3222 |  |  | 143 | 5720.675 | 8.6518 | 5.3381 |  |  | 143 | 5720.675 | 8.6518 | 5.3355
144 | 5760.68 | 8.6588 | 5.3194 |  |  | 144 | 5760.68 | 8.6588 | 5.3353 |  |  | 144 | 5760.68 | 8.6588 | 5.3327
145 | 5800.685 | 8.6657 | 5.3166 |  |  | 145 | 5800.685 | 8.6657 | 5.3326 |  |  | 145 | 5800.685 | 8.6657 | 5.3299
146 | 5840.689 | 8.6726 | 5.3139 |  |  | 146 | 5840.689 | 8.6726 | 5.3299 |  |  | 146 | 5840.689 | 8.6726 | 5.3272
147 | 5880.694 | 8.6794 | 5.3112 |  |  | 147 | 5880.694 | 8.6794 | 5.3273 |  |  | 147 | 5880.694 | 8.6794 | 5.3245
148 | 5920.699 | 8.6862 | 5.3086 |  |  | 148 | 5920.699 | 8.6862 | 5.3246 |  |  | 148 | 5920.699 | 8.6862 | 5.3219
149 | 5960.704 | 8.6929 | 5.3059 |  |  | 149 | 5960.704 | 8.6929 | 5.322 |  |  | 149 | 5960.704 | 8.6929 | 5.3192
150 | 6000.708 | 8.6996 | 5.3033 |  |  | 150 | 6000.708 | 8.6996 | 5.3194 |  |  | 150 | 6000.708 | 8.6996 | 5.3166
151 | 6040.713 | 8.7063 | 5.3007 |  |  | 151 | 6040.713 | 8.7063 | 5.3168 |  |  | 151 | 6040.713 | 8.7063 | 5.314
152 | 6080.718 | 8.7129 | 5.2982 |  |  | 152 | 6080.718 | 8.7129 | 5.3143 |  |  | 152 | 6080.718 | 8.7129 | 5.3115
153 | 6120.722 | 8.7194 | 5.2957 |  |  | 153 | 6120.722 | 8.7194 | 5.3118 |  |  | 153 | 6120.722 | 8.7194 | 5.3089
154 | 6160.727 | 8.726 | 5.2931 |  |  | 154 | 6160.727 | 8.726 | 5.3093 |  |  | 154 | 6160.727 | 8.726 | 5.3064
155 | 6200.732 | 8.7324 | 5.2907 |  |  | 155 | 6200.732 | 8.7324 | 5.3069 |  |  | 155 | 6200.732 | 8.7324 | 5.304
156 | 6240.737 | 8.7389 | 5.2882 |  |  | 156 | 6240.737 | 8.7389 | 5.3044 |  |  | 156 | 6240.737 | 8.7389 | 5.3015
157 | 6280.741 | 8.7452 | 5.2858 |  |  | 157 | 6280.741 | 8.7452 | 5.302 |  |  | 157 | 6280.741 | 8.7452 | 5.2991
158 | 6320.746 | 8.7516 | 5.2834 |  |  | 158 | 6320.746 | 8.7516 | 5.2996 |  |  | 158 | 6320.746 | 8.7516 | 5.2967
159 | 6360.751 | 8.7579 | 5.281 |  |  | 159 | 6360.751 | 8.7579 | 5.2973 |  |  | 159 | 6360.751 | 8.7579 | 5.2943
160 | 6400.755 | 8.7642 | 5.2786 |  |  | 160 | 6400.755 | 8.7642 | 5.2949 |  |  | 160 | 6400.755 | 8.7642 | 5.2919
161 | 6440.76 | 8.7704 | 5.2763 |  |  | 161 | 6440.76 | 8.7704 | 5.2926 |  |  | 161 | 6440.76 | 8.7704 | 5.2896
162 | 6480.765 | 8.7766 | 5.274 |  |  | 162 | 6480.765 | 8.7766 | 5.2903 |  |  | 162 | 6480.765 | 8.7766 | 5.2873
163 | 6520.77 | 8.7827 | 5.2717 |  |  | 163 | 6520.77 | 8.7827 | 5.288 |  |  | 163 | 6520.77 | 8.7827 | 5.285
164 | 6560.774 | 8.7889 | 5.2694 |  |  | 164 | 6560.774 | 8.7889 | 5.2858 |  |  | 164 | 6560.774 | 8.7889 | 5.2827
165 | 6600.779 | 8.7949 | 5.2672 |  |  | 165 | 6600.779 | 8.7949 | 5.2836 |  |  | 165 | 6600.779 | 8.7949 | 5.2805
166 | 6640.784 | 8.801 | 5.265 |  |  | 166 | 6640.784 | 8.801 | 5.2814 |  |  | 166 | 6640.784 | 8.801 | 5.2782
167 | 6680.789 | 8.807 | 5.2628 |  |  | 167 | 6680.789 | 8.807 | 5.2792 |  |  | 167 | 6680.789 | 8.807 | 5.276
168 | 6720.793 | 8.813 | 5.2606 |  |  | 168 | 6720.793 | 8.813 | 5.277 |  |  | 168 | 6720.793 | 8.813 | 5.2739
169 | 6760.798 | 8.8189 | 5.2584 |  |  | 169 | 6760.798 | 8.8189 | 5.2749 |  |  | 169 | 6760.798 | 8.8189 | 5.2717
170 | 6800.803 | 8.8248 | 5.2563 |  |  | 170 | 6800.803 | 8.8248 | 5.2727 |  |  | 170 | 6800.803 | 8.8248 | 5.2696
171 | 6840.807 | 8.8307 | 5.2542 |  |  | 171 | 6840.807 | 8.8307 | 5.2706 |  |  | 171 | 6840.807 | 8.8307 | 5.2675
172 | 6880.812 | 8.8365 | 5.2521 |  |  | 172 | 6880.812 | 8.8365 | 5.2686 |  |  | 172 | 6880.812 | 8.8365 | 5.2654
173 | 6920.817 | 8.8423 | 5.25 |  |  | 173 | 6920.817 | 8.8423 | 5.2665 |  |  | 173 | 6920.817 | 8.8423 | 5.2633
174 | 6960.822 | 8.8481 | 5.2479 |  |  | 174 | 6960.822 | 8.8481 | 5.2645 |  |  | 174 | 6960.822 | 8.8481 | 5.2612
175 | 7000.826 | 8.8538 | 5.2459 |  |  | 175 | 7000.826 | 8.8538 | 5.2624 |  |  | 175 | 7000.826 | 8.8538 | 5.2592
176 | 7040.831 | 8.8595 | 5.2439 |  |  | 176 | 7040.831 | 8.8595 | 5.2604 |  |  | 176 | 7040.831 | 8.8595 | 5.2572
177 | 7080.836 | 8.8651 | 5.2419 |  |  | 177 | 7080.836 | 8.8651 | 5.2584 |  |  | 177 | 7080.836 | 8.8651 | 5.2552
178 | 7120.84 | 8.8708 | 5.2399 |  |  | 178 | 7120.84 | 8.8708 | 5.2565 |  |  | 178 | 7120.84 | 8.8708 | 5.2532
179 | 7160.845 | 8.8764 | 5.2379 |  |  | 179 | 7160.845 | 8.8764 | 5.2545 |  |  | 179 | 7160.845 | 8.8764 | 5.2512
180 | 7200.85 | 8.882 | 5.236 |  |  | 180 | 7200.85 | 8.882 | 5.2526 |  |  | 180 | 7200.85 | 8.882 | 5.2493
181 | 7240.855 | 8.8875 | 5.2341 |  |  | 181 | 7240.855 | 8.8875 | 5.2507 |  |  | 181 | 7240.855 | 8.8875 | 5.2474
182 | 7280.859 | 8.893 | 5.2322 |  |  | 182 | 7280.859 | 8.893 | 5.2488 |  |  | 182 | 7280.859 | 8.893 | 5.2454
183 | 7320.864 | 8.8985 | 5.2303 |  |  | 183 | 7320.864 | 8.8985 | 5.2469 |  |  | 183 | 7320.864 | 8.8985 | 5.2436
184 | 7360.869 | 8.9039 | 5.2284 |  |  | 184 | 7360.869 | 8.9039 | 5.2451 |  |  | 184 | 7360.869 | 8.9039 | 5.2417
185 | 7400.874 | 8.9094 | 5.2266 |  |  | 185 | 7400.874 | 8.9094 | 5.2432 |  |  | 185 | 7400.874 | 8.9094 | 5.2398
186 | 7440.878 | 8.9147 | 5.2247 |  |  | 186 | 7440.878 | 8.9147 | 5.2414 |  |  | 186 | 7440.878 | 8.9147 | 5.238
187 | 7480.883 | 8.9201 | 5.2229 |  |  | 187 | 7480.883 | 8.9201 | 5.2396 |  |  | 187 | 7480.883 | 8.9201 | 5.2362
188 | 7520.888 | 8.9254 | 5.2211 |  |  | 188 | 7520.888 | 8.9254 | 5.2378 |  |  | 188 | 7520.888 | 8.9254 | 5.2344
189 | 7560.892 | 8.9307 | 5.2193 |  |  | 189 | 7560.892 | 8.9307 | 5.236 |  |  | 189 | 7560.892 | 8.9307 | 5.2326
190 | 7600.897 | 8.936 | 5.2175 |  |  | 190 | 7600.897 | 8.936 | 5.2343 |  |  | 190 | 7600.897 | 8.936 | 5.2308
191 | 7640.902 | 8.9413 | 5.2158 |  |  | 191 | 7640.902 | 8.9413 | 5.2325 |  |  | 191 | 7640.902 | 8.9413 | 5.2291
192 | 7680.907 | 8.9465 | 5.2141 |  |  | 192 | 7680.907 | 8.9465 | 5.2308 |  |  | 192 | 7680.907 | 8.9465 | 5.2273
193 | 7720.911 | 8.9517 | 5.2123 |  |  | 193 | 7720.911 | 8.9517 | 5.2291 |  |  | 193 | 7720.911 | 8.9517 | 5.2256
194 | 7760.916 | 8.9569 | 5.2106 |  |  | 194 | 7760.916 | 8.9569 | 5.2274 |  |  | 194 | 7760.916 | 8.9569 | 5.2239
195 | 7800.921 | 8.962 | 5.2089 |  |  | 195 | 7800.921 | 8.962 | 5.2257 |  |  | 195 | 7800.921 | 8.962 | 5.2222
196 | 7840.925 | 8.9671 | 5.2073 |  |  | 196 | 7840.925 | 8.9671 | 5.2241 |  |  | 196 | 7840.925 | 8.9671 | 5.2205
197 | 7880.93 | 8.9722 | 5.2056 |  |  | 197 | 7880.93 | 8.9722 | 5.2224 |  |  | 197 | 7880.93 | 8.9722 | 5.2189
198 | 7920.935 | 8.9773 | 5.204 |  |  | 198 | 7920.935 | 8.9773 | 5.2208 |  |  | 198 | 7920.935 | 8.9773 | 5.2172
199 | 7960.94 | 8.9823 | 5.2023 |  |  | 199 | 7960.94 | 8.9823 | 5.2192 |  |  | 199 | 7960.94 | 8.9823 | 5.2156
200 | 8000.944 | 8.9873 | 5.2007 |  |  | 200 | 8000.944 | 8.9873 | 5.2176 |  |  | 200 | 8000.944 | 8.9873 | 5.214
201 | 8040.949 | 8.9923 | 5.1991 |  |  | 201 | 8040.949 | 8.9923 | 5.216 |  |  | 201 | 8040.949 | 8.9923 | 5.2124
202 | 8080.954 | 8.9973 | 5.1975 |  |  | 202 | 8080.954 | 8.9973 | 5.2144 |  |  | 202 | 8080.954 | 8.9973 | 5.2108
203 | 8120.958 | 9.0022 | 5.196 |  |  | 203 | 8120.958 | 9.0022 | 5.2129 |  |  | 203 | 8120.958 | 9.0022 | 5.2092
204 | 8160.963 | 9.0071 | 5.1944 |  |  | 204 | 8160.963 | 9.0071 | 5.2113 |  |  | 204 | 8160.963 | 9.0071 | 5.2077
205 | 8200.968 | 9.012 | 5.1929 |  |  | 205 | 8200.968 | 9.012 | 5.2098 |  |  | 205 | 8200.968 | 9.012 | 5.2061
206 | 8240.973 | 9.0169 | 5.1913 |  |  | 206 | 8240.973 | 9.0169 | 5.2083 |  |  | 206 | 8240.973 | 9.0169 | 5.2046
207 | 8280.977 | 9.0217 | 5.1898 |  |  | 207 | 8280.977 | 9.0217 | 5.2068 |  |  | 207 | 8280.977 | 9.0217 | 5.2031
208 | 8320.982 | 9.0265 | 5.1883 |  |  | 208 | 8320.982 | 9.0265 | 5.2053 |  |  | 208 | 8320.982 | 9.0265 | 5.2016
209 | 8360.987 | 9.0313 | 5.1868 |  |  | 209 | 8360.987 | 9.0313 | 5.2038 |  |  | 209 | 8360.987 | 9.0313 | 5.2001
210 | 8400.992 | 9.0361 | 5.1853 |  |  | 210 | 8400.992 | 9.0361 | 5.2023 |  |  | 210 | 8400.992 | 9.0361 | 5.1986
211 | 8440.996 | 9.0409 | 5.1839 |  |  | 211 | 8440.996 | 9.0409 | 5.2009 |  |  | 211 | 8440.996 | 9.0409 | 5.1972
212 | 8481.001 | 9.0456 | 5.1824 |  |  | 212 | 8481.001 | 9.0456 | 5.1994 |  |  | 212 | 8481.001 | 9.0456 | 5.1957
213 | 8521.006 | 9.0503 | 5.181 |  |  | 213 | 8521.006 | 9.0503 | 5.198 |  |  | 213 | 8521.006 | 9.0503 | 5.1943
214 | 8561.01 | 9.055 | 5.1796 |  |  | 214 | 8561.01 | 9.055 | 5.1966 |  |  | 214 | 8561.01 | 9.055 | 5.1929
215 | 8601.015 | 9.0596 | 5.1782 |  |  | 215 | 8601.015 | 9.0596 | 5.1952 |  |  | 215 | 8601.015 | 9.0596 | 5.1914
216 | 8641.02 | 9.0643 | 5.1768 |  |  | 216 | 8641.02 | 9.0643 | 5.1938 |  |  | 216 | 8641.02 | 9.0643 | 5.19
217 | 8681.025 | 9.0689 | 5.1754 |  |  | 217 | 8681.025 | 9.0689 | 5.1924 |  |  | 217 | 8681.025 | 9.0689 | 5.1887
218 | 8721.029 | 9.0735 | 5.174 |  |  | 218 | 8721.029 | 9.0735 | 5.1911 |  |  | 218 | 8721.029 | 9.0735 | 5.1873
219 | 8761.034 | 9.0781 | 5.1726 |  |  | 219 | 8761.034 | 9.0781 | 5.1897 |  |  | 219 | 8761.034 | 9.0781 | 5.1859
220 | 8801.039 | 9.0826 | 5.1713 |  |  | 220 | 8801.039 | 9.0826 | 5.1884 |  |  | 220 | 8801.039 | 9.0826 | 5.1846
221 | 8841.043 | 9.0872 | 5.1699 |  |  | 221 | 8841.043 | 9.0872 | 5.187 |  |  | 221 | 8841.043 | 9.0872 | 5.1832
222 | 8881.048 | 9.0917 | 5.1686 |  |  | 222 | 8881.048 | 9.0917 | 5.1857 |  |  | 222 | 8881.048 | 9.0917 | 5.1819
223 | 8921.053 | 9.0962 | 5.1673 |  |  | 223 | 8921.053 | 9.0962 | 5.1844 |  |  | 223 | 8921.053 | 9.0962 | 5.1806
224 | 8961.058 | 9.1006 | 5.166 |  |  | 224 | 8961.058 | 9.1006 | 5.1831 |  |  | 224 | 8961.058 | 9.1006 | 5.1793
225 | 9001.062 | 9.1051 | 5.1647 |  |  | 225 | 9001.062 | 9.1051 | 5.1818 |  |  | 225 | 9001.062 | 9.1051 | 5.178
226 | 9041.067 | 9.1095 | 5.1634 |  |  | 226 | 9041.067 | 9.1095 | 5.1806 |  |  | 226 | 9041.067 | 9.1095 | 5.1767
227 | 9081.072 | 9.1139 | 5.1621 |  |  | 227 | 9081.072 | 9.1139 | 5.1793 |  |  | 227 | 9081.072 | 9.1139 | 5.1754
228 | 9121.077 | 9.1183 | 5.1609 |  |  | 228 | 9121.077 | 9.1183 | 5.178 |  |  | 228 | 9121.077 | 9.1183 | 5.1742
229 | 9161.081 | 9.1227 | 5.1596 |  |  | 229 | 9161.081 | 9.1227 | 5.1768 |  |  | 229 | 9161.081 | 9.1227 | 5.1729
230 | 9201.086 | 9.1271 | 5.1584 |  |  | 230 | 9201.086 | 9.1271 | 5.1756 |  |  | 230 | 9201.086 | 9.1271 | 5.1717
231 | 9241.091 | 9.1314 | 5.1572 |  |  | 231 | 9241.091 | 9.1314 | 5.1743 |  |  | 231 | 9241.091 | 9.1314 | 5.1705
232 | 9281.095 | 9.1357 | 5.156 |  |  | 232 | 9281.095 | 9.1357 | 5.1731 |  |  | 232 | 9281.095 | 9.1357 | 5.1693
233 | 9321.1 | 9.14 | 5.1547 |  |  | 233 | 9321.1 | 9.14 | 5.1719 |  |  | 233 | 9321.1 | 9.14 | 5.1681
234 | 9361.105 | 9.1443 | 5.1536 |  |  | 234 | 9361.105 | 9.1443 | 5.1708 |  |  | 234 | 9361.105 | 9.1443 | 5.1669
235 | 9401.11 | 9.1486 | 5.1524 |  |  | 235 | 9401.11 | 9.1486 | 5.1696 |  |  | 235 | 9401.11 | 9.1486 | 5.1657
236 | 9441.114 | 9.1528 | 5.1512 |  |  | 236 | 9441.114 | 9.1528 | 5.1684 |  |  | 236 | 9441.114 | 9.1528 | 5.1645
237 | 9481.119 | 9.1571 | 5.15 |  |  | 237 | 9481.119 | 9.1571 | 5.1672 |  |  | 237 | 9481.119 | 9.1571 | 5.1633
238 | 9521.124 | 9.1613 | 5.1489 |  |  | 238 | 9521.124 | 9.1613 | 5.1661 |  |  | 238 | 9521.124 | 9.1613 | 5.1622
239 | 9561.128 | 9.1655 | 5.1477 |  |  | 239 | 9561.128 | 9.1655 | 5.165 |  |  | 239 | 9561.128 | 9.1655 | 5.161
240 | 9601.133 | 9.1696 | 5.1466 |  |  | 240 | 9601.133 | 9.1696 | 5.1638 |  |  | 240 | 9601.133 | 9.1696 | 5.1599
241 | 9641.138 | 9.1738 | 5.1455 |  |  | 241 | 9641.138 | 9.1738 | 5.1627 |  |  | 241 | 9641.138 | 9.1738 | 5.1588
242 | 9681.143 | 9.1779 | 5.1443 |  |  | 242 | 9681.143 | 9.1779 | 5.1616 |  |  | 242 | 9681.143 | 9.1779 | 5.1577
243 | 9721.147 | 9.1821 | 5.1432 |  |  | 243 | 9721.147 | 9.1821 | 5.1605 |  |  | 243 | 9721.147 | 9.1821 | 5.1565
244 | 9761.152 | 9.1862 | 5.1421 |  |  | 244 | 9761.152 | 9.1862 | 5.1594 |  |  | 244 | 9761.152 | 9.1862 | 5.1554
245 | 9801.157 | 9.1903 | 5.141 |  |  | 245 | 9801.157 | 9.1903 | 5.1583 |  |  | 245 | 9801.157 | 9.1903 | 5.1544
246 | 9841.162 | 9.1943 | 5.14 |  |  | 246 | 9841.162 | 9.1943 | 5.1572 |  |  | 246 | 9841.162 | 9.1943 | 5.1533
247 | 9881.166 | 9.1984 | 5.1389 |  |  | 247 | 9881.166 | 9.1984 | 5.1562 |  |  | 247 | 9881.166 | 9.1984 | 5.1522
248 | 9921.171 | 9.2024 | 5.1378 |  |  | 248 | 9921.171 | 9.2024 | 5.1551 |  |  | 248 | 9921.171 | 9.2024 | 5.1511
249 | 9961.176 | 9.2065 | 5.1368 |  |  | 249 | 9961.176 | 9.2065 | 5.1541 |  |  | 249 | 9961.176 | 9.2065 | 5.1501
250 | 10001.18 | 9.2105 | 5.1357 |  |  | 250 | 10001.18 | 9.2105 | 5.153 |  |  | 250 | 10001.18 | 9.2105 | 5.1491
251 | 10041.185 | 9.2145 | 5.1347 |  |  | 251 | 10041.185 | 9.2145 | 5.152 |  |  | 251 | 10041.185 | 9.2145 | 5.148
252 | 10081.19 | 9.2184 | 5.1337 |  |  | 252 | 10081.19 | 9.2184 | 5.151 |  |  | 252 | 10081.19 | 9.2184 | 5.147
253 | 10121.195 | 9.2224 | 5.1326 |  |  | 253 | 10121.195 | 9.2224 | 5.15 |  |  | 253 | 10121.195 | 9.2224 | 5.146
254 | 10161.199 | 9.2263 | 5.1316 |  |  | 254 | 10161.199 | 9.2263 | 5.149 |  |  | 254 | 10161.199 | 9.2263 | 5.145
255 | 10201.204 | 9.2303 | 5.1306 |  |  | 255 | 10201.204 | 9.2303 | 5.148 |  |  | 255 | 10201.204 | 9.2303 | 5.144
256 | 10241.209 | 9.2342 | 5.1296 |  |  | 256 | 10241.209 | 9.2342 | 5.147 |  |  | 256 | 10241.209 | 9.2342 | 5.143
257 | 10281.213 | 9.2381 | 5.1286 |  |  | 257 | 10281.213 | 9.2381 | 5.146 |  |  | 257 | 10281.213 | 9.2381 | 5.142
258 | 10321.218 | 9.242 | 5.1277 |  |  | 258 | 10321.218 | 9.242 | 5.145 |  |  | 258 | 10321.218 | 9.242 | 5.141
259 | 10361.223 | 9.2458 | 5.1267 |  |  | 259 | 10361.223 | 9.2458 | 5.144 |  |  | 259 | 10361.223 | 9.2458 | 5.14
260 | 10401.228 | 9.2497 | 5.1257 |  |  | 260 | 10401.228 | 9.2497 | 5.1431 |  |  | 260 | 10401.228 | 9.2497 | 5.1391
261 | 10441.232 | 9.2535 | 5.1248 |  |  | 261 | 10441.232 | 9.2535 | 5.1421 |  |  | 261 | 10441.232 | 9.2535 | 5.1381
262 | 10481.237 | 9.2573 | 5.1238 |  |  | 262 | 10481.237 | 9.2573 | 5.1412 |  |  | 262 | 10481.237 | 9.2573 | 5.1372
263 | 10521.242 | 9.2612 | 5.1229 |  |  | 263 | 10521.242 | 9.2612 | 5.1403 |  |  | 263 | 10521.242 | 9.2612 | 5.1362
264 | 10561.247 | 9.2649 | 5.122 |  |  | 264 | 10561.247 | 9.2649 | 5.1393 |  |  | 264 | 10561.247 | 9.2649 | 5.1353
265 | 10601.251 | 9.2687 | 5.121 |  |  | 265 | 10601.251 | 9.2687 | 5.1384 |  |  | 265 | 10601.251 | 9.2687 | 5.1344
266 | 10641.256 | 9.2725 | 5.1201 |  |  | 266 | 10641.256 | 9.2725 | 5.1375 |  |  | 266 | 10641.256 | 9.2725 | 5.1335
267 | 10681.261 | 9.2762 | 5.1192 |  |  | 267 | 10681.261 | 9.2762 | 5.1366 |  |  | 267 | 10681.261 | 9.2762 | 5.1326
268 | 10721.265 | 9.28 | 5.1183 |  |  | 268 | 10721.265 | 9.28 | 5.1357 |  |  | 268 | 10721.265 | 9.28 | 5.1317
269 | 10761.27 | 9.2837 | 5.1174 |  |  | 269 | 10761.27 | 9.2837 | 5.1348 |  |  | 269 | 10761.27 | 9.2837 | 5.1308
270 | 10801.275 | 9.2874 | 5.1165 |  |  | 270 | 10801.275 | 9.2874 | 5.1339 |  |  | 270 | 10801.275 | 9.2874 | 5.1299
271 | 10841.28 | 9.2911 | 5.1156 |  |  | 271 | 10841.28 | 9.2911 | 5.1331 |  |  | 271 | 10841.28 | 9.2911 | 5.129
272 | 10881.284 | 9.2948 | 5.1148 |  |  | 272 | 10881.284 | 9.2948 | 5.1322 |  |  | 272 | 10881.284 | 9.2948 | 5.1281
273 | 10921.289 | 9.2985 | 5.1139 |  |  | 273 | 10921.289 | 9.2985 | 5.1313 |  |  | 273 | 10921.289 | 9.2985 | 5.1273
274 | 10961.294 | 9.3021 | 5.1131 |  |  | 274 | 10961.294 | 9.3021 | 5.1305 |  |  | 274 | 10961.294 | 9.3021 | 5.1264
275 | 11001.298 | 9.3058 | 5.1122 |  |  | 275 | 11001.298 | 9.3058 | 5.1296 |  |  | 275 | 11001.298 | 9.3058 | 5.1256
276 | 11041.303 | 9.3094 | 5.1114 |  |  | 276 | 11041.303 | 9.3094 | 5.1288 |  |  | 276 | 11041.303 | 9.3094 | 5.1247
277 | 11081.308 | 9.313 | 5.1105 |  |  | 277 | 11081.308 | 9.313 | 5.1279 |  |  | 277 | 11081.308 | 9.313 | 5.1239
278 | 11121.313 | 9.3166 | 5.1097 |  |  | 278 | 11121.313 | 9.3166 | 5.1271 |  |  | 278 | 11121.313 | 9.3166 | 5.1231
279 | 11161.317 | 9.3202 | 5.1089 |  |  | 279 | 11161.317 | 9.3202 | 5.1263 |  |  | 279 | 11161.317 | 9.3202 | 5.1222
280 | 11201.322 | 9.3238 | 5.108 |  |  | 280 | 11201.322 | 9.3238 | 5.1255 |  |  | 280 | 11201.322 | 9.3238 | 5.1214
281 | 11241.327 | 9.3274 | 5.1072 |  |  | 281 | 11241.327 | 9.3274 | 5.1247 |  |  | 281 | 11241.327 | 9.3274 | 5.1206
282 | 11281.331 | 9.3309 | 5.1064 |  |  | 282 | 11281.331 | 9.3309 | 5.1239 |  |  | 282 | 11281.331 | 9.3309 | 5.1198
283 | 11321.336 | 9.3344 | 5.1056 |  |  | 283 | 11321.336 | 9.3344 | 5.1231 |  |  | 283 | 11321.336 | 9.3344 | 5.119
284 | 11361.341 | 9.338 | 5.1048 |  |  | 284 | 11361.341 | 9.338 | 5.1223 |  |  | 284 | 11361.341 | 9.338 | 5.1182
285 | 11401.346 | 9.3415 | 5.104 |  |  | 285 | 11401.346 | 9.3415 | 5.1215 |  |  | 285 | 11401.346 | 9.3415 | 5.1174
286 | 11441.35 | 9.345 | 5.1033 |  |  | 286 | 11441.35 | 9.345 | 5.1207 |  |  | 286 | 11441.35 | 9.345 | 5.1167
287 | 11481.355 | 9.3485 | 5.1025 |  |  | 287 | 11481.355 | 9.3485 | 5.12 |  |  | 287 | 11481.355 | 9.3485 | 5.1159
288 | 11521.36 | 9.352 | 5.1017 |  |  | 288 | 11521.36 | 9.352 | 5.1192 |  |  | 288 | 11521.36 | 9.352 | 5.1151
289 | 11561.365 | 9.3554 | 5.101 |  |  | 289 | 11561.365 | 9.3554 | 5.1184 |  |  | 289 | 11561.365 | 9.3554 | 5.1144
290 | 11601.369 | 9.3589 | 5.1002 |  |  | 290 | 11601.369 | 9.3589 | 5.1177 |  |  | 290 | 11601.369 | 9.3589 | 5.1136
291 | 11641.374 | 9.3623 | 5.0995 |  |  | 291 | 11641.374 | 9.3623 | 5.1169 |  |  | 291 | 11641.374 | 9.3623 | 5.1129
292 | 11681.379 | 9.3658 | 5.0987 |  |  | 292 | 11681.379 | 9.3658 | 5.1162 |  |  | 292 | 11681.379 | 9.3658 | 5.1121
293 | 11721.383 | 9.3692 | 5.098 |  |  | 293 | 11721.383 | 9.3692 | 5.1155 |  |  | 293 | 11721.383 | 9.3692 | 5.1114
294 | 11761.388 | 9.3726 | 5.0972 |  |  | 294 | 11761.388 | 9.3726 | 5.1147 |  |  | 294 | 11761.388 | 9.3726 | 5.1107
295 | 11801.393 | 9.376 | 5.0965 |  |  | 295 | 11801.393 | 9.376 | 5.114 |  |  | 295 | 11801.393 | 9.376 | 5.1099
296 | 11841.398 | 9.3794 | 5.0958 |  |  | 296 | 11841.398 | 9.3794 | 5.1133 |  |  | 296 | 11841.398 | 9.3794 | 5.1092
297 | 11881.402 | 9.3827 | 5.0951 |  |  | 297 | 11881.402 | 9.3827 | 5.1126 |  |  | 297 | 11881.402 | 9.3827 | 5.1085
298 | 11921.407 | 9.3861 | 5.0944 |  |  | 298 | 11921.407 | 9.3861 | 5.1119 |  |  | 298 | 11921.407 | 9.3861 | 5.1078
299 | 11961.412 | 9.3894 | 5.0937 |  |  | 299 | 11961.412 | 9.3894 | 5.1112 |  |  | 299 | 11961.412 | 9.3894 | 5.1071
300 | 12001.416 | 9.3928 | 5.093 |  |  | 300 | 12001.416 | 9.3928 | 5.1105 |  |  | 300 | 12001.416 | 9.3928 | 5.1064
301 | 12041.421 | 9.3961 | 5.0923 |  |  | 301 | 12041.421 | 9.3961 | 5.1098 |  |  | 301 | 12041.421 | 9.3961 | 5.1057
302 | 12081.426 | 9.3994 | 5.0916 |  |  | 302 | 12081.426 | 9.3994 | 5.1091 |  |  | 302 | 12081.426 | 9.3994 | 5.105
303 | 12121.431 | 9.4027 | 5.0909 |  |  | 303 | 12121.431 | 9.4027 | 5.1084 |  |  | 303 | 12121.431 | 9.4027 | 5.1043
304 | 12161.435 | 9.406 | 5.0902 |  |  | 304 | 12161.435 | 9.406 | 5.1077 |  |  | 304 | 12161.435 | 9.406 | 5.1037
305 | 12201.44 | 9.4093 | 5.0896 |  |  | 305 | 12201.44 | 9.4093 | 5.1071 |  |  | 305 | 12201.44 | 9.4093 | 5.103
306 | 12241.445 | 9.4126 | 5.0889 |  |  | 306 | 12241.445 | 9.4126 | 5.1064 |  |  | 306 | 12241.445 | 9.4126 | 5.1023
307 | 12281.45 | 9.4158 | 5.0882 |  |  | 307 | 12281.45 | 9.4158 | 5.1057 |  |  | 307 | 12281.45 | 9.4158 | 5.1017
308 | 12321.454 | 9.4191 | 5.0876 |  |  | 308 | 12321.454 | 9.4191 | 5.1051 |  |  | 308 | 12321.454 | 9.4191 | 5.101
309 | 12361.459 | 9.4223 | 5.0869 |  |  | 309 | 12361.459 | 9.4223 | 5.1044 |  |  | 309 | 12361.459 | 9.4223 | 5.1004
310 | 12401.464 | 9.4256 | 5.0863 |  |  | 310 | 12401.464 | 9.4256 | 5.1038 |  |  | 310 | 12401.464 | 9.4256 | 5.0997
311 | 12441.468 | 9.4288 | 5.0856 |  |  | 311 | 12441.468 | 9.4288 | 5.1031 |  |  | 311 | 12441.468 | 9.4288 | 5.0991
312 | 12481.473 | 9.432 | 5.085 |  |  | 312 | 12481.473 | 9.432 | 5.1025 |  |  | 312 | 12481.473 | 9.432 | 5.0984
313 | 12521.478 | 9.4352 | 5.0844 |  |  | 313 | 12521.478 | 9.4352 | 5.1019 |  |  | 313 | 12521.478 | 9.4352 | 5.0978
314 | 12561.483 | 9.4384 | 5.0837 |  |  | 314 | 12561.483 | 9.4384 | 5.1012 |  |  | 314 | 12561.483 | 9.4384 | 5.0972
315 | 12601.487 | 9.4416 | 5.0831 |  |  | 315 | 12601.487 | 9.4416 | 5.1006 |  |  | 315 | 12601.487 | 9.4416 | 5.0966
316 | 12641.492 | 9.4447 | 5.0825 |  |  | 316 | 12641.492 | 9.4447 | 5.1 |  |  | 316 | 12641.492 | 9.4447 | 5.096
317 | 12681.497 | 9.4479 | 5.0819 |  |  | 317 | 12681.497 | 9.4479 | 5.0994 |  |  | 317 | 12681.497 | 9.4479 | 5.0953
318 | 12721.501 | 9.451 | 5.0813 |  |  | 318 | 12721.501 | 9.451 | 5.0988 |  |  | 318 | 12721.501 | 9.451 | 5.0947
319 | 12761.506 | 9.4542 | 5.0807 |  |  | 319 | 12761.506 | 9.4542 | 5.0982 |  |  | 319 | 12761.506 | 9.4542 | 5.0941
320 | 12801.511 | 9.4573 | 5.0801 |  |  | 320 | 12801.511 | 9.4573 | 5.0976 |  |  | 320 | 12801.511 | 9.4573 | 5.0935
321 | 12841.516 | 9.4604 | 5.0795 |  |  | 321 | 12841.516 | 9.4604 | 5.097 |  |  | 321 | 12841.516 | 9.4604 | 5.0929
322 | 12881.52 | 9.4635 | 5.0789 |  |  | 322 | 12881.52 | 9.4635 | 5.0964 |  |  | 322 | 12881.52 | 9.4635 | 5.0924
323 | 12921.525 | 9.4666 | 5.0783 |  |  | 323 | 12921.525 | 9.4666 | 5.0958 |  |  | 323 | 12921.525 | 9.4666 | 5.0918
324 | 12961.53 | 9.4697 | 5.0777 |  |  | 324 | 12961.53 | 9.4697 | 5.0952 |  |  | 324 | 12961.53 | 9.4697 | 5.0912
325 | 13001.535 | 9.4728 | 5.0771 |  |  | 325 | 13001.535 | 9.4728 | 5.0946 |  |  | 325 | 13001.535 | 9.4728 | 5.0906
326 | 13041.539 | 9.4759 | 5.0766 |  |  | 326 | 13041.539 | 9.4759 | 5.0941 |  |  | 326 | 13041.539 | 9.4759 | 5.09
327 | 13081.544 | 9.479 | 5.076 |  |  | 327 | 13081.544 | 9.479 | 5.0935 |  |  | 327 | 13081.544 | 9.479 | 5.0895
328 | 13121.549 | 9.482 | 5.0754 |  |  | 328 | 13121.549 | 9.482 | 5.0929 |  |  | 328 | 13121.549 | 9.482 | 5.0889
329 | 13161.553 | 9.4851 | 5.0749 |  |  | 329 | 13161.553 | 9.4851 | 5.0924 |  |  | 329 | 13161.553 | 9.4851 | 5.0884
330 | 13201.558 | 9.4881 | 5.0743 |  |  | 330 | 13201.558 | 9.4881 | 5.0918 |  |  | 330 | 13201.558 | 9.4881 | 5.0878
331 | 13241.563 | 9.4911 | 5.0738 |  |  | 331 | 13241.563 | 9.4911 | 5.0913 |  |  | 331 | 13241.563 | 9.4911 | 5.0873
332 | 13281.568 | 9.4941 | 5.0732 |  |  | 332 | 13281.568 | 9.4941 | 5.0907 |  |  | 332 | 13281.568 | 9.4941 | 5.0867
333 | 13321.572 | 9.4971 | 5.0727 |  |  | 333 | 13321.572 | 9.4971 | 5.0902 |  |  | 333 | 13321.572 | 9.4971 | 5.0862
334 | 13361.577 | 9.5001 | 5.0721 |  |  | 334 | 13361.577 | 9.5001 | 5.0896 |  |  | 334 | 13361.577 | 9.5001 | 5.0856
335 | 13401.582 | 9.5031 | 5.0716 |  |  | 335 | 13401.582 | 9.5031 | 5.0891 |  |  | 335 | 13401.582 | 9.5031 | 5.0851
336 | 13441.586 | 9.5061 | 5.0711 |  |  | 336 | 13441.586 | 9.5061 | 5.0886 |  |  | 336 | 13441.586 | 9.5061 | 5.0846
337 | 13481.591 | 9.5091 | 5.0705 |  |  | 337 | 13481.591 | 9.5091 | 5.088 |  |  | 337 | 13481.591 | 9.5091 | 5.084
338 | 13521.596 | 9.512 | 5.07 |  |  | 338 | 13521.596 | 9.512 | 5.0875 |  |  | 338 | 13521.596 | 9.512 | 5.0835
339 | 13561.601 | 9.515 | 5.0695 |  |  | 339 | 13561.601 | 9.515 | 5.087 |  |  | 339 | 13561.601 | 9.515 | 5.083
340 | 13601.605 | 9.5179 | 5.069 |  |  | 340 | 13601.605 | 9.5179 | 5.0865 |  |  | 340 | 13601.605 | 9.5179 | 5.0825
341 | 13641.61 | 9.5209 | 5.0685 |  |  | 341 | 13641.61 | 9.5209 | 5.086 |  |  | 341 | 13641.61 | 9.5209 | 5.082
342 | 13681.615 | 9.5238 | 5.0679 |  |  | 342 | 13681.615 | 9.5238 | 5.0855 |  |  | 342 | 13681.615 | 9.5238 | 5.0815
343 | 13721.62 | 9.5267 | 5.0674 |  |  | 343 | 13721.62 | 9.5267 | 5.085 |  |  | 343 | 13721.62 | 9.5267 | 5.081
344 | 13761.624 | 9.5296 | 5.0669 |  |  | 344 | 13761.624 | 9.5296 | 5.0844 |  |  | 344 | 13761.624 | 9.5296 | 5.0805
345 | 13801.629 | 9.5325 | 5.0664 |  |  | 345 | 13801.629 | 9.5325 | 5.0839 |  |  | 345 | 13801.629 | 9.5325 | 5.08
346 | 13841.634 | 9.5354 | 5.0659 |  |  | 346 | 13841.634 | 9.5354 | 5.0835 |  |  | 346 | 13841.634 | 9.5354 | 5.0795
347 | 13881.638 | 9.5383 | 5.0655 |  |  | 347 | 13881.638 | 9.5383 | 5.083 |  |  | 347 | 13881.638 | 9.5383 | 5.079
348 | 13921.643 | 9.5412 | 5.065 |  |  | 348 | 13921.643 | 9.5412 | 5.0825 |  |  | 348 | 13921.643 | 9.5412 | 5.0785
349 | 13961.648 | 9.5441 | 5.0645 |  |  | 349 | 13961.648 | 9.5441 | 5.082 |  |  | 349 | 13961.648 | 9.5441 | 5.078
350 | 14001.653 | 9.5469 | 5.064 |  |  | 350 | 14001.653 | 9.5469 | 5.0815 |  |  | 350 | 14001.653 | 9.5469 | 5.0775
351 | 14041.657 | 9.5498 | 5.0635 |  |  | 351 | 14041.657 | 9.5498 | 5.081 |  |  | 351 | 14041.657 | 9.5498 | 5.0771
352 | 14081.662 | 9.5526 | 5.0631 |  |  | 352 | 14081.662 | 9.5526 | 5.0806 |  |  | 352 | 14081.662 | 9.5526 | 5.0766
353 | 14121.667 | 9.5555 | 5.0626 |  |  | 353 | 14121.667 | 9.5555 | 5.0801 |  |  | 353 | 14121.667 | 9.5555 | 5.0761
354 | 14161.671 | 9.5583 | 5.0621 |  |  | 354 | 14161.671 | 9.5583 | 5.0796 |  |  | 354 | 14161.671 | 9.5583 | 5.0757
355 | 14201.676 | 9.5611 | 5.0617 |  |  | 355 | 14201.676 | 9.5611 | 5.0792 |  |  | 355 | 14201.676 | 9.5611 | 5.0752
356 | 14241.681 | 9.5639 | 5.0612 |  |  | 356 | 14241.681 | 9.5639 | 5.0787 |  |  | 356 | 14241.681 | 9.5639 | 5.0747
357 | 14281.686 | 9.5667 | 5.0607 |  |  | 357 | 14281.686 | 9.5667 | 5.0782 |  |  | 357 | 14281.686 | 9.5667 | 5.0743
358 | 14321.69 | 9.5695 | 5.0603 |  |  | 358 | 14321.69 | 9.5695 | 5.0778 |  |  | 358 | 14321.69 | 9.5695 | 5.0738
359 | 14361.695 | 9.5723 | 5.0598 |  |  | 359 | 14361.695 | 9.5723 | 5.0773 |  |  | 359 | 14361.695 | 9.5723 | 5.0734
360 | 14401.7 | 9.5751 | 5.0594 |  |  | 360 | 14401.7 | 9.5751 | 5.0769 |  |  | 360 | 14401.7 | 9.5751 | 5.0729
361 | 14441.705 | 9.5779 | 5.0589 |  |  | 361 | 14441.705 | 9.5779 | 5.0764 |  |  | 361 | 14441.705 | 9.5779 | 5.0725
362 | 14481.709 | 9.5806 | 5.0585 |  |  | 362 | 14481.709 | 9.5806 | 5.076 |  |  | 362 | 14481.709 | 9.5806 | 5.0721
363 | 14521.714 | 9.5834 | 5.0581 |  |  | 363 | 14521.714 | 9.5834 | 5.0755 |  |  | 363 | 14521.714 | 9.5834 | 5.0716
364 | 14561.719 | 9.5862 | 5.0576 |  |  | 364 | 14561.719 | 9.5862 | 5.0751 |  |  | 364 | 14561.719 | 9.5862 | 5.0712
365 | 14601.723 | 9.5889 | 5.0572 |  |  | 365 | 14601.723 | 9.5889 | 5.0747 |  |  | 365 | 14601.723 | 9.5889 | 5.0708
366 | 14641.728 | 9.5916 | 5.0568 |  |  | 366 | 14641.728 | 9.5916 | 5.0742 |  |  | 366 | 14641.728 | 9.5916 | 5.0703
367 | 14681.733 | 9.5944 | 5.0563 |  |  | 367 | 14681.733 | 9.5944 | 5.0738 |  |  | 367 | 14681.733 | 9.5944 | 5.0699
368 | 14721.738 | 9.5971 | 5.0559 |  |  | 368 | 14721.738 | 9.5971 | 5.0734 |  |  | 368 | 14721.738 | 9.5971 | 5.0695
369 | 14761.742 | 9.5998 | 5.0555 |  |  | 369 | 14761.742 | 9.5998 | 5.073 |  |  | 369 | 14761.742 | 9.5998 | 5.0691
370 | 14801.747 | 9.6025 | 5.0551 |  |  | 370 | 14801.747 | 9.6025 | 5.0726 |  |  | 370 | 14801.747 | 9.6025 | 5.0687
371 | 14841.752 | 9.6052 | 5.0547 |  |  | 371 | 14841.752 | 9.6052 | 5.0721 |  |  | 371 | 14841.752 | 9.6052 | 5.0683
372 | 14881.756 | 9.6079 | 5.0543 |  |  | 372 | 14881.756 | 9.6079 | 5.0717 |  |  | 372 | 14881.756 | 9.6079 | 5.0678
373 | 14921.761 | 9.6106 | 5.0539 |  |  | 373 | 14921.761 | 9.6106 | 5.0713 |  |  | 373 | 14921.761 | 9.6106 | 5.0674
374 | 14961.766 | 9.6133 | 5.0534 |  |  | 374 | 14961.766 | 9.6133 | 5.0709 |  |  | 374 | 14961.766 | 9.6133 | 5.067
375 | 15001.771 | 9.6159 | 5.053 |  |  | 375 | 15001.771 | 9.6159 | 5.0705 |  |  | 375 | 15001.771 | 9.6159 | 5.0666
376 | 15041.775 | 9.6186 | 5.0526 |  |  | 376 | 15041.775 | 9.6186 | 5.0701 |  |  | 376 | 15041.775 | 9.6186 | 5.0662
377 | 15081.78 | 9.6212 | 5.0523 |  |  | 377 | 15081.78 | 9.6212 | 5.0697 |  |  | 377 | 15081.78 | 9.6212 | 5.0658
378 | 15121.785 | 9.6239 | 5.0519 |  |  | 378 | 15121.785 | 9.6239 | 5.0693 |  |  | 378 | 15121.785 | 9.6239 | 5.0654
379 | 15161.789 | 9.6265 | 5.0515 |  |  | 379 | 15161.789 | 9.6265 | 5.0689 |  |  | 379 | 15161.789 | 9.6265 | 5.0651
380 | 15201.794 | 9.6292 | 5.0511 |  |  | 380 | 15201.794 | 9.6292 | 5.0685 |  |  | 380 | 15201.794 | 9.6292 | 5.0647
381 | 15241.799 | 9.6318 | 5.0507 |  |  | 381 | 15241.799 | 9.6318 | 5.0681 |  |  | 381 | 15241.799 | 9.6318 | 5.0643
382 | 15281.804 | 9.6344 | 5.0503 |  |  | 382 | 15281.804 | 9.6344 | 5.0677 |  |  | 382 | 15281.804 | 9.6344 | 5.0639
383 | 15321.808 | 9.637 | 5.0499 |  |  | 383 | 15321.808 | 9.637 | 5.0674 |  |  | 383 | 15321.808 | 9.637 | 5.0635
384 | 15361.813 | 9.6396 | 5.0495 |  |  | 384 | 15361.813 | 9.6396 | 5.067 |  |  | 384 | 15361.813 | 9.6396 | 5.0631
385 | 15401.818 | 9.6422 | 5.0492 |  |  | 385 | 15401.818 | 9.6422 | 5.0666 |  |  | 385 | 15401.818 | 9.6422 | 5.0628
386 | 15441.823 | 9.6448 | 5.0488 |  |  | 386 | 15441.823 | 9.6448 | 5.0662 |  |  | 386 | 15441.823 | 9.6448 | 5.0624
387 | 15481.827 | 9.6474 | 5.0484 |  |  | 387 | 15481.827 | 9.6474 | 5.0659 |  |  | 387 | 15481.827 | 9.6474 | 5.062
388 | 15521.832 | 9.65 | 5.0481 |  |  | 388 | 15521.832 | 9.65 | 5.0655 |  |  | 388 | 15521.832 | 9.65 | 5.0617
389 | 15561.837 | 9.6526 | 5.0477 |  |  | 389 | 15561.837 | 9.6526 | 5.0651 |  |  | 389 | 15561.837 | 9.6526 | 5.0613
390 | 15601.841 | 9.6551 | 5.0473 |  |  | 390 | 15601.841 | 9.6551 | 5.0648 |  |  | 390 | 15601.841 | 9.6551 | 5.0609
391 | 15641.846 | 9.6577 | 5.047 |  |  | 391 | 15641.846 | 9.6577 | 5.0644 |  |  | 391 | 15641.846 | 9.6577 | 5.0606
392 | 15681.851 | 9.6603 | 5.0466 |  |  | 392 | 15681.851 | 9.6603 | 5.064 |  |  | 392 | 15681.851 | 9.6603 | 5.0602
393 | 15721.856 | 9.6628 | 5.0463 |  |  | 393 | 15721.856 | 9.6628 | 5.0637 |  |  | 393 | 15721.856 | 9.6628 | 5.0599
394 | 15761.86 | 9.6653 | 5.0459 |  |  | 394 | 15761.86 | 9.6653 | 5.0633 |  |  | 394 | 15761.86 | 9.6653 | 5.0595
395 | 15801.865 | 9.6679 | 5.0456 |  |  | 395 | 15801.865 | 9.6679 | 5.063 |  |  | 395 | 15801.865 | 9.6679 | 5.0592
396 | 15841.87 | 9.6704 | 5.0452 |  |  | 396 | 15841.87 | 9.6704 | 5.0626 |  |  | 396 | 15841.87 | 9.6704 | 5.0588
397 | 15881.874 | 9.6729 | 5.0449 |  |  | 397 | 15881.874 | 9.6729 | 5.0623 |  |  | 397 | 15881.874 | 9.6729 | 5.0585
398 | 15921.879 | 9.6754 | 5.0445 |  |  | 398 | 15921.879 | 9.6754 | 5.0619 |  |  | 398 | 15921.879 | 9.6754 | 5.0581
399 | 15961.884 | 9.678 | 5.0442 |  |  | 399 | 15961.884 | 9.678 | 5.0616 |  |  | 399 | 15961.884 | 9.678 | 5.0578
400 | 16001.889 | 9.6805 | 5.0438 |  |  | 400 | 16001.889 | 9.6805 | 5.0612 |  |  | 400 | 16001.889 | 9.6805 | 5.0575
401 | 16041.893 | 9.683 | 5.0435 |  |  | 401 | 16041.893 | 9.683 | 5.0609 |  |  | 401 | 16041.893 | 9.683 | 5.0571
402 | 16081.898 | 9.6854 | 5.0432 |  |  | 402 | 16081.898 | 9.6854 | 5.0605 |  |  | 402 | 16081.898 | 9.6854 | 5.0568
403 | 16121.903 | 9.6879 | 5.0428 |  |  | 403 | 16121.903 | 9.6879 | 5.0602 |  |  | 403 | 16121.903 | 9.6879 | 5.0565
404 | 16161.908 | 9.6904 | 5.0425 |  |  | 404 | 16161.908 | 9.6904 | 5.0599 |  |  | 404 | 16161.908 | 9.6904 | 5.0561
405 | 16201.912 | 9.6929 | 5.0422 |  |  | 405 | 16201.912 | 9.6929 | 5.0595 |  |  | 405 | 16201.912 | 9.6929 | 5.0558
406 | 16241.917 | 9.6954 | 5.0418 |  |  | 406 | 16241.917 | 9.6954 | 5.0592 |  |  | 406 | 16241.917 | 9.6954 | 5.0555
407 | 16281.922 | 9.6978 | 5.0415 |  |  | 407 | 16281.922 | 9.6978 | 5.0589 |  |  | 407 | 16281.922 | 9.6978 | 5.0552
408 | 16321.926 | 9.7003 | 5.0412 |  |  | 408 | 16321.926 | 9.7003 | 5.0586 |  |  | 408 | 16321.926 | 9.7003 | 5.0548
409 | 16361.931 | 9.7027 | 5.0409 |  |  | 409 | 16361.931 | 9.7027 | 5.0582 |  |  | 409 | 16361.931 | 9.7027 | 5.0545
410 | 16401.936 | 9.7052 | 5.0406 |  |  | 410 | 16401.936 | 9.7052 | 5.0579 |  |  | 410 | 16401.936 | 9.7052 | 5.0542
411 | 16441.941 | 9.7076 | 5.0402 |  |  | 411 | 16441.941 | 9.7076 | 5.0576 |  |  | 411 | 16441.941 | 9.7076 | 5.0539
412 | 16481.945 | 9.71 | 5.0399 |  |  | 412 | 16481.945 | 9.71 | 5.0573 |  |  | 412 | 16481.945 | 9.71 | 5.0536
413 | 16521.95 | 9.7124 | 5.0396 |  |  | 413 | 16521.95 | 9.7124 | 5.057 |  |  | 413 | 16521.95 | 9.7124 | 5.0533
414 | 16561.955 | 9.7149 | 5.0393 |  |  | 414 | 16561.955 | 9.7149 | 5.0566 |  |  | 414 | 16561.955 | 9.7149 | 5.053
415 | 16601.959 | 9.7173 | 5.039 |  |  | 415 | 16601.959 | 9.7173 | 5.0563 |  |  | 415 | 16601.959 | 9.7173 | 5.0526
416 | 16641.964 | 9.7197 | 5.0387 |  |  | 416 | 16641.964 | 9.7197 | 5.056 |  |  | 416 | 16641.964 | 9.7197 | 5.0523
417 | 16681.969 | 9.7221 | 5.0384 |  |  | 417 | 16681.969 | 9.7221 | 5.0557 |  |  | 417 | 16681.969 | 9.7221 | 5.052
418 | 16721.974 | 9.7245 | 5.0381 |  |  | 418 | 16721.974 | 9.7245 | 5.0554 |  |  | 418 | 16721.974 | 9.7245 | 5.0517
419 | 16761.978 | 9.7269 | 5.0378 |  |  | 419 | 16761.978 | 9.7269 | 5.0551 |  |  | 419 | 16761.978 | 9.7269 | 5.0514
420 | 16801.983 | 9.7293 | 5.0375 |  |  | 420 | 16801.983 | 9.7293 | 5.0548 |  |  | 420 | 16801.983 | 9.7293 | 5.0511
421 | 16841.988 | 9.7316 | 5.0372 |  |  | 421 | 16841.988 | 9.7316 | 5.0545 |  |  | 421 | 16841.988 | 9.7316 | 5.0508
422 | 16881.993 | 9.734 | 5.0369 |  |  | 422 | 16881.993 | 9.734 | 5.0542 |  |  | 422 | 16881.993 | 9.734 | 5.0506
423 | 16921.997 | 9.7364 | 5.0366 |  |  | 423 | 16921.997 | 9.7364 | 5.0539 |  |  | 423 | 16921.997 | 9.7364 | 5.0503
424 | 16962.002 | 9.7387 | 5.0363 |  |  | 424 | 16962.002 | 9.7387 | 5.0536 |  |  | 424 | 16962.002 | 9.7387 | 5.05
425 | 17002.007 | 9.7411 | 5.036 |  |  | 425 | 17002.007 | 9.7411 | 5.0533 |  |  | 425 | 17002.007 | 9.7411 | 5.0497
426 | 17042.011 | 9.7434 | 5.0357 |  |  | 426 | 17042.011 | 9.7434 | 5.053 |  |  | 426 | 17042.011 | 9.7434 | 5.0494
427 | 17082.016 | 9.7458 | 5.0355 |  |  | 427 | 17082.016 | 9.7458 | 5.0527 |  |  | 427 | 17082.016 | 9.7458 | 5.0491
428 | 17122.021 | 9.7481 | 5.0352 |  |  | 428 | 17122.021 | 9.7481 | 5.0525 |  |  | 428 | 17122.021 | 9.7481 | 5.0488
429 | 17162.026 | 9.7505 | 5.0349 |  |  | 429 | 17162.026 | 9.7505 | 5.0522 |  |  | 429 | 17162.026 | 9.7505 | 5.0485
430 | 17202.03 | 9.7528 | 5.0346 |  |  | 430 | 17202.03 | 9.7528 | 5.0519 |  |  | 430 | 17202.03 | 9.7528 | 5.0483
431 | 17242.035 | 9.7551 | 5.0343 |  |  | 431 | 17242.035 | 9.7551 | 5.0516 |  |  | 431 | 17242.035 | 9.7551 | 5.048
432 | 17282.04 | 9.7574 | 5.0341 |  |  | 432 | 17282.04 | 9.7574 | 5.0513 |  |  | 432 | 17282.04 | 9.7574 | 5.0477
433 | 17322.044 | 9.7597 | 5.0338 |  |  | 433 | 17322.044 | 9.7597 | 5.051 |  |  | 433 | 17322.044 | 9.7597 | 5.0474
434 | 17362.049 | 9.762 | 5.0335 |  |  | 434 | 17362.049 | 9.762 | 5.0508 |  |  | 434 | 17362.049 | 9.762 | 5.0472
435 | 17402.054 | 9.7643 | 5.0332 |  |  | 435 | 17402.054 | 9.7643 | 5.0505 |  |  | 435 | 17402.054 | 9.7643 | 5.0469
436 | 17442.059 | 9.7666 | 5.033 |  |  | 436 | 17442.059 | 9.7666 | 5.0502 |  |  | 436 | 17442.059 | 9.7666 | 5.0466
437 | 17482.063 | 9.7689 | 5.0327 |  |  | 437 | 17482.063 | 9.7689 | 5.0499 |  |  | 437 | 17482.063 | 9.7689 | 5.0464
438 | 17522.068 | 9.7712 | 5.0324 |  |  | 438 | 17522.068 | 9.7712 | 5.0497 |  |  | 438 | 17522.068 | 9.7712 | 5.0461
439 | 17562.073 | 9.7735 | 5.0322 |  |  | 439 | 17562.073 | 9.7735 | 5.0494 |  |  | 439 | 17562.073 | 9.7735 | 5.0458
440 | 17602.078 | 9.7758 | 5.0319 |  |  | 440 | 17602.078 | 9.7758 | 5.0491 |  |  | 440 | 17602.078 | 9.7758 | 5.0456
441 | 17642.082 | 9.778 | 5.0316 |  |  | 441 | 17642.082 | 9.778 | 5.0489 |  |  | 441 | 17642.082 | 9.778 | 5.0453
442 | 17682.087 | 9.7803 | 5.0314 |  |  | 442 | 17682.087 | 9.7803 | 5.0486 |  |  | 442 | 17682.087 | 9.7803 | 5.0451
443 | 17722.092 | 9.7826 | 5.0311 |  |  | 443 | 17722.092 | 9.7826 | 5.0483 |  |  | 443 | 17722.092 | 9.7826 | 5.0448
444 | 17762.096 | 9.7848 | 5.0309 |  |  | 444 | 17762.096 | 9.7848 | 5.0481 |  |  | 444 | 17762.096 | 9.7848 | 5.0445
445 | 17802.101 | 9.7871 | 5.0306 |  |  | 445 | 17802.101 | 9.7871 | 5.0478 |  |  | 445 | 17802.101 | 9.7871 | 5.0443
446 | 17842.106 | 9.7893 | 5.0304 |  |  | 446 | 17842.106 | 9.7893 | 5.0476 |  |  | 446 | 17842.106 | 9.7893 | 5.044
447 | 17882.111 | 9.7916 | 5.0301 |  |  | 447 | 17882.111 | 9.7916 | 5.0473 |  |  | 447 | 17882.111 | 9.7916 | 5.0438
448 | 17922.115 | 9.7938 | 5.0299 |  |  | 448 | 17922.115 | 9.7938 | 5.047 |  |  | 448 | 17922.115 | 9.7938 | 5.0435
449 | 17962.12 | 9.796 | 5.0296 |  |  | 449 | 17962.12 | 9.796 | 5.0468 |  |  | 449 | 17962.12 | 9.796 | 5.0433
450 | 18002.125 | 9.7982 | 5.0294 |  |  | 450 | 18002.125 | 9.7982 | 5.0465 |  |  | 450 | 18002.125 | 9.7982 | 5.043
451 | 18042.129 | 9.8005 | 5.0291 |  |  | 451 | 18042.129 | 9.8005 | 5.0463 |  |  | 451 | 18042.129 | 9.8005 | 5.0428
452 | 18082.134 | 9.8027 | 5.0289 |  |  | 452 | 18082.134 | 9.8027 | 5.046 |  |  | 452 | 18082.134 | 9.8027 | 5.0425
453 | 18122.139 | 9.8049 | 5.0286 |  |  | 453 | 18122.139 | 9.8049 | 5.0458 |  |  | 453 | 18122.139 | 9.8049 | 5.0423
454 | 18162.144 | 9.8071 | 5.0284 |  |  | 454 | 18162.144 | 9.8071 | 5.0455 |  |  | 454 | 18162.144 | 9.8071 | 5.0421
455 | 18202.148 | 9.8093 | 5.0281 |  |  | 455 | 18202.148 | 9.8093 | 5.0453 |  |  | 455 | 18202.148 | 9.8093 | 5.0418
456 | 18242.153 | 9.8115 | 5.0279 |  |  | 456 | 18242.153 | 9.8115 | 5.0451 |  |  | 456 | 18242.153 | 9.8115 | 5.0416
457 | 18282.158 | 9.8137 | 5.0277 |  |  | 457 | 18282.158 | 9.8137 | 5.0448 |  |  | 457 | 18282.158 | 9.8137 | 5.0413
458 | 18322.163 | 9.8159 | 5.0274 |  |  | 458 | 18322.163 | 9.8159 | 5.0446 |  |  | 458 | 18322.163 | 9.8159 | 5.0411
459 | 18362.167 | 9.818 | 5.0272 |  |  | 459 | 18362.167 | 9.818 | 5.0443 |  |  | 459 | 18362.167 | 9.818 | 5.0409
460 | 18402.172 | 9.8202 | 5.027 |  |  | 460 | 18402.172 | 9.8202 | 5.0441 |  |  | 460 | 18402.172 | 9.8202 | 5.0406
461 | 18442.177 | 9.8224 | 5.0267 |  |  | 461 | 18442.177 | 9.8224 | 5.0439 |  |  | 461 | 18442.177 | 9.8224 | 5.0404
462 | 18482.181 | 9.8246 | 5.0265 |  |  | 462 | 18482.181 | 9.8246 | 5.0436 |  |  | 462 | 18482.181 | 9.8246 | 5.0402
463 | 18522.186 | 9.8267 | 5.0263 |  |  | 463 | 18522.186 | 9.8267 | 5.0434 |  |  | 463 | 18522.186 | 9.8267 | 5.0399
464 | 18562.191 | 9.8289 | 5.026 |  |  | 464 | 18562.191 | 9.8289 | 5.0431 |  |  | 464 | 18562.191 | 9.8289 | 5.0397
465 | 18602.196 | 9.831 | 5.0258 |  |  | 465 | 18602.196 | 9.831 | 5.0429 |  |  | 465 | 18602.196 | 9.831 | 5.0395
466 | 18642.2 | 9.8332 | 5.0256 |  |  | 466 | 18642.2 | 9.8332 | 5.0427 |  |  | 466 | 18642.2 | 9.8332 | 5.0393
467 | 18682.205 | 9.8353 | 5.0254 |  |  | 467 | 18682.205 | 9.8353 | 5.0425 |  |  | 467 | 18682.205 | 9.8353 | 5.039
468 | 18722.21 | 9.8375 | 5.0251 |  |  | 468 | 18722.21 | 9.8375 | 5.0422 |  |  | 468 | 18722.21 | 9.8375 | 5.0388
469 | 18762.214 | 9.8396 | 5.0249 |  |  | 469 | 18762.214 | 9.8396 | 5.042 |  |  | 469 | 18762.214 | 9.8396 | 5.0386
470 | 18802.219 | 9.8417 | 5.0247 |  |  | 470 | 18802.219 | 9.8417 | 5.0418 |  |  | 470 | 18802.219 | 9.8417 | 5.0384
471 | 18842.224 | 9.8439 | 5.0245 |  |  | 471 | 18842.224 | 9.8439 | 5.0415 |  |  | 471 | 18842.224 | 9.8439 | 5.0382
472 | 18882.229 | 9.846 | 5.0243 |  |  | 472 | 18882.229 | 9.846 | 5.0413 |  |  | 472 | 18882.229 | 9.846 | 5.0379
473 | 18922.233 | 9.8481 | 5.024 |  |  | 473 | 18922.233 | 9.8481 | 5.0411 |  |  | 473 | 18922.233 | 9.8481 | 5.0377
474 | 18962.238 | 9.8502 | 5.0238 |  |  | 474 | 18962.238 | 9.8502 | 5.0409 |  |  | 474 | 18962.238 | 9.8502 | 5.0375
475 | 19002.243 | 9.8523 | 5.0236 |  |  | 475 | 19002.243 | 9.8523 | 5.0407 |  |  | 475 | 19002.243 | 9.8523 | 5.0373
476 | 19042.247 | 9.8544 | 5.0234 |  |  | 476 | 19042.247 | 9.8544 | 5.0404 |  |  | 476 | 19042.247 | 9.8544 | 5.0371
477 | 19082.252 | 9.8565 | 5.0232 |  |  | 477 | 19082.252 | 9.8565 | 5.0402 |  |  | 477 | 19082.252 | 9.8565 | 5.0369
478 | 19122.257 | 9.8586 | 5.023 |  |  | 478 | 19122.257 | 9.8586 | 5.04 |  |  | 478 | 19122.257 | 9.8586 | 5.0367
479 | 19162.262 | 9.8607 | 5.0228 |  |  | 479 | 19162.262 | 9.8607 | 5.0398 |  |  | 479 | 19162.262 | 9.8607 | 5.0364
480 | 19202.266 | 9.8628 | 5.0225 |  |  | 480 | 19202.266 | 9.8628 | 5.0396 |  |  | 480 | 19202.266 | 9.8628 | 5.0362
481 | 19242.271 | 9.8649 | 5.0223 |  |  | 481 | 19242.271 | 9.8649 | 5.0394 |  |  | 481 | 19242.271 | 9.8649 | 5.036
482 | 19282.276 | 9.8669 | 5.0221 |  |  | 482 | 19282.276 | 9.8669 | 5.0392 |  |  | 482 | 19282.276 | 9.8669 | 5.0358
483 | 19322.281 | 9.869 | 5.0219 |  |  | 483 | 19322.281 | 9.869 | 5.0389 |  |  | 483 | 19322.281 | 9.869 | 5.0356
484 | 19362.285 | 9.8711 | 5.0217 |  |  | 484 | 19362.285 | 9.8711 | 5.0387 |  |  | 484 | 19362.285 | 9.8711 | 5.0354
485 | 19402.29 | 9.8731 | 5.0215 |  |  | 485 | 19402.29 | 9.8731 | 5.0385 |  |  | 485 | 19402.29 | 9.8731 | 5.0352
486 | 19442.295 | 9.8752 | 5.0213 |  |  | 486 | 19442.295 | 9.8752 | 5.0383 |  |  | 486 | 19442.295 | 9.8752 | 5.035
487 | 19482.299 | 9.8773 | 5.0211 |  |  | 487 | 19482.299 | 9.8773 | 5.0381 |  |  | 487 | 19482.299 | 9.8773 | 5.0348
488 | 19522.304 | 9.8793 | 5.0209 |  |  | 488 | 19522.304 | 9.8793 | 5.0379 |  |  | 488 | 19522.304 | 9.8793 | 5.0346
489 | 19562.309 | 9.8814 | 5.0207 |  |  | 489 | 19562.309 | 9.8814 | 5.0377 |  |  | 489 | 19562.309 | 9.8814 | 5.0344
490 | 19602.314 | 9.8834 | 5.0205 |  |  | 490 | 19602.314 | 9.8834 | 5.0375 |  |  | 490 | 19602.314 | 9.8834 | 5.0342
491 | 19642.318 | 9.8854 | 5.0203 |  |  | 491 | 19642.318 | 9.8854 | 5.0373 |  |  | 491 | 19642.318 | 9.8854 | 5.034
492 | 19682.323 | 9.8875 | 5.0201 |  |  | 492 | 19682.323 | 9.8875 | 5.0371 |  |  | 492 | 19682.323 | 9.8875 | 5.0338
493 | 19722.328 | 9.8895 | 5.0199 |  |  | 493 | 19722.328 | 9.8895 | 5.0369 |  |  | 493 | 19722.328 | 9.8895 | 5.0336
494 | 19762.332 | 9.8915 | 5.0197 |  |  | 494 | 19762.332 | 9.8915 | 5.0367 |  |  | 494 | 19762.332 | 9.8915 | 5.0334
495 | 19802.337 | 9.8936 | 5.0195 |  |  | 495 | 19802.337 | 9.8936 | 5.0365 |  |  | 495 | 19802.337 | 9.8936 | 5.0332
496 | 19842.342 | 9.8956 | 5.0194 |  |  | 496 | 19842.342 | 9.8956 | 5.0363 |  |  | 496 | 19842.342 | 9.8956 | 5.033
497 | 19882.347 | 9.8976 | 5.0192 |  |  | 497 | 19882.347 | 9.8976 | 5.0361 |  |  | 497 | 19882.347 | 9.8976 | 5.0328
498 | 19922.351 | 9.8996 | 5.019 |  |  | 498 | 19922.351 | 9.8996 | 5.0359 |  |  | 498 | 19922.351 | 9.8996 | 5.0327
499 | 19962.356 | 9.9016 | 5.0188 |  |  | 499 | 19962.356 | 9.9016 | 5.0357 |  |  | 499 | 19962.356 | 9.9016 | 5.0325
500 | 20002.361 | 9.9036 | 5.0186 |  |  | 500 | 20002.361 | 9.9036 | 5.0355 |  |  | 500 | 20002.361 | 9.9036 | 5.0323
501 | 20042.366 | 9.9056 | 5.0184 |  |  | 501 | 20042.366 | 9.9056 | 5.0353 |  |  | 501 | 20042.366 | 9.9056 | 5.0321
502 | 20082.37 | 9.9076 | 5.0182 |  |  | 502 | 20082.37 | 9.9076 | 5.0351 |  |  | 502 | 20082.37 | 9.9076 | 5.0319
503 | 20122.375 | 9.9096 | 5.018 |  |  | 503 | 20122.375 | 9.9096 | 5.0349 |  |  | 503 | 20122.375 | 9.9096 | 5.0317
504 | 20162.38 | 9.9116 | 5.0179 |  |  | 504 | 20162.38 | 9.9116 | 5.0347 |  |  | 504 | 20162.38 | 9.9116 | 5.0315
505 | 20202.384 | 9.9136 | 5.0177 |  |  | 505 | 20202.384 | 9.9136 | 5.0346 |  |  | 505 | 20202.384 | 9.9136 | 5.0313
506 | 20242.389 | 9.9155 | 5.0175 |  |  | 506 | 20242.389 | 9.9155 | 5.0344 |  |  | 506 | 20242.389 | 9.9155 | 5.0312
507 | 20282.394 | 9.9175 | 5.0173 |  |  | 507 | 20282.394 | 9.9175 | 5.0342 |  |  | 507 | 20282.394 | 9.9175 | 5.031
508 | 20322.399 | 9.9195 | 5.0171 |  |  | 508 | 20322.399 | 9.9195 | 5.034 |  |  | 508 | 20322.399 | 9.9195 | 5.0308
509 | 20362.403 | 9.9214 | 5.0169 |  |  | 509 | 20362.403 | 9.9214 | 5.0338 |  |  | 509 | 20362.403 | 9.9214 | 5.0306
510 | 20402.408 | 9.9234 | 5.0168 |  |  | 510 | 20402.408 | 9.9234 | 5.0336 |  |  | 510 | 20402.408 | 9.9234 | 5.0304
511 | 20442.413 | 9.9254 | 5.0166 |  |  | 511 | 20442.413 | 9.9254 | 5.0334 |  |  | 511 | 20442.413 | 9.9254 | 5.0303
512 | 20482.417 | 9.9273 | 5.0164 |  |  | 512 | 20482.417 | 9.9273 | 5.0333 |  |  | 512 | 20482.417 | 9.9273 | 5.0301
513 | 20522.422 | 9.9293 | 5.0162 |  |  | 513 | 20522.422 | 9.9293 | 5.0331 |  |  | 513 | 20522.422 | 9.9293 | 5.0299
514 | 20562.427 | 9.9312 | 5.0161 |  |  | 514 | 20562.427 | 9.9312 | 5.0329 |  |  | 514 | 20562.427 | 9.9312 | 5.0297
515 | 20602.432 | 9.9332 | 5.0159 |  |  | 515 | 20602.432 | 9.9332 | 5.0327 |  |  | 515 | 20602.432 | 9.9332 | 5.0296
516 | 20642.436 | 9.9351 | 5.0157 |  |  | 516 | 20642.436 | 9.9351 | 5.0325 |  |  | 516 | 20642.436 | 9.9351 | 5.0294
517 | 20682.441 | 9.937 | 5.0155 |  |  | 517 | 20682.441 | 9.937 | 5.0324 |  |  | 517 | 20682.441 | 9.937 | 5.0292
518 | 20722.446 | 9.939 | 5.0154 |  |  | 518 | 20722.446 | 9.939 | 5.0322 |  |  | 518 | 20722.446 | 9.939 | 5.029
519 | 20762.451 | 9.9409 | 5.0152 |  |  | 519 | 20762.451 | 9.9409 | 5.032 |  |  | 519 | 20762.451 | 9.9409 | 5.0289
520 | 20802.455 | 9.9428 | 5.015 |  |  | 520 | 20802.455 | 9.9428 | 5.0318 |  |  | 520 | 20802.455 | 9.9428 | 5.0287
521 | 20842.46 | 9.9447 | 5.0149 |  |  | 521 | 20842.46 | 9.9447 | 5.0316 |  |  | 521 | 20842.46 | 9.9447 | 5.0285
522 | 20882.465 | 9.9467 | 5.0147 |  |  | 522 | 20882.465 | 9.9467 | 5.0315 |  |  | 522 | 20882.465 | 9.9467 | 5.0284
523 | 20922.469 | 9.9486 | 5.0145 |  |  | 523 | 20922.469 | 9.9486 | 5.0313 |  |  | 523 | 20922.469 | 9.9486 | 5.0282
524 | 20962.474 | 9.9505 | 5.0143 |  |  | 524 | 20962.474 | 9.9505 | 5.0311 |  |  | 524 | 20962.474 | 9.9505 | 5.028
525 | 21002.479 | 9.9524 | 5.0142 |  |  | 525 | 21002.479 | 9.9524 | 5.0309 |  |  | 525 | 21002.479 | 9.9524 | 5.0278
526 | 21042.484 | 9.9543 | 5.014 |  |  | 526 | 21042.484 | 9.9543 | 5.0308 |  |  | 526 | 21042.484 | 9.9543 | 5.0277
527 | 21082.488 | 9.9562 | 5.0139 |  |  | 527 | 21082.488 | 9.9562 | 5.0306 |  |  | 527 | 21082.488 | 9.9562 | 5.0275
528 | 21122.493 | 9.9581 | 5.0137 |  |  | 528 | 21122.493 | 9.9581 | 5.0304 |  |  | 528 | 21122.493 | 9.9581 | 5.0274
529 | 21162.498 | 9.96 | 5.0135 |  |  | 529 | 21162.498 | 9.96 | 5.0303 |  |  | 529 | 21162.498 | 9.96 | 5.0272
530 | 21202.502 | 9.9619 | 5.0134 |  |  | 530 | 21202.502 | 9.9619 | 5.0301 |  |  | 530 | 21202.502 | 9.9619 | 5.027
531 | 21242.507 | 9.9638 | 5.0132 |  |  | 531 | 21242.507 | 9.9638 | 5.0299 |  |  | 531 | 21242.507 | 9.9638 | 5.0269
532 | 21282.512 | 9.9656 | 5.013 |  |  | 532 | 21282.512 | 9.9656 | 5.0298 |  |  | 532 | 21282.512 | 9.9656 | 5.0267
533 | 21322.517 | 9.9675 | 5.0129 |  |  | 533 | 21322.517 | 9.9675 | 5.0296 |  |  | 533 | 21322.517 | 9.9675 | 5.0265
534 | 21362.521 | 9.9694 | 5.0127 |  |  | 534 | 21362.521 | 9.9694 | 5.0294 |  |  | 534 | 21362.521 | 9.9694 | 5.0264
535 | 21402.526 | 9.9713 | 5.0126 |  |  | 535 | 21402.526 | 9.9713 | 5.0293 |  |  | 535 | 21402.526 | 9.9713 | 5.0262
536 | 21442.531 | 9.9731 | 5.0124 |  |  | 536 | 21442.531 | 9.9731 | 5.0291 |  |  | 536 | 21442.531 | 9.9731 | 5.0261
537 | 21482.536 | 9.975 | 5.0122 |  |  | 537 | 21482.536 | 9.975 | 5.0289 |  |  | 537 | 21482.536 | 9.975 | 5.0259
538 | 21522.54 | 9.9769 | 5.0121 |  |  | 538 | 21522.54 | 9.9769 | 5.0288 |  |  | 538 | 21522.54 | 9.9769 | 5.0257
539 | 21562.545 | 9.9787 | 5.0119 |  |  | 539 | 21562.545 | 9.9787 | 5.0286 |  |  | 539 | 21562.545 | 9.9787 | 5.0256
540 | 21602.55 | 9.9806 | 5.0118 |  |  | 540 | 21602.55 | 9.9806 | 5.0284 |  |  | 540 | 21602.55 | 9.9806 | 5.0254
541 | 21642.554 | 9.9824 | 5.0116 |  |  | 541 | 21642.554 | 9.9824 | 5.0283 |  |  | 541 | 21642.554 | 9.9824 | 5.0253
542 | 21682.559 | 9.9843 | 5.0115 |  |  | 542 | 21682.559 | 9.9843 | 5.0281 |  |  | 542 | 21682.559 | 9.9843 | 5.0251
543 | 21722.564 | 9.9861 | 5.0113 |  |  | 543 | 21722.564 | 9.9861 | 5.028 |  |  | 543 | 21722.564 | 9.9861 | 5.025
544 | 21762.569 | 9.9879 | 5.0112 |  |  | 544 | 21762.569 | 9.9879 | 5.0278 |  |  | 544 | 21762.569 | 9.9879 | 5.0248
545 | 21802.573 | 9.9898 | 5.011 |  |  | 545 | 21802.573 | 9.9898 | 5.0276 |  |  | 545 | 21802.573 | 9.9898 | 5.0247
546 | 21842.578 | 9.9916 | 5.0109 |  |  | 546 | 21842.578 | 9.9916 | 5.0275 |  |  | 546 | 21842.578 | 9.9916 | 5.0245
547 | 21882.583 | 9.9934 | 5.0107 |  |  | 547 | 21882.583 | 9.9934 | 5.0273 |  |  | 547 | 21882.583 | 9.9934 | 5.0243
548 | 21922.587 | 9.9953 | 5.0106 |  |  | 548 | 21922.587 | 9.9953 | 5.0272 |  |  | 548 | 21922.587 | 9.9953 | 5.0242
549 | 21962.592 | 9.9971 | 5.0104 |  |  | 549 | 21962.592 | 9.9971 | 5.027 |  |  | 549 | 21962.592 | 9.9971 | 5.024
550 | 22002.597 | 9.9989 | 5.0103 |  |  | 550 | 22002.597 | 9.9989 | 5.0269 |  |  | 550 | 22002.597 | 9.9989 | 5.0239
551 | 22042.602 | 10.0007 | 5.0101 |  |  | 551 | 22042.602 | 10.0007 | 5.0267 |  |  | 551 | 22042.602 | 10.0007 | 5.0237
552 | 22082.606 | 10.0025 | 5.01 |  |  | 552 | 22082.606 | 10.0025 | 5.0265 |  |  | 552 | 22082.606 | 10.0025 | 5.0236
553 | 22122.611 | 10.0044 | 5.0098 |  |  | 553 | 22122.611 | 10.0044 | 5.0264 |  |  | 553 | 22122.611 | 10.0044 | 5.0234
554 | 22162.616 | 10.0062 | 5.0097 |  |  | 554 | 22162.616 | 10.0062 | 5.0262 |  |  | 554 | 22162.616 | 10.0062 | 5.0233
555 | 22202.621 | 10.008 | 5.0095 |  |  | 555 | 22202.621 | 10.008 | 5.0261 |  |  | 555 | 22202.621 | 10.008 | 5.0232
556 | 22242.625 | 10.0098 | 5.0094 |  |  | 556 | 22242.625 | 10.0098 | 5.0259 |  |  | 556 | 22242.625 | 10.0098 | 5.023
557 | 22282.63 | 10.0116 | 5.0092 |  |  | 557 | 22282.63 | 10.0116 | 5.0258 |  |  | 557 | 22282.63 | 10.0116 | 5.0229
558 | 22322.635 | 10.0134 | 5.0091 |  |  | 558 | 22322.635 | 10.0134 | 5.0256 |  |  | 558 | 22322.635 | 10.0134 | 5.0227
559 | 22362.639 | 10.0151 | 5.0089 |  |  | 559 | 22362.639 | 10.0151 | 5.0255 |  |  | 559 | 22362.639 | 10.0151 | 5.0226
560 | 22402.644 | 10.0169 | 5.0088 |  |  | 560 | 22402.644 | 10.0169 | 5.0253 |  |  | 560 | 22402.644 | 10.0169 | 5.0224
561 | 22442.649 | 10.0187 | 5.0087 |  |  | 561 | 22442.649 | 10.0187 | 5.0252 |  |  | 561 | 22442.649 | 10.0187 | 5.0223
562 | 22482.654 | 10.0205 | 5.0085 |  |  | 562 | 22482.654 | 10.0205 | 5.025 |  |  | 562 | 22482.654 | 10.0205 | 5.0221
563 | 22522.658 | 10.0223 | 5.0084 |  |  | 563 | 22522.658 | 10.0223 | 5.0249 |  |  | 563 | 22522.658 | 10.0223 | 5.022
564 | 22562.663 | 10.0241 | 5.0082 |  |  | 564 | 22562.663 | 10.0241 | 5.0247 |  |  | 564 | 22562.663 | 10.0241 | 5.0218
565 | 22602.668 | 10.0258 | 5.0081 |  |  | 565 | 22602.668 | 10.0258 | 5.0246 |  |  | 565 | 22602.668 | 10.0258 | 5.0217
566 | 22642.672 | 10.0276 | 5.0079 |  |  | 566 | 22642.672 | 10.0276 | 5.0244 |  |  | 566 | 22642.672 | 10.0276 | 5.0216
567 | 22682.677 | 10.0294 | 5.0078 |  |  | 567 | 22682.677 | 10.0294 | 5.0243 |  |  | 567 | 22682.677 | 10.0294 | 5.0214
568 | 22722.682 | 10.0311 | 5.0077 |  |  | 568 | 22722.682 | 10.0311 | 5.0242 |  |  | 568 | 22722.682 | 10.0311 | 5.0213
569 | 22762.687 | 10.0329 | 5.0075 |  |  | 569 | 22762.687 | 10.0329 | 5.024 |  |  | 569 | 22762.687 | 10.0329 | 5.0211
570 | 22802.691 | 10.0346 | 5.0074 |  |  | 570 | 22802.691 | 10.0346 | 5.0239 |  |  | 570 | 22802.691 | 10.0346 | 5.021
571 | 22842.696 | 10.0364 | 5.0073 |  |  | 571 | 22842.696 | 10.0364 | 5.0237 |  |  | 571 | 22842.696 | 10.0364 | 5.0209
572 | 22882.701 | 10.0381 | 5.0071 |  |  | 572 | 22882.701 | 10.0381 | 5.0236 |  |  | 572 | 22882.701 | 10.0381 | 5.0207
573 | 22922.705 | 10.0399 | 5.007 |  |  | 573 | 22922.705 | 10.0399 | 5.0234 |  |  | 573 | 22922.705 | 10.0399 | 5.0206
574 | 22962.71 | 10.0416 | 5.0068 |  |  | 574 | 22962.71 | 10.0416 | 5.0233 |  |  | 574 | 22962.71 | 10.0416 | 5.0205
575 | 23002.715 | 10.0434 | 5.0067 |  |  | 575 | 23002.715 | 10.0434 | 5.0231 |  |  | 575 | 23002.715 | 10.0434 | 5.0203
576 | 23042.72 | 10.0451 | 5.0066 |  |  | 576 | 23042.72 | 10.0451 | 5.023 |  |  | 576 | 23042.72 | 10.0451 | 5.0202
577 | 23082.724 | 10.0468 | 5.0064 |  |  | 577 | 23082.724 | 10.0468 | 5.0229 |  |  | 577 | 23082.724 | 10.0468 | 5.02
578 | 23122.729 | 10.0486 | 5.0063 |  |  | 578 | 23122.729 | 10.0486 | 5.0227 |  |  | 578 | 23122.729 | 10.0486 | 5.0199
579 | 23162.734 | 10.0503 | 5.0062 |  |  | 579 | 23162.734 | 10.0503 | 5.0226 |  |  | 579 | 23162.734 | 10.0503 | 5.0198
580 | 23202.739 | 10.052 | 5.006 |  |  | 580 | 23202.739 | 10.052 | 5.0224 |  |  | 580 | 23202.739 | 10.052 | 5.0196
581 | 23242.743 | 10.0537 | 5.0059 |  |  | 581 | 23242.743 | 10.0537 | 5.0223 |  |  | 581 | 23242.743 | 10.0537 | 5.0195
582 | 23282.748 | 10.0555 | 5.0058 |  |  | 582 | 23282.748 | 10.0555 | 5.0222 |  |  | 582 | 23282.748 | 10.0555 | 5.0194
583 | 23322.753 | 10.0572 | 5.0056 |  |  | 583 | 23322.753 | 10.0572 | 5.022 |  |  | 583 | 23322.753 | 10.0572 | 5.0192
584 | 23362.757 | 10.0589 | 5.0055 |  |  | 584 | 23362.757 | 10.0589 | 5.0219 |  |  | 584 | 23362.757 | 10.0589 | 5.0191
585 | 23402.762 | 10.0606 | 5.0054 |  |  | 585 | 23402.762 | 10.0606 | 5.0218 |  |  | 585 | 23402.762 | 10.0606 | 5.019
586 | 23442.767 | 10.0623 | 5.0053 |  |  | 586 | 23442.767 | 10.0623 | 5.0216 |  |  | 586 | 23442.767 | 10.0623 | 5.0188
587 | 23482.772 | 10.064 | 5.0051 |  |  | 587 | 23482.772 | 10.064 | 5.0215 |  |  | 587 | 23482.772 | 10.064 | 5.0187
588 | 23522.776 | 10.0657 | 5.005 |  |  | 588 | 23522.776 | 10.0657 | 5.0213 |  |  | 588 | 23522.776 | 10.0657 | 5.0186
589 | 23562.781 | 10.0674 | 5.0049 |  |  | 589 | 23562.781 | 10.0674 | 5.0212 |  |  | 589 | 23562.781 | 10.0674 | 5.0184
590 | 23602.786 | 10.0691 | 5.0047 |  |  | 590 | 23602.786 | 10.0691 | 5.0211 |  |  | 590 | 23602.786 | 10.0691 | 5.0183
591 | 23642.79 | 10.0708 | 5.0046 |  |  | 591 | 23642.79 | 10.0708 | 5.0209 |  |  | 591 | 23642.79 | 10.0708 | 5.0182
592 | 23682.795 | 10.0725 | 5.0045 |  |  | 592 | 23682.795 | 10.0725 | 5.0208 |  |  | 592 | 23682.795 | 10.0725 | 5.0181
593 | 23722.8 | 10.0742 | 5.0044 |  |  | 593 | 23722.8 | 10.0742 | 5.0207 |  |  | 593 | 23722.8 | 10.0742 | 5.0179
594 | 23762.805 | 10.0759 | 5.0042 |  |  | 594 | 23762.805 | 10.0759 | 5.0205 |  |  | 594 | 23762.805 | 10.0759 | 5.0178
595 | 23802.809 | 10.0776 | 5.0041 |  |  | 595 | 23802.809 | 10.0776 | 5.0204 |  |  | 595 | 23802.809 | 10.0776 | 5.0177
596 | 23842.814 | 10.0792 | 5.004 |  |  | 596 | 23842.814 | 10.0792 | 5.0203 |  |  | 596 | 23842.814 | 10.0792 | 5.0175
597 | 23882.819 | 10.0809 | 5.0039 |  |  | 597 | 23882.819 | 10.0809 | 5.0201 |  |  | 597 | 23882.819 | 10.0809 | 5.0174
598 | 23922.824 | 10.0826 | 5.0037 |  |  | 598 | 23922.824 | 10.0826 | 5.02 |  |  | 598 | 23922.824 | 10.0826 | 5.0173
599 | 23962.828 | 10.0843 | 5.0036 |  |  | 599 | 23962.828 | 10.0843 | 5.0199 |  |  | 599 | 23962.828 | 10.0843 | 5.0172
600 | 24002.833 | 10.0859 | 5.0035 |  |  | 600 | 24002.833 | 10.0859 | 5.0197 |  |  | 600 | 24002.833 | 10.0859 | 5.017
601 | 24042.838 | 10.0876 | 5.0034 |  |  | 601 | 24042.838 | 10.0876 | 5.0196 |  |  | 601 | 24042.838 | 10.0876 | 5.0169
602 | 24082.842 | 10.0893 | 5.0032 |  |  | 602 | 24082.842 | 10.0893 | 5.0195 |  |  | 602 | 24082.842 | 10.0893 | 5.0168
603 | 24122.847 | 10.0909 | 5.0031 |  |  | 603 | 24122.847 | 10.0909 | 5.0194 |  |  | 603 | 24122.847 | 10.0909 | 5.0167
604 | 24162.852 | 10.0926 | 5.003 |  |  | 604 | 24162.852 | 10.0926 | 5.0192 |  |  | 604 | 24162.852 | 10.0926 | 5.0165
605 | 24202.857 | 10.0942 | 5.0029 |  |  | 605 | 24202.857 | 10.0942 | 5.0191 |  |  | 605 | 24202.857 | 10.0942 | 5.0164
606 | 24242.861 | 10.0959 | 5.0027 |  |  | 606 | 24242.861 | 10.0959 | 5.019 |  |  | 606 | 24242.861 | 10.0959 | 5.0163
607 | 24282.866 | 10.0975 | 5.0026 |  |  | 607 | 24282.866 | 10.0975 | 5.0188 |  |  | 607 | 24282.866 | 10.0975 | 5.0162
608 | 24322.871 | 10.0992 | 5.0025 |  |  | 608 | 24322.871 | 10.0992 | 5.0187 |  |  | 608 | 24322.871 | 10.0992 | 5.016
609 | 24362.875 | 10.1008 | 5.0024 |  |  | 609 | 24362.875 | 10.1008 | 5.0186 |  |  | 609 | 24362.875 | 10.1008 | 5.0159
610 | 24402.88 | 10.1025 | 5.0023 |  |  | 610 | 24402.88 | 10.1025 | 5.0185 |  |  | 610 | 24402.88 | 10.1025 | 5.0158
611 | 24442.885 | 10.1041 | 5.0021 |  |  | 611 | 24442.885 | 10.1041 | 5.0183 |  |  | 611 | 24442.885 | 10.1041 | 5.0157
612 | 24482.89 | 10.1057 | 5.002 |  |  | 612 | 24482.89 | 10.1057 | 5.0182 |  |  | 612 | 24482.89 | 10.1057 | 5.0155
613 | 24522.894 | 10.1074 | 5.0019 |  |  | 613 | 24522.894 | 10.1074 | 5.0181 |  |  | 613 | 24522.894 | 10.1074 | 5.0154
614 | 24562.899 | 10.109 | 5.0018 |  |  | 614 | 24562.899 | 10.109 | 5.0179 |  |  | 614 | 24562.899 | 10.109 | 5.0153
615 | 24602.904 | 10.1106 | 5.0017 |  |  | 615 | 24602.904 | 10.1106 | 5.0178 |  |  | 615 | 24602.904 | 10.1106 | 5.0152
616 | 24642.909 | 10.1122 | 5.0015 |  |  | 616 | 24642.909 | 10.1122 | 5.0177 |  |  | 616 | 24642.909 | 10.1122 | 5.0151
617 | 24682.913 | 10.1139 | 5.0014 |  |  | 617 | 24682.913 | 10.1139 | 5.0176 |  |  | 617 | 24682.913 | 10.1139 | 5.0149
618 | 24722.918 | 10.1155 | 5.0013 |  |  | 618 | 24722.918 | 10.1155 | 5.0174 |  |  | 618 | 24722.918 | 10.1155 | 5.0148
619 | 24762.923 | 10.1171 | 5.0012 |  |  | 619 | 24762.923 | 10.1171 | 5.0173 |  |  | 619 | 24762.923 | 10.1171 | 5.0147
620 | 24802.927 | 10.1187 | 5.0011 |  |  | 620 | 24802.927 | 10.1187 | 5.0172 |  |  | 620 | 24802.927 | 10.1187 | 5.0146
621 | 24842.932 | 10.1203 | 5.001 |  |  | 621 | 24842.932 | 10.1203 | 5.0171 |  |  | 621 | 24842.932 | 10.1203 | 5.0145
622 | 24882.937 | 10.1219 | 5.0008 |  |  | 622 | 24882.937 | 10.1219 | 5.0169 |  |  | 622 | 24882.937 | 10.1219 | 5.0143
623 | 24922.942 | 10.1235 | 5.0007 |  |  | 623 | 24922.942 | 10.1235 | 5.0168 |  |  | 623 | 24922.942 | 10.1235 | 5.0142
624 | 24962.946 | 10.1251 | 5.0006 |  |  | 624 | 24962.946 | 10.1251 | 5.0167 |  |  | 624 | 24962.946 | 10.1251 | 5.0141
625 | 25002.951 | 10.1267 | 5.0005 |  |  | 625 | 25002.951 | 10.1267 | 5.0166 |  |  | 625 | 25002.951 | 10.1267 | 5.014
626 | 25042.956 | 10.1283 | 5.0004 |  |  | 626 | 25042.956 | 10.1283 | 5.0165 |  |  | 626 | 25042.956 | 10.1283 | 5.0139
627 | 25082.96 | 10.1299 | 5.0003 |  |  | 627 | 25082.96 | 10.1299 | 5.0163 |  |  | 627 | 25082.96 | 10.1299 | 5.0138
628 | 25122.965 | 10.1315 | 5.0002 |  |  | 628 | 25122.965 | 10.1315 | 5.0162 |  |  | 628 | 25122.965 | 10.1315 | 5.0136
629 | 25162.97 | 10.1331 | 5 |  |  | 629 | 25162.97 | 10.1331 | 5.0161 |  |  | 629 | 25162.97 | 10.1331 | 5.0135
630 | 25202.975 | 10.1347 | 4.9999 |  |  | 630 | 25202.975 | 10.1347 | 5.016 |  |  | 630 | 25202.975 | 10.1347 | 5.0134
631 | 25242.979 | 10.1363 | 4.9998 |  |  | 631 | 25242.979 | 10.1363 | 5.0158 |  |  | 631 | 25242.979 | 10.1363 | 5.0133
632 | 25282.984 | 10.1379 | 4.9997 |  |  | 632 | 25282.984 | 10.1379 | 5.0157 |  |  | 632 | 25282.984 | 10.1379 | 5.0132
633 | 25322.989 | 10.1395 | 4.9996 |  |  | 633 | 25322.989 | 10.1395 | 5.0156 |  |  | 633 | 25322.989 | 10.1395 | 5.0131
634 | 25362.994 | 10.141 | 4.9995 |  |  | 634 | 25362.994 | 10.141 | 5.0155 |  |  | 634 | 25362.994 | 10.141 | 5.0129
635 | 25402.998 | 10.1426 | 4.9994 |  |  | 635 | 25402.998 | 10.1426 | 5.0154 |  |  | 635 | 25402.998 | 10.1426 | 5.0128
636 | 25443.003 | 10.1442 | 4.9992 |  |  | 636 | 25443.003 | 10.1442 | 5.0152 |  |  | 636 | 25443.003 | 10.1442 | 5.0127
637 | 25483.008 | 10.1458 | 4.9991 |  |  | 637 | 25483.008 | 10.1458 | 5.0151 |  |  | 637 | 25483.008 | 10.1458 | 5.0126
638 | 25523.012 | 10.1473 | 4.999 |  |  | 638 | 25523.012 | 10.1473 | 5.015 |  |  | 638 | 25523.012 | 10.1473 | 5.0125
639 | 25563.017 | 10.1489 | 4.9989 |  |  | 639 | 25563.017 | 10.1489 | 5.0149 |  |  | 639 | 25563.017 | 10.1489 | 5.0124
640 | 25603.022 | 10.1505 | 4.9988 |  |  | 640 | 25603.022 | 10.1505 | 5.0148 |  |  | 640 | 25603.022 | 10.1505 | 5.0122
641 | 25643.027 | 10.152 | 4.9987 |  |  | 641 | 25643.027 | 10.152 | 5.0147 |  |  | 641 | 25643.027 | 10.152 | 5.0121
642 | 25683.031 | 10.1536 | 4.9986 |  |  | 642 | 25683.031 | 10.1536 | 5.0145 |  |  | 642 | 25683.031 | 10.1536 | 5.012
643 | 25723.036 | 10.1551 | 4.9985 |  |  | 643 | 25723.036 | 10.1551 | 5.0144 |  |  | 643 | 25723.036 | 10.1551 | 5.0119
644 | 25763.041 | 10.1567 | 4.9984 |  |  | 644 | 25763.041 | 10.1567 | 5.0143 |  |  | 644 | 25763.041 | 10.1567 | 5.0118
645 | 25803.045 | 10.1582 | 4.9983 |  |  | 645 | 25803.045 | 10.1582 | 5.0142 |  |  | 645 | 25803.045 | 10.1582 | 5.0117
646 | 25843.05 | 10.1598 | 4.9981 |  |  | 646 | 25843.05 | 10.1598 | 5.0141 |  |  | 646 | 25843.05 | 10.1598 | 5.0116
647 | 25883.055 | 10.1613 | 4.998 |  |  | 647 | 25883.055 | 10.1613 | 5.0139 |  |  | 647 | 25883.055 | 10.1613 | 5.0115
648 | 25923.06 | 10.1629 | 4.9979 |  |  | 648 | 25923.06 | 10.1629 | 5.0138 |  |  | 648 | 25923.06 | 10.1629 | 5.0113
649 | 25963.064 | 10.1644 | 4.9978 |  |  | 649 | 25963.064 | 10.1644 | 5.0137 |  |  | 649 | 25963.064 | 10.1644 | 5.0112
650 | 26003.069 | 10.166 | 4.9977 |  |  | 650 | 26003.069 | 10.166 | 5.0136 |  |  | 650 | 26003.069 | 10.166 | 5.0111
651 | 26043.074 | 10.1675 | 4.9976 |  |  | 651 | 26043.074 | 10.1675 | 5.0135 |  |  | 651 | 26043.074 | 10.1675 | 5.011
652 | 26083.079 | 10.169 | 4.9975 |  |  | 652 | 26083.079 | 10.169 | 5.0134 |  |  | 652 | 26083.079 | 10.169 | 5.0109
653 | 26123.083 | 10.1706 | 4.9974 |  |  | 653 | 26123.083 | 10.1706 | 5.0133 |  |  | 653 | 26123.083 | 10.1706 | 5.0108
654 | 26163.088 | 10.1721 | 4.9973 |  |  | 654 | 26163.088 | 10.1721 | 5.0131 |  |  | 654 | 26163.088 | 10.1721 | 5.0107
655 | 26203.093 | 10.1736 | 4.9972 |  |  | 655 | 26203.093 | 10.1736 | 5.013 |  |  | 655 | 26203.093 | 10.1736 | 5.0106
656 | 26243.097 | 10.1752 | 4.9971 |  |  | 656 | 26243.097 | 10.1752 | 5.0129 |  |  | 656 | 26243.097 | 10.1752 | 5.0105
657 | 26283.102 | 10.1767 | 4.997 |  |  | 657 | 26283.102 | 10.1767 | 5.0128 |  |  | 657 | 26283.102 | 10.1767 | 5.0103
658 | 26323.107 | 10.1782 | 4.9968 |  |  | 658 | 26323.107 | 10.1782 | 5.0127 |  |  | 658 | 26323.107 | 10.1782 | 5.0102
659 | 26363.112 | 10.1797 | 4.9967 |  |  | 659 | 26363.112 | 10.1797 | 5.0126 |  |  | 659 | 26363.112 | 10.1797 | 5.0101
660 | 26403.116 | 10.1812 | 4.9966 |  |  | 660 | 26403.116 | 10.1812 | 5.0124 |  |  | 660 | 26403.116 | 10.1812 | 5.01
661 | 26443.121 | 10.1828 | 4.9965 |  |  | 661 | 26443.121 | 10.1828 | 5.0123 |  |  | 661 | 26443.121 | 10.1828 | 5.0099
662 | 26483.126 | 10.1843 | 4.9964 |  |  | 662 | 26483.126 | 10.1843 | 5.0122 |  |  | 662 | 26483.126 | 10.1843 | 5.0098
663 | 26523.13 | 10.1858 | 4.9963 |  |  | 663 | 26523.13 | 10.1858 | 5.0121 |  |  | 663 | 26523.13 | 10.1858 | 5.0097
664 | 26563.135 | 10.1873 | 4.9962 |  |  | 664 | 26563.135 | 10.1873 | 5.012 |  |  | 664 | 26563.135 | 10.1873 | 5.0096
665 | 26603.14 | 10.1888 | 4.9961 |  |  | 665 | 26603.14 | 10.1888 | 5.0119 |  |  | 665 | 26603.14 | 10.1888 | 5.0095
666 | 26643.145 | 10.1903 | 4.996 |  |  | 666 | 26643.145 | 10.1903 | 5.0118 |  |  | 666 | 26643.145 | 10.1903 | 5.0094
667 | 26683.149 | 10.1918 | 4.9959 |  |  | 667 | 26683.149 | 10.1918 | 5.0117 |  |  | 667 | 26683.149 | 10.1918 | 5.0093
668 | 26723.154 | 10.1933 | 4.9958 |  |  | 668 | 26723.154 | 10.1933 | 5.0115 |  |  | 668 | 26723.154 | 10.1933 | 5.0091
669 | 26763.159 | 10.1948 | 4.9957 |  |  | 669 | 26763.159 | 10.1948 | 5.0114 |  |  | 669 | 26763.159 | 10.1948 | 5.009
670 | 26803.163 | 10.1963 | 4.9956 |  |  | 670 | 26803.163 | 10.1963 | 5.0113 |  |  | 670 | 26803.163 | 10.1963 | 5.0089
671 | 26843.168 | 10.1978 | 4.9955 |  |  | 671 | 26843.168 | 10.1978 | 5.0112 |  |  | 671 | 26843.168 | 10.1978 | 5.0088
672 | 26883.173 | 10.1993 | 4.9954 |  |  | 672 | 26883.173 | 10.1993 | 5.0111 |  |  | 672 | 26883.173 | 10.1993 | 5.0087
673 | 26923.178 | 10.2007 | 4.9953 |  |  | 673 | 26923.178 | 10.2007 | 5.011 |  |  | 673 | 26923.178 | 10.2007 | 5.0086
674 | 26963.182 | 10.2022 | 4.9952 |  |  | 674 | 26963.182 | 10.2022 | 5.0109 |  |  | 674 | 26963.182 | 10.2022 | 5.0085
675 | 27003.187 | 10.2037 | 4.9951 |  |  | 675 | 27003.187 | 10.2037 | 5.0108 |  |  | 675 | 27003.187 | 10.2037 | 5.0084
676 | 27043.192 | 10.2052 | 4.995 |  |  | 676 | 27043.192 | 10.2052 | 5.0107 |  |  | 676 | 27043.192 | 10.2052 | 5.0083
677 | 27083.197 | 10.2067 | 4.9949 |  |  | 677 | 27083.197 | 10.2067 | 5.0105 |  |  | 677 | 27083.197 | 10.2067 | 5.0082
678 | 27123.201 | 10.2081 | 4.9948 |  |  | 678 | 27123.201 | 10.2081 | 5.0104 |  |  | 678 | 27123.201 | 10.2081 | 5.0081
679 | 27163.206 | 10.2096 | 4.9947 |  |  | 679 | 27163.206 | 10.2096 | 5.0103 |  |  | 679 | 27163.206 | 10.2096 | 5.008
680 | 27203.211 | 10.2111 | 4.9945 |  |  | 680 | 27203.211 | 10.2111 | 5.0102 |  |  | 680 | 27203.211 | 10.2111 | 5.0079
681 | 27243.215 | 10.2126 | 4.9944 |  |  | 681 | 27243.215 | 10.2126 | 5.0101 |  |  | 681 | 27243.215 | 10.2126 | 5.0078
682 | 27283.22 | 10.214 | 4.9943 |  |  | 682 | 27283.22 | 10.214 | 5.01 |  |  | 682 | 27283.22 | 10.214 | 5.0076
683 | 27323.225 | 10.2155 | 4.9942 |  |  | 683 | 27323.225 | 10.2155 | 5.0099 |  |  | 683 | 27323.225 | 10.2155 | 5.0075
684 | 27363.23 | 10.217 | 4.9941 |  |  | 684 | 27363.23 | 10.217 | 5.0098 |  |  | 684 | 27363.23 | 10.217 | 5.0074
685 | 27403.234 | 10.2184 | 4.994 |  |  | 685 | 27403.234 | 10.2184 | 5.0097 |  |  | 685 | 27403.234 | 10.2184 | 5.0073
686 | 27443.239 | 10.2199 | 4.9939 |  |  | 686 | 27443.239 | 10.2199 | 5.0096 |  |  | 686 | 27443.239 | 10.2199 | 5.0072
687 | 27483.244 | 10.2213 | 4.9938 |  |  | 687 | 27483.244 | 10.2213 | 5.0094 |  |  | 687 | 27483.244 | 10.2213 | 5.0071
688 | 27523.248 | 10.2228 | 4.9937 |  |  | 688 | 27523.248 | 10.2228 | 5.0093 |  |  | 688 | 27523.248 | 10.2228 | 5.007
689 | 27563.253 | 10.2242 | 4.9936 |  |  | 689 | 27563.253 | 10.2242 | 5.0092 |  |  | 689 | 27563.253 | 10.2242 | 5.0069
690 | 27603.258 | 10.2257 | 4.9935 |  |  | 690 | 27603.258 | 10.2257 | 5.0091 |  |  | 690 | 27603.258 | 10.2257 | 5.0068
691 | 27643.263 | 10.2271 | 4.9934 |  |  | 691 | 27643.263 | 10.2271 | 5.009 |  |  | 691 | 27643.263 | 10.2271 | 5.0067
692 | 27683.267 | 10.2286 | 4.9933 |  |  | 692 | 27683.267 | 10.2286 | 5.0089 |  |  | 692 | 27683.267 | 10.2286 | 5.0066
693 | 27723.272 | 10.23 | 4.9932 |  |  | 693 | 27723.272 | 10.23 | 5.0088 |  |  | 693 | 27723.272 | 10.23 | 5.0065
694 | 27763.277 | 10.2315 | 4.9931 |  |  | 694 | 27763.277 | 10.2315 | 5.0087 |  |  | 694 | 27763.277 | 10.2315 | 5.0064
695 | 27803.282 | 10.2329 | 4.993 |  |  | 695 | 27803.282 | 10.2329 | 5.0086 |  |  | 695 | 27803.282 | 10.2329 | 5.0063
696 | 27843.286 | 10.2343 | 4.9929 |  |  | 696 | 27843.286 | 10.2343 | 5.0085 |  |  | 696 | 27843.286 | 10.2343 | 5.0062
697 | 27883.291 | 10.2358 | 4.9928 |  |  | 697 | 27883.291 | 10.2358 | 5.0084 |  |  | 697 | 27883.291 | 10.2358 | 5.0061
698 | 27923.296 | 10.2372 | 4.9927 |  |  | 698 | 27923.296 | 10.2372 | 5.0083 |  |  | 698 | 27923.296 | 10.2372 | 5.006
699 | 27963.3 | 10.2386 | 4.9926 |  |  | 699 | 27963.3 | 10.2386 | 5.0082 |  |  | 699 | 27963.3 | 10.2386 | 5.0059
700 | 28003.305 | 10.2401 | 4.9925 |  |  | 700 | 28003.305 | 10.2401 | 5.008 |  |  | 700 | 28003.305 | 10.2401 | 5.0058
701 | 28043.31 | 10.2415 | 4.9924 |  |  | 701 | 28043.31 | 10.2415 | 5.0079 |  |  | 701 | 28043.31 | 10.2415 | 5.0057
702 | 28083.315 | 10.2429 | 4.9923 |  |  | 702 | 28083.315 | 10.2429 | 5.0078 |  |  | 702 | 28083.315 | 10.2429 | 5.0056
703 | 28123.319 | 10.2444 | 4.9922 |  |  | 703 | 28123.319 | 10.2444 | 5.0077 |  |  | 703 | 28123.319 | 10.2444 | 5.0055
704 | 28163.324 | 10.2458 | 4.9921 |  |  | 704 | 28163.324 | 10.2458 | 5.0076 |  |  | 704 | 28163.324 | 10.2458 | 5.0054
705 | 28203.329 | 10.2472 | 4.992 |  |  | 705 | 28203.329 | 10.2472 | 5.0075 |  |  | 705 | 28203.329 | 10.2472 | 5.0053
706 | 28243.333 | 10.2486 | 4.9919 |  |  | 706 | 28243.333 | 10.2486 | 5.0074 |  |  | 706 | 28243.333 | 10.2486 | 5.0052
707 | 28283.338 | 10.25 | 4.9918 |  |  | 707 | 28283.338 | 10.25 | 5.0073 |  |  | 707 | 28283.338 | 10.25 | 5.0051
708 | 28323.343 | 10.2514 | 4.9917 |  |  | 708 | 28323.343 | 10.2514 | 5.0072 |  |  | 708 | 28323.343 | 10.2514 | 5.0049
709 | 28363.348 | 10.2529 | 4.9916 |  |  | 709 | 28363.348 | 10.2529 | 5.0071 |  |  | 709 | 28363.348 | 10.2529 | 5.0048
710 | 28403.352 | 10.2543 | 4.9915 |  |  | 710 | 28403.352 | 10.2543 | 5.007 |  |  | 710 | 28403.352 | 10.2543 | 5.0047
711 | 28443.357 | 10.2557 | 4.9915 |  |  | 711 | 28443.357 | 10.2557 | 5.0069 |  |  | 711 | 28443.357 | 10.2557 | 5.0046
712 | 28483.362 | 10.2571 | 4.9914 |  |  | 712 | 28483.362 | 10.2571 | 5.0068 |  |  | 712 | 28483.362 | 10.2571 | 5.0045
713 | 28523.367 | 10.2585 | 4.9913 |  |  | 713 | 28523.367 | 10.2585 | 5.0067 |  |  | 713 | 28523.367 | 10.2585 | 5.0044
714 | 28563.371 | 10.2599 | 4.9912 |  |  | 714 | 28563.371 | 10.2599 | 5.0066 |  |  | 714 | 28563.371 | 10.2599 | 5.0043
715 | 28603.376 | 10.2613 | 4.9911 |  |  | 715 | 28603.376 | 10.2613 | 5.0065 |  |  | 715 | 28603.376 | 10.2613 | 5.0042
716 | 28643.381 | 10.2627 | 4.991 |  |  | 716 | 28643.381 | 10.2627 | 5.0063 |  |  | 716 | 28643.381 | 10.2627 | 5.0041
717 | 28683.385 | 10.2641 | 4.9909 |  |  | 717 | 28683.385 | 10.2641 | 5.0062 |  |  | 717 | 28683.385 | 10.2641 | 5.004
718 | 28723.39 | 10.2655 | 4.9908 |  |  | 718 | 28723.39 | 10.2655 | 5.0061 |  |  | 718 | 28723.39 | 10.2655 | 5.0039
719 | 28763.395 | 10.2669 | 4.9907 |  |  | 719 | 28763.395 | 10.2669 | 5.006 |  |  | 719 | 28763.395 | 10.2669 | 5.0038
720 | 28803.4 | 10.2682 | 4.9906 |  |  | 720 | 28803.4 | 10.2682 | 5.0059 |  |  | 720 | 28803.4 | 10.2682 | 5.0037
721 | 28843.404 | 10.2696 | 4.9905 |  |  | 721 | 28843.404 | 10.2696 | 5.0058 |  |  | 721 | 28843.404 | 10.2696 | 5.0036
722 | 28883.409 | 10.271 | 4.9904 |  |  | 722 | 28883.409 | 10.271 | 5.0057 |  |  | 722 | 28883.409 | 10.271 | 5.0035
723 | 28923.414 | 10.2724 | 4.9903 |  |  | 723 | 28923.414 | 10.2724 | 5.0056 |  |  | 723 | 28923.414 | 10.2724 | 5.0034
724 | 28963.418 | 10.2738 | 4.9902 |  |  | 724 | 28963.418 | 10.2738 | 5.0055 |  |  | 724 | 28963.418 | 10.2738 | 5.0033
725 | 29003.423 | 10.2752 | 4.9901 |  |  | 725 | 29003.423 | 10.2752 | 5.0054 |  |  | 725 | 29003.423 | 10.2752 | 5.0032
726 | 29043.428 | 10.2765 | 4.99 |  |  | 726 | 29043.428 | 10.2765 | 5.0053 |  |  | 726 | 29043.428 | 10.2765 | 5.0031
727 | 29083.433 | 10.2779 | 4.9899 |  |  | 727 | 29083.433 | 10.2779 | 5.0052 |  |  | 727 | 29083.433 | 10.2779 | 5.003
728 | 29123.437 | 10.2793 | 4.9898 |  |  | 728 | 29123.437 | 10.2793 | 5.0051 |  |  | 728 | 29123.437 | 10.2793 | 5.0029
729 | 29163.442 | 10.2807 | 4.9897 |  |  | 729 | 29163.442 | 10.2807 | 5.005 |  |  | 729 | 29163.442 | 10.2807 | 5.0028
730 | 29203.447 | 10.282 | 4.9896 |  |  | 730 | 29203.447 | 10.282 | 5.0049 |  |  | 730 | 29203.447 | 10.282 | 5.0027
731 | 29243.452 | 10.2834 | 4.9895 |  |  | 731 | 29243.452 | 10.2834 | 5.0048 |  |  | 731 | 29243.452 | 10.2834 | 5.0026
732 | 29283.456 | 10.2848 | 4.9894 |  |  | 732 | 29283.456 | 10.2848 | 5.0047 |  |  | 732 | 29283.456 | 10.2848 | 5.0025
733 | 29323.461 | 10.2861 | 4.9893 |  |  | 733 | 29323.461 | 10.2861 | 5.0046 |  |  | 733 | 29323.461 | 10.2861 | 5.0024
734 | 29363.466 | 10.2875 | 4.9892 |  |  | 734 | 29363.466 | 10.2875 | 5.0045 |  |  | 734 | 29363.466 | 10.2875 | 5.0023
735 | 29403.47 | 10.2889 | 4.9891 |  |  | 735 | 29403.47 | 10.2889 | 5.0044 |  |  | 735 | 29403.47 | 10.2889 | 5.0022
736 | 29443.475 | 10.2902 | 4.989 |  |  | 736 | 29443.475 | 10.2902 | 5.0043 |  |  | 736 | 29443.475 | 10.2902 | 5.0021
737 | 29483.48 | 10.2916 | 4.9889 |  |  | 737 | 29483.48 | 10.2916 | 5.0042 |  |  | 737 | 29483.48 | 10.2916 | 5.002
738 | 29523.485 | 10.2929 | 4.9888 |  |  | 738 | 29523.485 | 10.2929 | 5.0041 |  |  | 738 | 29523.485 | 10.2929 | 5.0019
739 | 29563.489 | 10.2943 | 4.9888 |  |  | 739 | 29563.489 | 10.2943 | 5.004 |  |  | 739 | 29563.489 | 10.2943 | 5.0018
740 | 29603.494 | 10.2956 | 4.9887 |  |  | 740 | 29603.494 | 10.2956 | 5.0038 |  |  | 740 | 29603.494 | 10.2956 | 5.0017
741 | 29643.499 | 10.297 | 4.9886 |  |  | 741 | 29643.499 | 10.297 | 5.0037 |  |  | 741 | 29643.499 | 10.297 | 5.0016
742 | 29683.503 | 10.2983 | 4.9885 |  |  | 742 | 29683.503 | 10.2983 | 5.0036 |  |  | 742 | 29683.503 | 10.2983 | 5.0015
743 | 29723.508 | 10.2997 | 4.9884 |  |  | 743 | 29723.508 | 10.2997 | 5.0035 |  |  | 743 | 29723.508 | 10.2997 | 5.0014
744 | 29763.513 | 10.301 | 4.9883 |  |  | 744 | 29763.513 | 10.301 | 5.0034 |  |  | 744 | 29763.513 | 10.301 | 5.0013
745 | 29803.518 | 10.3024 | 4.9882 |  |  | 745 | 29803.518 | 10.3024 | 5.0033 |  |  | 745 | 29803.518 | 10.3024 | 5.0012
746 | 29843.522 | 10.3037 | 4.9881 |  |  | 746 | 29843.522 | 10.3037 | 5.0032 |  |  | 746 | 29843.522 | 10.3037 | 5.0011
747 | 29883.527 | 10.3051 | 4.988 |  |  | 747 | 29883.527 | 10.3051 | 5.0031 |  |  | 747 | 29883.527 | 10.3051 | 5.001
748 | 29923.532 | 10.3064 | 4.9879 |  |  | 748 | 29923.532 | 10.3064 | 5.003 |  |  | 748 | 29923.532 | 10.3064 | 5.0009
749 | 29963.536 | 10.3077 | 4.9878 |  |  | 749 | 29963.536 | 10.3077 | 5.0029 |  |  | 749 | 29963.536 | 10.3077 | 5.0008
750 | 30003.541 | 10.3091 | 4.9877 |  |  | 750 | 30003.541 | 10.3091 | 5.0028 |  |  | 750 | 30003.541 | 10.3091 | 5.0007
751 | 30043.546 | 10.3104 | 4.9876 |  |  | 751 | 30043.546 | 10.3104 | 5.0027 |  |  | 751 | 30043.546 | 10.3104 | 5.0006
752 | 30083.551 | 10.3117 | 4.9875 |  |  | 752 | 30083.551 | 10.3117 | 5.0026 |  |  | 752 | 30083.551 | 10.3117 | 5.0005
753 | 30123.555 | 10.3131 | 4.9874 |  |  | 753 | 30123.555 | 10.3131 | 5.0025 |  |  | 753 | 30123.555 | 10.3131 | 5.0004
754 | 30163.56 | 10.3144 | 4.9873 |  |  | 754 | 30163.56 | 10.3144 | 5.0024 |  |  | 754 | 30163.56 | 10.3144 | 5.0003
755 | 30203.565 | 10.3157 | 4.9872 |  |  | 755 | 30203.565 | 10.3157 | 5.0023 |  |  | 755 | 30203.565 | 10.3157 | 5.0002
756 | 30243.57 | 10.317 | 4.9871 |  |  | 756 | 30243.57 | 10.317 | 5.0022 |  |  | 756 | 30243.57 | 10.317 | 5.0001
757 | 30283.574 | 10.3184 | 4.9871 |  |  | 757 | 30283.574 | 10.3184 | 5.0021 |  |  | 757 | 30283.574 | 10.3184 | 5
758 | 30323.579 | 10.3197 | 4.987 |  |  | 758 | 30323.579 | 10.3197 | 5.002 |  |  | 758 | 30323.579 | 10.3197 | 4.9999
759 | 30363.584 | 10.321 | 4.9869 |  |  | 759 | 30363.584 | 10.321 | 5.0019 |  |  | 759 | 30363.584 | 10.321 | 4.9998
760 | 30403.588 | 10.3223 | 4.9868 |  |  | 760 | 30403.588 | 10.3223 | 5.0018 |  |  | 760 | 30403.588 | 10.3223 | 4.9997
761 | 30443.593 | 10.3236 | 4.9867 |  |  | 761 | 30443.593 | 10.3236 | 5.0017 |  |  | 761 | 30443.593 | 10.3236 | 4.9996
762 | 30483.598 | 10.3249 | 4.9866 |  |  | 762 | 30483.598 | 10.3249 | 5.0016 |  |  | 762 | 30483.598 | 10.3249 | 4.9995
763 | 30523.603 | 10.3263 | 4.9865 |  |  | 763 | 30523.603 | 10.3263 | 5.0015 |  |  | 763 | 30523.603 | 10.3263 | 4.9994
764 | 30563.607 | 10.3276 | 4.9864 |  |  | 764 | 30563.607 | 10.3276 | 5.0014 |  |  | 764 | 30563.607 | 10.3276 | 4.9993
765 | 30603.612 | 10.3289 | 4.9863 |  |  | 765 | 30603.612 | 10.3289 | 5.0013 |  |  | 765 | 30603.612 | 10.3289 | 4.9993
766 | 30643.617 | 10.3302 | 4.9862 |  |  | 766 | 30643.617 | 10.3302 | 5.0012 |  |  | 766 | 30643.617 | 10.3302 | 4.9992
767 | 30683.621 | 10.3315 | 4.9861 |  |  | 767 | 30683.621 | 10.3315 | 5.0011 |  |  | 767 | 30683.621 | 10.3315 | 4.9991
768 | 30723.626 | 10.3328 | 4.986 |  |  | 768 | 30723.626 | 10.3328 | 5.001 |  |  | 768 | 30723.626 | 10.3328 | 4.999
769 | 30763.631 | 10.3341 | 4.9859 |  |  | 769 | 30763.631 | 10.3341 | 5.0009 |  |  | 769 | 30763.631 | 10.3341 | 4.9989
770 | 30803.636 | 10.3354 | 4.9858 |  |  | 770 | 30803.636 | 10.3354 | 5.0008 |  |  | 770 | 30803.636 | 10.3354 | 4.9988
771 | 30843.64 | 10.3367 | 4.9857 |  |  | 771 | 30843.64 | 10.3367 | 5.0007 |  |  | 771 | 30843.64 | 10.3367 | 4.9987
772 | 30883.645 | 10.338 | 4.9856 |  |  | 772 | 30883.645 | 10.338 | 5.0006 |  |  | 772 | 30883.645 | 10.338 | 4.9986
773 | 30923.65 | 10.3393 | 4.9856 |  |  | 773 | 30923.65 | 10.3393 | 5.0005 |  |  | 773 | 30923.65 | 10.3393 | 4.9985
774 | 30963.655 | 10.3406 | 4.9855 |  |  | 774 | 30963.655 | 10.3406 | 5.0004 |  |  | 774 | 30963.655 | 10.3406 | 4.9984
775 | 31003.659 | 10.3419 | 4.9854 |  |  | 775 | 31003.659 | 10.3419 | 5.0003 |  |  | 775 | 31003.659 | 10.3419 | 4.9983
776 | 31043.664 | 10.3431 | 4.9853 |  |  | 776 | 31043.664 | 10.3431 | 5.0002 |  |  | 776 | 31043.664 | 10.3431 | 4.9982
777 | 31083.669 | 10.3444 | 4.9852 |  |  | 777 | 31083.669 | 10.3444 | 5.0001 |  |  | 777 | 31083.669 | 10.3444 | 4.9981
778 | 31123.673 | 10.3457 | 4.9851 |  |  | 778 | 31123.673 | 10.3457 | 5 |  |  | 778 | 31123.673 | 10.3457 | 4.998
779 | 31163.678 | 10.347 | 4.985 |  |  | 779 | 31163.678 | 10.347 | 4.9999 |  |  | 779 | 31163.678 | 10.347 | 4.9979
780 | 31203.683 | 10.3483 | 4.9849 |  |  | 780 | 31203.683 | 10.3483 | 4.9998 |  |  | 780 | 31203.683 | 10.3483 | 4.9978
781 | 31243.688 | 10.3496 | 4.9848 |  |  | 781 | 31243.688 | 10.3496 | 4.9997 |  |  | 781 | 31243.688 | 10.3496 | 4.9977
782 | 31283.692 | 10.3509 | 4.9847 |  |  | 782 | 31283.692 | 10.3509 | 4.9996 |  |  | 782 | 31283.692 | 10.3509 | 4.9976
783 | 31323.697 | 10.3521 | 4.9846 |  |  | 783 | 31323.697 | 10.3521 | 4.9995 |  |  | 783 | 31323.697 | 10.3521 | 4.9975
784 | 31363.702 | 10.3534 | 4.9845 |  |  | 784 | 31363.702 | 10.3534 | 4.9994 |  |  | 784 | 31363.702 | 10.3534 | 4.9974
785 | 31403.706 | 10.3547 | 4.9844 |  |  | 785 | 31403.706 | 10.3547 | 4.9993 |  |  | 785 | 31403.706 | 10.3547 | 4.9973
786 | 31443.711 | 10.356 | 4.9844 |  |  | 786 | 31443.711 | 10.356 | 4.9992 |  |  | 786 | 31443.711 | 10.356 | 4.9972
787 | 31483.716 | 10.3572 | 4.9843 |  |  | 787 | 31483.716 | 10.3572 | 4.9991 |  |  | 787 | 31483.716 | 10.3572 | 4.9971
788 | 31523.721 | 10.3585 | 4.9842 |  |  | 788 | 31523.721 | 10.3585 | 4.999 |  |  | 788 | 31523.721 | 10.3585 | 4.997
789 | 31563.725 | 10.3598 | 4.9841 |  |  | 789 | 31563.725 | 10.3598 | 4.9989 |  |  | 789 | 31563.725 | 10.3598 | 4.9969
790 | 31603.73 | 10.361 | 4.984 |  |  | 790 | 31603.73 | 10.361 | 4.9988 |  |  | 790 | 31603.73 | 10.361 | 4.9968
791 | 31643.735 | 10.3623 | 4.9839 |  |  | 791 | 31643.735 | 10.3623 | 4.9987 |  |  | 791 | 31643.735 | 10.3623 | 4.9967
792 | 31683.74 | 10.3636 | 4.9838 |  |  | 792 | 31683.74 | 10.3636 | 4.9986 |  |  | 792 | 31683.74 | 10.3636 | 4.9966
793 | 31723.744 | 10.3648 | 4.9837 |  |  | 793 | 31723.744 | 10.3648 | 4.9985 |  |  | 793 | 31723.744 | 10.3648 | 4.9965
794 | 31763.749 | 10.3661 | 4.9836 |  |  | 794 | 31763.749 | 10.3661 | 4.9984 |  |  | 794 | 31763.749 | 10.3661 | 4.9964
795 | 31803.754 | 10.3673 | 4.9835 |  |  | 795 | 31803.754 | 10.3673 | 4.9983 |  |  | 795 | 31803.754 | 10.3673 | 4.9963
796 | 31843.758 | 10.3686 | 4.9834 |  |  | 796 | 31843.758 | 10.3686 | 4.9982 |  |  | 796 | 31843.758 | 10.3686 | 4.9962
797 | 31883.763 | 10.3699 | 4.9833 |  |  | 797 | 31883.763 | 10.3699 | 4.9981 |  |  | 797 | 31883.763 | 10.3699 | 4.9961
798 | 31923.768 | 10.3711 | 4.9832 |  |  | 798 | 31923.768 | 10.3711 | 4.998 |  |  | 798 | 31923.768 | 10.3711 | 4.996
799 | 31963.773 | 10.3724 | 4.9831 |  |  | 799 | 31963.773 | 10.3724 | 4.9979 |  |  | 799 | 31963.773 | 10.3724 | 4.9959
800 | 32003.777 | 10.3736 | 4.9831 |  |  | 800 | 32003.777 | 10.3736 | 4.9978 |  |  | 800 | 32003.777 | 10.3736 | 4.9958
801 | 32043.782 | 10.3749 | 4.983 |  |  | 801 | 32043.782 | 10.3749 | 4.9977 |  |  | 801 | 32043.782 | 10.3749 | 4.9957
802 | 32083.787 | 10.3761 | 4.9829 |  |  | 802 | 32083.787 | 10.3761 | 4.9976 |  |  | 802 | 32083.787 | 10.3761 | 4.9956
803 | 32123.791 | 10.3774 | 4.9828 |  |  | 803 | 32123.791 | 10.3774 | 4.9975 |  |  | 803 | 32123.791 | 10.3774 | 4.9955
804 | 32163.796 | 10.3786 | 4.9827 |  |  | 804 | 32163.796 | 10.3786 | 4.9974 |  |  | 804 | 32163.796 | 10.3786 | 4.9954
805 | 32203.801 | 10.3798 | 4.9826 |  |  | 805 | 32203.801 | 10.3798 | 4.9973 |  |  | 805 | 32203.801 | 10.3798 | 4.9953
806 | 32243.806 | 10.3811 | 4.9825 |  |  | 806 | 32243.806 | 10.3811 | 4.9972 |  |  | 806 | 32243.806 | 10.3811 | 4.9952
807 | 32283.81 | 10.3823 | 4.9824 |  |  | 807 | 32283.81 | 10.3823 | 4.9971 |  |  | 807 | 32283.81 | 10.3823 | 4.9951
808 | 32323.815 | 10.3836 | 4.9823 |  |  | 808 | 32323.815 | 10.3836 | 4.997 |  |  | 808 | 32323.815 | 10.3836 | 4.995
809 | 32363.82 | 10.3848 | 4.9822 |  |  | 809 | 32363.82 | 10.3848 | 4.9969 |  |  | 809 | 32363.82 | 10.3848 | 4.9949
810 | 32403.825 | 10.386 | 4.9821 |  |  | 810 | 32403.825 | 10.386 | 4.9968 |  |  | 810 | 32403.825 | 10.386 | 4.9949
811 | 32443.829 | 10.3873 | 4.982 |  |  | 811 | 32443.829 | 10.3873 | 4.9967 |  |  | 811 | 32443.829 | 10.3873 | 4.9948
812 | 32483.834 | 10.3885 | 4.982 |  |  | 812 | 32483.834 | 10.3885 | 4.9966 |  |  | 812 | 32483.834 | 10.3885 | 4.9947
813 | 32523.839 | 10.3897 | 4.9819 |  |  | 813 | 32523.839 | 10.3897 | 4.9965 |  |  | 813 | 32523.839 | 10.3897 | 4.9946
814 | 32563.843 | 10.391 | 4.9818 |  |  | 814 | 32563.843 | 10.391 | 4.9964 |  |  | 814 | 32563.843 | 10.391 | 4.9945
815 | 32603.848 | 10.3922 | 4.9817 |  |  | 815 | 32603.848 | 10.3922 | 4.9963 |  |  | 815 | 32603.848 | 10.3922 | 4.9944
816 | 32643.853 | 10.3934 | 4.9816 |  |  | 816 | 32643.853 | 10.3934 | 4.9962 |  |  | 816 | 32643.853 | 10.3934 | 4.9943
817 | 32683.858 | 10.3946 | 4.9815 |  |  | 817 | 32683.858 | 10.3946 | 4.9961 |  |  | 817 | 32683.858 | 10.3946 | 4.9942
818 | 32723.862 | 10.3959 | 4.9814 |  |  | 818 | 32723.862 | 10.3959 | 4.996 |  |  | 818 | 32723.862 | 10.3959 | 4.9941
819 | 32763.867 | 10.3971 | 4.9813 |  |  | 819 | 32763.867 | 10.3971 | 4.9959 |  |  | 819 | 32763.867 | 10.3971 | 4.994
820 | 32803.872 | 10.3983 | 4.9812 |  |  | 820 | 32803.872 | 10.3983 | 4.9958 |  |  | 820 | 32803.872 | 10.3983 | 4.9939
821 | 32843.876 | 10.3995 | 4.9811 |  |  | 821 | 32843.876 | 10.3995 | 4.9957 |  |  | 821 | 32843.876 | 10.3995 | 4.9938
822 | 32883.881 | 10.4007 | 4.981 |  |  | 822 | 32883.881 | 10.4007 | 4.9956 |  |  | 822 | 32883.881 | 10.4007 | 4.9937
823 | 32923.886 | 10.402 | 4.9809 |  |  | 823 | 32923.886 | 10.402 | 4.9955 |  |  | 823 | 32923.886 | 10.402 | 4.9936
824 | 32963.891 | 10.4032 | 4.9809 |  |  | 824 | 32963.891 | 10.4032 | 4.9954 |  |  | 824 | 32963.891 | 10.4032 | 4.9935
825 | 33003.895 | 10.4044 | 4.9808 |  |  | 825 | 33003.895 | 10.4044 | 4.9953 |  |  | 825 | 33003.895 | 10.4044 | 4.9934
826 | 33043.9 | 10.4056 | 4.9807 |  |  | 826 | 33043.9 | 10.4056 | 4.9952 |  |  | 826 | 33043.9 | 10.4056 | 4.9933
827 | 33083.905 | 10.4068 | 4.9806 |  |  | 827 | 33083.905 | 10.4068 | 4.9951 |  |  | 827 | 33083.905 | 10.4068 | 4.9932
828 | 33123.91 | 10.408 | 4.9805 |  |  | 828 | 33123.91 | 10.408 | 4.995 |  |  | 828 | 33123.91 | 10.408 | 4.9931
829 | 33163.914 | 10.4092 | 4.9804 |  |  | 829 | 33163.914 | 10.4092 | 4.9949 |  |  | 829 | 33163.914 | 10.4092 | 4.993
830 | 33203.919 | 10.4104 | 4.9803 |  |  | 830 | 33203.919 | 10.4104 | 4.9948 |  |  | 830 | 33203.919 | 10.4104 | 4.9929
831 | 33243.924 | 10.4116 | 4.9802 |  |  | 831 | 33243.924 | 10.4116 | 4.9947 |  |  | 831 | 33243.924 | 10.4116 | 4.9928
832 | 33283.928 | 10.4128 | 4.9801 |  |  | 832 | 33283.928 | 10.4128 | 4.9946 |  |  | 832 | 33283.928 | 10.4128 | 4.9927
833 | 33323.933 | 10.414 | 4.98 |  |  | 833 | 33323.933 | 10.414 | 4.9945 |  |  | 833 | 33323.933 | 10.414 | 4.9926
834 | 33363.938 | 10.4152 | 4.9799 |  |  | 834 | 33363.938 | 10.4152 | 4.9944 |  |  | 834 | 33363.938 | 10.4152 | 4.9925
835 | 33403.943 | 10.4164 | 4.9798 |  |  | 835 | 33403.943 | 10.4164 | 4.9943 |  |  | 835 | 33403.943 | 10.4164 | 4.9924
836 | 33443.947 | 10.4176 | 4.9797 |  |  | 836 | 33443.947 | 10.4176 | 4.9942 |  |  | 836 | 33443.947 | 10.4176 | 4.9923
837 | 33483.952 | 10.4188 | 4.9797 |  |  | 837 | 33483.952 | 10.4188 | 4.9941 |  |  | 837 | 33483.952 | 10.4188 | 4.9922
838 | 33523.957 | 10.42 | 4.9796 |  |  | 838 | 33523.957 | 10.42 | 4.994 |  |  | 838 | 33523.957 | 10.42 | 4.9921
839 | 33563.961 | 10.4212 | 4.9795 |  |  | 839 | 33563.961 | 10.4212 | 4.9939 |  |  | 839 | 33563.961 | 10.4212 | 4.992
840 | 33603.966 | 10.4224 | 4.9794 |  |  | 840 | 33603.966 | 10.4224 | 4.9938 |  |  | 840 | 33603.966 | 10.4224 | 4.9919
841 | 33643.971 | 10.4236 | 4.9793 |  |  | 841 | 33643.971 | 10.4236 | 4.9937 |  |  | 841 | 33643.971 | 10.4236 | 4.9918
842 | 33683.976 | 10.4248 | 4.9792 |  |  | 842 | 33683.976 | 10.4248 | 4.9936 |  |  | 842 | 33683.976 | 10.4248 | 4.9917
843 | 33723.98 | 10.426 | 4.9791 |  |  | 843 | 33723.98 | 10.426 | 4.9935 |  |  | 843 | 33723.98 | 10.426 | 4.9916
844 | 33763.985 | 10.4271 | 4.979 |  |  | 844 | 33763.985 | 10.4271 | 4.9934 |  |  | 844 | 33763.985 | 10.4271 | 4.9915
845 | 33803.99 | 10.4283 | 4.9789 |  |  | 845 | 33803.99 | 10.4283 | 4.9933 |  |  | 845 | 33803.99 | 10.4283 | 4.9914
846 | 33843.994 | 10.4295 | 4.9788 |  |  | 846 | 33843.994 | 10.4295 | 4.9932 |  |  | 846 | 33843.994 | 10.4295 | 4.9913
847 | 33883.999 | 10.4307 | 4.9787 |  |  | 847 | 33883.999 | 10.4307 | 4.9931 |  |  | 847 | 33883.999 | 10.4307 | 4.9912
848 | 33924.004 | 10.4319 | 4.9786 |  |  | 848 | 33924.004 | 10.4319 | 4.993 |  |  | 848 | 33924.004 | 10.4319 | 4.9911
849 | 33964.009 | 10.4331 | 4.9786 |  |  | 849 | 33964.009 | 10.4331 | 4.9929 |  |  | 849 | 33964.009 | 10.4331 | 4.991
850 | 34004.013 | 10.4342 | 4.9785 |  |  | 850 | 34004.013 | 10.4342 | 4.9928 |  |  | 850 | 34004.013 | 10.4342 | 4.9909
851 | 34044.018 | 10.4354 | 4.9784 |  |  | 851 | 34044.018 | 10.4354 | 4.9927 |  |  | 851 | 34044.018 | 10.4354 | 4.9909
852 | 34084.023 | 10.4366 | 4.9783 |  |  | 852 | 34084.023 | 10.4366 | 4.9926 |  |  | 852 | 34084.023 | 10.4366 | 4.9908
853 | 34124.028 | 10.4378 | 4.9782 |  |  | 853 | 34124.028 | 10.4378 | 4.9925 |  |  | 853 | 34124.028 | 10.4378 | 4.9907
854 | 34164.032 | 10.4389 | 4.9781 |  |  | 854 | 34164.032 | 10.4389 | 4.9924 |  |  | 854 | 34164.032 | 10.4389 | 4.9906
855 | 34204.037 | 10.4401 | 4.978 |  |  | 855 | 34204.037 | 10.4401 | 4.9923 |  |  | 855 | 34204.037 | 10.4401 | 4.9905
856 | 34244.042 | 10.4413 | 4.9779 |  |  | 856 | 34244.042 | 10.4413 | 4.9922 |  |  | 856 | 34244.042 | 10.4413 | 4.9904
857 | 34284.046 | 10.4424 | 4.9778 |  |  | 857 | 34284.046 | 10.4424 | 4.9921 |  |  | 857 | 34284.046 | 10.4424 | 4.9903
858 | 34324.051 | 10.4436 | 4.9777 |  |  | 858 | 34324.051 | 10.4436 | 4.992 |  |  | 858 | 34324.051 | 10.4436 | 4.9902
859 | 34364.056 | 10.4448 | 4.9776 |  |  | 859 | 34364.056 | 10.4448 | 4.9919 |  |  | 859 | 34364.056 | 10.4448 | 4.9901
860 | 34404.061 | 10.4459 | 4.9775 |  |  | 860 | 34404.061 | 10.4459 | 4.9918 |  |  | 860 | 34404.061 | 10.4459 | 4.99
861 | 34444.065 | 10.4471 | 4.9775 |  |  | 861 | 34444.065 | 10.4471 | 4.9917 |  |  | 861 | 34444.065 | 10.4471 | 4.9899
862 | 34484.07 | 10.4483 | 4.9774 |  |  | 862 | 34484.07 | 10.4483 | 4.9916 |  |  | 862 | 34484.07 | 10.4483 | 4.9898
863 | 34524.075 | 10.4494 | 4.9773 |  |  | 863 | 34524.075 | 10.4494 | 4.9915 |  |  | 863 | 34524.075 | 10.4494 | 4.9897
864 | 34564.079 | 10.4506 | 4.9772 |  |  | 864 | 34564.079 | 10.4506 | 4.9914 |  |  | 864 | 34564.079 | 10.4506 | 4.9896
865 | 34604.084 | 10.4517 | 4.9771 |  |  | 865 | 34604.084 | 10.4517 | 4.9913 |  |  | 865 | 34604.084 | 10.4517 | 4.9895
866 | 34644.089 | 10.4529 | 4.977 |  |  | 866 | 34644.089 | 10.4529 | 4.9912 |  |  | 866 | 34644.089 | 10.4529 | 4.9894
867 | 34684.094 | 10.454 | 4.9769 |  |  | 867 | 34684.094 | 10.454 | 4.9911 |  |  | 867 | 34684.094 | 10.454 | 4.9893
868 | 34724.098 | 10.4552 | 4.9768 |  |  | 868 | 34724.098 | 10.4552 | 4.991 |  |  | 868 | 34724.098 | 10.4552 | 4.9892
869 | 34764.103 | 10.4563 | 4.9767 |  |  | 869 | 34764.103 | 10.4563 | 4.9909 |  |  | 869 | 34764.103 | 10.4563 | 4.9891
870 | 34804.108 | 10.4575 | 4.9766 |  |  | 870 | 34804.108 | 10.4575 | 4.9908 |  |  | 870 | 34804.108 | 10.4575 | 4.989
871 | 34844.113 | 10.4586 | 4.9765 |  |  | 871 | 34844.113 | 10.4586 | 4.9907 |  |  | 871 | 34844.113 | 10.4586 | 4.9889
872 | 34884.117 | 10.4598 | 4.9764 |  |  | 872 | 34884.117 | 10.4598 | 4.9906 |  |  | 872 | 34884.117 | 10.4598 | 4.9888
873 | 34924.122 | 10.4609 | 4.9763 |  |  | 873 | 34924.122 | 10.4609 | 4.9905 |  |  | 873 | 34924.122 | 10.4609 | 4.9887
874 | 34964.127 | 10.4621 | 4.9763 |  |  | 874 | 34964.127 | 10.4621 | 4.9904 |  |  | 874 | 34964.127 | 10.4621 | 4.9886
875 | 35004.131 | 10.4632 | 4.9762 |  |  | 875 | 35004.131 | 10.4632 | 4.9903 |  |  | 875 | 35004.131 | 10.4632 | 4.9885
876 | 35044.136 | 10.4644 | 4.9761 |  |  | 876 | 35044.136 | 10.4644 | 4.9902 |  |  | 876 | 35044.136 | 10.4644 | 4.9884
877 | 35084.141 | 10.4655 | 4.976 |  |  | 877 | 35084.141 | 10.4655 | 4.9901 |  |  | 877 | 35084.141 | 10.4655 | 4.9883
878 | 35124.146 | 10.4666 | 4.9759 |  |  | 878 | 35124.146 | 10.4666 | 4.99 |  |  | 878 | 35124.146 | 10.4666 | 4.9882
879 | 35164.15 | 10.4678 | 4.9758 |  |  | 879 | 35164.15 | 10.4678 | 4.9899 |  |  | 879 | 35164.15 | 10.4678 | 4.9881
880 | 35204.155 | 10.4689 | 4.9757 |  |  | 880 | 35204.155 | 10.4689 | 4.9898 |  |  | 880 | 35204.155 | 10.4689 | 4.988
881 | 35244.16 | 10.4701 | 4.9756 |  |  | 881 | 35244.16 | 10.4701 | 4.9897 |  |  | 881 | 35244.16 | 10.4701 | 4.9879
882 | 35284.164 | 10.4712 | 4.9755 |  |  | 882 | 35284.164 | 10.4712 | 4.9896 |  |  | 882 | 35284.164 | 10.4712 | 4.9878
883 | 35324.169 | 10.4723 | 4.9754 |  |  | 883 | 35324.169 | 10.4723 | 4.9895 |  |  | 883 | 35324.169 | 10.4723 | 4.9877
884 | 35364.174 | 10.4735 | 4.9753 |  |  | 884 | 35364.174 | 10.4735 | 4.9894 |  |  | 884 | 35364.174 | 10.4735 | 4.9876
885 | 35404.179 | 10.4746 | 4.9752 |  |  | 885 | 35404.179 | 10.4746 | 4.9893 |  |  | 885 | 35404.179 | 10.4746 | 4.9875
886 | 35444.183 | 10.4757 | 4.9751 |  |  | 886 | 35444.183 | 10.4757 | 4.9892 |  |  | 886 | 35444.183 | 10.4757 | 4.9874
887 | 35484.188 | 10.4768 | 4.9751 |  |  | 887 | 35484.188 | 10.4768 | 4.9891 |  |  | 887 | 35484.188 | 10.4768 | 4.9873
888 | 35524.193 | 10.478 | 4.975 |  |  | 888 | 35524.193 | 10.478 | 4.9889 |  |  | 888 | 35524.193 | 10.478 | 4.9872
889 | 35564.198 | 10.4791 | 4.9749 |  |  | 889 | 35564.198 | 10.4791 | 4.9888 |  |  | 889 | 35564.198 | 10.4791 | 4.9871
890 | 35604.202 | 10.4802 | 4.9748 |  |  | 890 | 35604.202 | 10.4802 | 4.9887 |  |  | 890 | 35604.202 | 10.4802 | 4.987
891 | 35644.207 | 10.4813 | 4.9747 |  |  | 891 | 35644.207 | 10.4813 | 4.9886 |  |  | 891 | 35644.207 | 10.4813 | 4.9869
892 | 35684.212 | 10.4825 | 4.9746 |  |  | 892 | 35684.212 | 10.4825 | 4.9885 |  |  | 892 | 35684.212 | 10.4825 | 4.9868
893 | 35724.216 | 10.4836 | 4.9745 |  |  | 893 | 35724.216 | 10.4836 | 4.9884 |  |  | 893 | 35724.216 | 10.4836 | 4.9867
894 | 35764.221 | 10.4847 | 4.9744 |  |  | 894 | 35764.221 | 10.4847 | 4.9883 |  |  | 894 | 35764.221 | 10.4847 | 4.9866
895 | 35804.226 | 10.4858 | 4.9743 |  |  | 895 | 35804.226 | 10.4858 | 4.9882 |  |  | 895 | 35804.226 | 10.4858 | 4.9865
896 | 35844.231 | 10.4869 | 4.9742 |  |  | 896 | 35844.231 | 10.4869 | 4.9881 |  |  | 896 | 35844.231 | 10.4869 | 4.9864
897 | 35884.235 | 10.4881 | 4.9741 |  |  | 897 | 35884.235 | 10.4881 | 4.988 |  |  | 897 | 35884.235 | 10.4881 | 4.9863
898 | 35924.24 | 10.4892 | 4.974 |  |  | 898 | 35924.24 | 10.4892 | 4.9879 |  |  | 898 | 35924.24 | 10.4892 | 4.9862
899 | 35964.245 | 10.4903 | 4.9739 |  |  | 899 | 35964.245 | 10.4903 | 4.9878 |  |  | 899 | 35964.245 | 10.4903 | 4.9861
900 | 36004.249 | 10.4914 | 4.9739 |  |  | 900 | 36004.249 | 10.4914 | 4.9877 |  |  | 900 | 36004.249 | 10.4914 | 4.986
901 | 36044.254 | 10.4925 | 4.9738 |  |  | 901 | 36044.254 | 10.4925 | 4.9876 |  |  | 901 | 36044.254 | 10.4925 | 4.9859
902 | 36084.259 | 10.4936 | 4.9737 |  |  | 902 | 36084.259 | 10.4936 | 4.9875 |  |  | 902 | 36084.259 | 10.4936 | 4.9858
903 | 36124.264 | 10.4947 | 4.9736 |  |  | 903 | 36124.264 | 10.4947 | 4.9874 |  |  | 903 | 36124.264 | 10.4947 | 4.9857
904 | 36164.268 | 10.4958 | 4.9735 |  |  | 904 | 36164.268 | 10.4958 | 4.9873 |  |  | 904 | 36164.268 | 10.4958 | 4.9856
905 | 36204.273 | 10.4969 | 4.9734 |  |  | 905 | 36204.273 | 10.4969 | 4.9872 |  |  | 905 | 36204.273 | 10.4969 | 4.9855
906 | 36244.278 | 10.498 | 4.9733 |  |  | 906 | 36244.278 | 10.498 | 4.9871 |  |  | 906 | 36244.278 | 10.498 | 4.9854
907 | 36284.283 | 10.4991 | 4.9732 |  |  | 907 | 36284.283 | 10.4991 | 4.987 |  |  | 907 | 36284.283 | 10.4991 | 4.9853
908 | 36324.287 | 10.5002 | 4.9731 |  |  | 908 | 36324.287 | 10.5002 | 4.9869 |  |  | 908 | 36324.287 | 10.5002 | 4.9852
909 | 36364.292 | 10.5013 | 4.973 |  |  | 909 | 36364.292 | 10.5013 | 4.9868 |  |  | 909 | 36364.292 | 10.5013 | 4.9851
910 | 36404.297 | 10.5024 | 4.9729 |  |  | 910 | 36404.297 | 10.5024 | 4.9867 |  |  | 910 | 36404.297 | 10.5024 | 4.985
911 | 36444.301 | 10.5035 | 4.9728 |  |  | 911 | 36444.301 | 10.5035 | 4.9866 |  |  | 911 | 36444.301 | 10.5035 | 4.9849
912 | 36484.306 | 10.5046 | 4.9727 |  |  | 912 | 36484.306 | 10.5046 | 4.9865 |  |  | 912 | 36484.306 | 10.5046 | 4.9848
913 | 36524.311 | 10.5057 | 4.9726 |  |  | 913 | 36524.311 | 10.5057 | 4.9864 |  |  | 913 | 36524.311 | 10.5057 | 4.9847
914 | 36564.316 | 10.5068 | 4.9726 |  |  | 914 | 36564.316 | 10.5068 | 4.9863 |  |  | 914 | 36564.316 | 10.5068 | 4.9846
915 | 36604.32 | 10.5079 | 4.9725 |  |  | 915 | 36604.32 | 10.5079 | 4.9862 |  |  | 915 | 36604.32 | 10.5079 | 4.9845
916 | 36644.325 | 10.509 | 4.9724 |  |  | 916 | 36644.325 | 10.509 | 4.9861 |  |  | 916 | 36644.325 | 10.509 | 4.9844
917 | 36684.33 | 10.5101 | 4.9723 |  |  | 917 | 36684.33 | 10.5101 | 4.986 |  |  | 917 | 36684.33 | 10.5101 | 4.9843
918 | 36724.334 | 10.5112 | 4.9722 |  |  | 918 | 36724.334 | 10.5112 | 4.9859 |  |  | 918 | 36724.334 | 10.5112 | 4.9842
919 | 36764.339 | 10.5123 | 4.9721 |  |  | 919 | 36764.339 | 10.5123 | 4.9858 |  |  | 919 | 36764.339 | 10.5123 | 4.9841
920 | 36804.344 | 10.5134 | 4.972 |  |  | 920 | 36804.344 | 10.5134 | 4.9857 |  |  | 920 | 36804.344 | 10.5134 | 4.984
921 | 36844.349 | 10.5145 | 4.9719 |  |  | 921 | 36844.349 | 10.5145 | 4.9856 |  |  | 921 | 36844.349 | 10.5145 | 4.9839
922 | 36884.353 | 10.5155 | 4.9718 |  |  | 922 | 36884.353 | 10.5155 | 4.9855 |  |  | 922 | 36884.353 | 10.5155 | 4.9838
923 | 36924.358 | 10.5166 | 4.9717 |  |  | 923 | 36924.358 | 10.5166 | 4.9854 |  |  | 923 | 36924.358 | 10.5166 | 4.9837
924 | 36964.363 | 10.5177 | 4.9716 |  |  | 924 | 36964.363 | 10.5177 | 4.9853 |  |  | 924 | 36964.363 | 10.5177 | 4.9836
925 | 37004.368 | 10.5188 | 4.9715 |  |  | 925 | 37004.368 | 10.5188 | 4.9852 |  |  | 925 | 37004.368 | 10.5188 | 4.9835
926 | 37044.372 | 10.5199 | 4.9714 |  |  | 926 | 37044.372 | 10.5199 | 4.9851 |  |  | 926 | 37044.372 | 10.5199 | 4.9834
927 | 37084.377 | 10.521 | 4.9713 |  |  | 927 | 37084.377 | 10.521 | 4.985 |  |  | 927 | 37084.377 | 10.521 | 4.9833
928 | 37124.382 | 10.522 | 4.9712 |  |  | 928 | 37124.382 | 10.522 | 4.9849 |  |  | 928 | 37124.382 | 10.522 | 4.9832
929 | 37164.386 | 10.5231 | 4.9712 |  |  | 929 | 37164.386 | 10.5231 | 4.9848 |  |  | 929 | 37164.386 | 10.5231 | 4.9831
930 | 37204.391 | 10.5242 | 4.9711 |  |  | 930 | 37204.391 | 10.5242 | 4.9847 |  |  | 930 | 37204.391 | 10.5242 | 4.983
931 | 37244.396 | 10.5253 | 4.971 |  |  | 931 | 37244.396 | 10.5253 | 4.9846 |  |  | 931 | 37244.396 | 10.5253 | 4.9829
932 | 37284.401 | 10.5263 | 4.9709 |  |  | 932 | 37284.401 | 10.5263 | 4.9845 |  |  | 932 | 37284.401 | 10.5263 | 4.9828
933 | 37324.405 | 10.5274 | 4.9708 |  |  | 933 | 37324.405 | 10.5274 | 4.9844 |  |  | 933 | 37324.405 | 10.5274 | 4.9827
934 | 37364.41 | 10.5285 | 4.9707 |  |  | 934 | 37364.41 | 10.5285 | 4.9843 |  |  | 934 | 37364.41 | 10.5285 | 4.9826
935 | 37404.415 | 10.5295 | 4.9706 |  |  | 935 | 37404.415 | 10.5295 | 4.9842 |  |  | 935 | 37404.415 | 10.5295 | 4.9825
936 | 37444.419 | 10.5306 | 4.9705 |  |  | 936 | 37444.419 | 10.5306 | 4.9841 |  |  | 936 | 37444.419 | 10.5306 | 4.9824
937 | 37484.424 | 10.5317 | 4.9704 |  |  | 937 | 37484.424 | 10.5317 | 4.984 |  |  | 937 | 37484.424 | 10.5317 | 4.9823
938 | 37524.429 | 10.5327 | 4.9703 |  |  | 938 | 37524.429 | 10.5327 | 4.9839 |  |  | 938 | 37524.429 | 10.5327 | 4.9822
939 | 37564.434 | 10.5338 | 4.9702 |  |  | 939 | 37564.434 | 10.5338 | 4.9838 |  |  | 939 | 37564.434 | 10.5338 | 4.9821
940 | 37604.438 | 10.5349 | 4.9701 |  |  | 940 | 37604.438 | 10.5349 | 4.9837 |  |  | 940 | 37604.438 | 10.5349 | 4.982
941 | 37644.443 | 10.5359 | 4.97 |  |  | 941 | 37644.443 | 10.5359 | 4.9836 |  |  | 941 | 37644.443 | 10.5359 | 4.9819
942 | 37684.448 | 10.537 | 4.9699 |  |  | 942 | 37684.448 | 10.537 | 4.9835 |  |  | 942 | 37684.448 | 10.537 | 4.9818
943 | 37724.452 | 10.5381 | 4.9698 |  |  | 943 | 37724.452 | 10.5381 | 4.9834 |  |  | 943 | 37724.452 | 10.5381 | 4.9817
944 | 37764.457 | 10.5391 | 4.9697 |  |  | 944 | 37764.457 | 10.5391 | 4.9833 |  |  | 944 | 37764.457 | 10.5391 | 4.9816
945 | 37804.462 | 10.5402 | 4.9697 |  |  | 945 | 37804.462 | 10.5402 | 4.9832 |  |  | 945 | 37804.462 | 10.5402 | 4.9815
946 | 37844.467 | 10.5412 | 4.9696 |  |  | 946 | 37844.467 | 10.5412 | 4.9831 |  |  | 946 | 37844.467 | 10.5412 | 4.9814
947 | 37884.471 | 10.5423 | 4.9695 |  |  | 947 | 37884.471 | 10.5423 | 4.983 |  |  | 947 | 37884.471 | 10.5423 | 4.9813
948 | 37924.476 | 10.5434 | 4.9694 |  |  | 948 | 37924.476 | 10.5434 | 4.9829 |  |  | 948 | 37924.476 | 10.5434 | 4.9812
949 | 37964.481 | 10.5444 | 4.9693 |  |  | 949 | 37964.481 | 10.5444 | 4.9828 |  |  | 949 | 37964.481 | 10.5444 | 4.9811
950 | 38004.486 | 10.5455 | 4.9692 |  |  | 950 | 38004.486 | 10.5455 | 4.9827 |  |  | 950 | 38004.486 | 10.5455 | 4.981
951 | 38044.49 | 10.5465 | 4.9691 |  |  | 951 | 38044.49 | 10.5465 | 4.9826 |  |  | 951 | 38044.49 | 10.5465 | 4.9809
952 | 38084.495 | 10.5476 | 4.969 |  |  | 952 | 38084.495 | 10.5476 | 4.9825 |  |  | 952 | 38084.495 | 10.5476 | 4.9808
953 | 38124.5 | 10.5486 | 4.9689 |  |  | 953 | 38124.5 | 10.5486 | 4.9823 |  |  | 953 | 38124.5 | 10.5486 | 4.9807
954 | 38164.504 | 10.5497 | 4.9688 |  |  | 954 | 38164.504 | 10.5497 | 4.9822 |  |  | 954 | 38164.504 | 10.5497 | 4.9806
955 | 38204.509 | 10.5507 | 4.9687 |  |  | 955 | 38204.509 | 10.5507 | 4.9821 |  |  | 955 | 38204.509 | 10.5507 | 4.9805
956 | 38244.514 | 10.5518 | 4.9686 |  |  | 956 | 38244.514 | 10.5518 | 4.982 |  |  | 956 | 38244.514 | 10.5518 | 4.9804
957 | 38284.519 | 10.5528 | 4.9685 |  |  | 957 | 38284.519 | 10.5528 | 4.9819 |  |  | 957 | 38284.519 | 10.5528 | 4.9803
958 | 38324.523 | 10.5538 | 4.9684 |  |  | 958 | 38324.523 | 10.5538 | 4.9818 |  |  | 958 | 38324.523 | 10.5538 | 4.9802
959 | 38364.528 | 10.5549 | 4.9683 |  |  | 959 | 38364.528 | 10.5549 | 4.9817 |  |  | 959 | 38364.528 | 10.5549 | 4.9801
960 | 38404.533 | 10.5559 | 4.9682 |  |  | 960 | 38404.533 | 10.5559 | 4.9816 |  |  | 960 | 38404.533 | 10.5559 | 4.98
961 | 38444.537 | 10.557 | 4.9681 |  |  | 961 | 38444.537 | 10.557 | 4.9815 |  |  | 961 | 38444.537 | 10.557 | 4.9799
962 | 38484.542 | 10.558 | 4.9681 |  |  | 962 | 38484.542 | 10.558 | 4.9814 |  |  | 962 | 38484.542 | 10.558 | 4.9798
963 | 38524.547 | 10.5591 | 4.968 |  |  | 963 | 38524.547 | 10.5591 | 4.9813 |  |  | 963 | 38524.547 | 10.5591 | 4.9797
964 | 38564.552 | 10.5601 | 4.9679 |  |  | 964 | 38564.552 | 10.5601 | 4.9812 |  |  | 964 | 38564.552 | 10.5601 | 4.9796
965 | 38604.556 | 10.5611 | 4.9678 |  |  | 965 | 38604.556 | 10.5611 | 4.9811 |  |  | 965 | 38604.556 | 10.5611 | 4.9795
966 | 38644.561 | 10.5622 | 4.9677 |  |  | 966 | 38644.561 | 10.5622 | 4.981 |  |  | 966 | 38644.561 | 10.5622 | 4.9794
967 | 38684.566 | 10.5632 | 4.9676 |  |  | 967 | 38684.566 | 10.5632 | 4.9809 |  |  | 967 | 38684.566 | 10.5632 | 4.9793
968 | 38724.571 | 10.5642 | 4.9675 |  |  | 968 | 38724.571 | 10.5642 | 4.9808 |  |  | 968 | 38724.571 | 10.5642 | 4.9792
969 | 38764.575 | 10.5653 | 4.9674 |  |  | 969 | 38764.575 | 10.5653 | 4.9807 |  |  | 969 | 38764.575 | 10.5653 | 4.9791
970 | 38804.58 | 10.5663 | 4.9673 |  |  | 970 | 38804.58 | 10.5663 | 4.9806 |  |  | 970 | 38804.58 | 10.5663 | 4.979
971 | 38844.585 | 10.5673 | 4.9672 |  |  | 971 | 38844.585 | 10.5673 | 4.9805 |  |  | 971 | 38844.585 | 10.5673 | 4.9789
972 | 38884.589 | 10.5684 | 4.9671 |  |  | 972 | 38884.589 | 10.5684 | 4.9804 |  |  | 972 | 38884.589 | 10.5684 | 4.9788
973 | 38924.594 | 10.5694 | 4.967 |  |  | 973 | 38924.594 | 10.5694 | 4.9803 |  |  | 973 | 38924.594 | 10.5694 | 4.9787
974 | 38964.599 | 10.5704 | 4.9669 |  |  | 974 | 38964.599 | 10.5704 | 4.9802 |  |  | 974 | 38964.599 | 10.5704 | 4.9786
975 | 39004.604 | 10.5714 | 4.9668 |  |  | 975 | 39004.604 | 10.5714 | 4.9801 |  |  | 975 | 39004.604 | 10.5714 | 4.9785
976 | 39044.608 | 10.5725 | 4.9667 |  |  | 976 | 39044.608 | 10.5725 | 4.98 |  |  | 976 | 39044.608 | 10.5725 | 4.9784
977 | 39084.613 | 10.5735 | 4.9666 |  |  | 977 | 39084.613 | 10.5735 | 4.9799 |  |  | 977 | 39084.613 | 10.5735 | 4.9783
978 | 39124.618 | 10.5745 | 4.9665 |  |  | 978 | 39124.618 | 10.5745 | 4.9798 |  |  | 978 | 39124.618 | 10.5745 | 4.9782
979 | 39164.622 | 10.5755 | 4.9664 |  |  | 979 | 39164.622 | 10.5755 | 4.9797 |  |  | 979 | 39164.622 | 10.5755 | 4.9781
980 | 39204.627 | 10.5766 | 4.9663 |  |  | 980 | 39204.627 | 10.5766 | 4.9796 |  |  | 980 | 39204.627 | 10.5766 | 4.978
981 | 39244.632 | 10.5776 | 4.9662 |  |  | 981 | 39244.632 | 10.5776 | 4.9795 |  |  | 981 | 39244.632 | 10.5776 | 4.9779
982 | 39284.637 | 10.5786 | 4.9662 |  |  | 982 | 39284.637 | 10.5786 | 4.9794 |  |  | 982 | 39284.637 | 10.5786 | 4.9778
983 | 39324.641 | 10.5796 | 4.9661 |  |  | 983 | 39324.641 | 10.5796 | 4.9793 |  |  | 983 | 39324.641 | 10.5796 | 4.9777
984 | 39364.646 | 10.5806 | 4.966 |  |  | 984 | 39364.646 | 10.5806 | 4.9792 |  |  | 984 | 39364.646 | 10.5806 | 4.9776
985 | 39404.651 | 10.5816 | 4.9659 |  |  | 985 | 39404.651 | 10.5816 | 4.979 |  |  | 985 | 39404.651 | 10.5816 | 4.9775
986 | 39444.656 | 10.5827 | 4.9658 |  |  | 986 | 39444.656 | 10.5827 | 4.9789 |  |  | 986 | 39444.656 | 10.5827 | 4.9774
987 | 39484.66 | 10.5837 | 4.9657 |  |  | 987 | 39484.66 | 10.5837 | 4.9788 |  |  | 987 | 39484.66 | 10.5837 | 4.9773
988 | 39524.665 | 10.5847 | 4.9656 |  |  | 988 | 39524.665 | 10.5847 | 4.9787 |  |  | 988 | 39524.665 | 10.5847 | 4.9772
989 | 39564.67 | 10.5857 | 4.9655 |  |  | 989 | 39564.67 | 10.5857 | 4.9786 |  |  | 989 | 39564.67 | 10.5857 | 4.9771
990 | 39604.674 | 10.5867 | 4.9654 |  |  | 990 | 39604.674 | 10.5867 | 4.9785 |  |  | 990 | 39604.674 | 10.5867 | 4.977

## Run4
MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR
File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run4\4 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run4\5 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run4\6 01 001.RAW
Cursor Location: |  |  | (179, 193) |  |  | Cursor Location: |  |  | (179, 193) |  |  | Cursor Location: |  |  | (178, 193)
Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15)
Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev
1 | 40.002 | 3.6889 | 7.7149 | 5.1106 | 0.3237 | 1 | 40.002 | 3.6889 | 7.7294 | 5.1255 | 0.3187 | 1 | 40.002 | 3.6889 | 7.7331 | 5.1288 | 0.32069
2 | 80.004 | 4.3821 | 7.3865 |  |  | 2 | 80.004 | 4.3821 | 7.3942 |  |  | 2 | 80.004 | 4.3821 | 7.4013
3 | 120.007 | 4.7875 | 7.1895 |  |  | 3 | 120.007 | 4.7875 | 7.2001 |  |  | 3 | 120.007 | 4.7875 | 7.207
4 | 160.009 | 5.0752 | 7.0483 |  |  | 4 | 160.009 | 5.0752 | 7.0604 |  |  | 4 | 160.009 | 5.0752 | 7.0668
5 | 200.011 | 5.2984 | 6.9381 |  |  | 5 | 200.011 | 5.2984 | 6.9505 |  |  | 5 | 200.011 | 5.2984 | 6.9565
6 | 240.013 | 5.4807 | 6.8477 |  |  | 6 | 240.013 | 5.4807 | 6.8597 |  |  | 6 | 240.013 | 5.4807 | 6.8654
7 | 280.016 | 5.6348 | 6.771 |  |  | 7 | 280.016 | 5.6348 | 6.7822 |  |  | 7 | 280.016 | 5.6348 | 6.7878
8 | 320.018 | 5.7684 | 6.7043 |  |  | 8 | 320.018 | 5.7684 | 6.7147 |  |  | 8 | 320.018 | 5.7684 | 6.7202
9 | 360.02 | 5.8862 | 6.6452 |  |  | 9 | 360.02 | 5.8862 | 6.6548 |  |  | 9 | 360.02 | 5.8862 | 6.6603
10 | 400.022 | 5.9915 | 6.5923 |  |  | 10 | 400.022 | 5.9915 | 6.601 |  |  | 10 | 400.022 | 5.9915 | 6.6066
11 | 440.025 | 6.0868 | 6.5442 |  |  | 11 | 440.025 | 6.0868 | 6.5523 |  |  | 11 | 440.025 | 6.0868 | 6.5579
12 | 480.027 | 6.1738 | 6.5003 |  |  | 12 | 480.027 | 6.1738 | 6.5076 |  |  | 12 | 480.027 | 6.1738 | 6.5134
13 | 520.029 | 6.2539 | 6.4597 |  |  | 13 | 520.029 | 6.2539 | 6.4665 |  |  | 13 | 520.029 | 6.2539 | 6.4725
14 | 560.031 | 6.328 | 6.4221 |  |  | 14 | 560.031 | 6.328 | 6.4284 |  |  | 14 | 560.031 | 6.328 | 6.4345
15 | 600.033 | 6.397 | 6.3869 |  |  | 15 | 600.033 | 6.397 | 6.3929 |  |  | 15 | 600.033 | 6.397 | 6.3991
16 | 640.036 | 6.4615 | 6.354 |  |  | 16 | 640.036 | 6.4615 | 6.3596 |  |  | 16 | 640.036 | 6.4615 | 6.366
17 | 680.038 | 6.5221 | 6.323 |  |  | 17 | 680.038 | 6.5221 | 6.3283 |  |  | 17 | 680.038 | 6.5221 | 6.3349
18 | 720.04 | 6.5793 | 6.2937 |  |  | 18 | 720.04 | 6.5793 | 6.2989 |  |  | 18 | 720.04 | 6.5793 | 6.3056
19 | 760.042 | 6.6334 | 6.266 |  |  | 19 | 760.042 | 6.6334 | 6.271 |  |  | 19 | 760.042 | 6.6334 | 6.2779
20 | 800.045 | 6.6847 | 6.2396 |  |  | 20 | 800.045 | 6.6847 | 6.2445 |  |  | 20 | 800.045 | 6.6847 | 6.2516
21 | 840.047 | 6.7335 | 6.2145 |  |  | 21 | 840.047 | 6.7335 | 6.2193 |  |  | 21 | 840.047 | 6.7335 | 6.2265
22 | 880.049 | 6.78 | 6.1905 |  |  | 22 | 880.049 | 6.78 | 6.1953 |  |  | 22 | 880.049 | 6.78 | 6.2027
23 | 920.051 | 6.8244 | 6.1676 |  |  | 23 | 920.051 | 6.8244 | 6.1724 |  |  | 23 | 920.051 | 6.8244 | 6.1799
24 | 960.053 | 6.867 | 6.1456 |  |  | 24 | 960.053 | 6.867 | 6.1505 |  |  | 24 | 960.053 | 6.867 | 6.1581
25 | 1000.056 | 6.9078 | 6.1245 |  |  | 25 | 1000.056 | 6.9078 | 6.1294 |  |  | 25 | 1000.056 | 6.9078 | 6.1372
26 | 1040.058 | 6.947 | 6.1042 |  |  | 26 | 1040.058 | 6.947 | 6.1092 |  |  | 26 | 1040.058 | 6.947 | 6.1171
27 | 1080.06 | 6.9848 | 6.0847 |  |  | 27 | 1080.06 | 6.9848 | 6.0898 |  |  | 27 | 1080.06 | 6.9848 | 6.0978
28 | 1120.062 | 7.0211 | 6.0659 |  |  | 28 | 1120.062 | 7.0211 | 6.0711 |  |  | 28 | 1120.062 | 7.0211 | 6.0792
29 | 1160.065 | 7.0562 | 6.0477 |  |  | 29 | 1160.065 | 7.0562 | 6.053 |  |  | 29 | 1160.065 | 7.0562 | 6.0613
30 | 1200.067 | 7.0901 | 6.0301 |  |  | 30 | 1200.067 | 7.0901 | 6.0356 |  |  | 30 | 1200.067 | 7.0901 | 6.0439
31 | 1240.069 | 7.1229 | 6.0131 |  |  | 31 | 1240.069 | 7.1229 | 6.0187 |  |  | 31 | 1240.069 | 7.1229 | 6.0272
32 | 1280.071 | 7.1547 | 5.9967 |  |  | 32 | 1280.071 | 7.1547 | 6.0024 |  |  | 32 | 1280.071 | 7.1547 | 6.011
33 | 1320.074 | 7.1854 | 5.9807 |  |  | 33 | 1320.074 | 7.1854 | 5.9867 |  |  | 33 | 1320.074 | 7.1854 | 5.9953
34 | 1360.076 | 7.2153 | 5.9653 |  |  | 34 | 1360.076 | 7.2153 | 5.9713 |  |  | 34 | 1360.076 | 7.2153 | 5.9801
35 | 1400.078 | 7.2443 | 5.9502 |  |  | 35 | 1400.078 | 7.2443 | 5.9565 |  |  | 35 | 1400.078 | 7.2443 | 5.9653
36 | 1440.08 | 7.2725 | 5.9357 |  |  | 36 | 1440.08 | 7.2725 | 5.9421 |  |  | 36 | 1440.08 | 7.2725 | 5.9509
37 | 1480.082 | 7.2999 | 5.9215 |  |  | 37 | 1480.082 | 7.2999 | 5.9281 |  |  | 37 | 1480.082 | 7.2999 | 5.937
38 | 1520.085 | 7.3265 | 5.9077 |  |  | 38 | 1520.085 | 7.3265 | 5.9144 |  |  | 38 | 1520.085 | 7.3265 | 5.9234
39 | 1560.087 | 7.3525 | 5.8943 |  |  | 39 | 1560.087 | 7.3525 | 5.9012 |  |  | 39 | 1560.087 | 7.3525 | 5.9102
40 | 1600.089 | 7.3778 | 5.8812 |  |  | 40 | 1600.089 | 7.3778 | 5.8882 |  |  | 40 | 1600.089 | 7.3778 | 5.8974
41 | 1640.091 | 7.4025 | 5.8684 |  |  | 41 | 1640.091 | 7.4025 | 5.8757 |  |  | 41 | 1640.091 | 7.4025 | 5.8848
42 | 1680.094 | 7.4266 | 5.856 |  |  | 42 | 1680.094 | 7.4266 | 5.8634 |  |  | 42 | 1680.094 | 7.4266 | 5.8726
43 | 1720.096 | 7.4501 | 5.8439 |  |  | 43 | 1720.096 | 7.4501 | 5.8514 |  |  | 43 | 1720.096 | 7.4501 | 5.8607
44 | 1760.098 | 7.4731 | 5.8321 |  |  | 44 | 1760.098 | 7.4731 | 5.8397 |  |  | 44 | 1760.098 | 7.4731 | 5.8491
45 | 1800.1 | 7.4956 | 5.8205 |  |  | 45 | 1800.1 | 7.4956 | 5.8283 |  |  | 45 | 1800.1 | 7.4956 | 5.8377
46 | 1840.102 | 7.5176 | 5.8093 |  |  | 46 | 1840.102 | 7.5176 | 5.8172 |  |  | 46 | 1840.102 | 7.5176 | 5.8266
47 | 1880.105 | 7.5391 | 5.7982 |  |  | 47 | 1880.105 | 7.5391 | 5.8063 |  |  | 47 | 1880.105 | 7.5391 | 5.8157
48 | 1920.107 | 7.5601 | 5.7875 |  |  | 48 | 1920.107 | 7.5601 | 5.7957 |  |  | 48 | 1920.107 | 7.5601 | 5.8051
49 | 1960.109 | 7.5808 | 5.7769 |  |  | 49 | 1960.109 | 7.5808 | 5.7852 |  |  | 49 | 1960.109 | 7.5808 | 5.7947
50 | 2000.111 | 7.601 | 5.7666 |  |  | 50 | 2000.111 | 7.601 | 5.7751 |  |  | 50 | 2000.111 | 7.601 | 5.7846
51 | 2040.114 | 7.6208 | 5.7565 |  |  | 51 | 2040.114 | 7.6208 | 5.7651 |  |  | 51 | 2040.114 | 7.6208 | 5.7746
52 | 2080.116 | 7.6402 | 5.7467 |  |  | 52 | 2080.116 | 7.6402 | 5.7553 |  |  | 52 | 2080.116 | 7.6402 | 5.7649
53 | 2120.118 | 7.6592 | 5.737 |  |  | 53 | 2120.118 | 7.6592 | 5.7457 |  |  | 53 | 2120.118 | 7.6592 | 5.7553
54 | 2160.12 | 7.6779 | 5.7275 |  |  | 54 | 2160.12 | 7.6779 | 5.7364 |  |  | 54 | 2160.12 | 7.6779 | 5.746
55 | 2200.123 | 7.6963 | 5.7183 |  |  | 55 | 2200.123 | 7.6963 | 5.7272 |  |  | 55 | 2200.123 | 7.6963 | 5.7368
56 | 2240.125 | 7.7143 | 5.7092 |  |  | 56 | 2240.125 | 7.7143 | 5.7182 |  |  | 56 | 2240.125 | 7.7143 | 5.7278
57 | 2280.127 | 7.732 | 5.7002 |  |  | 57 | 2280.127 | 7.732 | 5.7093 |  |  | 57 | 2280.127 | 7.732 | 5.719
58 | 2320.129 | 7.7494 | 5.6915 |  |  | 58 | 2320.129 | 7.7494 | 5.7006 |  |  | 58 | 2320.129 | 7.7494 | 5.7103
59 | 2360.131 | 7.7665 | 5.6829 |  |  | 59 | 2360.131 | 7.7665 | 5.6921 |  |  | 59 | 2360.131 | 7.7665 | 5.7018
60 | 2400.134 | 7.7833 | 5.6745 |  |  | 60 | 2400.134 | 7.7833 | 5.6838 |  |  | 60 | 2400.134 | 7.7833 | 5.6934
61 | 2440.136 | 7.7998 | 5.6663 |  |  | 61 | 2440.136 | 7.7998 | 5.6756 |  |  | 61 | 2440.136 | 7.7998 | 5.6852
62 | 2480.138 | 7.8161 | 5.6582 |  |  | 62 | 2480.138 | 7.8161 | 5.6675 |  |  | 62 | 2480.138 | 7.8161 | 5.6772
63 | 2520.14 | 7.8321 | 5.6502 |  |  | 63 | 2520.14 | 7.8321 | 5.6596 |  |  | 63 | 2520.14 | 7.8321 | 5.6693
64 | 2560.143 | 7.8478 | 5.6424 |  |  | 64 | 2560.143 | 7.8478 | 5.6518 |  |  | 64 | 2560.143 | 7.8478 | 5.6615
65 | 2600.145 | 7.8633 | 5.6347 |  |  | 65 | 2600.145 | 7.8633 | 5.6442 |  |  | 65 | 2600.145 | 7.8633 | 5.6538
66 | 2640.147 | 7.8786 | 5.6272 |  |  | 66 | 2640.147 | 7.8786 | 5.6367 |  |  | 66 | 2640.147 | 7.8786 | 5.6463
67 | 2680.149 | 7.8936 | 5.6198 |  |  | 67 | 2680.149 | 7.8936 | 5.6293 |  |  | 67 | 2680.149 | 7.8936 | 5.6389
68 | 2720.151 | 7.9084 | 5.6125 |  |  | 68 | 2720.151 | 7.9084 | 5.622 |  |  | 68 | 2720.151 | 7.9084 | 5.6317
69 | 2760.154 | 7.923 | 5.6054 |  |  | 69 | 2760.154 | 7.923 | 5.6149 |  |  | 69 | 2760.154 | 7.923 | 5.6245
70 | 2800.156 | 7.9374 | 5.5983 |  |  | 70 | 2800.156 | 7.9374 | 5.6079 |  |  | 70 | 2800.156 | 7.9374 | 5.6175
71 | 2840.158 | 7.9516 | 5.5914 |  |  | 71 | 2840.158 | 7.9516 | 5.601 |  |  | 71 | 2840.158 | 7.9516 | 5.6106
72 | 2880.16 | 7.9656 | 5.5846 |  |  | 72 | 2880.16 | 7.9656 | 5.5941 |  |  | 72 | 2880.16 | 7.9656 | 5.6037
73 | 2920.163 | 7.9794 | 5.5779 |  |  | 73 | 2920.163 | 7.9794 | 5.5875 |  |  | 73 | 2920.163 | 7.9794 | 5.597
74 | 2960.165 | 7.993 | 5.5714 |  |  | 74 | 2960.165 | 7.993 | 5.5809 |  |  | 74 | 2960.165 | 7.993 | 5.5904
75 | 3000.167 | 8.0064 | 5.5649 |  |  | 75 | 3000.167 | 8.0064 | 5.5744 |  |  | 75 | 3000.167 | 8.0064 | 5.5839
76 | 3040.169 | 8.0197 | 5.5585 |  |  | 76 | 3040.169 | 8.0197 | 5.568 |  |  | 76 | 3040.169 | 8.0197 | 5.5775
77 | 3080.172 | 8.0327 | 5.5523 |  |  | 77 | 3080.172 | 8.0327 | 5.5617 |  |  | 77 | 3080.172 | 8.0327 | 5.5712
78 | 3120.174 | 8.0456 | 5.5461 |  |  | 78 | 3120.174 | 8.0456 | 5.5555 |  |  | 78 | 3120.174 | 8.0456 | 5.565
79 | 3160.176 | 8.0584 | 5.54 |  |  | 79 | 3160.176 | 8.0584 | 5.5494 |  |  | 79 | 3160.176 | 8.0584 | 5.5589
80 | 3200.178 | 8.071 | 5.5341 |  |  | 80 | 3200.178 | 8.071 | 5.5434 |  |  | 80 | 3200.178 | 8.071 | 5.5529
81 | 3240.18 | 8.0834 | 5.5282 |  |  | 81 | 3240.18 | 8.0834 | 5.5375 |  |  | 81 | 3240.18 | 8.0834 | 5.5469
82 | 3280.183 | 8.0957 | 5.5224 |  |  | 82 | 3280.183 | 8.0957 | 5.5316 |  |  | 82 | 3280.183 | 8.0957 | 5.541
83 | 3320.185 | 8.1078 | 5.5167 |  |  | 83 | 3320.185 | 8.1078 | 5.5259 |  |  | 83 | 3320.185 | 8.1078 | 5.5353
84 | 3360.187 | 8.1198 | 5.5111 |  |  | 84 | 3360.187 | 8.1198 | 5.5202 |  |  | 84 | 3360.187 | 8.1198 | 5.5296
85 | 3400.189 | 8.1316 | 5.5055 |  |  | 85 | 3400.189 | 8.1316 | 5.5146 |  |  | 85 | 3400.189 | 8.1316 | 5.524
86 | 3440.192 | 8.1433 | 5.5001 |  |  | 86 | 3440.192 | 8.1433 | 5.5091 |  |  | 86 | 3440.192 | 8.1433 | 5.5184
87 | 3480.194 | 8.1548 | 5.4947 |  |  | 87 | 3480.194 | 8.1548 | 5.5037 |  |  | 87 | 3480.194 | 8.1548 | 5.513
88 | 3520.196 | 8.1663 | 5.4894 |  |  | 88 | 3520.196 | 8.1663 | 5.4983 |  |  | 88 | 3520.196 | 8.1663 | 5.5076
89 | 3560.198 | 8.1776 | 5.4842 |  |  | 89 | 3560.198 | 8.1776 | 5.493 |  |  | 89 | 3560.198 | 8.1776 | 5.5023
90 | 3600.201 | 8.1887 | 5.4791 |  |  | 90 | 3600.201 | 8.1887 | 5.4878 |  |  | 90 | 3600.201 | 8.1887 | 5.497
91 | 3640.203 | 8.1998 | 5.474 |  |  | 91 | 3640.203 | 8.1998 | 5.4827 |  |  | 91 | 3640.203 | 8.1998 | 5.4919
92 | 3680.205 | 8.2107 | 5.469 |  |  | 92 | 3680.205 | 8.2107 | 5.4776 |  |  | 92 | 3680.205 | 8.2107 | 5.4868
93 | 3720.207 | 8.2215 | 5.4641 |  |  | 93 | 3720.207 | 8.2215 | 5.4726 |  |  | 93 | 3720.207 | 8.2215 | 5.4817
94 | 3760.209 | 8.2322 | 5.4592 |  |  | 94 | 3760.209 | 8.2322 | 5.4676 |  |  | 94 | 3760.209 | 8.2322 | 5.4768
95 | 3800.212 | 8.2428 | 5.4544 |  |  | 95 | 3800.212 | 8.2428 | 5.4628 |  |  | 95 | 3800.212 | 8.2428 | 5.4719
96 | 3840.214 | 8.2533 | 5.4497 |  |  | 96 | 3840.214 | 8.2533 | 5.458 |  |  | 96 | 3840.214 | 8.2533 | 5.467
97 | 3880.216 | 8.2636 | 5.445 |  |  | 97 | 3880.216 | 8.2636 | 5.4532 |  |  | 97 | 3880.216 | 8.2636 | 5.4623
98 | 3920.218 | 8.2739 | 5.4404 |  |  | 98 | 3920.218 | 8.2739 | 5.4485 |  |  | 98 | 3920.218 | 8.2739 | 5.4576
99 | 3960.221 | 8.2841 | 5.4359 |  |  | 99 | 3960.221 | 8.2841 | 5.4439 |  |  | 99 | 3960.221 | 8.2841 | 5.4529
100 | 4000.223 | 8.2941 | 5.4314 |  |  | 100 | 4000.223 | 8.2941 | 5.4393 |  |  | 100 | 4000.223 | 8.2941 | 5.4483
101 | 4040.225 | 8.3041 | 5.427 |  |  | 101 | 4040.225 | 8.3041 | 5.4348 |  |  | 101 | 4040.225 | 8.3041 | 5.4438
102 | 4080.227 | 8.3139 | 5.4227 |  |  | 102 | 4080.227 | 8.3139 | 5.4304 |  |  | 102 | 4080.227 | 8.3139 | 5.4393
103 | 4120.229 | 8.3237 | 5.4184 |  |  | 103 | 4120.229 | 8.3237 | 5.426 |  |  | 103 | 4120.229 | 8.3237 | 5.4349
104 | 4160.232 | 8.3333 | 5.4142 |  |  | 104 | 4160.232 | 8.3333 | 5.4216 |  |  | 104 | 4160.232 | 8.3333 | 5.4305
105 | 4200.234 | 8.3429 | 5.41 |  |  | 105 | 4200.234 | 8.3429 | 5.4174 |  |  | 105 | 4200.234 | 8.3429 | 5.4262
106 | 4240.236 | 8.3524 | 5.4058 |  |  | 106 | 4240.236 | 8.3524 | 5.4131 |  |  | 106 | 4240.236 | 8.3524 | 5.4219
107 | 4280.238 | 8.3618 | 5.4018 |  |  | 107 | 4280.238 | 8.3618 | 5.4089 |  |  | 107 | 4280.238 | 8.3618 | 5.4177
108 | 4320.241 | 8.3711 | 5.3977 |  |  | 108 | 4320.241 | 8.3711 | 5.4048 |  |  | 108 | 4320.241 | 8.3711 | 5.4136
109 | 4360.243 | 8.3803 | 5.3938 |  |  | 109 | 4360.243 | 8.3803 | 5.4007 |  |  | 109 | 4360.243 | 8.3803 | 5.4095
110 | 4400.245 | 8.3894 | 5.3899 |  |  | 110 | 4400.245 | 8.3894 | 5.3967 |  |  | 110 | 4400.245 | 8.3894 | 5.4054
111 | 4440.247 | 8.3985 | 5.386 |  |  | 111 | 4440.247 | 8.3985 | 5.3927 |  |  | 111 | 4440.247 | 8.3985 | 5.4014
112 | 4480.25 | 8.4074 | 5.3822 |  |  | 112 | 4480.25 | 8.4074 | 5.3888 |  |  | 112 | 4480.25 | 8.4074 | 5.3974
113 | 4520.252 | 8.4163 | 5.3784 |  |  | 113 | 4520.252 | 8.4163 | 5.3849 |  |  | 113 | 4520.252 | 8.4163 | 5.3935
114 | 4560.254 | 8.4251 | 5.3747 |  |  | 114 | 4560.254 | 8.4251 | 5.3811 |  |  | 114 | 4560.254 | 8.4251 | 5.3896
115 | 4600.256 | 8.4339 | 5.371 |  |  | 115 | 4600.256 | 8.4339 | 5.3773 |  |  | 115 | 4600.256 | 8.4339 | 5.3858
116 | 4640.258 | 8.4425 | 5.3674 |  |  | 116 | 4640.258 | 8.4425 | 5.3735 |  |  | 116 | 4640.258 | 8.4425 | 5.382
117 | 4680.261 | 8.4511 | 5.3638 |  |  | 117 | 4680.261 | 8.4511 | 5.3698 |  |  | 117 | 4680.261 | 8.4511 | 5.3783
118 | 4720.263 | 8.4596 | 5.3602 |  |  | 118 | 4720.263 | 8.4596 | 5.3661 |  |  | 118 | 4720.263 | 8.4596 | 5.3746
119 | 4760.265 | 8.4681 | 5.3567 |  |  | 119 | 4760.265 | 8.4681 | 5.3625 |  |  | 119 | 4760.265 | 8.4681 | 5.371
120 | 4800.267 | 8.4764 | 5.3533 |  |  | 120 | 4800.267 | 8.4764 | 5.3589 |  |  | 120 | 4800.267 | 8.4764 | 5.3674
121 | 4840.27 | 8.4847 | 5.3499 |  |  | 121 | 4840.27 | 8.4847 | 5.3554 |  |  | 121 | 4840.27 | 8.4847 | 5.3638
122 | 4880.272 | 8.493 | 5.3465 |  |  | 122 | 4880.272 | 8.493 | 5.3519 |  |  | 122 | 4880.272 | 8.493 | 5.3603
123 | 4920.274 | 8.5011 | 5.3432 |  |  | 123 | 4920.274 | 8.5011 | 5.3485 |  |  | 123 | 4920.274 | 8.5011 | 5.3568
124 | 4960.276 | 8.5092 | 5.3399 |  |  | 124 | 4960.276 | 8.5092 | 5.345 |  |  | 124 | 4960.276 | 8.5092 | 5.3533
125 | 5000.278 | 8.5172 | 5.3366 |  |  | 125 | 5000.278 | 8.5172 | 5.3417 |  |  | 125 | 5000.278 | 8.5172 | 5.3499
126 | 5040.281 | 8.5252 | 5.3334 |  |  | 126 | 5040.281 | 8.5252 | 5.3383 |  |  | 126 | 5040.281 | 8.5252 | 5.3466
127 | 5080.283 | 8.5331 | 5.3302 |  |  | 127 | 5080.283 | 8.5331 | 5.335 |  |  | 127 | 5080.283 | 8.5331 | 5.3432
128 | 5120.285 | 8.541 | 5.3271 |  |  | 128 | 5120.285 | 8.541 | 5.3318 |  |  | 128 | 5120.285 | 8.541 | 5.3399
129 | 5160.287 | 8.5487 | 5.324 |  |  | 129 | 5160.287 | 8.5487 | 5.3285 |  |  | 129 | 5160.287 | 8.5487 | 5.3367
130 | 5200.29 | 8.5565 | 5.3209 |  |  | 130 | 5200.29 | 8.5565 | 5.3253 |  |  | 130 | 5200.29 | 8.5565 | 5.3335
131 | 5240.292 | 8.5641 | 5.3179 |  |  | 131 | 5240.292 | 8.5641 | 5.3222 |  |  | 131 | 5240.292 | 8.5641 | 5.3303
132 | 5280.294 | 8.5717 | 5.3149 |  |  | 132 | 5280.294 | 8.5717 | 5.319 |  |  | 132 | 5280.294 | 8.5717 | 5.3271
133 | 5320.296 | 8.5793 | 5.3119 |  |  | 133 | 5320.296 | 8.5793 | 5.316 |  |  | 133 | 5320.296 | 8.5793 | 5.324
134 | 5360.299 | 8.5868 | 5.309 |  |  | 134 | 5360.299 | 8.5868 | 5.3129 |  |  | 134 | 5360.299 | 8.5868 | 5.3209
135 | 5400.301 | 8.5942 | 5.3061 |  |  | 135 | 5400.301 | 8.5942 | 5.3099 |  |  | 135 | 5400.301 | 8.5942 | 5.3179
136 | 5440.303 | 8.6016 | 5.3032 |  |  | 136 | 5440.303 | 8.6016 | 5.3069 |  |  | 136 | 5440.303 | 8.6016 | 5.3148
137 | 5480.305 | 8.6089 | 5.3004 |  |  | 137 | 5480.305 | 8.6089 | 5.3039 |  |  | 137 | 5480.305 | 8.6089 | 5.3119
138 | 5520.307 | 8.6162 | 5.2976 |  |  | 138 | 5520.307 | 8.6162 | 5.301 |  |  | 138 | 5520.307 | 8.6162 | 5.3089
139 | 5560.31 | 8.6234 | 5.2948 |  |  | 139 | 5560.31 | 8.6234 | 5.2981 |  |  | 139 | 5560.31 | 8.6234 | 5.306
140 | 5600.312 | 8.6306 | 5.2921 |  |  | 140 | 5600.312 | 8.6306 | 5.2953 |  |  | 140 | 5600.312 | 8.6306 | 5.3031
141 | 5640.314 | 8.6377 | 5.2894 |  |  | 141 | 5640.314 | 8.6377 | 5.2924 |  |  | 141 | 5640.314 | 8.6377 | 5.3002
142 | 5680.316 | 8.6448 | 5.2867 |  |  | 142 | 5680.316 | 8.6448 | 5.2896 |  |  | 142 | 5680.316 | 8.6448 | 5.2974
143 | 5720.319 | 8.6518 | 5.284 |  |  | 143 | 5720.319 | 8.6518 | 5.2869 |  |  | 143 | 5720.319 | 8.6518 | 5.2946
144 | 5760.321 | 8.6587 | 5.2814 |  |  | 144 | 5760.321 | 8.6587 | 5.2841 |  |  | 144 | 5760.321 | 8.6587 | 5.2918
145 | 5800.323 | 8.6657 | 5.2788 |  |  | 145 | 5800.323 | 8.6657 | 5.2814 |  |  | 145 | 5800.323 | 8.6657 | 5.2891
146 | 5840.325 | 8.6725 | 5.2763 |  |  | 146 | 5840.325 | 8.6725 | 5.2787 |  |  | 146 | 5840.325 | 8.6725 | 5.2864
147 | 5880.328 | 8.6794 | 5.2737 |  |  | 147 | 5880.328 | 8.6794 | 5.2761 |  |  | 147 | 5880.328 | 8.6794 | 5.2837
148 | 5920.33 | 8.6861 | 5.2712 |  |  | 148 | 5920.33 | 8.6861 | 5.2734 |  |  | 148 | 5920.33 | 8.6861 | 5.2811
149 | 5960.332 | 8.6929 | 5.2687 |  |  | 149 | 5960.332 | 8.6929 | 5.2708 |  |  | 149 | 5960.332 | 8.6929 | 5.2784
150 | 6000.334 | 8.6996 | 5.2663 |  |  | 150 | 6000.334 | 8.6996 | 5.2683 |  |  | 150 | 6000.334 | 8.6996 | 5.2758
151 | 6040.336 | 8.7062 | 5.2639 |  |  | 151 | 6040.336 | 8.7062 | 5.2657 |  |  | 151 | 6040.336 | 8.7062 | 5.2733
152 | 6080.339 | 8.7128 | 5.2615 |  |  | 152 | 6080.339 | 8.7128 | 5.2632 |  |  | 152 | 6080.339 | 8.7128 | 5.2707
153 | 6120.341 | 8.7194 | 5.2591 |  |  | 153 | 6120.341 | 8.7194 | 5.2607 |  |  | 153 | 6120.341 | 8.7194 | 5.2682
154 | 6160.343 | 8.7259 | 5.2567 |  |  | 154 | 6160.343 | 8.7259 | 5.2582 |  |  | 154 | 6160.343 | 8.7259 | 5.2657
155 | 6200.345 | 8.7324 | 5.2544 |  |  | 155 | 6200.345 | 8.7324 | 5.2558 |  |  | 155 | 6200.345 | 8.7324 | 5.2632
156 | 6240.348 | 8.7388 | 5.2521 |  |  | 156 | 6240.348 | 8.7388 | 5.2534 |  |  | 156 | 6240.348 | 8.7388 | 5.2608
157 | 6280.35 | 8.7452 | 5.2498 |  |  | 157 | 6280.35 | 8.7452 | 5.251 |  |  | 157 | 6280.35 | 8.7452 | 5.2584
158 | 6320.352 | 8.7515 | 5.2476 |  |  | 158 | 6320.352 | 8.7515 | 5.2486 |  |  | 158 | 6320.352 | 8.7515 | 5.256
159 | 6360.354 | 8.7578 | 5.2453 |  |  | 159 | 6360.354 | 8.7578 | 5.2463 |  |  | 159 | 6360.354 | 8.7578 | 5.2536
160 | 6400.356 | 8.7641 | 5.2431 |  |  | 160 | 6400.356 | 8.7641 | 5.2439 |  |  | 160 | 6400.356 | 8.7641 | 5.2513
161 | 6440.359 | 8.7703 | 5.2409 |  |  | 161 | 6440.359 | 8.7703 | 5.2416 |  |  | 161 | 6440.359 | 8.7703 | 5.2489
162 | 6480.361 | 8.7765 | 5.2388 |  |  | 162 | 6480.361 | 8.7765 | 5.2394 |  |  | 162 | 6480.361 | 8.7765 | 5.2466
163 | 6520.363 | 8.7827 | 5.2366 |  |  | 163 | 6520.363 | 8.7827 | 5.2371 |  |  | 163 | 6520.363 | 8.7827 | 5.2444
164 | 6560.365 | 8.7888 | 5.2345 |  |  | 164 | 6560.365 | 8.7888 | 5.2349 |  |  | 164 | 6560.365 | 8.7888 | 5.2421
165 | 6600.368 | 8.7949 | 5.2324 |  |  | 165 | 6600.368 | 8.7949 | 5.2327 |  |  | 165 | 6600.368 | 8.7949 | 5.2399
166 | 6640.37 | 8.8009 | 5.2304 |  |  | 166 | 6640.37 | 8.8009 | 5.2305 |  |  | 166 | 6640.37 | 8.8009 | 5.2377
167 | 6680.372 | 8.8069 | 5.2283 |  |  | 167 | 6680.372 | 8.8069 | 5.2283 |  |  | 167 | 6680.372 | 8.8069 | 5.2355
168 | 6720.374 | 8.8129 | 5.2263 |  |  | 168 | 6720.374 | 8.8129 | 5.2262 |  |  | 168 | 6720.374 | 8.8129 | 5.2333
169 | 6760.377 | 8.8188 | 5.2243 |  |  | 169 | 6760.377 | 8.8188 | 5.2241 |  |  | 169 | 6760.377 | 8.8188 | 5.2312
170 | 6800.379 | 8.8247 | 5.2223 |  |  | 170 | 6800.379 | 8.8247 | 5.222 |  |  | 170 | 6800.379 | 8.8247 | 5.229
171 | 6840.381 | 8.8306 | 5.2203 |  |  | 171 | 6840.381 | 8.8306 | 5.2199 |  |  | 171 | 6840.381 | 8.8306 | 5.2269
172 | 6880.383 | 8.8364 | 5.2183 |  |  | 172 | 6880.383 | 8.8364 | 5.2178 |  |  | 172 | 6880.383 | 8.8364 | 5.2249
173 | 6920.385 | 8.8422 | 5.2164 |  |  | 173 | 6920.385 | 8.8422 | 5.2158 |  |  | 173 | 6920.385 | 8.8422 | 5.2228
174 | 6960.388 | 8.848 | 5.2145 |  |  | 174 | 6960.388 | 8.848 | 5.2138 |  |  | 174 | 6960.388 | 8.848 | 5.2208
175 | 7000.39 | 8.8537 | 5.2126 |  |  | 175 | 7000.39 | 8.8537 | 5.2118 |  |  | 175 | 7000.39 | 8.8537 | 5.2187
176 | 7040.392 | 8.8594 | 5.2107 |  |  | 176 | 7040.392 | 8.8594 | 5.2098 |  |  | 176 | 7040.392 | 8.8594 | 5.2167
177 | 7080.394 | 8.8651 | 5.2089 |  |  | 177 | 7080.394 | 8.8651 | 5.2078 |  |  | 177 | 7080.394 | 8.8651 | 5.2148
178 | 7120.397 | 8.8707 | 5.207 |  |  | 178 | 7120.397 | 8.8707 | 5.2059 |  |  | 178 | 7120.397 | 8.8707 | 5.2128
179 | 7160.399 | 8.8763 | 5.2052 |  |  | 179 | 7160.399 | 8.8763 | 5.204 |  |  | 179 | 7160.399 | 8.8763 | 5.2109
180 | 7200.401 | 8.8819 | 5.2034 |  |  | 180 | 7200.401 | 8.8819 | 5.2021 |  |  | 180 | 7200.401 | 8.8819 | 5.2089
181 | 7240.403 | 8.8874 | 5.2016 |  |  | 181 | 7240.403 | 8.8874 | 5.2002 |  |  | 181 | 7240.403 | 8.8874 | 5.207
182 | 7280.405 | 8.8929 | 5.1998 |  |  | 182 | 7280.405 | 8.8929 | 5.1983 |  |  | 182 | 7280.405 | 8.8929 | 5.2051
183 | 7320.408 | 8.8984 | 5.1981 |  |  | 183 | 7320.408 | 8.8984 | 5.1965 |  |  | 183 | 7320.408 | 8.8984 | 5.2033
184 | 7360.41 | 8.9039 | 5.1963 |  |  | 184 | 7360.41 | 8.9039 | 5.1946 |  |  | 184 | 7360.41 | 8.9039 | 5.2014
185 | 7400.412 | 8.9093 | 5.1946 |  |  | 185 | 7400.412 | 8.9093 | 5.1928 |  |  | 185 | 7400.412 | 8.9093 | 5.1996
186 | 7440.414 | 8.9147 | 5.1929 |  |  | 186 | 7440.414 | 8.9147 | 5.191 |  |  | 186 | 7440.414 | 8.9147 | 5.1978
187 | 7480.417 | 8.92 | 5.1912 |  |  | 187 | 7480.417 | 8.92 | 5.1893 |  |  | 187 | 7480.417 | 8.92 | 5.196
188 | 7520.419 | 8.9254 | 5.1896 |  |  | 188 | 7520.419 | 8.9254 | 5.1875 |  |  | 188 | 7520.419 | 8.9254 | 5.1942
189 | 7560.421 | 8.9307 | 5.1879 |  |  | 189 | 7560.421 | 8.9307 | 5.1858 |  |  | 189 | 7560.421 | 8.9307 | 5.1924
190 | 7600.423 | 8.936 | 5.1863 |  |  | 190 | 7600.423 | 8.936 | 5.184 |  |  | 190 | 7600.423 | 8.936 | 5.1907
191 | 7640.426 | 8.9412 | 5.1846 |  |  | 191 | 7640.426 | 8.9412 | 5.1823 |  |  | 191 | 7640.426 | 8.9412 | 5.1889
192 | 7680.428 | 8.9464 | 5.183 |  |  | 192 | 7680.428 | 8.9464 | 5.1806 |  |  | 192 | 7680.428 | 8.9464 | 5.1872
193 | 7720.43 | 8.9516 | 5.1814 |  |  | 193 | 7720.43 | 8.9516 | 5.1789 |  |  | 193 | 7720.43 | 8.9516 | 5.1855
194 | 7760.432 | 8.9568 | 5.1798 |  |  | 194 | 7760.432 | 8.9568 | 5.1773 |  |  | 194 | 7760.432 | 8.9568 | 5.1838
195 | 7800.434 | 8.9619 | 5.1783 |  |  | 195 | 7800.434 | 8.9619 | 5.1756 |  |  | 195 | 7800.434 | 8.9619 | 5.1822
196 | 7840.437 | 8.967 | 5.1767 |  |  | 196 | 7840.437 | 8.967 | 5.174 |  |  | 196 | 7840.437 | 8.967 | 5.1805
197 | 7880.439 | 8.9721 | 5.1752 |  |  | 197 | 7880.439 | 8.9721 | 5.1724 |  |  | 197 | 7880.439 | 8.9721 | 5.1789
198 | 7920.441 | 8.9772 | 5.1737 |  |  | 198 | 7920.441 | 8.9772 | 5.1708 |  |  | 198 | 7920.441 | 8.9772 | 5.1773
199 | 7960.443 | 8.9822 | 5.1722 |  |  | 199 | 7960.443 | 8.9822 | 5.1692 |  |  | 199 | 7960.443 | 8.9822 | 5.1757
200 | 8000.446 | 8.9873 | 5.1707 |  |  | 200 | 8000.446 | 8.9873 | 5.1676 |  |  | 200 | 8000.446 | 8.9873 | 5.1741
201 | 8040.448 | 8.9922 | 5.1692 |  |  | 201 | 8040.448 | 8.9922 | 5.1661 |  |  | 201 | 8040.448 | 8.9922 | 5.1725
202 | 8080.45 | 8.9972 | 5.1677 |  |  | 202 | 8080.45 | 8.9972 | 5.1645 |  |  | 202 | 8080.45 | 8.9972 | 5.1709
203 | 8120.452 | 9.0021 | 5.1663 |  |  | 203 | 8120.452 | 9.0021 | 5.163 |  |  | 203 | 8120.452 | 9.0021 | 5.1694
204 | 8160.454 | 9.0071 | 5.1648 |  |  | 204 | 8160.454 | 9.0071 | 5.1615 |  |  | 204 | 8160.454 | 9.0071 | 5.1679
205 | 8200.457 | 9.0119 | 5.1634 |  |  | 205 | 8200.457 | 9.0119 | 5.16 |  |  | 205 | 8200.457 | 9.0119 | 5.1663
206 | 8240.459 | 9.0168 | 5.162 |  |  | 206 | 8240.459 | 9.0168 | 5.1585 |  |  | 206 | 8240.459 | 9.0168 | 5.1648
207 | 8280.461 | 9.0217 | 5.1606 |  |  | 207 | 8280.461 | 9.0217 | 5.157 |  |  | 207 | 8280.461 | 9.0217 | 5.1633
208 | 8320.463 | 9.0265 | 5.1592 |  |  | 208 | 8320.463 | 9.0265 | 5.1556 |  |  | 208 | 8320.463 | 9.0265 | 5.1619
209 | 8360.466 | 9.0313 | 5.1578 |  |  | 209 | 8360.466 | 9.0313 | 5.1541 |  |  | 209 | 8360.466 | 9.0313 | 5.1604
210 | 8400.468 | 9.036 | 5.1564 |  |  | 210 | 8400.468 | 9.036 | 5.1527 |  |  | 210 | 8400.468 | 9.036 | 5.159
211 | 8440.47 | 9.0408 | 5.1551 |  |  | 211 | 8440.47 | 9.0408 | 5.1513 |  |  | 211 | 8440.47 | 9.0408 | 5.1575
212 | 8480.472 | 9.0455 | 5.1537 |  |  | 212 | 8480.472 | 9.0455 | 5.1499 |  |  | 212 | 8480.472 | 9.0455 | 5.1561
213 | 8520.475 | 9.0502 | 5.1524 |  |  | 213 | 8520.475 | 9.0502 | 5.1485 |  |  | 213 | 8520.475 | 9.0502 | 5.1547
214 | 8560.477 | 9.0549 | 5.1511 |  |  | 214 | 8560.477 | 9.0549 | 5.1471 |  |  | 214 | 8560.477 | 9.0549 | 5.1533
215 | 8600.479 | 9.0596 | 5.1498 |  |  | 215 | 8600.479 | 9.0596 | 5.1457 |  |  | 215 | 8600.479 | 9.0596 | 5.1519
216 | 8640.481 | 9.0642 | 5.1485 |  |  | 216 | 8640.481 | 9.0642 | 5.1444 |  |  | 216 | 8640.481 | 9.0642 | 5.1505
217 | 8680.483 | 9.0688 | 5.1472 |  |  | 217 | 8680.483 | 9.0688 | 5.143 |  |  | 217 | 8680.483 | 9.0688 | 5.1492
218 | 8720.486 | 9.0734 | 5.1459 |  |  | 218 | 8720.486 | 9.0734 | 5.1417 |  |  | 218 | 8720.486 | 9.0734 | 5.1478
219 | 8760.488 | 9.078 | 5.1447 |  |  | 219 | 8760.488 | 9.078 | 5.1404 |  |  | 219 | 8760.488 | 9.078 | 5.1465
220 | 8800.49 | 9.0826 | 5.1434 |  |  | 220 | 8800.49 | 9.0826 | 5.1391 |  |  | 220 | 8800.49 | 9.0826 | 5.1452
221 | 8840.492 | 9.0871 | 5.1422 |  |  | 221 | 8840.492 | 9.0871 | 5.1378 |  |  | 221 | 8840.492 | 9.0871 | 5.1439
222 | 8880.495 | 9.0916 | 5.141 |  |  | 222 | 8880.495 | 9.0916 | 5.1365 |  |  | 222 | 8880.495 | 9.0916 | 5.1426
223 | 8920.497 | 9.0961 | 5.1397 |  |  | 223 | 8920.497 | 9.0961 | 5.1352 |  |  | 223 | 8920.497 | 9.0961 | 5.1413
224 | 8960.499 | 9.1006 | 5.1385 |  |  | 224 | 8960.499 | 9.1006 | 5.134 |  |  | 224 | 8960.499 | 9.1006 | 5.14
225 | 9000.501 | 9.105 | 5.1373 |  |  | 225 | 9000.501 | 9.105 | 5.1327 |  |  | 225 | 9000.501 | 9.105 | 5.1387
226 | 9040.504 | 9.1095 | 5.1361 |  |  | 226 | 9040.504 | 9.1095 | 5.1315 |  |  | 226 | 9040.504 | 9.1095 | 5.1375
227 | 9080.506 | 9.1139 | 5.135 |  |  | 227 | 9080.506 | 9.1139 | 5.1302 |  |  | 227 | 9080.506 | 9.1139 | 5.1362
228 | 9120.508 | 9.1183 | 5.1338 |  |  | 228 | 9120.508 | 9.1183 | 5.129 |  |  | 228 | 9120.508 | 9.1183 | 5.135
229 | 9160.51 | 9.1227 | 5.1326 |  |  | 229 | 9160.51 | 9.1227 | 5.1278 |  |  | 229 | 9160.51 | 9.1227 | 5.1338
230 | 9200.512 | 9.127 | 5.1315 |  |  | 230 | 9200.512 | 9.127 | 5.1266 |  |  | 230 | 9200.512 | 9.127 | 5.1326
231 | 9240.515 | 9.1314 | 5.1303 |  |  | 231 | 9240.515 | 9.1314 | 5.1255 |  |  | 231 | 9240.515 | 9.1314 | 5.1314
232 | 9280.517 | 9.1357 | 5.1292 |  |  | 232 | 9280.517 | 9.1357 | 5.1243 |  |  | 232 | 9280.517 | 9.1357 | 5.1302
233 | 9320.519 | 9.14 | 5.1281 |  |  | 233 | 9320.519 | 9.14 | 5.1231 |  |  | 233 | 9320.519 | 9.14 | 5.129
234 | 9360.521 | 9.1443 | 5.127 |  |  | 234 | 9360.521 | 9.1443 | 5.122 |  |  | 234 | 9360.521 | 9.1443 | 5.1278
235 | 9400.524 | 9.1485 | 5.1259 |  |  | 235 | 9400.524 | 9.1485 | 5.1208 |  |  | 235 | 9400.524 | 9.1485 | 5.1267
236 | 9440.526 | 9.1528 | 5.1248 |  |  | 236 | 9440.526 | 9.1528 | 5.1197 |  |  | 236 | 9440.526 | 9.1528 | 5.1255
237 | 9480.528 | 9.157 | 5.1237 |  |  | 237 | 9480.528 | 9.157 | 5.1186 |  |  | 237 | 9480.528 | 9.157 | 5.1244
238 | 9520.53 | 9.1612 | 5.1226 |  |  | 238 | 9520.53 | 9.1612 | 5.1175 |  |  | 238 | 9520.53 | 9.1612 | 5.1233
239 | 9560.532 | 9.1654 | 5.1216 |  |  | 239 | 9560.532 | 9.1654 | 5.1164 |  |  | 239 | 9560.532 | 9.1654 | 5.1222
240 | 9600.535 | 9.1696 | 5.1205 |  |  | 240 | 9600.535 | 9.1696 | 5.1153 |  |  | 240 | 9600.535 | 9.1696 | 5.1211
241 | 9640.537 | 9.1737 | 5.1195 |  |  | 241 | 9640.537 | 9.1737 | 5.1142 |  |  | 241 | 9640.537 | 9.1737 | 5.12
242 | 9680.539 | 9.1779 | 5.1184 |  |  | 242 | 9680.539 | 9.1779 | 5.1131 |  |  | 242 | 9680.539 | 9.1779 | 5.1189
243 | 9720.541 | 9.182 | 5.1174 |  |  | 243 | 9720.541 | 9.182 | 5.112 |  |  | 243 | 9720.541 | 9.182 | 5.1178
244 | 9760.544 | 9.1861 | 5.1164 |  |  | 244 | 9760.544 | 9.1861 | 5.111 |  |  | 244 | 9760.544 | 9.1861 | 5.1167
245 | 9800.546 | 9.1902 | 5.1153 |  |  | 245 | 9800.546 | 9.1902 | 5.1099 |  |  | 245 | 9800.546 | 9.1902 | 5.1157
246 | 9840.548 | 9.1943 | 5.1143 |  |  | 246 | 9840.548 | 9.1943 | 5.1089 |  |  | 246 | 9840.548 | 9.1943 | 5.1146
247 | 9880.55 | 9.1983 | 5.1133 |  |  | 247 | 9880.55 | 9.1983 | 5.1079 |  |  | 247 | 9880.55 | 9.1983 | 5.1136
248 | 9920.553 | 9.2024 | 5.1123 |  |  | 248 | 9920.553 | 9.2024 | 5.1069 |  |  | 248 | 9920.553 | 9.2024 | 5.1126
249 | 9960.555 | 9.2064 | 5.1114 |  |  | 249 | 9960.555 | 9.2064 | 5.1059 |  |  | 249 | 9960.555 | 9.2064 | 5.1115
250 | 10000.557 | 9.2104 | 5.1104 |  |  | 250 | 10000.557 | 9.2104 | 5.1049 |  |  | 250 | 10000.557 | 9.2104 | 5.1105
251 | 10040.559 | 9.2144 | 5.1094 |  |  | 251 | 10040.559 | 9.2144 | 5.1039 |  |  | 251 | 10040.559 | 9.2144 | 5.1095
252 | 10080.561 | 9.2184 | 5.1085 |  |  | 252 | 10080.561 | 9.2184 | 5.1029 |  |  | 252 | 10080.561 | 9.2184 | 5.1085
253 | 10120.564 | 9.2223 | 5.1075 |  |  | 253 | 10120.564 | 9.2223 | 5.1019 |  |  | 253 | 10120.564 | 9.2223 | 5.1075
254 | 10160.566 | 9.2263 | 5.1066 |  |  | 254 | 10160.566 | 9.2263 | 5.1009 |  |  | 254 | 10160.566 | 9.2263 | 5.1066
255 | 10200.568 | 9.2302 | 5.1056 |  |  | 255 | 10200.568 | 9.2302 | 5.1 |  |  | 255 | 10200.568 | 9.2302 | 5.1056
256 | 10240.57 | 9.2341 | 5.1047 |  |  | 256 | 10240.57 | 9.2341 | 5.099 |  |  | 256 | 10240.57 | 9.2341 | 5.1046
257 | 10280.573 | 9.238 | 5.1038 |  |  | 257 | 10280.573 | 9.238 | 5.0981 |  |  | 257 | 10280.573 | 9.238 | 5.1037
258 | 10320.575 | 9.2419 | 5.1028 |  |  | 258 | 10320.575 | 9.2419 | 5.0972 |  |  | 258 | 10320.575 | 9.2419 | 5.1027
259 | 10360.577 | 9.2458 | 5.1019 |  |  | 259 | 10360.577 | 9.2458 | 5.0962 |  |  | 259 | 10360.577 | 9.2458 | 5.1018
260 | 10400.579 | 9.2496 | 5.101 |  |  | 260 | 10400.579 | 9.2496 | 5.0953 |  |  | 260 | 10400.579 | 9.2496 | 5.1009
261 | 10440.581 | 9.2535 | 5.1001 |  |  | 261 | 10440.581 | 9.2535 | 5.0944 |  |  | 261 | 10440.581 | 9.2535 | 5.0999
262 | 10480.584 | 9.2573 | 5.0993 |  |  | 262 | 10480.584 | 9.2573 | 5.0935 |  |  | 262 | 10480.584 | 9.2573 | 5.099
263 | 10520.586 | 9.2611 | 5.0984 |  |  | 263 | 10520.586 | 9.2611 | 5.0926 |  |  | 263 | 10520.586 | 9.2611 | 5.0981
264 | 10560.588 | 9.2649 | 5.0975 |  |  | 264 | 10560.588 | 9.2649 | 5.0917 |  |  | 264 | 10560.588 | 9.2649 | 5.0972
265 | 10600.59 | 9.2687 | 5.0966 |  |  | 265 | 10600.59 | 9.2687 | 5.0909 |  |  | 265 | 10600.59 | 9.2687 | 5.0963
266 | 10640.593 | 9.2724 | 5.0958 |  |  | 266 | 10640.593 | 9.2724 | 5.09 |  |  | 266 | 10640.593 | 9.2724 | 5.0954
267 | 10680.595 | 9.2762 | 5.0949 |  |  | 267 | 10680.595 | 9.2762 | 5.0891 |  |  | 267 | 10680.595 | 9.2762 | 5.0946
268 | 10720.597 | 9.2799 | 5.0941 |  |  | 268 | 10720.597 | 9.2799 | 5.0883 |  |  | 268 | 10720.597 | 9.2799 | 5.0937
269 | 10760.599 | 9.2836 | 5.0932 |  |  | 269 | 10760.599 | 9.2836 | 5.0874 |  |  | 269 | 10760.599 | 9.2836 | 5.0928
270 | 10800.602 | 9.2874 | 5.0924 |  |  | 270 | 10800.602 | 9.2874 | 5.0866 |  |  | 270 | 10800.602 | 9.2874 | 5.092
271 | 10840.604 | 9.2911 | 5.0916 |  |  | 271 | 10840.604 | 9.2911 | 5.0857 |  |  | 271 | 10840.604 | 9.2911 | 5.0911
272 | 10880.606 | 9.2947 | 5.0907 |  |  | 272 | 10880.606 | 9.2947 | 5.0849 |  |  | 272 | 10880.606 | 9.2947 | 5.0903
273 | 10920.608 | 9.2984 | 5.0899 |  |  | 273 | 10920.608 | 9.2984 | 5.0841 |  |  | 273 | 10920.608 | 9.2984 | 5.0895
274 | 10960.61 | 9.3021 | 5.0891 |  |  | 274 | 10960.61 | 9.3021 | 5.0833 |  |  | 274 | 10960.61 | 9.3021 | 5.0887
275 | 11000.613 | 9.3057 | 5.0883 |  |  | 275 | 11000.613 | 9.3057 | 5.0825 |  |  | 275 | 11000.613 | 9.3057 | 5.0878
276 | 11040.615 | 9.3093 | 5.0875 |  |  | 276 | 11040.615 | 9.3093 | 5.0817 |  |  | 276 | 11040.615 | 9.3093 | 5.087
277 | 11080.617 | 9.313 | 5.0867 |  |  | 277 | 11080.617 | 9.313 | 5.0809 |  |  | 277 | 11080.617 | 9.313 | 5.0862
278 | 11120.619 | 9.3166 | 5.0859 |  |  | 278 | 11120.619 | 9.3166 | 5.0801 |  |  | 278 | 11120.619 | 9.3166 | 5.0854
279 | 11160.622 | 9.3201 | 5.0851 |  |  | 279 | 11160.622 | 9.3201 | 5.0793 |  |  | 279 | 11160.622 | 9.3201 | 5.0846
280 | 11200.624 | 9.3237 | 5.0844 |  |  | 280 | 11200.624 | 9.3237 | 5.0786 |  |  | 280 | 11200.624 | 9.3237 | 5.0839
281 | 11240.626 | 9.3273 | 5.0836 |  |  | 281 | 11240.626 | 9.3273 | 5.0778 |  |  | 281 | 11240.626 | 9.3273 | 5.0831
282 | 11280.628 | 9.3308 | 5.0828 |  |  | 282 | 11280.628 | 9.3308 | 5.077 |  |  | 282 | 11280.628 | 9.3308 | 5.0823
283 | 11320.63 | 9.3344 | 5.0821 |  |  | 283 | 11320.63 | 9.3344 | 5.0763 |  |  | 283 | 11320.63 | 9.3344 | 5.0815
284 | 11360.633 | 9.3379 | 5.0813 |  |  | 284 | 11360.633 | 9.3379 | 5.0755 |  |  | 284 | 11360.633 | 9.3379 | 5.0808
285 | 11400.635 | 9.3414 | 5.0806 |  |  | 285 | 11400.635 | 9.3414 | 5.0748 |  |  | 285 | 11400.635 | 9.3414 | 5.08
286 | 11440.637 | 9.3449 | 5.0798 |  |  | 286 | 11440.637 | 9.3449 | 5.0741 |  |  | 286 | 11440.637 | 9.3449 | 5.0793
287 | 11480.639 | 9.3484 | 5.0791 |  |  | 287 | 11480.639 | 9.3484 | 5.0733 |  |  | 287 | 11480.639 | 9.3484 | 5.0786
288 | 11520.642 | 9.3519 | 5.0783 |  |  | 288 | 11520.642 | 9.3519 | 5.0726 |  |  | 288 | 11520.642 | 9.3519 | 5.0778
289 | 11560.644 | 9.3554 | 5.0776 |  |  | 289 | 11560.644 | 9.3554 | 5.0719 |  |  | 289 | 11560.644 | 9.3554 | 5.0771
290 | 11600.646 | 9.3588 | 5.0769 |  |  | 290 | 11600.646 | 9.3588 | 5.0712 |  |  | 290 | 11600.646 | 9.3588 | 5.0764
291 | 11640.648 | 9.3623 | 5.0762 |  |  | 291 | 11640.648 | 9.3623 | 5.0705 |  |  | 291 | 11640.648 | 9.3623 | 5.0757
292 | 11680.651 | 9.3657 | 5.0755 |  |  | 292 | 11680.651 | 9.3657 | 5.0698 |  |  | 292 | 11680.651 | 9.3657 | 5.0749
293 | 11720.653 | 9.3691 | 5.0748 |  |  | 293 | 11720.653 | 9.3691 | 5.0691 |  |  | 293 | 11720.653 | 9.3691 | 5.0742
294 | 11760.655 | 9.3725 | 5.0741 |  |  | 294 | 11760.655 | 9.3725 | 5.0684 |  |  | 294 | 11760.655 | 9.3725 | 5.0735
295 | 11800.657 | 9.3759 | 5.0734 |  |  | 295 | 11800.657 | 9.3759 | 5.0677 |  |  | 295 | 11800.657 | 9.3759 | 5.0729
296 | 11840.659 | 9.3793 | 5.0727 |  |  | 296 | 11840.659 | 9.3793 | 5.067 |  |  | 296 | 11840.659 | 9.3793 | 5.0722
297 | 11880.662 | 9.3827 | 5.072 |  |  | 297 | 11880.662 | 9.3827 | 5.0664 |  |  | 297 | 11880.662 | 9.3827 | 5.0715
298 | 11920.664 | 9.386 | 5.0713 |  |  | 298 | 11920.664 | 9.386 | 5.0657 |  |  | 298 | 11920.664 | 9.386 | 5.0708
299 | 11960.666 | 9.3894 | 5.0706 |  |  | 299 | 11960.666 | 9.3894 | 5.065 |  |  | 299 | 11960.666 | 9.3894 | 5.0702
300 | 12000.668 | 9.3927 | 5.0699 |  |  | 300 | 12000.668 | 9.3927 | 5.0644 |  |  | 300 | 12000.668 | 9.3927 | 5.0695
301 | 12040.671 | 9.396 | 5.0693 |  |  | 301 | 12040.671 | 9.396 | 5.0637 |  |  | 301 | 12040.671 | 9.396 | 5.0688
302 | 12080.673 | 9.3994 | 5.0686 |  |  | 302 | 12080.673 | 9.3994 | 5.0631 |  |  | 302 | 12080.673 | 9.3994 | 5.0682
303 | 12120.675 | 9.4027 | 5.068 |  |  | 303 | 12120.675 | 9.4027 | 5.0625 |  |  | 303 | 12120.675 | 9.4027 | 5.0675
304 | 12160.677 | 9.406 | 5.0673 |  |  | 304 | 12160.677 | 9.406 | 5.0618 |  |  | 304 | 12160.677 | 9.406 | 5.0669
305 | 12200.68 | 9.4092 | 5.0667 |  |  | 305 | 12200.68 | 9.4092 | 5.0612 |  |  | 305 | 12200.68 | 9.4092 | 5.0663
306 | 12240.682 | 9.4125 | 5.066 |  |  | 306 | 12240.682 | 9.4125 | 5.0606 |  |  | 306 | 12240.682 | 9.4125 | 5.0656
307 | 12280.684 | 9.4158 | 5.0654 |  |  | 307 | 12280.684 | 9.4158 | 5.06 |  |  | 307 | 12280.684 | 9.4158 | 5.065
308 | 12320.686 | 9.419 | 5.0647 |  |  | 308 | 12320.686 | 9.419 | 5.0594 |  |  | 308 | 12320.686 | 9.419 | 5.0644
309 | 12360.688 | 9.4223 | 5.0641 |  |  | 309 | 12360.688 | 9.4223 | 5.0588 |  |  | 309 | 12360.688 | 9.4223 | 5.0638
310 | 12400.691 | 9.4255 | 5.0635 |  |  | 310 | 12400.691 | 9.4255 | 5.0582 |  |  | 310 | 12400.691 | 9.4255 | 5.0631
311 | 12440.693 | 9.4287 | 5.0628 |  |  | 311 | 12440.693 | 9.4287 | 5.0576 |  |  | 311 | 12440.693 | 9.4287 | 5.0625
312 | 12480.695 | 9.4319 | 5.0622 |  |  | 312 | 12480.695 | 9.4319 | 5.057 |  |  | 312 | 12480.695 | 9.4319 | 5.0619
313 | 12520.697 | 9.4351 | 5.0616 |  |  | 313 | 12520.697 | 9.4351 | 5.0564 |  |  | 313 | 12520.697 | 9.4351 | 5.0613
314 | 12560.7 | 9.4383 | 5.061 |  |  | 314 | 12560.7 | 9.4383 | 5.0558 |  |  | 314 | 12560.7 | 9.4383 | 5.0608
315 | 12600.702 | 9.4415 | 5.0604 |  |  | 315 | 12600.702 | 9.4415 | 5.0552 |  |  | 315 | 12600.702 | 9.4415 | 5.0602
316 | 12640.704 | 9.4447 | 5.0598 |  |  | 316 | 12640.704 | 9.4447 | 5.0546 |  |  | 316 | 12640.704 | 9.4447 | 5.0596
317 | 12680.706 | 9.4478 | 5.0592 |  |  | 317 | 12680.706 | 9.4478 | 5.0541 |  |  | 317 | 12680.706 | 9.4478 | 5.059
318 | 12720.708 | 9.451 | 5.0586 |  |  | 318 | 12720.708 | 9.451 | 5.0535 |  |  | 318 | 12720.708 | 9.451 | 5.0584
319 | 12760.711 | 9.4541 | 5.058 |  |  | 319 | 12760.711 | 9.4541 | 5.053 |  |  | 319 | 12760.711 | 9.4541 | 5.0579
320 | 12800.713 | 9.4573 | 5.0574 |  |  | 320 | 12800.713 | 9.4573 | 5.0524 |  |  | 320 | 12800.713 | 9.4573 | 5.0573
321 | 12840.715 | 9.4604 | 5.0568 |  |  | 321 | 12840.715 | 9.4604 | 5.0519 |  |  | 321 | 12840.715 | 9.4604 | 5.0567
322 | 12880.717 | 9.4635 | 5.0562 |  |  | 322 | 12880.717 | 9.4635 | 5.0513 |  |  | 322 | 12880.717 | 9.4635 | 5.0562
323 | 12920.72 | 9.4666 | 5.0557 |  |  | 323 | 12920.72 | 9.4666 | 5.0508 |  |  | 323 | 12920.72 | 9.4666 | 5.0556
324 | 12960.722 | 9.4697 | 5.0551 |  |  | 324 | 12960.722 | 9.4697 | 5.0502 |  |  | 324 | 12960.722 | 9.4697 | 5.0551
325 | 13000.724 | 9.4728 | 5.0545 |  |  | 325 | 13000.724 | 9.4728 | 5.0497 |  |  | 325 | 13000.724 | 9.4728 | 5.0546
326 | 13040.726 | 9.4758 | 5.054 |  |  | 326 | 13040.726 | 9.4758 | 5.0492 |  |  | 326 | 13040.726 | 9.4758 | 5.054
327 | 13080.729 | 9.4789 | 5.0534 |  |  | 327 | 13080.729 | 9.4789 | 5.0486 |  |  | 327 | 13080.729 | 9.4789 | 5.0535
328 | 13120.731 | 9.4819 | 5.0528 |  |  | 328 | 13120.731 | 9.4819 | 5.0481 |  |  | 328 | 13120.731 | 9.4819 | 5.053
329 | 13160.733 | 9.485 | 5.0523 |  |  | 329 | 13160.733 | 9.485 | 5.0476 |  |  | 329 | 13160.733 | 9.485 | 5.0524
330 | 13200.735 | 9.488 | 5.0517 |  |  | 330 | 13200.735 | 9.488 | 5.0471 |  |  | 330 | 13200.735 | 9.488 | 5.0519
331 | 13240.737 | 9.4911 | 5.0512 |  |  | 331 | 13240.737 | 9.4911 | 5.0466 |  |  | 331 | 13240.737 | 9.4911 | 5.0514
332 | 13280.74 | 9.4941 | 5.0506 |  |  | 332 | 13280.74 | 9.4941 | 5.0461 |  |  | 332 | 13280.74 | 9.4941 | 5.0509
333 | 13320.742 | 9.4971 | 5.0501 |  |  | 333 | 13320.742 | 9.4971 | 5.0456 |  |  | 333 | 13320.742 | 9.4971 | 5.0504
334 | 13360.744 | 9.5001 | 5.0496 |  |  | 334 | 13360.744 | 9.5001 | 5.0451 |  |  | 334 | 13360.744 | 9.5001 | 5.0499
335 | 13400.746 | 9.5031 | 5.049 |  |  | 335 | 13400.746 | 9.5031 | 5.0446 |  |  | 335 | 13400.746 | 9.5031 | 5.0494
336 | 13440.749 | 9.506 | 5.0485 |  |  | 336 | 13440.749 | 9.506 | 5.0441 |  |  | 336 | 13440.749 | 9.506 | 5.0489
337 | 13480.751 | 9.509 | 5.048 |  |  | 337 | 13480.751 | 9.509 | 5.0436 |  |  | 337 | 13480.751 | 9.509 | 5.0484
338 | 13520.753 | 9.512 | 5.0474 |  |  | 338 | 13520.753 | 9.512 | 5.0432 |  |  | 338 | 13520.753 | 9.512 | 5.0479
339 | 13560.755 | 9.5149 | 5.0469 |  |  | 339 | 13560.755 | 9.5149 | 5.0427 |  |  | 339 | 13560.755 | 9.5149 | 5.0474
340 | 13600.757 | 9.5179 | 5.0464 |  |  | 340 | 13600.757 | 9.5179 | 5.0422 |  |  | 340 | 13600.757 | 9.5179 | 5.0469
341 | 13640.76 | 9.5208 | 5.0459 |  |  | 341 | 13640.76 | 9.5208 | 5.0417 |  |  | 341 | 13640.76 | 9.5208 | 5.0465
342 | 13680.762 | 9.5237 | 5.0454 |  |  | 342 | 13680.762 | 9.5237 | 5.0413 |  |  | 342 | 13680.762 | 9.5237 | 5.046
343 | 13720.764 | 9.5267 | 5.0449 |  |  | 343 | 13720.764 | 9.5267 | 5.0408 |  |  | 343 | 13720.764 | 9.5267 | 5.0455
344 | 13760.766 | 9.5296 | 5.0444 |  |  | 344 | 13760.766 | 9.5296 | 5.0404 |  |  | 344 | 13760.766 | 9.5296 | 5.0451
345 | 13800.769 | 9.5325 | 5.0438 |  |  | 345 | 13800.769 | 9.5325 | 5.0399 |  |  | 345 | 13800.769 | 9.5325 | 5.0446
346 | 13840.771 | 9.5354 | 5.0433 |  |  | 346 | 13840.771 | 9.5354 | 5.0395 |  |  | 346 | 13840.771 | 9.5354 | 5.0441
347 | 13880.773 | 9.5383 | 5.0429 |  |  | 347 | 13880.773 | 9.5383 | 5.039 |  |  | 347 | 13880.773 | 9.5383 | 5.0437
348 | 13920.775 | 9.5411 | 5.0424 |  |  | 348 | 13920.775 | 9.5411 | 5.0386 |  |  | 348 | 13920.775 | 9.5411 | 5.0432
349 | 13960.778 | 9.544 | 5.0419 |  |  | 349 | 13960.778 | 9.544 | 5.0381 |  |  | 349 | 13960.778 | 9.544 | 5.0428
350 | 14000.78 | 9.5469 | 5.0414 |  |  | 350 | 14000.78 | 9.5469 | 5.0377 |  |  | 350 | 14000.78 | 9.5469 | 5.0423
351 | 14040.782 | 9.5497 | 5.0409 |  |  | 351 | 14040.782 | 9.5497 | 5.0373 |  |  | 351 | 14040.782 | 9.5497 | 5.0419
352 | 14080.784 | 9.5526 | 5.0404 |  |  | 352 | 14080.784 | 9.5526 | 5.0368 |  |  | 352 | 14080.784 | 9.5526 | 5.0415
353 | 14120.786 | 9.5554 | 5.0399 |  |  | 353 | 14120.786 | 9.5554 | 5.0364 |  |  | 353 | 14120.786 | 9.5554 | 5.041
354 | 14160.789 | 9.5582 | 5.0395 |  |  | 354 | 14160.789 | 9.5582 | 5.036 |  |  | 354 | 14160.789 | 9.5582 | 5.0406
355 | 14200.791 | 9.5611 | 5.039 |  |  | 355 | 14200.791 | 9.5611 | 5.0356 |  |  | 355 | 14200.791 | 9.5611 | 5.0402
356 | 14240.793 | 9.5639 | 5.0385 |  |  | 356 | 14240.793 | 9.5639 | 5.0351 |  |  | 356 | 14240.793 | 9.5639 | 5.0397
357 | 14280.795 | 9.5667 | 5.038 |  |  | 357 | 14280.795 | 9.5667 | 5.0347 |  |  | 357 | 14280.795 | 9.5667 | 5.0393
358 | 14320.798 | 9.5695 | 5.0376 |  |  | 358 | 14320.798 | 9.5695 | 5.0343 |  |  | 358 | 14320.798 | 9.5695 | 5.0389
359 | 14360.8 | 9.5723 | 5.0371 |  |  | 359 | 14360.8 | 9.5723 | 5.0339 |  |  | 359 | 14360.8 | 9.5723 | 5.0385
360 | 14400.802 | 9.575 | 5.0367 |  |  | 360 | 14400.802 | 9.575 | 5.0335 |  |  | 360 | 14400.802 | 9.575 | 5.0381
361 | 14440.804 | 9.5778 | 5.0362 |  |  | 361 | 14440.804 | 9.5778 | 5.0331 |  |  | 361 | 14440.804 | 9.5778 | 5.0377
362 | 14480.807 | 9.5806 | 5.0357 |  |  | 362 | 14480.807 | 9.5806 | 5.0327 |  |  | 362 | 14480.807 | 9.5806 | 5.0373
363 | 14520.809 | 9.5833 | 5.0353 |  |  | 363 | 14520.809 | 9.5833 | 5.0323 |  |  | 363 | 14520.809 | 9.5833 | 5.0369
364 | 14560.811 | 9.5861 | 5.0348 |  |  | 364 | 14560.811 | 9.5861 | 5.0319 |  |  | 364 | 14560.811 | 9.5861 | 5.0365
365 | 14600.813 | 9.5888 | 5.0344 |  |  | 365 | 14600.813 | 9.5888 | 5.0316 |  |  | 365 | 14600.813 | 9.5888 | 5.0361
366 | 14640.815 | 9.5916 | 5.034 |  |  | 366 | 14640.815 | 9.5916 | 5.0312 |  |  | 366 | 14640.815 | 9.5916 | 5.0357
367 | 14680.818 | 9.5943 | 5.0335 |  |  | 367 | 14680.818 | 9.5943 | 5.0308 |  |  | 367 | 14680.818 | 9.5943 | 5.0353
368 | 14720.82 | 9.597 | 5.0331 |  |  | 368 | 14720.82 | 9.597 | 5.0304 |  |  | 368 | 14720.82 | 9.597 | 5.0349
369 | 14760.822 | 9.5997 | 5.0326 |  |  | 369 | 14760.822 | 9.5997 | 5.03 |  |  | 369 | 14760.822 | 9.5997 | 5.0345
370 | 14800.824 | 9.6024 | 5.0322 |  |  | 370 | 14800.824 | 9.6024 | 5.0297 |  |  | 370 | 14800.824 | 9.6024 | 5.0341
371 | 14840.827 | 9.6051 | 5.0318 |  |  | 371 | 14840.827 | 9.6051 | 5.0293 |  |  | 371 | 14840.827 | 9.6051 | 5.0338
372 | 14880.829 | 9.6078 | 5.0313 |  |  | 372 | 14880.829 | 9.6078 | 5.0289 |  |  | 372 | 14880.829 | 9.6078 | 5.0334
373 | 14920.831 | 9.6105 | 5.0309 |  |  | 373 | 14920.831 | 9.6105 | 5.0286 |  |  | 373 | 14920.831 | 9.6105 | 5.033
374 | 14960.833 | 9.6132 | 5.0305 |  |  | 374 | 14960.833 | 9.6132 | 5.0282 |  |  | 374 | 14960.833 | 9.6132 | 5.0327
375 | 15000.835 | 9.6159 | 5.0301 |  |  | 375 | 15000.835 | 9.6159 | 5.0278 |  |  | 375 | 15000.835 | 9.6159 | 5.0323
376 | 15040.838 | 9.6185 | 5.0296 |  |  | 376 | 15040.838 | 9.6185 | 5.0275 |  |  | 376 | 15040.838 | 9.6185 | 5.0319
377 | 15080.84 | 9.6212 | 5.0292 |  |  | 377 | 15080.84 | 9.6212 | 5.0271 |  |  | 377 | 15080.84 | 9.6212 | 5.0316
378 | 15120.842 | 9.6238 | 5.0288 |  |  | 378 | 15120.842 | 9.6238 | 5.0268 |  |  | 378 | 15120.842 | 9.6238 | 5.0312
379 | 15160.844 | 9.6265 | 5.0284 |  |  | 379 | 15160.844 | 9.6265 | 5.0264 |  |  | 379 | 15160.844 | 9.6265 | 5.0308
380 | 15200.847 | 9.6291 | 5.028 |  |  | 380 | 15200.847 | 9.6291 | 5.0261 |  |  | 380 | 15200.847 | 9.6291 | 5.0305
381 | 15240.849 | 9.6317 | 5.0276 |  |  | 381 | 15240.849 | 9.6317 | 5.0257 |  |  | 381 | 15240.849 | 9.6317 | 5.0301
382 | 15280.851 | 9.6344 | 5.0272 |  |  | 382 | 15280.851 | 9.6344 | 5.0254 |  |  | 382 | 15280.851 | 9.6344 | 5.0298
383 | 15320.853 | 9.637 | 5.0268 |  |  | 383 | 15320.853 | 9.637 | 5.0251 |  |  | 383 | 15320.853 | 9.637 | 5.0294
384 | 15360.856 | 9.6396 | 5.0264 |  |  | 384 | 15360.856 | 9.6396 | 5.0247 |  |  | 384 | 15360.856 | 9.6396 | 5.0291
385 | 15400.858 | 9.6422 | 5.026 |  |  | 385 | 15400.858 | 9.6422 | 5.0244 |  |  | 385 | 15400.858 | 9.6422 | 5.0288
386 | 15440.86 | 9.6448 | 5.0256 |  |  | 386 | 15440.86 | 9.6448 | 5.0241 |  |  | 386 | 15440.86 | 9.6448 | 5.0284
387 | 15480.862 | 9.6474 | 5.0252 |  |  | 387 | 15480.862 | 9.6474 | 5.0237 |  |  | 387 | 15480.862 | 9.6474 | 5.0281
388 | 15520.864 | 9.6499 | 5.0248 |  |  | 388 | 15520.864 | 9.6499 | 5.0234 |  |  | 388 | 15520.864 | 9.6499 | 5.0278
389 | 15560.867 | 9.6525 | 5.0244 |  |  | 389 | 15560.867 | 9.6525 | 5.0231 |  |  | 389 | 15560.867 | 9.6525 | 5.0274
390 | 15600.869 | 9.6551 | 5.024 |  |  | 390 | 15600.869 | 9.6551 | 5.0228 |  |  | 390 | 15600.869 | 9.6551 | 5.0271
391 | 15640.871 | 9.6576 | 5.0236 |  |  | 391 | 15640.871 | 9.6576 | 5.0225 |  |  | 391 | 15640.871 | 9.6576 | 5.0268
392 | 15680.873 | 9.6602 | 5.0232 |  |  | 392 | 15680.873 | 9.6602 | 5.0221 |  |  | 392 | 15680.873 | 9.6602 | 5.0265
393 | 15720.876 | 9.6627 | 5.0228 |  |  | 393 | 15720.876 | 9.6627 | 5.0218 |  |  | 393 | 15720.876 | 9.6627 | 5.0261
394 | 15760.878 | 9.6653 | 5.0225 |  |  | 394 | 15760.878 | 9.6653 | 5.0215 |  |  | 394 | 15760.878 | 9.6653 | 5.0258
395 | 15800.88 | 9.6678 | 5.0221 |  |  | 395 | 15800.88 | 9.6678 | 5.0212 |  |  | 395 | 15800.88 | 9.6678 | 5.0255
396 | 15840.882 | 9.6703 | 5.0217 |  |  | 396 | 15840.882 | 9.6703 | 5.0209 |  |  | 396 | 15840.882 | 9.6703 | 5.0252
397 | 15880.884 | 9.6729 | 5.0213 |  |  | 397 | 15880.884 | 9.6729 | 5.0206 |  |  | 397 | 15880.884 | 9.6729 | 5.0249
398 | 15920.887 | 9.6754 | 5.021 |  |  | 398 | 15920.887 | 9.6754 | 5.0203 |  |  | 398 | 15920.887 | 9.6754 | 5.0246
399 | 15960.889 | 9.6779 | 5.0206 |  |  | 399 | 15960.889 | 9.6779 | 5.02 |  |  | 399 | 15960.889 | 9.6779 | 5.0243
400 | 16000.891 | 9.6804 | 5.0202 |  |  | 400 | 16000.891 | 9.6804 | 5.0197 |  |  | 400 | 16000.891 | 9.6804 | 5.024
401 | 16040.893 | 9.6829 | 5.0198 |  |  | 401 | 16040.893 | 9.6829 | 5.0194 |  |  | 401 | 16040.893 | 9.6829 | 5.0237
402 | 16080.896 | 9.6854 | 5.0195 |  |  | 402 | 16080.896 | 9.6854 | 5.0191 |  |  | 402 | 16080.896 | 9.6854 | 5.0234
403 | 16120.898 | 9.6879 | 5.0191 |  |  | 403 | 16120.898 | 9.6879 | 5.0188 |  |  | 403 | 16120.898 | 9.6879 | 5.0231
404 | 16160.9 | 9.6904 | 5.0188 |  |  | 404 | 16160.9 | 9.6904 | 5.0185 |  |  | 404 | 16160.9 | 9.6904 | 5.0228
405 | 16200.902 | 9.6928 | 5.0184 |  |  | 405 | 16200.902 | 9.6928 | 5.0183 |  |  | 405 | 16200.902 | 9.6928 | 5.0225
406 | 16240.905 | 9.6953 | 5.018 |  |  | 406 | 16240.905 | 9.6953 | 5.018 |  |  | 406 | 16240.905 | 9.6953 | 5.0222
407 | 16280.907 | 9.6977 | 5.0177 |  |  | 407 | 16280.907 | 9.6977 | 5.0177 |  |  | 407 | 16280.907 | 9.6977 | 5.0219
408 | 16320.909 | 9.7002 | 5.0173 |  |  | 408 | 16320.909 | 9.7002 | 5.0174 |  |  | 408 | 16320.909 | 9.7002 | 5.0216
409 | 16360.911 | 9.7027 | 5.017 |  |  | 409 | 16360.911 | 9.7027 | 5.0171 |  |  | 409 | 16360.911 | 9.7027 | 5.0213
410 | 16400.913 | 9.7051 | 5.0166 |  |  | 410 | 16400.913 | 9.7051 | 5.0169 |  |  | 410 | 16400.913 | 9.7051 | 5.021
411 | 16440.916 | 9.7075 | 5.0163 |  |  | 411 | 16440.916 | 9.7075 | 5.0166 |  |  | 411 | 16440.916 | 9.7075 | 5.0207
412 | 16480.918 | 9.71 | 5.0159 |  |  | 412 | 16480.918 | 9.71 | 5.0163 |  |  | 412 | 16480.918 | 9.71 | 5.0205
413 | 16520.92 | 9.7124 | 5.0156 |  |  | 413 | 16520.92 | 9.7124 | 5.016 |  |  | 413 | 16520.92 | 9.7124 | 5.0202
414 | 16560.922 | 9.7148 | 5.0152 |  |  | 414 | 16560.922 | 9.7148 | 5.0158 |  |  | 414 | 16560.922 | 9.7148 | 5.0199
415 | 16600.925 | 9.7172 | 5.0149 |  |  | 415 | 16600.925 | 9.7172 | 5.0155 |  |  | 415 | 16600.925 | 9.7172 | 5.0196
416 | 16640.927 | 9.7196 | 5.0145 |  |  | 416 | 16640.927 | 9.7196 | 5.0153 |  |  | 416 | 16640.927 | 9.7196 | 5.0194
417 | 16680.929 | 9.722 | 5.0142 |  |  | 417 | 16680.929 | 9.722 | 5.015 |  |  | 417 | 16680.929 | 9.722 | 5.0191
418 | 16720.931 | 9.7244 | 5.0139 |  |  | 418 | 16720.931 | 9.7244 | 5.0147 |  |  | 418 | 16720.931 | 9.7244 | 5.0188
419 | 16760.933 | 9.7268 | 5.0135 |  |  | 419 | 16760.933 | 9.7268 | 5.0145 |  |  | 419 | 16760.933 | 9.7268 | 5.0186
420 | 16800.936 | 9.7292 | 5.0132 |  |  | 420 | 16800.936 | 9.7292 | 5.0142 |  |  | 420 | 16800.936 | 9.7292 | 5.0183
421 | 16840.938 | 9.7316 | 5.0129 |  |  | 421 | 16840.938 | 9.7316 | 5.014 |  |  | 421 | 16840.938 | 9.7316 | 5.018
422 | 16880.94 | 9.7339 | 5.0125 |  |  | 422 | 16880.94 | 9.7339 | 5.0137 |  |  | 422 | 16880.94 | 9.7339 | 5.0178
423 | 16920.942 | 9.7363 | 5.0122 |  |  | 423 | 16920.942 | 9.7363 | 5.0135 |  |  | 423 | 16920.942 | 9.7363 | 5.0175
424 | 16960.945 | 9.7387 | 5.0119 |  |  | 424 | 16960.945 | 9.7387 | 5.0132 |  |  | 424 | 16960.945 | 9.7387 | 5.0173
425 | 17000.947 | 9.741 | 5.0115 |  |  | 425 | 17000.947 | 9.741 | 5.013 |  |  | 425 | 17000.947 | 9.741 | 5.017
426 | 17040.949 | 9.7434 | 5.0112 |  |  | 426 | 17040.949 | 9.7434 | 5.0127 |  |  | 426 | 17040.949 | 9.7434 | 5.0168
427 | 17080.951 | 9.7457 | 5.0109 |  |  | 427 | 17080.951 | 9.7457 | 5.0125 |  |  | 427 | 17080.951 | 9.7457 | 5.0165
428 | 17120.954 | 9.7481 | 5.0106 |  |  | 428 | 17120.954 | 9.7481 | 5.0122 |  |  | 428 | 17120.954 | 9.7481 | 5.0163
429 | 17160.956 | 9.7504 | 5.0102 |  |  | 429 | 17160.956 | 9.7504 | 5.012 |  |  | 429 | 17160.956 | 9.7504 | 5.016
430 | 17200.958 | 9.7527 | 5.0099 |  |  | 430 | 17200.958 | 9.7527 | 5.0118 |  |  | 430 | 17200.958 | 9.7527 | 5.0158
431 | 17240.96 | 9.755 | 5.0096 |  |  | 431 | 17240.96 | 9.755 | 5.0115 |  |  | 431 | 17240.96 | 9.755 | 5.0155
432 | 17280.962 | 9.7574 | 5.0093 |  |  | 432 | 17280.962 | 9.7574 | 5.0113 |  |  | 432 | 17280.962 | 9.7574 | 5.0153
433 | 17320.965 | 9.7597 | 5.009 |  |  | 433 | 17320.965 | 9.7597 | 5.0111 |  |  | 433 | 17320.965 | 9.7597 | 5.015
434 | 17360.967 | 9.762 | 5.0087 |  |  | 434 | 17360.967 | 9.762 | 5.0108 |  |  | 434 | 17360.967 | 9.762 | 5.0148
435 | 17400.969 | 9.7643 | 5.0083 |  |  | 435 | 17400.969 | 9.7643 | 5.0106 |  |  | 435 | 17400.969 | 9.7643 | 5.0146
436 | 17440.971 | 9.7666 | 5.008 |  |  | 436 | 17440.971 | 9.7666 | 5.0104 |  |  | 436 | 17440.971 | 9.7666 | 5.0143
437 | 17480.974 | 9.7689 | 5.0077 |  |  | 437 | 17480.974 | 9.7689 | 5.0101 |  |  | 437 | 17480.974 | 9.7689 | 5.0141
438 | 17520.976 | 9.7712 | 5.0074 |  |  | 438 | 17520.976 | 9.7712 | 5.0099 |  |  | 438 | 17520.976 | 9.7712 | 5.0139
439 | 17560.978 | 9.7734 | 5.0071 |  |  | 439 | 17560.978 | 9.7734 | 5.0097 |  |  | 439 | 17560.978 | 9.7734 | 5.0136
440 | 17600.98 | 9.7757 | 5.0068 |  |  | 440 | 17600.98 | 9.7757 | 5.0095 |  |  | 440 | 17600.98 | 9.7757 | 5.0134
441 | 17640.983 | 9.778 | 5.0065 |  |  | 441 | 17640.983 | 9.778 | 5.0092 |  |  | 441 | 17640.983 | 9.778 | 5.0132
442 | 17680.985 | 9.7802 | 5.0062 |  |  | 442 | 17680.985 | 9.7802 | 5.009 |  |  | 442 | 17680.985 | 9.7802 | 5.0129
443 | 17720.987 | 9.7825 | 5.0059 |  |  | 443 | 17720.987 | 9.7825 | 5.0088 |  |  | 443 | 17720.987 | 9.7825 | 5.0127
444 | 17760.989 | 9.7848 | 5.0056 |  |  | 444 | 17760.989 | 9.7848 | 5.0086 |  |  | 444 | 17760.989 | 9.7848 | 5.0125
445 | 17800.991 | 9.787 | 5.0053 |  |  | 445 | 17800.991 | 9.787 | 5.0084 |  |  | 445 | 17800.991 | 9.787 | 5.0123
446 | 17840.994 | 9.7893 | 5.005 |  |  | 446 | 17840.994 | 9.7893 | 5.0082 |  |  | 446 | 17840.994 | 9.7893 | 5.012
447 | 17880.996 | 9.7915 | 5.0047 |  |  | 447 | 17880.996 | 9.7915 | 5.008 |  |  | 447 | 17880.996 | 9.7915 | 5.0118
448 | 17920.998 | 9.7937 | 5.0044 |  |  | 448 | 17920.998 | 9.7937 | 5.0078 |  |  | 448 | 17920.998 | 9.7937 | 5.0116
449 | 17961 | 9.796 | 5.0041 |  |  | 449 | 17961 | 9.796 | 5.0075 |  |  | 449 | 17961 | 9.796 | 5.0114
450 | 18001.003 | 9.7982 | 5.0038 |  |  | 450 | 18001.003 | 9.7982 | 5.0073 |  |  | 450 | 18001.003 | 9.7982 | 5.0112
451 | 18041.005 | 9.8004 | 5.0035 |  |  | 451 | 18041.005 | 9.8004 | 5.0071 |  |  | 451 | 18041.005 | 9.8004 | 5.011
452 | 18081.007 | 9.8026 | 5.0032 |  |  | 452 | 18081.007 | 9.8026 | 5.0069 |  |  | 452 | 18081.007 | 9.8026 | 5.0108
453 | 18121.009 | 9.8048 | 5.0029 |  |  | 453 | 18121.009 | 9.8048 | 5.0067 |  |  | 453 | 18121.009 | 9.8048 | 5.0105
454 | 18161.011 | 9.807 | 5.0027 |  |  | 454 | 18161.011 | 9.807 | 5.0065 |  |  | 454 | 18161.011 | 9.807 | 5.0103
455 | 18201.014 | 9.8092 | 5.0024 |  |  | 455 | 18201.014 | 9.8092 | 5.0063 |  |  | 455 | 18201.014 | 9.8092 | 5.0101
456 | 18241.016 | 9.8114 | 5.0021 |  |  | 456 | 18241.016 | 9.8114 | 5.0061 |  |  | 456 | 18241.016 | 9.8114 | 5.0099
457 | 18281.018 | 9.8136 | 5.0018 |  |  | 457 | 18281.018 | 9.8136 | 5.0059 |  |  | 457 | 18281.018 | 9.8136 | 5.0097
458 | 18321.02 | 9.8158 | 5.0015 |  |  | 458 | 18321.02 | 9.8158 | 5.0057 |  |  | 458 | 18321.02 | 9.8158 | 5.0095
459 | 18361.023 | 9.818 | 5.0012 |  |  | 459 | 18361.023 | 9.818 | 5.0055 |  |  | 459 | 18361.023 | 9.818 | 5.0093
460 | 18401.025 | 9.8202 | 5.001 |  |  | 460 | 18401.025 | 9.8202 | 5.0053 |  |  | 460 | 18401.025 | 9.8202 | 5.0091
461 | 18441.027 | 9.8223 | 5.0007 |  |  | 461 | 18441.027 | 9.8223 | 5.0052 |  |  | 461 | 18441.027 | 9.8223 | 5.0089
462 | 18481.029 | 9.8245 | 5.0004 |  |  | 462 | 18481.029 | 9.8245 | 5.005 |  |  | 462 | 18481.029 | 9.8245 | 5.0087
463 | 18521.032 | 9.8267 | 5.0001 |  |  | 463 | 18521.032 | 9.8267 | 5.0048 |  |  | 463 | 18521.032 | 9.8267 | 5.0085
464 | 18561.034 | 9.8288 | 4.9998 |  |  | 464 | 18561.034 | 9.8288 | 5.0046 |  |  | 464 | 18561.034 | 9.8288 | 5.0083
465 | 18601.036 | 9.831 | 4.9996 |  |  | 465 | 18601.036 | 9.831 | 5.0044 |  |  | 465 | 18601.036 | 9.831 | 5.0081
466 | 18641.038 | 9.8331 | 4.9993 |  |  | 466 | 18641.038 | 9.8331 | 5.0042 |  |  | 466 | 18641.038 | 9.8331 | 5.0079
467 | 18681.04 | 9.8353 | 4.999 |  |  | 467 | 18681.04 | 9.8353 | 5.004 |  |  | 467 | 18681.04 | 9.8353 | 5.0077
468 | 18721.043 | 9.8374 | 4.9987 |  |  | 468 | 18721.043 | 9.8374 | 5.0038 |  |  | 468 | 18721.043 | 9.8374 | 5.0075
469 | 18761.045 | 9.8395 | 4.9985 |  |  | 469 | 18761.045 | 9.8395 | 5.0037 |  |  | 469 | 18761.045 | 9.8395 | 5.0074
470 | 18801.047 | 9.8417 | 4.9982 |  |  | 470 | 18801.047 | 9.8417 | 5.0035 |  |  | 470 | 18801.047 | 9.8417 | 5.0072
471 | 18841.049 | 9.8438 | 4.9979 |  |  | 471 | 18841.049 | 9.8438 | 5.0033 |  |  | 471 | 18841.049 | 9.8438 | 5.007
472 | 18881.052 | 9.8459 | 4.9977 |  |  | 472 | 18881.052 | 9.8459 | 5.0031 |  |  | 472 | 18881.052 | 9.8459 | 5.0068
473 | 18921.054 | 9.848 | 4.9974 |  |  | 473 | 18921.054 | 9.848 | 5.003 |  |  | 473 | 18921.054 | 9.848 | 5.0066
474 | 18961.056 | 9.8501 | 4.9971 |  |  | 474 | 18961.056 | 9.8501 | 5.0028 |  |  | 474 | 18961.056 | 9.8501 | 5.0064
475 | 19001.058 | 9.8522 | 4.9969 |  |  | 475 | 19001.058 | 9.8522 | 5.0026 |  |  | 475 | 19001.058 | 9.8522 | 5.0062
476 | 19041.06 | 9.8544 | 4.9966 |  |  | 476 | 19041.06 | 9.8544 | 5.0024 |  |  | 476 | 19041.06 | 9.8544 | 5.0061
477 | 19081.063 | 9.8565 | 4.9963 |  |  | 477 | 19081.063 | 9.8565 | 5.0023 |  |  | 477 | 19081.063 | 9.8565 | 5.0059
478 | 19121.065 | 9.8585 | 4.9961 |  |  | 478 | 19121.065 | 9.8585 | 5.0021 |  |  | 478 | 19121.065 | 9.8585 | 5.0057
479 | 19161.067 | 9.8606 | 4.9958 |  |  | 479 | 19161.067 | 9.8606 | 5.0019 |  |  | 479 | 19161.067 | 9.8606 | 5.0055
480 | 19201.069 | 9.8627 | 4.9956 |  |  | 480 | 19201.069 | 9.8627 | 5.0017 |  |  | 480 | 19201.069 | 9.8627 | 5.0053
481 | 19241.072 | 9.8648 | 4.9953 |  |  | 481 | 19241.072 | 9.8648 | 5.0016 |  |  | 481 | 19241.072 | 9.8648 | 5.0052
482 | 19281.074 | 9.8669 | 4.9951 |  |  | 482 | 19281.074 | 9.8669 | 5.0014 |  |  | 482 | 19281.074 | 9.8669 | 5.005
483 | 19321.076 | 9.869 | 4.9948 |  |  | 483 | 19321.076 | 9.869 | 5.0012 |  |  | 483 | 19321.076 | 9.869 | 5.0048
484 | 19361.078 | 9.871 | 4.9945 |  |  | 484 | 19361.078 | 9.871 | 5.0011 |  |  | 484 | 19361.078 | 9.871 | 5.0046
485 | 19401.081 | 9.8731 | 4.9943 |  |  | 485 | 19401.081 | 9.8731 | 5.0009 |  |  | 485 | 19401.081 | 9.8731 | 5.0045
486 | 19441.083 | 9.8751 | 4.994 |  |  | 486 | 19441.083 | 9.8751 | 5.0008 |  |  | 486 | 19441.083 | 9.8751 | 5.0043
487 | 19481.085 | 9.8772 | 4.9938 |  |  | 487 | 19481.085 | 9.8772 | 5.0006 |  |  | 487 | 19481.085 | 9.8772 | 5.0041
488 | 19521.087 | 9.8793 | 4.9935 |  |  | 488 | 19521.087 | 9.8793 | 5.0004 |  |  | 488 | 19521.087 | 9.8793 | 5.004
489 | 19561.089 | 9.8813 | 4.9933 |  |  | 489 | 19561.089 | 9.8813 | 5.0003 |  |  | 489 | 19561.089 | 9.8813 | 5.0038
490 | 19601.092 | 9.8833 | 4.993 |  |  | 490 | 19601.092 | 9.8833 | 5.0001 |  |  | 490 | 19601.092 | 9.8833 | 5.0036
491 | 19641.094 | 9.8854 | 4.9928 |  |  | 491 | 19641.094 | 9.8854 | 5 |  |  | 491 | 19641.094 | 9.8854 | 5.0035
492 | 19681.096 | 9.8874 | 4.9925 |  |  | 492 | 19681.096 | 9.8874 | 4.9998 |  |  | 492 | 19681.096 | 9.8874 | 5.0033
493 | 19721.098 | 9.8894 | 4.9923 |  |  | 493 | 19721.098 | 9.8894 | 4.9996 |  |  | 493 | 19721.098 | 9.8894 | 5.0031
494 | 19761.101 | 9.8915 | 4.992 |  |  | 494 | 19761.101 | 9.8915 | 4.9995 |  |  | 494 | 19761.101 | 9.8915 | 5.003
495 | 19801.103 | 9.8935 | 4.9918 |  |  | 495 | 19801.103 | 9.8935 | 4.9993 |  |  | 495 | 19801.103 | 9.8935 | 5.0028
496 | 19841.105 | 9.8955 | 4.9915 |  |  | 496 | 19841.105 | 9.8955 | 4.9992 |  |  | 496 | 19841.105 | 9.8955 | 5.0026
497 | 19881.107 | 9.8975 | 4.9913 |  |  | 497 | 19881.107 | 9.8975 | 4.999 |  |  | 497 | 19881.107 | 9.8975 | 5.0025
498 | 19921.109 | 9.8995 | 4.9911 |  |  | 498 | 19921.109 | 9.8995 | 4.9989 |  |  | 498 | 19921.109 | 9.8995 | 5.0023
499 | 19961.112 | 9.9015 | 4.9908 |  |  | 499 | 19961.112 | 9.9015 | 4.9987 |  |  | 499 | 19961.112 | 9.9015 | 5.0022
500 | 20001.114 | 9.9035 | 4.9906 |  |  | 500 | 20001.114 | 9.9035 | 4.9986 |  |  | 500 | 20001.114 | 9.9035 | 5.002
501 | 20041.116 | 9.9055 | 4.9903 |  |  | 501 | 20041.116 | 9.9055 | 4.9984 |  |  | 501 | 20041.116 | 9.9055 | 5.0018
502 | 20081.118 | 9.9075 | 4.9901 |  |  | 502 | 20081.118 | 9.9075 | 4.9983 |  |  | 502 | 20081.118 | 9.9075 | 5.0017
503 | 20121.121 | 9.9095 | 4.9899 |  |  | 503 | 20121.121 | 9.9095 | 4.9982 |  |  | 503 | 20121.121 | 9.9095 | 5.0015
504 | 20161.123 | 9.9115 | 4.9896 |  |  | 504 | 20161.123 | 9.9115 | 4.998 |  |  | 504 | 20161.123 | 9.9115 | 5.0014
505 | 20201.125 | 9.9135 | 4.9894 |  |  | 505 | 20201.125 | 9.9135 | 4.9979 |  |  | 505 | 20201.125 | 9.9135 | 5.0012
506 | 20241.127 | 9.9155 | 4.9891 |  |  | 506 | 20241.127 | 9.9155 | 4.9977 |  |  | 506 | 20241.127 | 9.9155 | 5.0011
507 | 20281.13 | 9.9174 | 4.9889 |  |  | 507 | 20281.13 | 9.9174 | 4.9976 |  |  | 507 | 20281.13 | 9.9174 | 5.0009
508 | 20321.132 | 9.9194 | 4.9887 |  |  | 508 | 20321.132 | 9.9194 | 4.9974 |  |  | 508 | 20321.132 | 9.9194 | 5.0008
509 | 20361.134 | 9.9214 | 4.9884 |  |  | 509 | 20361.134 | 9.9214 | 4.9973 |  |  | 509 | 20361.134 | 9.9214 | 5.0006
510 | 20401.136 | 9.9233 | 4.9882 |  |  | 510 | 20401.136 | 9.9233 | 4.9972 |  |  | 510 | 20401.136 | 9.9233 | 5.0005
511 | 20441.138 | 9.9253 | 4.988 |  |  | 511 | 20441.138 | 9.9253 | 4.997 |  |  | 511 | 20441.138 | 9.9253 | 5.0003
512 | 20481.141 | 9.9273 | 4.9877 |  |  | 512 | 20481.141 | 9.9273 | 4.9969 |  |  | 512 | 20481.141 | 9.9273 | 5.0002
513 | 20521.143 | 9.9292 | 4.9875 |  |  | 513 | 20521.143 | 9.9292 | 4.9967 |  |  | 513 | 20521.143 | 9.9292 | 5
514 | 20561.145 | 9.9312 | 4.9873 |  |  | 514 | 20561.145 | 9.9312 | 4.9966 |  |  | 514 | 20561.145 | 9.9312 | 4.9999
515 | 20601.147 | 9.9331 | 4.9871 |  |  | 515 | 20601.147 | 9.9331 | 4.9965 |  |  | 515 | 20601.147 | 9.9331 | 4.9997
516 | 20641.15 | 9.935 | 4.9868 |  |  | 516 | 20641.15 | 9.935 | 4.9963 |  |  | 516 | 20641.15 | 9.935 | 4.9996
517 | 20681.152 | 9.937 | 4.9866 |  |  | 517 | 20681.152 | 9.937 | 4.9962 |  |  | 517 | 20681.152 | 9.937 | 4.9995
518 | 20721.154 | 9.9389 | 4.9864 |  |  | 518 | 20721.154 | 9.9389 | 4.9961 |  |  | 518 | 20721.154 | 9.9389 | 4.9993
519 | 20761.156 | 9.9408 | 4.9861 |  |  | 519 | 20761.156 | 9.9408 | 4.9959 |  |  | 519 | 20761.156 | 9.9408 | 4.9992
520 | 20801.159 | 9.9428 | 4.9859 |  |  | 520 | 20801.159 | 9.9428 | 4.9958 |  |  | 520 | 20801.159 | 9.9428 | 4.999
521 | 20841.161 | 9.9447 | 4.9857 |  |  | 521 | 20841.161 | 9.9447 | 4.9957 |  |  | 521 | 20841.161 | 9.9447 | 4.9989
522 | 20881.163 | 9.9466 | 4.9855 |  |  | 522 | 20881.163 | 9.9466 | 4.9955 |  |  | 522 | 20881.163 | 9.9466 | 4.9988
523 | 20921.165 | 9.9485 | 4.9852 |  |  | 523 | 20921.165 | 9.9485 | 4.9954 |  |  | 523 | 20921.165 | 9.9485 | 4.9986
524 | 20961.167 | 9.9504 | 4.985 |  |  | 524 | 20961.167 | 9.9504 | 4.9953 |  |  | 524 | 20961.167 | 9.9504 | 4.9985
525 | 21001.17 | 9.9523 | 4.9848 |  |  | 525 | 21001.17 | 9.9523 | 4.9952 |  |  | 525 | 21001.17 | 9.9523 | 4.9983
526 | 21041.172 | 9.9542 | 4.9846 |  |  | 526 | 21041.172 | 9.9542 | 4.995 |  |  | 526 | 21041.172 | 9.9542 | 4.9982
527 | 21081.174 | 9.9561 | 4.9844 |  |  | 527 | 21081.174 | 9.9561 | 4.9949 |  |  | 527 | 21081.174 | 9.9561 | 4.9981
528 | 21121.176 | 9.958 | 4.9841 |  |  | 528 | 21121.176 | 9.958 | 4.9948 |  |  | 528 | 21121.176 | 9.958 | 4.9979
529 | 21161.179 | 9.9599 | 4.9839 |  |  | 529 | 21161.179 | 9.9599 | 4.9947 |  |  | 529 | 21161.179 | 9.9599 | 4.9978
530 | 21201.181 | 9.9618 | 4.9837 |  |  | 530 | 21201.181 | 9.9618 | 4.9945 |  |  | 530 | 21201.181 | 9.9618 | 4.9977
531 | 21241.183 | 9.9637 | 4.9835 |  |  | 531 | 21241.183 | 9.9637 | 4.9944 |  |  | 531 | 21241.183 | 9.9637 | 4.9975
532 | 21281.185 | 9.9656 | 4.9833 |  |  | 532 | 21281.185 | 9.9656 | 4.9943 |  |  | 532 | 21281.185 | 9.9656 | 4.9974
533 | 21321.187 | 9.9675 | 4.9831 |  |  | 533 | 21321.187 | 9.9675 | 4.9942 |  |  | 533 | 21321.187 | 9.9675 | 4.9973
534 | 21361.19 | 9.9693 | 4.9828 |  |  | 534 | 21361.19 | 9.9693 | 4.994 |  |  | 534 | 21361.19 | 9.9693 | 4.9971
535 | 21401.192 | 9.9712 | 4.9826 |  |  | 535 | 21401.192 | 9.9712 | 4.9939 |  |  | 535 | 21401.192 | 9.9712 | 4.997
536 | 21441.194 | 9.9731 | 4.9824 |  |  | 536 | 21441.194 | 9.9731 | 4.9938 |  |  | 536 | 21441.194 | 9.9731 | 4.9969
537 | 21481.196 | 9.9749 | 4.9822 |  |  | 537 | 21481.196 | 9.9749 | 4.9937 |  |  | 537 | 21481.196 | 9.9749 | 4.9967
538 | 21521.199 | 9.9768 | 4.982 |  |  | 538 | 21521.199 | 9.9768 | 4.9936 |  |  | 538 | 21521.199 | 9.9768 | 4.9966
539 | 21561.201 | 9.9787 | 4.9818 |  |  | 539 | 21561.201 | 9.9787 | 4.9934 |  |  | 539 | 21561.201 | 9.9787 | 4.9965
540 | 21601.203 | 9.9805 | 4.9816 |  |  | 540 | 21601.203 | 9.9805 | 4.9933 |  |  | 540 | 21601.203 | 9.9805 | 4.9964
541 | 21641.205 | 9.9824 | 4.9813 |  |  | 541 | 21641.205 | 9.9824 | 4.9932 |  |  | 541 | 21641.205 | 9.9824 | 4.9962
542 | 21681.208 | 9.9842 | 4.9811 |  |  | 542 | 21681.208 | 9.9842 | 4.9931 |  |  | 542 | 21681.208 | 9.9842 | 4.9961
543 | 21721.21 | 9.986 | 4.9809 |  |  | 543 | 21721.21 | 9.986 | 4.993 |  |  | 543 | 21721.21 | 9.986 | 4.996
544 | 21761.212 | 9.9879 | 4.9807 |  |  | 544 | 21761.212 | 9.9879 | 4.9929 |  |  | 544 | 21761.212 | 9.9879 | 4.9959
545 | 21801.214 | 9.9897 | 4.9805 |  |  | 545 | 21801.214 | 9.9897 | 4.9927 |  |  | 545 | 21801.214 | 9.9897 | 4.9957
546 | 21841.216 | 9.9916 | 4.9803 |  |  | 546 | 21841.216 | 9.9916 | 4.9926 |  |  | 546 | 21841.216 | 9.9916 | 4.9956
547 | 21881.219 | 9.9934 | 4.9801 |  |  | 547 | 21881.219 | 9.9934 | 4.9925 |  |  | 547 | 21881.219 | 9.9934 | 4.9955
548 | 21921.221 | 9.9952 | 4.9799 |  |  | 548 | 21921.221 | 9.9952 | 4.9924 |  |  | 548 | 21921.221 | 9.9952 | 4.9954
549 | 21961.223 | 9.997 | 4.9797 |  |  | 549 | 21961.223 | 9.997 | 4.9923 |  |  | 549 | 21961.223 | 9.997 | 4.9952
550 | 22001.225 | 9.9989 | 4.9795 |  |  | 550 | 22001.225 | 9.9989 | 4.9922 |  |  | 550 | 22001.225 | 9.9989 | 4.9951
551 | 22041.228 | 10.0007 | 4.9793 |  |  | 551 | 22041.228 | 10.0007 | 4.9921 |  |  | 551 | 22041.228 | 10.0007 | 4.995
552 | 22081.23 | 10.0025 | 4.9791 |  |  | 552 | 22081.23 | 10.0025 | 4.992 |  |  | 552 | 22081.23 | 10.0025 | 4.9949
553 | 22121.232 | 10.0043 | 4.9789 |  |  | 553 | 22121.232 | 10.0043 | 4.9919 |  |  | 553 | 22121.232 | 10.0043 | 4.9948
554 | 22161.234 | 10.0061 | 4.9786 |  |  | 554 | 22161.234 | 10.0061 | 4.9917 |  |  | 554 | 22161.234 | 10.0061 | 4.9946
555 | 22201.236 | 10.0079 | 4.9784 |  |  | 555 | 22201.236 | 10.0079 | 4.9916 |  |  | 555 | 22201.236 | 10.0079 | 4.9945
556 | 22241.239 | 10.0097 | 4.9782 |  |  | 556 | 22241.239 | 10.0097 | 4.9915 |  |  | 556 | 22241.239 | 10.0097 | 4.9944
557 | 22281.241 | 10.0115 | 4.978 |  |  | 557 | 22281.241 | 10.0115 | 4.9914 |  |  | 557 | 22281.241 | 10.0115 | 4.9943
558 | 22321.243 | 10.0133 | 4.9778 |  |  | 558 | 22321.243 | 10.0133 | 4.9913 |  |  | 558 | 22321.243 | 10.0133 | 4.9942
559 | 22361.245 | 10.0151 | 4.9776 |  |  | 559 | 22361.245 | 10.0151 | 4.9912 |  |  | 559 | 22361.245 | 10.0151 | 4.9941
560 | 22401.248 | 10.0169 | 4.9774 |  |  | 560 | 22401.248 | 10.0169 | 4.9911 |  |  | 560 | 22401.248 | 10.0169 | 4.9939
561 | 22441.25 | 10.0187 | 4.9772 |  |  | 561 | 22441.25 | 10.0187 | 4.991 |  |  | 561 | 22441.25 | 10.0187 | 4.9938
562 | 22481.252 | 10.0204 | 4.977 |  |  | 562 | 22481.252 | 10.0204 | 4.9909 |  |  | 562 | 22481.252 | 10.0204 | 4.9937
563 | 22521.254 | 10.0222 | 4.9768 |  |  | 563 | 22521.254 | 10.0222 | 4.9908 |  |  | 563 | 22521.254 | 10.0222 | 4.9936
564 | 22561.257 | 10.024 | 4.9766 |  |  | 564 | 22561.257 | 10.024 | 4.9907 |  |  | 564 | 22561.257 | 10.024 | 4.9935
565 | 22601.259 | 10.0258 | 4.9764 |  |  | 565 | 22601.259 | 10.0258 | 4.9906 |  |  | 565 | 22601.259 | 10.0258 | 4.9934
566 | 22641.261 | 10.0275 | 4.9762 |  |  | 566 | 22641.261 | 10.0275 | 4.9905 |  |  | 566 | 22641.261 | 10.0275 | 4.9933
567 | 22681.263 | 10.0293 | 4.9761 |  |  | 567 | 22681.263 | 10.0293 | 4.9904 |  |  | 567 | 22681.263 | 10.0293 | 4.9931
568 | 22721.265 | 10.0311 | 4.9759 |  |  | 568 | 22721.265 | 10.0311 | 4.9903 |  |  | 568 | 22721.265 | 10.0311 | 4.993
569 | 22761.268 | 10.0328 | 4.9757 |  |  | 569 | 22761.268 | 10.0328 | 4.9902 |  |  | 569 | 22761.268 | 10.0328 | 4.9929
570 | 22801.27 | 10.0346 | 4.9755 |  |  | 570 | 22801.27 | 10.0346 | 4.9901 |  |  | 570 | 22801.27 | 10.0346 | 4.9928
571 | 22841.272 | 10.0363 | 4.9753 |  |  | 571 | 22841.272 | 10.0363 | 4.99 |  |  | 571 | 22841.272 | 10.0363 | 4.9927
572 | 22881.274 | 10.0381 | 4.9751 |  |  | 572 | 22881.274 | 10.0381 | 4.9899 |  |  | 572 | 22881.274 | 10.0381 | 4.9926
573 | 22921.277 | 10.0398 | 4.9749 |  |  | 573 | 22921.277 | 10.0398 | 4.9898 |  |  | 573 | 22921.277 | 10.0398 | 4.9925
574 | 22961.279 | 10.0416 | 4.9747 |  |  | 574 | 22961.279 | 10.0416 | 4.9897 |  |  | 574 | 22961.279 | 10.0416 | 4.9924
575 | 23001.281 | 10.0433 | 4.9745 |  |  | 575 | 23001.281 | 10.0433 | 4.9896 |  |  | 575 | 23001.281 | 10.0433 | 4.9923
576 | 23041.283 | 10.045 | 4.9743 |  |  | 576 | 23041.283 | 10.045 | 4.9895 |  |  | 576 | 23041.283 | 10.045 | 4.9922
577 | 23081.286 | 10.0468 | 4.9741 |  |  | 577 | 23081.286 | 10.0468 | 4.9894 |  |  | 577 | 23081.286 | 10.0468 | 4.992
578 | 23121.288 | 10.0485 | 4.9739 |  |  | 578 | 23121.288 | 10.0485 | 4.9893 |  |  | 578 | 23121.288 | 10.0485 | 4.9919
579 | 23161.29 | 10.0502 | 4.9737 |  |  | 579 | 23161.29 | 10.0502 | 4.9892 |  |  | 579 | 23161.29 | 10.0502 | 4.9918
580 | 23201.292 | 10.052 | 4.9735 |  |  | 580 | 23201.292 | 10.052 | 4.9891 |  |  | 580 | 23201.292 | 10.052 | 4.9917
581 | 23241.294 | 10.0537 | 4.9734 |  |  | 581 | 23241.294 | 10.0537 | 4.989 |  |  | 581 | 23241.294 | 10.0537 | 4.9916
582 | 23281.297 | 10.0554 | 4.9732 |  |  | 582 | 23281.297 | 10.0554 | 4.9889 |  |  | 582 | 23281.297 | 10.0554 | 4.9915
583 | 23321.299 | 10.0571 | 4.973 |  |  | 583 | 23321.299 | 10.0571 | 4.9888 |  |  | 583 | 23321.299 | 10.0571 | 4.9914
584 | 23361.301 | 10.0588 | 4.9728 |  |  | 584 | 23361.301 | 10.0588 | 4.9887 |  |  | 584 | 23361.301 | 10.0588 | 4.9913
585 | 23401.303 | 10.0605 | 4.9726 |  |  | 585 | 23401.303 | 10.0605 | 4.9886 |  |  | 585 | 23401.303 | 10.0605 | 4.9912
586 | 23441.306 | 10.0623 | 4.9724 |  |  | 586 | 23441.306 | 10.0623 | 4.9885 |  |  | 586 | 23441.306 | 10.0623 | 4.9911
587 | 23481.308 | 10.064 | 4.9722 |  |  | 587 | 23481.308 | 10.064 | 4.9884 |  |  | 587 | 23481.308 | 10.064 | 4.991
588 | 23521.31 | 10.0657 | 4.972 |  |  | 588 | 23521.31 | 10.0657 | 4.9883 |  |  | 588 | 23521.31 | 10.0657 | 4.9909
589 | 23561.312 | 10.0674 | 4.9719 |  |  | 589 | 23561.312 | 10.0674 | 4.9883 |  |  | 589 | 23561.312 | 10.0674 | 4.9908
590 | 23601.314 | 10.0691 | 4.9717 |  |  | 590 | 23601.314 | 10.0691 | 4.9882 |  |  | 590 | 23601.314 | 10.0691 | 4.9907
591 | 23641.317 | 10.0708 | 4.9715 |  |  | 591 | 23641.317 | 10.0708 | 4.9881 |  |  | 591 | 23641.317 | 10.0708 | 4.9906
592 | 23681.319 | 10.0724 | 4.9713 |  |  | 592 | 23681.319 | 10.0724 | 4.988 |  |  | 592 | 23681.319 | 10.0724 | 4.9905
593 | 23721.321 | 10.0741 | 4.9711 |  |  | 593 | 23721.321 | 10.0741 | 4.9879 |  |  | 593 | 23721.321 | 10.0741 | 4.9904
594 | 23761.323 | 10.0758 | 4.9709 |  |  | 594 | 23761.323 | 10.0758 | 4.9878 |  |  | 594 | 23761.323 | 10.0758 | 4.9903
595 | 23801.326 | 10.0775 | 4.9707 |  |  | 595 | 23801.326 | 10.0775 | 4.9877 |  |  | 595 | 23801.326 | 10.0775 | 4.9902
596 | 23841.328 | 10.0792 | 4.9706 |  |  | 596 | 23841.328 | 10.0792 | 4.9876 |  |  | 596 | 23841.328 | 10.0792 | 4.9901
597 | 23881.33 | 10.0809 | 4.9704 |  |  | 597 | 23881.33 | 10.0809 | 4.9875 |  |  | 597 | 23881.33 | 10.0809 | 4.99
598 | 23921.332 | 10.0825 | 4.9702 |  |  | 598 | 23921.332 | 10.0825 | 4.9874 |  |  | 598 | 23921.332 | 10.0825 | 4.9899
599 | 23961.335 | 10.0842 | 4.97 |  |  | 599 | 23961.335 | 10.0842 | 4.9874 |  |  | 599 | 23961.335 | 10.0842 | 4.9898
600 | 24001.337 | 10.0859 | 4.9698 |  |  | 600 | 24001.337 | 10.0859 | 4.9873 |  |  | 600 | 24001.337 | 10.0859 | 4.9897
601 | 24041.339 | 10.0875 | 4.9697 |  |  | 601 | 24041.339 | 10.0875 | 4.9872 |  |  | 601 | 24041.339 | 10.0875 | 4.9896
602 | 24081.341 | 10.0892 | 4.9695 |  |  | 602 | 24081.341 | 10.0892 | 4.9871 |  |  | 602 | 24081.341 | 10.0892 | 4.9895
603 | 24121.343 | 10.0909 | 4.9693 |  |  | 603 | 24121.343 | 10.0909 | 4.987 |  |  | 603 | 24121.343 | 10.0909 | 4.9894
604 | 24161.346 | 10.0925 | 4.9691 |  |  | 604 | 24161.346 | 10.0925 | 4.9869 |  |  | 604 | 24161.346 | 10.0925 | 4.9893
605 | 24201.348 | 10.0942 | 4.9689 |  |  | 605 | 24201.348 | 10.0942 | 4.9868 |  |  | 605 | 24201.348 | 10.0942 | 4.9892
606 | 24241.35 | 10.0958 | 4.9688 |  |  | 606 | 24241.35 | 10.0958 | 4.9868 |  |  | 606 | 24241.35 | 10.0958 | 4.9891
607 | 24281.352 | 10.0975 | 4.9686 |  |  | 607 | 24281.352 | 10.0975 | 4.9867 |  |  | 607 | 24281.352 | 10.0975 | 4.989
608 | 24321.355 | 10.0991 | 4.9684 |  |  | 608 | 24321.355 | 10.0991 | 4.9866 |  |  | 608 | 24321.355 | 10.0991 | 4.9889
609 | 24361.357 | 10.1008 | 4.9682 |  |  | 609 | 24361.357 | 10.1008 | 4.9865 |  |  | 609 | 24361.357 | 10.1008 | 4.9888
610 | 24401.359 | 10.1024 | 4.968 |  |  | 610 | 24401.359 | 10.1024 | 4.9864 |  |  | 610 | 24401.359 | 10.1024 | 4.9887
611 | 24441.361 | 10.104 | 4.9679 |  |  | 611 | 24441.361 | 10.104 | 4.9863 |  |  | 611 | 24441.361 | 10.104 | 4.9886
612 | 24481.363 | 10.1057 | 4.9677 |  |  | 612 | 24481.363 | 10.1057 | 4.9862 |  |  | 612 | 24481.363 | 10.1057 | 4.9885
613 | 24521.366 | 10.1073 | 4.9675 |  |  | 613 | 24521.366 | 10.1073 | 4.9862 |  |  | 613 | 24521.366 | 10.1073 | 4.9884
614 | 24561.368 | 10.1089 | 4.9673 |  |  | 614 | 24561.368 | 10.1089 | 4.9861 |  |  | 614 | 24561.368 | 10.1089 | 4.9883
615 | 24601.37 | 10.1106 | 4.9672 |  |  | 615 | 24601.37 | 10.1106 | 4.986 |  |  | 615 | 24601.37 | 10.1106 | 4.9882
616 | 24641.372 | 10.1122 | 4.967 |  |  | 616 | 24641.372 | 10.1122 | 4.9859 |  |  | 616 | 24641.372 | 10.1122 | 4.9881
617 | 24681.375 | 10.1138 | 4.9668 |  |  | 617 | 24681.375 | 10.1138 | 4.9858 |  |  | 617 | 24681.375 | 10.1138 | 4.9881
618 | 24721.377 | 10.1154 | 4.9666 |  |  | 618 | 24721.377 | 10.1154 | 4.9857 |  |  | 618 | 24721.377 | 10.1154 | 4.988
619 | 24761.379 | 10.117 | 4.9665 |  |  | 619 | 24761.379 | 10.117 | 4.9857 |  |  | 619 | 24761.379 | 10.117 | 4.9879
620 | 24801.381 | 10.1187 | 4.9663 |  |  | 620 | 24801.381 | 10.1187 | 4.9856 |  |  | 620 | 24801.381 | 10.1187 | 4.9878
621 | 24841.384 | 10.1203 | 4.9661 |  |  | 621 | 24841.384 | 10.1203 | 4.9855 |  |  | 621 | 24841.384 | 10.1203 | 4.9877
622 | 24881.386 | 10.1219 | 4.9659 |  |  | 622 | 24881.386 | 10.1219 | 4.9854 |  |  | 622 | 24881.386 | 10.1219 | 4.9876
623 | 24921.388 | 10.1235 | 4.9658 |  |  | 623 | 24921.388 | 10.1235 | 4.9853 |  |  | 623 | 24921.388 | 10.1235 | 4.9875
624 | 24961.39 | 10.1251 | 4.9656 |  |  | 624 | 24961.39 | 10.1251 | 4.9853 |  |  | 624 | 24961.39 | 10.1251 | 4.9874
625 | 25001.392 | 10.1267 | 4.9654 |  |  | 625 | 25001.392 | 10.1267 | 4.9852 |  |  | 625 | 25001.392 | 10.1267 | 4.9873
626 | 25041.395 | 10.1283 | 4.9653 |  |  | 626 | 25041.395 | 10.1283 | 4.9851 |  |  | 626 | 25041.395 | 10.1283 | 4.9872
627 | 25081.397 | 10.1299 | 4.9651 |  |  | 627 | 25081.397 | 10.1299 | 4.985 |  |  | 627 | 25081.397 | 10.1299 | 4.9871
628 | 25121.399 | 10.1315 | 4.9649 |  |  | 628 | 25121.399 | 10.1315 | 4.9849 |  |  | 628 | 25121.399 | 10.1315 | 4.987
629 | 25161.401 | 10.1331 | 4.9647 |  |  | 629 | 25161.401 | 10.1331 | 4.9849 |  |  | 629 | 25161.401 | 10.1331 | 4.987
630 | 25201.404 | 10.1347 | 4.9646 |  |  | 630 | 25201.404 | 10.1347 | 4.9848 |  |  | 630 | 25201.404 | 10.1347 | 4.9869
631 | 25241.406 | 10.1362 | 4.9644 |  |  | 631 | 25241.406 | 10.1362 | 4.9847 |  |  | 631 | 25241.406 | 10.1362 | 4.9868
632 | 25281.408 | 10.1378 | 4.9642 |  |  | 632 | 25281.408 | 10.1378 | 4.9846 |  |  | 632 | 25281.408 | 10.1378 | 4.9867
633 | 25321.41 | 10.1394 | 4.9641 |  |  | 633 | 25321.41 | 10.1394 | 4.9846 |  |  | 633 | 25321.41 | 10.1394 | 4.9866
634 | 25361.412 | 10.141 | 4.9639 |  |  | 634 | 25361.412 | 10.141 | 4.9845 |  |  | 634 | 25361.412 | 10.141 | 4.9865
635 | 25401.415 | 10.1426 | 4.9637 |  |  | 635 | 25401.415 | 10.1426 | 4.9844 |  |  | 635 | 25401.415 | 10.1426 | 4.9864
636 | 25441.417 | 10.1441 | 4.9636 |  |  | 636 | 25441.417 | 10.1441 | 4.9843 |  |  | 636 | 25441.417 | 10.1441 | 4.9863
637 | 25481.419 | 10.1457 | 4.9634 |  |  | 637 | 25481.419 | 10.1457 | 4.9843 |  |  | 637 | 25481.419 | 10.1457 | 4.9862
638 | 25521.421 | 10.1473 | 4.9632 |  |  | 638 | 25521.421 | 10.1473 | 4.9842 |  |  | 638 | 25521.421 | 10.1473 | 4.9862
639 | 25561.424 | 10.1488 | 4.9631 |  |  | 639 | 25561.424 | 10.1488 | 4.9841 |  |  | 639 | 25561.424 | 10.1488 | 4.9861
640 | 25601.426 | 10.1504 | 4.9629 |  |  | 640 | 25601.426 | 10.1504 | 4.984 |  |  | 640 | 25601.426 | 10.1504 | 4.986
641 | 25641.428 | 10.152 | 4.9627 |  |  | 641 | 25641.428 | 10.152 | 4.9839 |  |  | 641 | 25641.428 | 10.152 | 4.9859
642 | 25681.43 | 10.1535 | 4.9626 |  |  | 642 | 25681.43 | 10.1535 | 4.9839 |  |  | 642 | 25681.43 | 10.1535 | 4.9858
643 | 25721.433 | 10.1551 | 4.9624 |  |  | 643 | 25721.433 | 10.1551 | 4.9838 |  |  | 643 | 25721.433 | 10.1551 | 4.9857
644 | 25761.435 | 10.1566 | 4.9622 |  |  | 644 | 25761.435 | 10.1566 | 4.9837 |  |  | 644 | 25761.435 | 10.1566 | 4.9856
645 | 25801.437 | 10.1582 | 4.9621 |  |  | 645 | 25801.437 | 10.1582 | 4.9836 |  |  | 645 | 25801.437 | 10.1582 | 4.9855
646 | 25841.439 | 10.1597 | 4.9619 |  |  | 646 | 25841.439 | 10.1597 | 4.9836 |  |  | 646 | 25841.439 | 10.1597 | 4.9855
647 | 25881.441 | 10.1613 | 4.9617 |  |  | 647 | 25881.441 | 10.1613 | 4.9835 |  |  | 647 | 25881.441 | 10.1613 | 4.9854
648 | 25921.444 | 10.1628 | 4.9616 |  |  | 648 | 25921.444 | 10.1628 | 4.9834 |  |  | 648 | 25921.444 | 10.1628 | 4.9853
649 | 25961.446 | 10.1644 | 4.9614 |  |  | 649 | 25961.446 | 10.1644 | 4.9834 |  |  | 649 | 25961.446 | 10.1644 | 4.9852
650 | 26001.448 | 10.1659 | 4.9612 |  |  | 650 | 26001.448 | 10.1659 | 4.9833 |  |  | 650 | 26001.448 | 10.1659 | 4.9851
651 | 26041.45 | 10.1674 | 4.9611 |  |  | 651 | 26041.45 | 10.1674 | 4.9832 |  |  | 651 | 26041.45 | 10.1674 | 4.985
652 | 26081.453 | 10.169 | 4.9609 |  |  | 652 | 26081.453 | 10.169 | 4.9831 |  |  | 652 | 26081.453 | 10.169 | 4.9849
653 | 26121.455 | 10.1705 | 4.9608 |  |  | 653 | 26121.455 | 10.1705 | 4.9831 |  |  | 653 | 26121.455 | 10.1705 | 4.9849
654 | 26161.457 | 10.172 | 4.9606 |  |  | 654 | 26161.457 | 10.172 | 4.983 |  |  | 654 | 26161.457 | 10.172 | 4.9848
655 | 26201.459 | 10.1736 | 4.9604 |  |  | 655 | 26201.459 | 10.1736 | 4.9829 |  |  | 655 | 26201.459 | 10.1736 | 4.9847
656 | 26241.462 | 10.1751 | 4.9603 |  |  | 656 | 26241.462 | 10.1751 | 4.9828 |  |  | 656 | 26241.462 | 10.1751 | 4.9846
657 | 26281.464 | 10.1766 | 4.9601 |  |  | 657 | 26281.464 | 10.1766 | 4.9828 |  |  | 657 | 26281.464 | 10.1766 | 4.9845
658 | 26321.466 | 10.1781 | 4.9599 |  |  | 658 | 26321.466 | 10.1781 | 4.9827 |  |  | 658 | 26321.466 | 10.1781 | 4.9844
659 | 26361.468 | 10.1797 | 4.9598 |  |  | 659 | 26361.468 | 10.1797 | 4.9826 |  |  | 659 | 26361.468 | 10.1797 | 4.9843
660 | 26401.47 | 10.1812 | 4.9596 |  |  | 660 | 26401.47 | 10.1812 | 4.9826 |  |  | 660 | 26401.47 | 10.1812 | 4.9843
661 | 26441.473 | 10.1827 | 4.9595 |  |  | 661 | 26441.473 | 10.1827 | 4.9825 |  |  | 661 | 26441.473 | 10.1827 | 4.9842
662 | 26481.475 | 10.1842 | 4.9593 |  |  | 662 | 26481.475 | 10.1842 | 4.9824 |  |  | 662 | 26481.475 | 10.1842 | 4.9841
663 | 26521.477 | 10.1857 | 4.9591 |  |  | 663 | 26521.477 | 10.1857 | 4.9823 |  |  | 663 | 26521.477 | 10.1857 | 4.984
664 | 26561.479 | 10.1872 | 4.959 |  |  | 664 | 26561.479 | 10.1872 | 4.9823 |  |  | 664 | 26561.479 | 10.1872 | 4.9839
665 | 26601.482 | 10.1887 | 4.9588 |  |  | 665 | 26601.482 | 10.1887 | 4.9822 |  |  | 665 | 26601.482 | 10.1887 | 4.9838
666 | 26641.484 | 10.1902 | 4.9587 |  |  | 666 | 26641.484 | 10.1902 | 4.9821 |  |  | 666 | 26641.484 | 10.1902 | 4.9838
667 | 26681.486 | 10.1917 | 4.9585 |  |  | 667 | 26681.486 | 10.1917 | 4.9821 |  |  | 667 | 26681.486 | 10.1917 | 4.9837
668 | 26721.488 | 10.1932 | 4.9583 |  |  | 668 | 26721.488 | 10.1932 | 4.982 |  |  | 668 | 26721.488 | 10.1932 | 4.9836
669 | 26761.49 | 10.1947 | 4.9582 |  |  | 669 | 26761.49 | 10.1947 | 4.9819 |  |  | 669 | 26761.49 | 10.1947 | 4.9835
670 | 26801.493 | 10.1962 | 4.958 |  |  | 670 | 26801.493 | 10.1962 | 4.9818 |  |  | 670 | 26801.493 | 10.1962 | 4.9834
671 | 26841.495 | 10.1977 | 4.9579 |  |  | 671 | 26841.495 | 10.1977 | 4.9818 |  |  | 671 | 26841.495 | 10.1977 | 4.9834
672 | 26881.497 | 10.1992 | 4.9577 |  |  | 672 | 26881.497 | 10.1992 | 4.9817 |  |  | 672 | 26881.497 | 10.1992 | 4.9833
673 | 26921.499 | 10.2007 | 4.9576 |  |  | 673 | 26921.499 | 10.2007 | 4.9816 |  |  | 673 | 26921.499 | 10.2007 | 4.9832
674 | 26961.502 | 10.2022 | 4.9574 |  |  | 674 | 26961.502 | 10.2022 | 4.9816 |  |  | 674 | 26961.502 | 10.2022 | 4.9831
675 | 27001.504 | 10.2036 | 4.9572 |  |  | 675 | 27001.504 | 10.2036 | 4.9815 |  |  | 675 | 27001.504 | 10.2036 | 4.983
676 | 27041.506 | 10.2051 | 4.9571 |  |  | 676 | 27041.506 | 10.2051 | 4.9814 |  |  | 676 | 27041.506 | 10.2051 | 4.9829
677 | 27081.508 | 10.2066 | 4.9569 |  |  | 677 | 27081.508 | 10.2066 | 4.9814 |  |  | 677 | 27081.508 | 10.2066 | 4.9829
678 | 27121.511 | 10.2081 | 4.9568 |  |  | 678 | 27121.511 | 10.2081 | 4.9813 |  |  | 678 | 27121.511 | 10.2081 | 4.9828
679 | 27161.513 | 10.2096 | 4.9566 |  |  | 679 | 27161.513 | 10.2096 | 4.9812 |  |  | 679 | 27161.513 | 10.2096 | 4.9827
680 | 27201.515 | 10.211 | 4.9565 |  |  | 680 | 27201.515 | 10.211 | 4.9812 |  |  | 680 | 27201.515 | 10.211 | 4.9826
681 | 27241.517 | 10.2125 | 4.9563 |  |  | 681 | 27241.517 | 10.2125 | 4.9811 |  |  | 681 | 27241.517 | 10.2125 | 4.9825
682 | 27281.519 | 10.214 | 4.9562 |  |  | 682 | 27281.519 | 10.214 | 4.981 |  |  | 682 | 27281.519 | 10.214 | 4.9825
683 | 27321.522 | 10.2154 | 4.956 |  |  | 683 | 27321.522 | 10.2154 | 4.981 |  |  | 683 | 27321.522 | 10.2154 | 4.9824
684 | 27361.524 | 10.2169 | 4.9558 |  |  | 684 | 27361.524 | 10.2169 | 4.9809 |  |  | 684 | 27361.524 | 10.2169 | 4.9823
685 | 27401.526 | 10.2184 | 4.9557 |  |  | 685 | 27401.526 | 10.2184 | 4.9808 |  |  | 685 | 27401.526 | 10.2184 | 4.9822
686 | 27441.528 | 10.2198 | 4.9555 |  |  | 686 | 27441.528 | 10.2198 | 4.9808 |  |  | 686 | 27441.528 | 10.2198 | 4.9821
687 | 27481.531 | 10.2213 | 4.9554 |  |  | 687 | 27481.531 | 10.2213 | 4.9807 |  |  | 687 | 27481.531 | 10.2213 | 4.9821
688 | 27521.533 | 10.2227 | 4.9552 |  |  | 688 | 27521.533 | 10.2227 | 4.9806 |  |  | 688 | 27521.533 | 10.2227 | 4.982
689 | 27561.535 | 10.2242 | 4.9551 |  |  | 689 | 27561.535 | 10.2242 | 4.9806 |  |  | 689 | 27561.535 | 10.2242 | 4.9819
690 | 27601.537 | 10.2256 | 4.9549 |  |  | 690 | 27601.537 | 10.2256 | 4.9805 |  |  | 690 | 27601.537 | 10.2256 | 4.9818
691 | 27641.539 | 10.2271 | 4.9548 |  |  | 691 | 27641.539 | 10.2271 | 4.9804 |  |  | 691 | 27641.539 | 10.2271 | 4.9817
692 | 27681.542 | 10.2285 | 4.9546 |  |  | 692 | 27681.542 | 10.2285 | 4.9804 |  |  | 692 | 27681.542 | 10.2285 | 4.9817
693 | 27721.544 | 10.23 | 4.9545 |  |  | 693 | 27721.544 | 10.23 | 4.9803 |  |  | 693 | 27721.544 | 10.23 | 4.9816
694 | 27761.546 | 10.2314 | 4.9543 |  |  | 694 | 27761.546 | 10.2314 | 4.9802 |  |  | 694 | 27761.546 | 10.2314 | 4.9815
695 | 27801.548 | 10.2328 | 4.9542 |  |  | 695 | 27801.548 | 10.2328 | 4.9802 |  |  | 695 | 27801.548 | 10.2328 | 4.9814
696 | 27841.551 | 10.2343 | 4.954 |  |  | 696 | 27841.551 | 10.2343 | 4.9801 |  |  | 696 | 27841.551 | 10.2343 | 4.9814
697 | 27881.553 | 10.2357 | 4.9539 |  |  | 697 | 27881.553 | 10.2357 | 4.98 |  |  | 697 | 27881.553 | 10.2357 | 4.9813
698 | 27921.555 | 10.2372 | 4.9537 |  |  | 698 | 27921.555 | 10.2372 | 4.98 |  |  | 698 | 27921.555 | 10.2372 | 4.9812
699 | 27961.557 | 10.2386 | 4.9536 |  |  | 699 | 27961.557 | 10.2386 | 4.9799 |  |  | 699 | 27961.557 | 10.2386 | 4.9811
700 | 28001.56 | 10.24 | 4.9534 |  |  | 700 | 28001.56 | 10.24 | 4.9798 |  |  | 700 | 28001.56 | 10.24 | 4.981
701 | 28041.562 | 10.2414 | 4.9532 |  |  | 701 | 28041.562 | 10.2414 | 4.9798 |  |  | 701 | 28041.562 | 10.2414 | 4.981
702 | 28081.564 | 10.2429 | 4.9531 |  |  | 702 | 28081.564 | 10.2429 | 4.9797 |  |  | 702 | 28081.564 | 10.2429 | 4.9809
703 | 28121.566 | 10.2443 | 4.9529 |  |  | 703 | 28121.566 | 10.2443 | 4.9796 |  |  | 703 | 28121.566 | 10.2443 | 4.9808
704 | 28161.568 | 10.2457 | 4.9528 |  |  | 704 | 28161.568 | 10.2457 | 4.9796 |  |  | 704 | 28161.568 | 10.2457 | 4.9807
705 | 28201.571 | 10.2471 | 4.9526 |  |  | 705 | 28201.571 | 10.2471 | 4.9795 |  |  | 705 | 28201.571 | 10.2471 | 4.9806
706 | 28241.573 | 10.2486 | 4.9525 |  |  | 706 | 28241.573 | 10.2486 | 4.9794 |  |  | 706 | 28241.573 | 10.2486 | 4.9806
707 | 28281.575 | 10.25 | 4.9524 |  |  | 707 | 28281.575 | 10.25 | 4.9794 |  |  | 707 | 28281.575 | 10.25 | 4.9805
708 | 28321.577 | 10.2514 | 4.9522 |  |  | 708 | 28321.577 | 10.2514 | 4.9793 |  |  | 708 | 28321.577 | 10.2514 | 4.9804
709 | 28361.58 | 10.2528 | 4.9521 |  |  | 709 | 28361.58 | 10.2528 | 4.9792 |  |  | 709 | 28361.58 | 10.2528 | 4.9803
710 | 28401.582 | 10.2542 | 4.9519 |  |  | 710 | 28401.582 | 10.2542 | 4.9792 |  |  | 710 | 28401.582 | 10.2542 | 4.9803
711 | 28441.584 | 10.2556 | 4.9518 |  |  | 711 | 28441.584 | 10.2556 | 4.9791 |  |  | 711 | 28441.584 | 10.2556 | 4.9802
712 | 28481.586 | 10.257 | 4.9516 |  |  | 712 | 28481.586 | 10.257 | 4.9791 |  |  | 712 | 28481.586 | 10.257 | 4.9801
713 | 28521.588 | 10.2584 | 4.9515 |  |  | 713 | 28521.588 | 10.2584 | 4.979 |  |  | 713 | 28521.588 | 10.2584 | 4.98
714 | 28561.591 | 10.2598 | 4.9513 |  |  | 714 | 28561.591 | 10.2598 | 4.9789 |  |  | 714 | 28561.591 | 10.2598 | 4.9799
715 | 28601.593 | 10.2612 | 4.9512 |  |  | 715 | 28601.593 | 10.2612 | 4.9789 |  |  | 715 | 28601.593 | 10.2612 | 4.9799
716 | 28641.595 | 10.2626 | 4.951 |  |  | 716 | 28641.595 | 10.2626 | 4.9788 |  |  | 716 | 28641.595 | 10.2626 | 4.9798
717 | 28681.597 | 10.264 | 4.9509 |  |  | 717 | 28681.597 | 10.264 | 4.9787 |  |  | 717 | 28681.597 | 10.264 | 4.9797
718 | 28721.6 | 10.2654 | 4.9507 |  |  | 718 | 28721.6 | 10.2654 | 4.9787 |  |  | 718 | 28721.6 | 10.2654 | 4.9796
719 | 28761.602 | 10.2668 | 4.9506 |  |  | 719 | 28761.602 | 10.2668 | 4.9786 |  |  | 719 | 28761.602 | 10.2668 | 4.9796
720 | 28801.604 | 10.2682 | 4.9504 |  |  | 720 | 28801.604 | 10.2682 | 4.9785 |  |  | 720 | 28801.604 | 10.2682 | 4.9795
721 | 28841.606 | 10.2696 | 4.9503 |  |  | 721 | 28841.606 | 10.2696 | 4.9785 |  |  | 721 | 28841.606 | 10.2696 | 4.9794
722 | 28881.609 | 10.271 | 4.9501 |  |  | 722 | 28881.609 | 10.271 | 4.9784 |  |  | 722 | 28881.609 | 10.271 | 4.9793
723 | 28921.611 | 10.2723 | 4.95 |  |  | 723 | 28921.611 | 10.2723 | 4.9784 |  |  | 723 | 28921.611 | 10.2723 | 4.9793
724 | 28961.613 | 10.2737 | 4.9498 |  |  | 724 | 28961.613 | 10.2737 | 4.9783 |  |  | 724 | 28961.613 | 10.2737 | 4.9792
725 | 29001.615 | 10.2751 | 4.9497 |  |  | 725 | 29001.615 | 10.2751 | 4.9782 |  |  | 725 | 29001.615 | 10.2751 | 4.9791
726 | 29041.617 | 10.2765 | 4.9496 |  |  | 726 | 29041.617 | 10.2765 | 4.9782 |  |  | 726 | 29041.617 | 10.2765 | 4.979
727 | 29081.62 | 10.2779 | 4.9494 |  |  | 727 | 29081.62 | 10.2779 | 4.9781 |  |  | 727 | 29081.62 | 10.2779 | 4.9789
728 | 29121.622 | 10.2792 | 4.9493 |  |  | 728 | 29121.622 | 10.2792 | 4.978 |  |  | 728 | 29121.622 | 10.2792 | 4.9789
729 | 29161.624 | 10.2806 | 4.9491 |  |  | 729 | 29161.624 | 10.2806 | 4.978 |  |  | 729 | 29161.624 | 10.2806 | 4.9788
730 | 29201.626 | 10.282 | 4.949 |  |  | 730 | 29201.626 | 10.282 | 4.9779 |  |  | 730 | 29201.626 | 10.282 | 4.9787
731 | 29241.629 | 10.2833 | 4.9488 |  |  | 731 | 29241.629 | 10.2833 | 4.9778 |  |  | 731 | 29241.629 | 10.2833 | 4.9786
732 | 29281.631 | 10.2847 | 4.9487 |  |  | 732 | 29281.631 | 10.2847 | 4.9778 |  |  | 732 | 29281.631 | 10.2847 | 4.9786
733 | 29321.633 | 10.2861 | 4.9485 |  |  | 733 | 29321.633 | 10.2861 | 4.9777 |  |  | 733 | 29321.633 | 10.2861 | 4.9785
734 | 29361.635 | 10.2874 | 4.9484 |  |  | 734 | 29361.635 | 10.2874 | 4.9777 |  |  | 734 | 29361.635 | 10.2874 | 4.9784
735 | 29401.638 | 10.2888 | 4.9483 |  |  | 735 | 29401.638 | 10.2888 | 4.9776 |  |  | 735 | 29401.638 | 10.2888 | 4.9783
736 | 29441.64 | 10.2902 | 4.9481 |  |  | 736 | 29441.64 | 10.2902 | 4.9775 |  |  | 736 | 29441.64 | 10.2902 | 4.9783
737 | 29481.642 | 10.2915 | 4.948 |  |  | 737 | 29481.642 | 10.2915 | 4.9775 |  |  | 737 | 29481.642 | 10.2915 | 4.9782
738 | 29521.644 | 10.2929 | 4.9478 |  |  | 738 | 29521.644 | 10.2929 | 4.9774 |  |  | 738 | 29521.644 | 10.2929 | 4.9781
739 | 29561.646 | 10.2942 | 4.9477 |  |  | 739 | 29561.646 | 10.2942 | 4.9774 |  |  | 739 | 29561.646 | 10.2942 | 4.978
740 | 29601.649 | 10.2956 | 4.9475 |  |  | 740 | 29601.649 | 10.2956 | 4.9773 |  |  | 740 | 29601.649 | 10.2956 | 4.978
741 | 29641.651 | 10.2969 | 4.9474 |  |  | 741 | 29641.651 | 10.2969 | 4.9772 |  |  | 741 | 29641.651 | 10.2969 | 4.9779
742 | 29681.653 | 10.2983 | 4.9472 |  |  | 742 | 29681.653 | 10.2983 | 4.9772 |  |  | 742 | 29681.653 | 10.2983 | 4.9778
743 | 29721.655 | 10.2996 | 4.9471 |  |  | 743 | 29721.655 | 10.2996 | 4.9771 |  |  | 743 | 29721.655 | 10.2996 | 4.9777
744 | 29761.658 | 10.301 | 4.947 |  |  | 744 | 29761.658 | 10.301 | 4.977 |  |  | 744 | 29761.658 | 10.301 | 4.9777
745 | 29801.66 | 10.3023 | 4.9468 |  |  | 745 | 29801.66 | 10.3023 | 4.977 |  |  | 745 | 29801.66 | 10.3023 | 4.9776
746 | 29841.662 | 10.3037 | 4.9467 |  |  | 746 | 29841.662 | 10.3037 | 4.9769 |  |  | 746 | 29841.662 | 10.3037 | 4.9775
747 | 29881.664 | 10.305 | 4.9465 |  |  | 747 | 29881.664 | 10.305 | 4.9769 |  |  | 747 | 29881.664 | 10.305 | 4.9774
748 | 29921.666 | 10.3063 | 4.9464 |  |  | 748 | 29921.666 | 10.3063 | 4.9768 |  |  | 748 | 29921.666 | 10.3063 | 4.9774
749 | 29961.669 | 10.3077 | 4.9463 |  |  | 749 | 29961.669 | 10.3077 | 4.9767 |  |  | 749 | 29961.669 | 10.3077 | 4.9773
750 | 30001.671 | 10.309 | 4.9461 |  |  | 750 | 30001.671 | 10.309 | 4.9767 |  |  | 750 | 30001.671 | 10.309 | 4.9772
751 | 30041.673 | 10.3103 | 4.946 |  |  | 751 | 30041.673 | 10.3103 | 4.9766 |  |  | 751 | 30041.673 | 10.3103 | 4.9771
752 | 30081.675 | 10.3117 | 4.9458 |  |  | 752 | 30081.675 | 10.3117 | 4.9765 |  |  | 752 | 30081.675 | 10.3117 | 4.9771
753 | 30121.678 | 10.313 | 4.9457 |  |  | 753 | 30121.678 | 10.313 | 4.9765 |  |  | 753 | 30121.678 | 10.313 | 4.977
754 | 30161.68 | 10.3143 | 4.9456 |  |  | 754 | 30161.68 | 10.3143 | 4.9764 |  |  | 754 | 30161.68 | 10.3143 | 4.9769
755 | 30201.682 | 10.3157 | 4.9454 |  |  | 755 | 30201.682 | 10.3157 | 4.9764 |  |  | 755 | 30201.682 | 10.3157 | 4.9768
756 | 30241.684 | 10.317 | 4.9453 |  |  | 756 | 30241.684 | 10.317 | 4.9763 |  |  | 756 | 30241.684 | 10.317 | 4.9768
757 | 30281.687 | 10.3183 | 4.9451 |  |  | 757 | 30281.687 | 10.3183 | 4.9762 |  |  | 757 | 30281.687 | 10.3183 | 4.9767
758 | 30321.689 | 10.3196 | 4.945 |  |  | 758 | 30321.689 | 10.3196 | 4.9762 |  |  | 758 | 30321.689 | 10.3196 | 4.9766
759 | 30361.691 | 10.3209 | 4.9449 |  |  | 759 | 30361.691 | 10.3209 | 4.9761 |  |  | 759 | 30361.691 | 10.3209 | 4.9765
760 | 30401.693 | 10.3223 | 4.9447 |  |  | 760 | 30401.693 | 10.3223 | 4.9761 |  |  | 760 | 30401.693 | 10.3223 | 4.9765
761 | 30441.695 | 10.3236 | 4.9446 |  |  | 761 | 30441.695 | 10.3236 | 4.976 |  |  | 761 | 30441.695 | 10.3236 | 4.9764
762 | 30481.698 | 10.3249 | 4.9444 |  |  | 762 | 30481.698 | 10.3249 | 4.9759 |  |  | 762 | 30481.698 | 10.3249 | 4.9763
763 | 30521.7 | 10.3262 | 4.9443 |  |  | 763 | 30521.7 | 10.3262 | 4.9759 |  |  | 763 | 30521.7 | 10.3262 | 4.9762
764 | 30561.702 | 10.3275 | 4.9442 |  |  | 764 | 30561.702 | 10.3275 | 4.9758 |  |  | 764 | 30561.702 | 10.3275 | 4.9762
765 | 30601.704 | 10.3288 | 4.944 |  |  | 765 | 30601.704 | 10.3288 | 4.9758 |  |  | 765 | 30601.704 | 10.3288 | 4.9761
766 | 30641.707 | 10.3301 | 4.9439 |  |  | 766 | 30641.707 | 10.3301 | 4.9757 |  |  | 766 | 30641.707 | 10.3301 | 4.976
767 | 30681.709 | 10.3314 | 4.9437 |  |  | 767 | 30681.709 | 10.3314 | 4.9756 |  |  | 767 | 30681.709 | 10.3314 | 4.9759
768 | 30721.711 | 10.3327 | 4.9436 |  |  | 768 | 30721.711 | 10.3327 | 4.9756 |  |  | 768 | 30721.711 | 10.3327 | 4.9759
769 | 30761.713 | 10.334 | 4.9435 |  |  | 769 | 30761.713 | 10.334 | 4.9755 |  |  | 769 | 30761.713 | 10.334 | 4.9758
770 | 30801.715 | 10.3353 | 4.9433 |  |  | 770 | 30801.715 | 10.3353 | 4.9754 |  |  | 770 | 30801.715 | 10.3353 | 4.9757
771 | 30841.718 | 10.3366 | 4.9432 |  |  | 771 | 30841.718 | 10.3366 | 4.9754 |  |  | 771 | 30841.718 | 10.3366 | 4.9756
772 | 30881.72 | 10.3379 | 4.9431 |  |  | 772 | 30881.72 | 10.3379 | 4.9753 |  |  | 772 | 30881.72 | 10.3379 | 4.9756
773 | 30921.722 | 10.3392 | 4.9429 |  |  | 773 | 30921.722 | 10.3392 | 4.9753 |  |  | 773 | 30921.722 | 10.3392 | 4.9755
774 | 30961.724 | 10.3405 | 4.9428 |  |  | 774 | 30961.724 | 10.3405 | 4.9752 |  |  | 774 | 30961.724 | 10.3405 | 4.9754
775 | 31001.727 | 10.3418 | 4.9426 |  |  | 775 | 31001.727 | 10.3418 | 4.9751 |  |  | 775 | 31001.727 | 10.3418 | 4.9753
776 | 31041.729 | 10.3431 | 4.9425 |  |  | 776 | 31041.729 | 10.3431 | 4.9751 |  |  | 776 | 31041.729 | 10.3431 | 4.9753
777 | 31081.731 | 10.3444 | 4.9424 |  |  | 777 | 31081.731 | 10.3444 | 4.975 |  |  | 777 | 31081.731 | 10.3444 | 4.9752
778 | 31121.733 | 10.3457 | 4.9422 |  |  | 778 | 31121.733 | 10.3457 | 4.975 |  |  | 778 | 31121.733 | 10.3457 | 4.9751
779 | 31161.736 | 10.3469 | 4.9421 |  |  | 779 | 31161.736 | 10.3469 | 4.9749 |  |  | 779 | 31161.736 | 10.3469 | 4.975
780 | 31201.738 | 10.3482 | 4.942 |  |  | 780 | 31201.738 | 10.3482 | 4.9748 |  |  | 780 | 31201.738 | 10.3482 | 4.975
781 | 31241.74 | 10.3495 | 4.9418 |  |  | 781 | 31241.74 | 10.3495 | 4.9748 |  |  | 781 | 31241.74 | 10.3495 | 4.9749
782 | 31281.742 | 10.3508 | 4.9417 |  |  | 782 | 31281.742 | 10.3508 | 4.9747 |  |  | 782 | 31281.742 | 10.3508 | 4.9748
783 | 31321.744 | 10.3521 | 4.9415 |  |  | 783 | 31321.744 | 10.3521 | 4.9747 |  |  | 783 | 31321.744 | 10.3521 | 4.9747
784 | 31361.747 | 10.3533 | 4.9414 |  |  | 784 | 31361.747 | 10.3533 | 4.9746 |  |  | 784 | 31361.747 | 10.3533 | 4.9747
785 | 31401.749 | 10.3546 | 4.9413 |  |  | 785 | 31401.749 | 10.3546 | 4.9745 |  |  | 785 | 31401.749 | 10.3546 | 4.9746
786 | 31441.751 | 10.3559 | 4.9411 |  |  | 786 | 31441.751 | 10.3559 | 4.9745 |  |  | 786 | 31441.751 | 10.3559 | 4.9745
787 | 31481.753 | 10.3572 | 4.941 |  |  | 787 | 31481.753 | 10.3572 | 4.9744 |  |  | 787 | 31481.753 | 10.3572 | 4.9744
788 | 31521.756 | 10.3584 | 4.9409 |  |  | 788 | 31521.756 | 10.3584 | 4.9744 |  |  | 788 | 31521.756 | 10.3584 | 4.9744
789 | 31561.758 | 10.3597 | 4.9407 |  |  | 789 | 31561.758 | 10.3597 | 4.9743 |  |  | 789 | 31561.758 | 10.3597 | 4.9743
790 | 31601.76 | 10.361 | 4.9406 |  |  | 790 | 31601.76 | 10.361 | 4.9742 |  |  | 790 | 31601.76 | 10.361 | 4.9742
791 | 31641.762 | 10.3622 | 4.9405 |  |  | 791 | 31641.762 | 10.3622 | 4.9742 |  |  | 791 | 31641.762 | 10.3622 | 4.9741
792 | 31681.765 | 10.3635 | 4.9403 |  |  | 792 | 31681.765 | 10.3635 | 4.9741 |  |  | 792 | 31681.765 | 10.3635 | 4.9741
793 | 31721.767 | 10.3648 | 4.9402 |  |  | 793 | 31721.767 | 10.3648 | 4.9741 |  |  | 793 | 31721.767 | 10.3648 | 4.974
794 | 31761.769 | 10.366 | 4.9401 |  |  | 794 | 31761.769 | 10.366 | 4.974 |  |  | 794 | 31761.769 | 10.366 | 4.9739
795 | 31801.771 | 10.3673 | 4.9399 |  |  | 795 | 31801.771 | 10.3673 | 4.9739 |  |  | 795 | 31801.771 | 10.3673 | 4.9738
796 | 31841.773 | 10.3685 | 4.9398 |  |  | 796 | 31841.773 | 10.3685 | 4.9739 |  |  | 796 | 31841.773 | 10.3685 | 4.9738
797 | 31881.776 | 10.3698 | 4.9397 |  |  | 797 | 31881.776 | 10.3698 | 4.9738 |  |  | 797 | 31881.776 | 10.3698 | 4.9737
798 | 31921.778 | 10.371 | 4.9395 |  |  | 798 | 31921.778 | 10.371 | 4.9738 |  |  | 798 | 31921.778 | 10.371 | 4.9736
799 | 31961.78 | 10.3723 | 4.9394 |  |  | 799 | 31961.78 | 10.3723 | 4.9737 |  |  | 799 | 31961.78 | 10.3723 | 4.9735
800 | 32001.782 | 10.3735 | 4.9393 |  |  | 800 | 32001.782 | 10.3735 | 4.9736 |  |  | 800 | 32001.782 | 10.3735 | 4.9735
801 | 32041.785 | 10.3748 | 4.9391 |  |  | 801 | 32041.785 | 10.3748 | 4.9736 |  |  | 801 | 32041.785 | 10.3748 | 4.9734
802 | 32081.787 | 10.376 | 4.939 |  |  | 802 | 32081.787 | 10.376 | 4.9735 |  |  | 802 | 32081.787 | 10.376 | 4.9733
803 | 32121.789 | 10.3773 | 4.9389 |  |  | 803 | 32121.789 | 10.3773 | 4.9735 |  |  | 803 | 32121.789 | 10.3773 | 4.9732
804 | 32161.791 | 10.3785 | 4.9387 |  |  | 804 | 32161.791 | 10.3785 | 4.9734 |  |  | 804 | 32161.791 | 10.3785 | 4.9732
805 | 32201.793 | 10.3798 | 4.9386 |  |  | 805 | 32201.793 | 10.3798 | 4.9733 |  |  | 805 | 32201.793 | 10.3798 | 4.9731
806 | 32241.796 | 10.381 | 4.9385 |  |  | 806 | 32241.796 | 10.381 | 4.9733 |  |  | 806 | 32241.796 | 10.381 | 4.973
807 | 32281.798 | 10.3823 | 4.9383 |  |  | 807 | 32281.798 | 10.3823 | 4.9732 |  |  | 807 | 32281.798 | 10.3823 | 4.9729
808 | 32321.8 | 10.3835 | 4.9382 |  |  | 808 | 32321.8 | 10.3835 | 4.9732 |  |  | 808 | 32321.8 | 10.3835 | 4.9729
809 | 32361.802 | 10.3847 | 4.9381 |  |  | 809 | 32361.802 | 10.3847 | 4.9731 |  |  | 809 | 32361.802 | 10.3847 | 4.9728
810 | 32401.805 | 10.386 | 4.9379 |  |  | 810 | 32401.805 | 10.386 | 4.973 |  |  | 810 | 32401.805 | 10.386 | 4.9727
811 | 32441.807 | 10.3872 | 4.9378 |  |  | 811 | 32441.807 | 10.3872 | 4.973 |  |  | 811 | 32441.807 | 10.3872 | 4.9726
812 | 32481.809 | 10.3884 | 4.9377 |  |  | 812 | 32481.809 | 10.3884 | 4.9729 |  |  | 812 | 32481.809 | 10.3884 | 4.9726
813 | 32521.811 | 10.3897 | 4.9375 |  |  | 813 | 32521.811 | 10.3897 | 4.9728 |  |  | 813 | 32521.811 | 10.3897 | 4.9725
814 | 32561.814 | 10.3909 | 4.9374 |  |  | 814 | 32561.814 | 10.3909 | 4.9728 |  |  | 814 | 32561.814 | 10.3909 | 4.9724
815 | 32601.816 | 10.3921 | 4.9373 |  |  | 815 | 32601.816 | 10.3921 | 4.9727 |  |  | 815 | 32601.816 | 10.3921 | 4.9723
816 | 32641.818 | 10.3933 | 4.9371 |  |  | 816 | 32641.818 | 10.3933 | 4.9727 |  |  | 816 | 32641.818 | 10.3933 | 4.9723
817 | 32681.82 | 10.3946 | 4.937 |  |  | 817 | 32681.82 | 10.3946 | 4.9726 |  |  | 817 | 32681.82 | 10.3946 | 4.9722
818 | 32721.822 | 10.3958 | 4.9369 |  |  | 818 | 32721.822 | 10.3958 | 4.9725 |  |  | 818 | 32721.822 | 10.3958 | 4.9721
819 | 32761.825 | 10.397 | 4.9367 |  |  | 819 | 32761.825 | 10.397 | 4.9725 |  |  | 819 | 32761.825 | 10.397 | 4.972
820 | 32801.827 | 10.3982 | 4.9366 |  |  | 820 | 32801.827 | 10.3982 | 4.9724 |  |  | 820 | 32801.827 | 10.3982 | 4.972
821 | 32841.829 | 10.3995 | 4.9365 |  |  | 821 | 32841.829 | 10.3995 | 4.9724 |  |  | 821 | 32841.829 | 10.3995 | 4.9719
822 | 32881.831 | 10.4007 | 4.9364 |  |  | 822 | 32881.831 | 10.4007 | 4.9723 |  |  | 822 | 32881.831 | 10.4007 | 4.9718
823 | 32921.834 | 10.4019 | 4.9362 |  |  | 823 | 32921.834 | 10.4019 | 4.9722 |  |  | 823 | 32921.834 | 10.4019 | 4.9717
824 | 32961.836 | 10.4031 | 4.9361 |  |  | 824 | 32961.836 | 10.4031 | 4.9722 |  |  | 824 | 32961.836 | 10.4031 | 4.9717
825 | 33001.838 | 10.4043 | 4.936 |  |  | 825 | 33001.838 | 10.4043 | 4.9721 |  |  | 825 | 33001.838 | 10.4043 | 4.9716
826 | 33041.84 | 10.4055 | 4.9358 |  |  | 826 | 33041.84 | 10.4055 | 4.9721 |  |  | 826 | 33041.84 | 10.4055 | 4.9715
827 | 33081.842 | 10.4067 | 4.9357 |  |  | 827 | 33081.842 | 10.4067 | 4.972 |  |  | 827 | 33081.842 | 10.4067 | 4.9714
828 | 33121.845 | 10.4079 | 4.9356 |  |  | 828 | 33121.845 | 10.4079 | 4.9719 |  |  | 828 | 33121.845 | 10.4079 | 4.9714
829 | 33161.847 | 10.4092 | 4.9354 |  |  | 829 | 33161.847 | 10.4092 | 4.9719 |  |  | 829 | 33161.847 | 10.4092 | 4.9713
830 | 33201.849 | 10.4104 | 4.9353 |  |  | 830 | 33201.849 | 10.4104 | 4.9718 |  |  | 830 | 33201.849 | 10.4104 | 4.9712
831 | 33241.851 | 10.4116 | 4.9352 |  |  | 831 | 33241.851 | 10.4116 | 4.9718 |  |  | 831 | 33241.851 | 10.4116 | 4.9711
832 | 33281.854 | 10.4128 | 4.9351 |  |  | 832 | 33281.854 | 10.4128 | 4.9717 |  |  | 832 | 33281.854 | 10.4128 | 4.971
833 | 33321.856 | 10.414 | 4.9349 |  |  | 833 | 33321.856 | 10.414 | 4.9716 |  |  | 833 | 33321.856 | 10.414 | 4.971
834 | 33361.858 | 10.4152 | 4.9348 |  |  | 834 | 33361.858 | 10.4152 | 4.9716 |  |  | 834 | 33361.858 | 10.4152 | 4.9709
835 | 33401.86 | 10.4164 | 4.9347 |  |  | 835 | 33401.86 | 10.4164 | 4.9715 |  |  | 835 | 33401.86 | 10.4164 | 4.9708
836 | 33441.863 | 10.4176 | 4.9345 |  |  | 836 | 33441.863 | 10.4176 | 4.9715 |  |  | 836 | 33441.863 | 10.4176 | 4.9707
837 | 33481.865 | 10.4188 | 4.9344 |  |  | 837 | 33481.865 | 10.4188 | 4.9714 |  |  | 837 | 33481.865 | 10.4188 | 4.9707
838 | 33521.867 | 10.42 | 4.9343 |  |  | 838 | 33521.867 | 10.42 | 4.9713 |  |  | 838 | 33521.867 | 10.42 | 4.9706
839 | 33561.869 | 10.4211 | 4.9341 |  |  | 839 | 33561.869 | 10.4211 | 4.9713 |  |  | 839 | 33561.869 | 10.4211 | 4.9705
840 | 33601.871 | 10.4223 | 4.934 |  |  | 840 | 33601.871 | 10.4223 | 4.9712 |  |  | 840 | 33601.871 | 10.4223 | 4.9704
841 | 33641.874 | 10.4235 | 4.9339 |  |  | 841 | 33641.874 | 10.4235 | 4.9712 |  |  | 841 | 33641.874 | 10.4235 | 4.9704
842 | 33681.876 | 10.4247 | 4.9338 |  |  | 842 | 33681.876 | 10.4247 | 4.9711 |  |  | 842 | 33681.876 | 10.4247 | 4.9703
843 | 33721.878 | 10.4259 | 4.9336 |  |  | 843 | 33721.878 | 10.4259 | 4.971 |  |  | 843 | 33721.878 | 10.4259 | 4.9702
844 | 33761.88 | 10.4271 | 4.9335 |  |  | 844 | 33761.88 | 10.4271 | 4.971 |  |  | 844 | 33761.88 | 10.4271 | 4.9701
845 | 33801.883 | 10.4283 | 4.9334 |  |  | 845 | 33801.883 | 10.4283 | 4.9709 |  |  | 845 | 33801.883 | 10.4283 | 4.9701
846 | 33841.885 | 10.4295 | 4.9333 |  |  | 846 | 33841.885 | 10.4295 | 4.9709 |  |  | 846 | 33841.885 | 10.4295 | 4.97
847 | 33881.887 | 10.4306 | 4.9331 |  |  | 847 | 33881.887 | 10.4306 | 4.9708 |  |  | 847 | 33881.887 | 10.4306 | 4.9699
848 | 33921.889 | 10.4318 | 4.933 |  |  | 848 | 33921.889 | 10.4318 | 4.9707 |  |  | 848 | 33921.889 | 10.4318 | 4.9698
849 | 33961.891 | 10.433 | 4.9329 |  |  | 849 | 33961.891 | 10.433 | 4.9707 |  |  | 849 | 33961.891 | 10.433 | 4.9698
850 | 34001.894 | 10.4342 | 4.9327 |  |  | 850 | 34001.894 | 10.4342 | 4.9706 |  |  | 850 | 34001.894 | 10.4342 | 4.9697
851 | 34041.896 | 10.4353 | 4.9326 |  |  | 851 | 34041.896 | 10.4353 | 4.9705 |  |  | 851 | 34041.896 | 10.4353 | 4.9696
852 | 34081.898 | 10.4365 | 4.9325 |  |  | 852 | 34081.898 | 10.4365 | 4.9705 |  |  | 852 | 34081.898 | 10.4365 | 4.9695
853 | 34121.9 | 10.4377 | 4.9324 |  |  | 853 | 34121.9 | 10.4377 | 4.9704 |  |  | 853 | 34121.9 | 10.4377 | 4.9695
854 | 34161.903 | 10.4389 | 4.9322 |  |  | 854 | 34161.903 | 10.4389 | 4.9704 |  |  | 854 | 34161.903 | 10.4389 | 4.9694
855 | 34201.905 | 10.44 | 4.9321 |  |  | 855 | 34201.905 | 10.44 | 4.9703 |  |  | 855 | 34201.905 | 10.44 | 4.9693
856 | 34241.907 | 10.4412 | 4.932 |  |  | 856 | 34241.907 | 10.4412 | 4.9702 |  |  | 856 | 34241.907 | 10.4412 | 4.9692
857 | 34281.909 | 10.4424 | 4.9319 |  |  | 857 | 34281.909 | 10.4424 | 4.9702 |  |  | 857 | 34281.909 | 10.4424 | 4.9691
858 | 34321.912 | 10.4435 | 4.9317 |  |  | 858 | 34321.912 | 10.4435 | 4.9701 |  |  | 858 | 34321.912 | 10.4435 | 4.9691
859 | 34361.914 | 10.4447 | 4.9316 |  |  | 859 | 34361.914 | 10.4447 | 4.9701 |  |  | 859 | 34361.914 | 10.4447 | 4.969
860 | 34401.916 | 10.4459 | 4.9315 |  |  | 860 | 34401.916 | 10.4459 | 4.97 |  |  | 860 | 34401.916 | 10.4459 | 4.9689
861 | 34441.918 | 10.447 | 4.9313 |  |  | 861 | 34441.918 | 10.447 | 4.9699 |  |  | 861 | 34441.918 | 10.447 | 4.9688
862 | 34481.92 | 10.4482 | 4.9312 |  |  | 862 | 34481.92 | 10.4482 | 4.9699 |  |  | 862 | 34481.92 | 10.4482 | 4.9688
863 | 34521.923 | 10.4493 | 4.9311 |  |  | 863 | 34521.923 | 10.4493 | 4.9698 |  |  | 863 | 34521.923 | 10.4493 | 4.9687
864 | 34561.925 | 10.4505 | 4.931 |  |  | 864 | 34561.925 | 10.4505 | 4.9698 |  |  | 864 | 34561.925 | 10.4505 | 4.9686
865 | 34601.927 | 10.4517 | 4.9308 |  |  | 865 | 34601.927 | 10.4517 | 4.9697 |  |  | 865 | 34601.927 | 10.4517 | 4.9685
866 | 34641.929 | 10.4528 | 4.9307 |  |  | 866 | 34641.929 | 10.4528 | 4.9696 |  |  | 866 | 34641.929 | 10.4528 | 4.9685
867 | 34681.932 | 10.454 | 4.9306 |  |  | 867 | 34681.932 | 10.454 | 4.9696 |  |  | 867 | 34681.932 | 10.454 | 4.9684
868 | 34721.934 | 10.4551 | 4.9305 |  |  | 868 | 34721.934 | 10.4551 | 4.9695 |  |  | 868 | 34721.934 | 10.4551 | 4.9683
869 | 34761.936 | 10.4563 | 4.9303 |  |  | 869 | 34761.936 | 10.4563 | 4.9694 |  |  | 869 | 34761.936 | 10.4563 | 4.9682
870 | 34801.938 | 10.4574 | 4.9302 |  |  | 870 | 34801.938 | 10.4574 | 4.9694 |  |  | 870 | 34801.938 | 10.4574 | 4.9682
871 | 34841.941 | 10.4586 | 4.9301 |  |  | 871 | 34841.941 | 10.4586 | 4.9693 |  |  | 871 | 34841.941 | 10.4586 | 4.9681
872 | 34881.943 | 10.4597 | 4.93 |  |  | 872 | 34881.943 | 10.4597 | 4.9693 |  |  | 872 | 34881.943 | 10.4597 | 4.968
873 | 34921.945 | 10.4609 | 4.9298 |  |  | 873 | 34921.945 | 10.4609 | 4.9692 |  |  | 873 | 34921.945 | 10.4609 | 4.9679
874 | 34961.947 | 10.462 | 4.9297 |  |  | 874 | 34961.947 | 10.462 | 4.9691 |  |  | 874 | 34961.947 | 10.462 | 4.9678
875 | 35001.949 | 10.4632 | 4.9296 |  |  | 875 | 35001.949 | 10.4632 | 4.9691 |  |  | 875 | 35001.949 | 10.4632 | 4.9678
876 | 35041.952 | 10.4643 | 4.9295 |  |  | 876 | 35041.952 | 10.4643 | 4.969 |  |  | 876 | 35041.952 | 10.4643 | 4.9677
877 | 35081.954 | 10.4654 | 4.9293 |  |  | 877 | 35081.954 | 10.4654 | 4.969 |  |  | 877 | 35081.954 | 10.4654 | 4.9676
878 | 35121.956 | 10.4666 | 4.9292 |  |  | 878 | 35121.956 | 10.4666 | 4.9689 |  |  | 878 | 35121.956 | 10.4666 | 4.9675
879 | 35161.958 | 10.4677 | 4.9291 |  |  | 879 | 35161.958 | 10.4677 | 4.9688 |  |  | 879 | 35161.958 | 10.4677 | 4.9675
880 | 35201.961 | 10.4689 | 4.929 |  |  | 880 | 35201.961 | 10.4689 | 4.9688 |  |  | 880 | 35201.961 | 10.4689 | 4.9674
881 | 35241.963 | 10.47 | 4.9288 |  |  | 881 | 35241.963 | 10.47 | 4.9687 |  |  | 881 | 35241.963 | 10.47 | 4.9673
882 | 35281.965 | 10.4711 | 4.9287 |  |  | 882 | 35281.965 | 10.4711 | 4.9686 |  |  | 882 | 35281.965 | 10.4711 | 4.9672
883 | 35321.967 | 10.4723 | 4.9286 |  |  | 883 | 35321.967 | 10.4723 | 4.9686 |  |  | 883 | 35321.967 | 10.4723 | 4.9671
884 | 35361.969 | 10.4734 | 4.9285 |  |  | 884 | 35361.969 | 10.4734 | 4.9685 |  |  | 884 | 35361.969 | 10.4734 | 4.9671
885 | 35401.972 | 10.4745 | 4.9283 |  |  | 885 | 35401.972 | 10.4745 | 4.9685 |  |  | 885 | 35401.972 | 10.4745 | 4.967
886 | 35441.974 | 10.4757 | 4.9282 |  |  | 886 | 35441.974 | 10.4757 | 4.9684 |  |  | 886 | 35441.974 | 10.4757 | 4.9669
887 | 35481.976 | 10.4768 | 4.9281 |  |  | 887 | 35481.976 | 10.4768 | 4.9683 |  |  | 887 | 35481.976 | 10.4768 | 4.9668
888 | 35521.978 | 10.4779 | 4.928 |  |  | 888 | 35521.978 | 10.4779 | 4.9683 |  |  | 888 | 35521.978 | 10.4779 | 4.9668
889 | 35561.981 | 10.479 | 4.9279 |  |  | 889 | 35561.981 | 10.479 | 4.9682 |  |  | 889 | 35561.981 | 10.479 | 4.9667
890 | 35601.983 | 10.4802 | 4.9277 |  |  | 890 | 35601.983 | 10.4802 | 4.9681 |  |  | 890 | 35601.983 | 10.4802 | 4.9666
891 | 35641.985 | 10.4813 | 4.9276 |  |  | 891 | 35641.985 | 10.4813 | 4.9681 |  |  | 891 | 35641.985 | 10.4813 | 4.9665
892 | 35681.987 | 10.4824 | 4.9275 |  |  | 892 | 35681.987 | 10.4824 | 4.968 |  |  | 892 | 35681.987 | 10.4824 | 4.9664
893 | 35721.99 | 10.4835 | 4.9274 |  |  | 893 | 35721.99 | 10.4835 | 4.968 |  |  | 893 | 35721.99 | 10.4835 | 4.9664
894 | 35761.992 | 10.4846 | 4.9272 |  |  | 894 | 35761.992 | 10.4846 | 4.9679 |  |  | 894 | 35761.992 | 10.4846 | 4.9663
895 | 35801.994 | 10.4858 | 4.9271 |  |  | 895 | 35801.994 | 10.4858 | 4.9678 |  |  | 895 | 35801.994 | 10.4858 | 4.9662
896 | 35841.996 | 10.4869 | 4.927 |  |  | 896 | 35841.996 | 10.4869 | 4.9678 |  |  | 896 | 35841.996 | 10.4869 | 4.9661
897 | 35881.998 | 10.488 | 4.9269 |  |  | 897 | 35881.998 | 10.488 | 4.9677 |  |  | 897 | 35881.998 | 10.488 | 4.9661
898 | 35922.001 | 10.4891 | 4.9267 |  |  | 898 | 35922.001 | 10.4891 | 4.9677 |  |  | 898 | 35922.001 | 10.4891 | 4.966
899 | 35962.003 | 10.4902 | 4.9266 |  |  | 899 | 35962.003 | 10.4902 | 4.9676 |  |  | 899 | 35962.003 | 10.4902 | 4.9659
900 | 36002.005 | 10.4913 | 4.9265 |  |  | 900 | 36002.005 | 10.4913 | 4.9675 |  |  | 900 | 36002.005 | 10.4913 | 4.9658
901 | 36042.007 | 10.4924 | 4.9264 |  |  | 901 | 36042.007 | 10.4924 | 4.9675 |  |  | 901 | 36042.007 | 10.4924 | 4.9657
902 | 36082.01 | 10.4935 | 4.9263 |  |  | 902 | 36082.01 | 10.4935 | 4.9674 |  |  | 902 | 36082.01 | 10.4935 | 4.9657
903 | 36122.012 | 10.4947 | 4.9261 |  |  | 903 | 36122.012 | 10.4947 | 4.9673 |  |  | 903 | 36122.012 | 10.4947 | 4.9656
904 | 36162.014 | 10.4958 | 4.926 |  |  | 904 | 36162.014 | 10.4958 | 4.9673 |  |  | 904 | 36162.014 | 10.4958 | 4.9655
905 | 36202.016 | 10.4969 | 4.9259 |  |  | 905 | 36202.016 | 10.4969 | 4.9672 |  |  | 905 | 36202.016 | 10.4969 | 4.9654
906 | 36242.018 | 10.498 | 4.9258 |  |  | 906 | 36242.018 | 10.498 | 4.9672 |  |  | 906 | 36242.018 | 10.498 | 4.9653
907 | 36282.021 | 10.4991 | 4.9257 |  |  | 907 | 36282.021 | 10.4991 | 4.9671 |  |  | 907 | 36282.021 | 10.4991 | 4.9653
908 | 36322.023 | 10.5002 | 4.9255 |  |  | 908 | 36322.023 | 10.5002 | 4.967 |  |  | 908 | 36322.023 | 10.5002 | 4.9652
909 | 36362.025 | 10.5013 | 4.9254 |  |  | 909 | 36362.025 | 10.5013 | 4.967 |  |  | 909 | 36362.025 | 10.5013 | 4.9651
910 | 36402.027 | 10.5024 | 4.9253 |  |  | 910 | 36402.027 | 10.5024 | 4.9669 |  |  | 910 | 36402.027 | 10.5024 | 4.965
911 | 36442.03 | 10.5035 | 4.9252 |  |  | 911 | 36442.03 | 10.5035 | 4.9668 |  |  | 911 | 36442.03 | 10.5035 | 4.965
912 | 36482.032 | 10.5046 | 4.925 |  |  | 912 | 36482.032 | 10.5046 | 4.9668 |  |  | 912 | 36482.032 | 10.5046 | 4.9649
913 | 36522.034 | 10.5057 | 4.9249 |  |  | 913 | 36522.034 | 10.5057 | 4.9667 |  |  | 913 | 36522.034 | 10.5057 | 4.9648
914 | 36562.036 | 10.5068 | 4.9248 |  |  | 914 | 36562.036 | 10.5068 | 4.9666 |  |  | 914 | 36562.036 | 10.5068 | 4.9647
915 | 36602.039 | 10.5079 | 4.9247 |  |  | 915 | 36602.039 | 10.5079 | 4.9666 |  |  | 915 | 36602.039 | 10.5079 | 4.9646
916 | 36642.041 | 10.509 | 4.9246 |  |  | 916 | 36642.041 | 10.509 | 4.9665 |  |  | 916 | 36642.041 | 10.509 | 4.9646
917 | 36682.043 | 10.51 | 4.9244 |  |  | 917 | 36682.043 | 10.51 | 4.9665 |  |  | 917 | 36682.043 | 10.51 | 4.9645
918 | 36722.045 | 10.5111 | 4.9243 |  |  | 918 | 36722.045 | 10.5111 | 4.9664 |  |  | 918 | 36722.045 | 10.5111 | 4.9644
919 | 36762.047 | 10.5122 | 4.9242 |  |  | 919 | 36762.047 | 10.5122 | 4.9663 |  |  | 919 | 36762.047 | 10.5122 | 4.9643
920 | 36802.05 | 10.5133 | 4.9241 |  |  | 920 | 36802.05 | 10.5133 | 4.9663 |  |  | 920 | 36802.05 | 10.5133 | 4.9642
921 | 36842.052 | 10.5144 | 4.924 |  |  | 921 | 36842.052 | 10.5144 | 4.9662 |  |  | 921 | 36842.052 | 10.5144 | 4.9642
922 | 36882.054 | 10.5155 | 4.9238 |  |  | 922 | 36882.054 | 10.5155 | 4.9661 |  |  | 922 | 36882.054 | 10.5155 | 4.9641
923 | 36922.056 | 10.5166 | 4.9237 |  |  | 923 | 36922.056 | 10.5166 | 4.9661 |  |  | 923 | 36922.056 | 10.5166 | 4.964
924 | 36962.059 | 10.5176 | 4.9236 |  |  | 924 | 36962.059 | 10.5176 | 4.966 |  |  | 924 | 36962.059 | 10.5176 | 4.9639
925 | 37002.061 | 10.5187 | 4.9235 |  |  | 925 | 37002.061 | 10.5187 | 4.966 |  |  | 925 | 37002.061 | 10.5187 | 4.9638
926 | 37042.063 | 10.5198 | 4.9234 |  |  | 926 | 37042.063 | 10.5198 | 4.9659 |  |  | 926 | 37042.063 | 10.5198 | 4.9638
927 | 37082.065 | 10.5209 | 4.9232 |  |  | 927 | 37082.065 | 10.5209 | 4.9658 |  |  | 927 | 37082.065 | 10.5209 | 4.9637
928 | 37122.067 | 10.522 | 4.9231 |  |  | 928 | 37122.067 | 10.522 | 4.9658 |  |  | 928 | 37122.067 | 10.522 | 4.9636
929 | 37162.07 | 10.523 | 4.923 |  |  | 929 | 37162.07 | 10.523 | 4.9657 |  |  | 929 | 37162.07 | 10.523 | 4.9635
930 | 37202.072 | 10.5241 | 4.9229 |  |  | 930 | 37202.072 | 10.5241 | 4.9656 |  |  | 930 | 37202.072 | 10.5241 | 4.9634
931 | 37242.074 | 10.5252 | 4.9228 |  |  | 931 | 37242.074 | 10.5252 | 4.9656 |  |  | 931 | 37242.074 | 10.5252 | 4.9634
932 | 37282.076 | 10.5263 | 4.9226 |  |  | 932 | 37282.076 | 10.5263 | 4.9655 |  |  | 932 | 37282.076 | 10.5263 | 4.9633
933 | 37322.079 | 10.5273 | 4.9225 |  |  | 933 | 37322.079 | 10.5273 | 4.9654 |  |  | 933 | 37322.079 | 10.5273 | 4.9632
934 | 37362.081 | 10.5284 | 4.9224 |  |  | 934 | 37362.081 | 10.5284 | 4.9654 |  |  | 934 | 37362.081 | 10.5284 | 4.9631
935 | 37402.083 | 10.5295 | 4.9223 |  |  | 935 | 37402.083 | 10.5295 | 4.9653 |  |  | 935 | 37402.083 | 10.5295 | 4.963
936 | 37442.085 | 10.5306 | 4.9222 |  |  | 936 | 37442.085 | 10.5306 | 4.9652 |  |  | 936 | 37442.085 | 10.5306 | 4.963
937 | 37482.088 | 10.5316 | 4.9221 |  |  | 937 | 37482.088 | 10.5316 | 4.9652 |  |  | 937 | 37482.088 | 10.5316 | 4.9629
938 | 37522.09 | 10.5327 | 4.9219 |  |  | 938 | 37522.09 | 10.5327 | 4.9651 |  |  | 938 | 37522.09 | 10.5327 | 4.9628
939 | 37562.092 | 10.5338 | 4.9218 |  |  | 939 | 37562.092 | 10.5338 | 4.9651 |  |  | 939 | 37562.092 | 10.5338 | 4.9627
940 | 37602.094 | 10.5348 | 4.9217 |  |  | 940 | 37602.094 | 10.5348 | 4.965 |  |  | 940 | 37602.094 | 10.5348 | 4.9626
941 | 37642.096 | 10.5359 | 4.9216 |  |  | 941 | 37642.096 | 10.5359 | 4.9649 |  |  | 941 | 37642.096 | 10.5359 | 4.9626
942 | 37682.099 | 10.5369 | 4.9215 |  |  | 942 | 37682.099 | 10.5369 | 4.9649 |  |  | 942 | 37682.099 | 10.5369 | 4.9625
943 | 37722.101 | 10.538 | 4.9213 |  |  | 943 | 37722.101 | 10.538 | 4.9648 |  |  | 943 | 37722.101 | 10.538 | 4.9624
944 | 37762.103 | 10.5391 | 4.9212 |  |  | 944 | 37762.103 | 10.5391 | 4.9647 |  |  | 944 | 37762.103 | 10.5391 | 4.9623
945 | 37802.105 | 10.5401 | 4.9211 |  |  | 945 | 37802.105 | 10.5401 | 4.9647 |  |  | 945 | 37802.105 | 10.5401 | 4.9622
946 | 37842.108 | 10.5412 | 4.921 |  |  | 946 | 37842.108 | 10.5412 | 4.9646 |  |  | 946 | 37842.108 | 10.5412 | 4.9622
947 | 37882.11 | 10.5422 | 4.9209 |  |  | 947 | 37882.11 | 10.5422 | 4.9645 |  |  | 947 | 37882.11 | 10.5422 | 4.9621
948 | 37922.112 | 10.5433 | 4.9208 |  |  | 948 | 37922.112 | 10.5433 | 4.9645 |  |  | 948 | 37922.112 | 10.5433 | 4.962
949 | 37962.114 | 10.5443 | 4.9206 |  |  | 949 | 37962.114 | 10.5443 | 4.9644 |  |  | 949 | 37962.114 | 10.5443 | 4.9619
950 | 38002.117 | 10.5454 | 4.9205 |  |  | 950 | 38002.117 | 10.5454 | 4.9643 |  |  | 950 | 38002.117 | 10.5454 | 4.9618
951 | 38042.119 | 10.5464 | 4.9204 |  |  | 951 | 38042.119 | 10.5464 | 4.9643 |  |  | 951 | 38042.119 | 10.5464 | 4.9618
952 | 38082.121 | 10.5475 | 4.9203 |  |  | 952 | 38082.121 | 10.5475 | 4.9642 |  |  | 952 | 38082.121 | 10.5475 | 4.9617
953 | 38122.123 | 10.5486 | 4.9202 |  |  | 953 | 38122.123 | 10.5486 | 4.9642 |  |  | 953 | 38122.123 | 10.5486 | 4.9616
954 | 38162.125 | 10.5496 | 4.92 |  |  | 954 | 38162.125 | 10.5496 | 4.9641 |  |  | 954 | 38162.125 | 10.5496 | 4.9615
955 | 38202.128 | 10.5506 | 4.9199 |  |  | 955 | 38202.128 | 10.5506 | 4.964 |  |  | 955 | 38202.128 | 10.5506 | 4.9614
956 | 38242.13 | 10.5517 | 4.9198 |  |  | 956 | 38242.13 | 10.5517 | 4.964 |  |  | 956 | 38242.13 | 10.5517 | 4.9613
957 | 38282.132 | 10.5527 | 4.9197 |  |  | 957 | 38282.132 | 10.5527 | 4.9639 |  |  | 957 | 38282.132 | 10.5527 | 4.9613
958 | 38322.134 | 10.5538 | 4.9196 |  |  | 958 | 38322.134 | 10.5538 | 4.9638 |  |  | 958 | 38322.134 | 10.5538 | 4.9612
959 | 38362.137 | 10.5548 | 4.9195 |  |  | 959 | 38362.137 | 10.5548 | 4.9638 |  |  | 959 | 38362.137 | 10.5548 | 4.9611
960 | 38402.139 | 10.5559 | 4.9193 |  |  | 960 | 38402.139 | 10.5559 | 4.9637 |  |  | 960 | 38402.139 | 10.5559 | 4.961
961 | 38442.141 | 10.5569 | 4.9192 |  |  | 961 | 38442.141 | 10.5569 | 4.9636 |  |  | 961 | 38442.141 | 10.5569 | 4.9609
962 | 38482.143 | 10.5579 | 4.9191 |  |  | 962 | 38482.143 | 10.5579 | 4.9636 |  |  | 962 | 38482.143 | 10.5579 | 4.9609
963 | 38522.145 | 10.559 | 4.919 |  |  | 963 | 38522.145 | 10.559 | 4.9635 |  |  | 963 | 38522.145 | 10.559 | 4.9608
964 | 38562.148 | 10.56 | 4.9189 |  |  | 964 | 38562.148 | 10.56 | 4.9634 |  |  | 964 | 38562.148 | 10.56 | 4.9607
965 | 38602.15 | 10.5611 | 4.9188 |  |  | 965 | 38602.15 | 10.5611 | 4.9634 |  |  | 965 | 38602.15 | 10.5611 | 4.9606
966 | 38642.152 | 10.5621 | 4.9187 |  |  | 966 | 38642.152 | 10.5621 | 4.9633 |  |  | 966 | 38642.152 | 10.5621 | 4.9605
967 | 38682.154 | 10.5631 | 4.9185 |  |  | 967 | 38682.154 | 10.5631 | 4.9632 |  |  | 967 | 38682.154 | 10.5631 | 4.9604
968 | 38722.157 | 10.5642 | 4.9184 |  |  | 968 | 38722.157 | 10.5642 | 4.9632 |  |  | 968 | 38722.157 | 10.5642 | 4.9604
969 | 38762.159 | 10.5652 | 4.9183 |  |  | 969 | 38762.159 | 10.5652 | 4.9631 |  |  | 969 | 38762.159 | 10.5652 | 4.9603
970 | 38802.161 | 10.5662 | 4.9182 |  |  | 970 | 38802.161 | 10.5662 | 4.963 |  |  | 970 | 38802.161 | 10.5662 | 4.9602
971 | 38842.163 | 10.5673 | 4.9181 |  |  | 971 | 38842.163 | 10.5673 | 4.963 |  |  | 971 | 38842.163 | 10.5673 | 4.9601
972 | 38882.166 | 10.5683 | 4.918 |  |  | 972 | 38882.166 | 10.5683 | 4.9629 |  |  | 972 | 38882.166 | 10.5683 | 4.96
973 | 38922.168 | 10.5693 | 4.9178 |  |  | 973 | 38922.168 | 10.5693 | 4.9628 |  |  | 973 | 38922.168 | 10.5693 | 4.96
974 | 38962.17 | 10.5703 | 4.9177 |  |  | 974 | 38962.17 | 10.5703 | 4.9628 |  |  | 974 | 38962.17 | 10.5703 | 4.9599
975 | 39002.172 | 10.5714 | 4.9176 |  |  | 975 | 39002.172 | 10.5714 | 4.9627 |  |  | 975 | 39002.172 | 10.5714 | 4.9598
976 | 39042.174 | 10.5724 | 4.9175 |  |  | 976 | 39042.174 | 10.5724 | 4.9626 |  |  | 976 | 39042.174 | 10.5724 | 4.9597
977 | 39082.177 | 10.5734 | 4.9174 |  |  | 977 | 39082.177 | 10.5734 | 4.9626 |  |  | 977 | 39082.177 | 10.5734 | 4.9596
978 | 39122.179 | 10.5744 | 4.9173 |  |  | 978 | 39122.179 | 10.5744 | 4.9625 |  |  | 978 | 39122.179 | 10.5744 | 4.9595
979 | 39162.181 | 10.5755 | 4.9171 |  |  | 979 | 39162.181 | 10.5755 | 4.9624 |  |  | 979 | 39162.181 | 10.5755 | 4.9595
980 | 39202.183 | 10.5765 | 4.917 |  |  | 980 | 39202.183 | 10.5765 | 4.9624 |  |  | 980 | 39202.183 | 10.5765 | 4.9594
981 | 39242.186 | 10.5775 | 4.9169 |  |  | 981 | 39242.186 | 10.5775 | 4.9623 |  |  | 981 | 39242.186 | 10.5775 | 4.9593
982 | 39282.188 | 10.5785 | 4.9168 |  |  | 982 | 39282.188 | 10.5785 | 4.9623 |  |  | 982 | 39282.188 | 10.5785 | 4.9592
983 | 39322.19 | 10.5795 | 4.9167 |  |  | 983 | 39322.19 | 10.5795 | 4.9622 |  |  | 983 | 39322.19 | 10.5795 | 4.9591
984 | 39362.192 | 10.5806 | 4.9166 |  |  | 984 | 39362.192 | 10.5806 | 4.9621 |  |  | 984 | 39362.192 | 10.5806 | 4.959
985 | 39402.194 | 10.5816 | 4.9165 |  |  | 985 | 39402.194 | 10.5816 | 4.9621 |  |  | 985 | 39402.194 | 10.5816 | 4.959
986 | 39442.197 | 10.5826 | 4.9163 |  |  | 986 | 39442.197 | 10.5826 | 4.962 |  |  | 986 | 39442.197 | 10.5826 | 4.9589
987 | 39482.199 | 10.5836 | 4.9162 |  |  | 987 | 39482.199 | 10.5836 | 4.9619 |  |  | 987 | 39482.199 | 10.5836 | 4.9588
988 | 39522.201 | 10.5846 | 4.9161 |  |  | 988 | 39522.201 | 10.5846 | 4.9619 |  |  | 988 | 39522.201 | 10.5846 | 4.9587
989 | 39562.203 | 10.5856 | 4.916 |  |  | 989 | 39562.203 | 10.5856 | 4.9618 |  |  | 989 | 39562.203 | 10.5856 | 4.9586
990 | 39602.206 | 10.5866 | 4.9159 |  |  | 990 | 39602.206 | 10.5866 | 4.9617 |  |  | 990 | 39602.206 | 10.5866 | 4.9585

## Run5
MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR
File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run5\1 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run5\2 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run5\3 01 001.RAW
Cursor Location: |  |  | (178, 193) |  |  | Cursor Location: |  |  | (179, 193) |  |  | Cursor Location: |  |  | (179, 193)
Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15)
Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev
1 | 39.997 | 3.6888 | 7.7155 | 5.1123 | 0.32386 | 1 | 39.997 | 3.6888 | 7.717 | 5.1134 | 0.32548 | 1 | 39.997 | 3.6888 | 7.7211 | 5.116 | 0.32516
2 | 79.995 | 4.382 | 7.3853 |  |  | 2 | 79.995 | 4.382 | 7.3883 |  |  | 2 | 79.995 | 4.382 | 7.3907
3 | 119.992 | 4.7874 | 7.1939 |  |  | 3 | 119.992 | 4.7874 | 7.1972 |  |  | 3 | 119.992 | 4.7874 | 7.1993
4 | 159.989 | 5.0751 | 7.0558 |  |  | 4 | 159.989 | 5.0751 | 7.059 |  |  | 4 | 159.989 | 5.0751 | 7.0612
5 | 199.986 | 5.2982 | 6.9469 |  |  | 5 | 199.986 | 5.2982 | 6.95 |  |  | 5 | 199.986 | 5.2982 | 6.9523
6 | 239.984 | 5.4806 | 6.8567 |  |  | 6 | 239.984 | 5.4806 | 6.8598 |  |  | 6 | 239.984 | 5.4806 | 6.8622
7 | 279.981 | 5.6347 | 6.7797 |  |  | 7 | 279.981 | 5.6347 | 6.7828 |  |  | 7 | 279.981 | 5.6347 | 6.7853
8 | 319.978 | 5.7683 | 6.7126 |  |  | 8 | 319.978 | 5.7683 | 6.7156 |  |  | 8 | 319.978 | 5.7683 | 6.7181
9 | 359.975 | 5.886 | 6.653 |  |  | 9 | 359.975 | 5.886 | 6.656 |  |  | 9 | 359.975 | 5.886 | 6.6586
10 | 399.973 | 5.9914 | 6.5994 |  |  | 10 | 399.973 | 5.9914 | 6.6025 |  |  | 10 | 399.973 | 5.9914 | 6.6052
11 | 439.97 | 6.0867 | 6.5509 |  |  | 11 | 439.97 | 6.0867 | 6.554 |  |  | 11 | 439.97 | 6.0867 | 6.5566
12 | 479.967 | 6.1737 | 6.5064 |  |  | 12 | 479.967 | 6.1737 | 6.5096 |  |  | 12 | 479.967 | 6.1737 | 6.5122
13 | 519.964 | 6.2538 | 6.4655 |  |  | 13 | 519.964 | 6.2538 | 6.4687 |  |  | 13 | 519.964 | 6.2538 | 6.4713
14 | 559.962 | 6.3279 | 6.4275 |  |  | 14 | 559.962 | 6.3279 | 6.4308 |  |  | 14 | 559.962 | 6.3279 | 6.4334
15 | 599.959 | 6.3969 | 6.3921 |  |  | 15 | 599.959 | 6.3969 | 6.3954 |  |  | 15 | 599.959 | 6.3969 | 6.3981
16 | 639.956 | 6.4614 | 6.3589 |  |  | 16 | 639.956 | 6.4614 | 6.3624 |  |  | 16 | 639.956 | 6.4614 | 6.365
17 | 679.953 | 6.522 | 6.3278 |  |  | 17 | 679.953 | 6.522 | 6.3313 |  |  | 17 | 679.953 | 6.522 | 6.3339
18 | 719.951 | 6.5792 | 6.2984 |  |  | 18 | 719.951 | 6.5792 | 6.302 |  |  | 18 | 719.951 | 6.5792 | 6.3045
19 | 759.948 | 6.6332 | 6.2707 |  |  | 19 | 759.948 | 6.6332 | 6.2743 |  |  | 19 | 759.948 | 6.6332 | 6.2768
20 | 799.945 | 6.6845 | 6.2443 |  |  | 20 | 799.945 | 6.6845 | 6.248 |  |  | 20 | 799.945 | 6.6845 | 6.2505
21 | 839.942 | 6.7333 | 6.2192 |  |  | 21 | 839.942 | 6.7333 | 6.223 |  |  | 21 | 839.942 | 6.7333 | 6.2254
22 | 879.94 | 6.7799 | 6.1953 |  |  | 22 | 879.94 | 6.7799 | 6.1991 |  |  | 22 | 879.94 | 6.7799 | 6.2016
23 | 919.937 | 6.8243 | 6.1725 |  |  | 23 | 919.937 | 6.8243 | 6.1764 |  |  | 23 | 919.937 | 6.8243 | 6.1788
24 | 959.934 | 6.8669 | 6.1506 |  |  | 24 | 959.934 | 6.8669 | 6.1546 |  |  | 24 | 959.934 | 6.8669 | 6.1569
25 | 999.931 | 6.9077 | 6.1297 |  |  | 25 | 999.931 | 6.9077 | 6.1337 |  |  | 25 | 999.931 | 6.9077 | 6.136
26 | 1039.929 | 6.9469 | 6.1095 |  |  | 26 | 1039.929 | 6.9469 | 6.1136 |  |  | 26 | 1039.929 | 6.9469 | 6.1159
27 | 1079.926 | 6.9846 | 6.0902 |  |  | 27 | 1079.926 | 6.9846 | 6.0943 |  |  | 27 | 1079.926 | 6.9846 | 6.0966
28 | 1119.923 | 7.021 | 6.0715 |  |  | 28 | 1119.923 | 7.021 | 6.0757 |  |  | 28 | 1119.923 | 7.021 | 6.078
29 | 1159.92 | 7.0561 | 6.0535 |  |  | 29 | 1159.92 | 7.0561 | 6.0578 |  |  | 29 | 1159.92 | 7.0561 | 6.06
30 | 1199.918 | 7.09 | 6.0362 |  |  | 30 | 1199.918 | 7.09 | 6.0405 |  |  | 30 | 1199.918 | 7.09 | 6.0426
31 | 1239.915 | 7.1228 | 6.0194 |  |  | 31 | 1239.915 | 7.1228 | 6.0237 |  |  | 31 | 1239.915 | 7.1228 | 6.0259
32 | 1279.912 | 7.1545 | 6.0031 |  |  | 32 | 1279.912 | 7.1545 | 6.0075 |  |  | 32 | 1279.912 | 7.1545 | 6.0097
33 | 1319.909 | 7.1853 | 5.9874 |  |  | 33 | 1319.909 | 7.1853 | 5.9918 |  |  | 33 | 1319.909 | 7.1853 | 5.9939
34 | 1359.907 | 7.2152 | 5.9721 |  |  | 34 | 1359.907 | 7.2152 | 5.9766 |  |  | 34 | 1359.907 | 7.2152 | 5.9787
35 | 1399.904 | 7.2442 | 5.9573 |  |  | 35 | 1399.904 | 7.2442 | 5.9619 |  |  | 35 | 1399.904 | 7.2442 | 5.9639
36 | 1439.901 | 7.2723 | 5.943 |  |  | 36 | 1439.901 | 7.2723 | 5.9475 |  |  | 36 | 1439.901 | 7.2723 | 5.9496
37 | 1479.898 | 7.2997 | 5.929 |  |  | 37 | 1479.898 | 7.2997 | 5.9336 |  |  | 37 | 1479.898 | 7.2997 | 5.9356
38 | 1519.896 | 7.3264 | 5.9154 |  |  | 38 | 1519.896 | 7.3264 | 5.92 |  |  | 38 | 1519.896 | 7.3264 | 5.922
39 | 1559.893 | 7.3524 | 5.9022 |  |  | 39 | 1559.893 | 7.3524 | 5.9069 |  |  | 39 | 1559.893 | 7.3524 | 5.9088
40 | 1599.89 | 7.3777 | 5.8893 |  |  | 40 | 1599.89 | 7.3777 | 5.894 |  |  | 40 | 1599.89 | 7.3777 | 5.896
41 | 1639.887 | 7.4024 | 5.8767 |  |  | 41 | 1639.887 | 7.4024 | 5.8815 |  |  | 41 | 1639.887 | 7.4024 | 5.8834
42 | 1679.885 | 7.4265 | 5.8645 |  |  | 42 | 1679.885 | 7.4265 | 5.8693 |  |  | 42 | 1679.885 | 7.4265 | 5.8712
43 | 1719.882 | 7.45 | 5.8526 |  |  | 43 | 1719.882 | 7.45 | 5.8574 |  |  | 43 | 1719.882 | 7.45 | 5.8593
44 | 1759.879 | 7.473 | 5.8409 |  |  | 44 | 1759.879 | 7.473 | 5.8457 |  |  | 44 | 1759.879 | 7.473 | 5.8476
45 | 1799.876 | 7.4955 | 5.8295 |  |  | 45 | 1799.876 | 7.4955 | 5.8344 |  |  | 45 | 1799.876 | 7.4955 | 5.8363
46 | 1839.874 | 7.5175 | 5.8184 |  |  | 46 | 1839.874 | 7.5175 | 5.8233 |  |  | 46 | 1839.874 | 7.5175 | 5.8251
47 | 1879.871 | 7.539 | 5.8075 |  |  | 47 | 1879.871 | 7.539 | 5.8124 |  |  | 47 | 1879.871 | 7.539 | 5.8143
48 | 1919.868 | 7.56 | 5.7969 |  |  | 48 | 1919.868 | 7.56 | 5.8018 |  |  | 48 | 1919.868 | 7.56 | 5.8037
49 | 1959.865 | 7.5806 | 5.7865 |  |  | 49 | 1959.865 | 7.5806 | 5.7915 |  |  | 49 | 1959.865 | 7.5806 | 5.7933
50 | 1999.863 | 7.6008 | 5.7763 |  |  | 50 | 1999.863 | 7.6008 | 5.7813 |  |  | 50 | 1999.863 | 7.6008 | 5.7831
51 | 2039.86 | 7.6206 | 5.7663 |  |  | 51 | 2039.86 | 7.6206 | 5.7713 |  |  | 51 | 2039.86 | 7.6206 | 5.7732
52 | 2079.857 | 7.6401 | 5.7566 |  |  | 52 | 2079.857 | 7.6401 | 5.7616 |  |  | 52 | 2079.857 | 7.6401 | 5.7634
53 | 2119.854 | 7.6591 | 5.747 |  |  | 53 | 2119.854 | 7.6591 | 5.7521 |  |  | 53 | 2119.854 | 7.6591 | 5.7538
54 | 2159.852 | 7.6778 | 5.7376 |  |  | 54 | 2159.852 | 7.6778 | 5.7427 |  |  | 54 | 2159.852 | 7.6778 | 5.7445
55 | 2199.849 | 7.6961 | 5.7284 |  |  | 55 | 2199.849 | 7.6961 | 5.7335 |  |  | 55 | 2199.849 | 7.6961 | 5.7353
56 | 2239.846 | 7.7142 | 5.7194 |  |  | 56 | 2239.846 | 7.7142 | 5.7245 |  |  | 56 | 2239.846 | 7.7142 | 5.7263
57 | 2279.843 | 7.7319 | 5.7106 |  |  | 57 | 2279.843 | 7.7319 | 5.7157 |  |  | 57 | 2279.843 | 7.7319 | 5.7175
58 | 2319.841 | 7.7493 | 5.7019 |  |  | 58 | 2319.841 | 7.7493 | 5.707 |  |  | 58 | 2319.841 | 7.7493 | 5.7088
59 | 2359.838 | 7.7663 | 5.6934 |  |  | 59 | 2359.838 | 7.7663 | 5.6985 |  |  | 59 | 2359.838 | 7.7663 | 5.7003
60 | 2399.835 | 7.7832 | 5.685 |  |  | 60 | 2399.835 | 7.7832 | 5.6902 |  |  | 60 | 2399.835 | 7.7832 | 5.6919
61 | 2439.832 | 7.7997 | 5.6768 |  |  | 61 | 2439.832 | 7.7997 | 5.682 |  |  | 61 | 2439.832 | 7.7997 | 5.6837
62 | 2479.83 | 7.8159 | 5.6688 |  |  | 62 | 2479.83 | 7.8159 | 5.6739 |  |  | 62 | 2479.83 | 7.8159 | 5.6757
63 | 2519.827 | 7.8319 | 5.6608 |  |  | 63 | 2519.827 | 7.8319 | 5.666 |  |  | 63 | 2519.827 | 7.8319 | 5.6677
64 | 2559.824 | 7.8477 | 5.6531 |  |  | 64 | 2559.824 | 7.8477 | 5.6582 |  |  | 64 | 2559.824 | 7.8477 | 5.66
65 | 2599.821 | 7.8632 | 5.6454 |  |  | 65 | 2599.821 | 7.8632 | 5.6506 |  |  | 65 | 2599.821 | 7.8632 | 5.6523
66 | 2639.819 | 7.8785 | 5.6379 |  |  | 66 | 2639.819 | 7.8785 | 5.6431 |  |  | 66 | 2639.819 | 7.8785 | 5.6448
67 | 2679.816 | 7.8935 | 5.6305 |  |  | 67 | 2679.816 | 7.8935 | 5.6357 |  |  | 67 | 2679.816 | 7.8935 | 5.6374
68 | 2719.813 | 7.9083 | 5.6232 |  |  | 68 | 2719.813 | 7.9083 | 5.6284 |  |  | 68 | 2719.813 | 7.9083 | 5.6301
69 | 2759.81 | 7.9229 | 5.616 |  |  | 69 | 2759.81 | 7.9229 | 5.6213 |  |  | 69 | 2759.81 | 7.9229 | 5.623
70 | 2799.808 | 7.9373 | 5.609 |  |  | 70 | 2799.808 | 7.9373 | 5.6142 |  |  | 70 | 2799.808 | 7.9373 | 5.6159
71 | 2839.805 | 7.9515 | 5.6021 |  |  | 71 | 2839.805 | 7.9515 | 5.6073 |  |  | 71 | 2839.805 | 7.9515 | 5.609
72 | 2879.802 | 7.9655 | 5.5952 |  |  | 72 | 2879.802 | 7.9655 | 5.6005 |  |  | 72 | 2879.802 | 7.9655 | 5.6022
73 | 2919.799 | 7.9793 | 5.5885 |  |  | 73 | 2919.799 | 7.9793 | 5.5938 |  |  | 73 | 2919.799 | 7.9793 | 5.5955
74 | 2959.797 | 7.9929 | 5.5819 |  |  | 74 | 2959.797 | 7.9929 | 5.5871 |  |  | 74 | 2959.797 | 7.9929 | 5.5889
75 | 2999.794 | 8.0063 | 5.5754 |  |  | 75 | 2999.794 | 8.0063 | 5.5806 |  |  | 75 | 2999.794 | 8.0063 | 5.5824
76 | 3039.791 | 8.0195 | 5.569 |  |  | 76 | 3039.791 | 8.0195 | 5.5742 |  |  | 76 | 3039.791 | 8.0195 | 5.5759
77 | 3079.788 | 8.0326 | 5.5627 |  |  | 77 | 3079.788 | 8.0326 | 5.5679 |  |  | 77 | 3079.788 | 8.0326 | 5.5696
78 | 3119.786 | 8.0455 | 5.5564 |  |  | 78 | 3119.786 | 8.0455 | 5.5617 |  |  | 78 | 3119.786 | 8.0455 | 5.5634
79 | 3159.783 | 8.0583 | 5.5503 |  |  | 79 | 3159.783 | 8.0583 | 5.5556 |  |  | 79 | 3159.783 | 8.0583 | 5.5573
80 | 3199.78 | 8.0708 | 5.5443 |  |  | 80 | 3199.78 | 8.0708 | 5.5495 |  |  | 80 | 3199.78 | 8.0708 | 5.5512
81 | 3239.777 | 8.0833 | 5.5383 |  |  | 81 | 3239.777 | 8.0833 | 5.5436 |  |  | 81 | 3239.777 | 8.0833 | 5.5453
82 | 3279.775 | 8.0955 | 5.5324 |  |  | 82 | 3279.775 | 8.0955 | 5.5377 |  |  | 82 | 3279.775 | 8.0955 | 5.5394
83 | 3319.772 | 8.1077 | 5.5266 |  |  | 83 | 3319.772 | 8.1077 | 5.5319 |  |  | 83 | 3319.772 | 8.1077 | 5.5336
84 | 3359.769 | 8.1196 | 5.5209 |  |  | 84 | 3359.769 | 8.1196 | 5.5262 |  |  | 84 | 3359.769 | 8.1196 | 5.5279
85 | 3399.766 | 8.1315 | 5.5153 |  |  | 85 | 3399.766 | 8.1315 | 5.5206 |  |  | 85 | 3399.766 | 8.1315 | 5.5223
86 | 3439.764 | 8.1432 | 5.5098 |  |  | 86 | 3439.764 | 8.1432 | 5.515 |  |  | 86 | 3439.764 | 8.1432 | 5.5168
87 | 3479.761 | 8.1547 | 5.5043 |  |  | 87 | 3479.761 | 8.1547 | 5.5096 |  |  | 87 | 3479.761 | 8.1547 | 5.5113
88 | 3519.758 | 8.1661 | 5.4989 |  |  | 88 | 3519.758 | 8.1661 | 5.5042 |  |  | 88 | 3519.758 | 8.1661 | 5.5059
89 | 3559.755 | 8.1774 | 5.4936 |  |  | 89 | 3559.755 | 8.1774 | 5.4988 |  |  | 89 | 3559.755 | 8.1774 | 5.5006
90 | 3599.753 | 8.1886 | 5.4883 |  |  | 90 | 3599.753 | 8.1886 | 5.4936 |  |  | 90 | 3599.753 | 8.1886 | 5.4953
91 | 3639.75 | 8.1997 | 5.4831 |  |  | 91 | 3639.75 | 8.1997 | 5.4884 |  |  | 91 | 3639.75 | 8.1997 | 5.4901
92 | 3679.747 | 8.2106 | 5.478 |  |  | 92 | 3679.747 | 8.2106 | 5.4833 |  |  | 92 | 3679.747 | 8.2106 | 5.485
93 | 3719.744 | 8.2214 | 5.473 |  |  | 93 | 3719.744 | 8.2214 | 5.4782 |  |  | 93 | 3719.744 | 8.2214 | 5.48
94 | 3759.742 | 8.2321 | 5.468 |  |  | 94 | 3759.742 | 8.2321 | 5.4732 |  |  | 94 | 3759.742 | 8.2321 | 5.475
95 | 3799.739 | 8.2427 | 5.4631 |  |  | 95 | 3799.739 | 8.2427 | 5.4683 |  |  | 95 | 3799.739 | 8.2427 | 5.4701
96 | 3839.736 | 8.2532 | 5.4582 |  |  | 96 | 3839.736 | 8.2532 | 5.4635 |  |  | 96 | 3839.736 | 8.2532 | 5.4652
97 | 3879.733 | 8.2635 | 5.4534 |  |  | 97 | 3879.733 | 8.2635 | 5.4587 |  |  | 97 | 3879.733 | 8.2635 | 5.4605
98 | 3919.731 | 8.2738 | 5.4487 |  |  | 98 | 3919.731 | 8.2738 | 5.454 |  |  | 98 | 3919.731 | 8.2738 | 5.4557
99 | 3959.728 | 8.2839 | 5.444 |  |  | 99 | 3959.728 | 8.2839 | 5.4493 |  |  | 99 | 3959.728 | 8.2839 | 5.4511
100 | 3999.725 | 8.294 | 5.4394 |  |  | 100 | 3999.725 | 8.294 | 5.4447 |  |  | 100 | 3999.725 | 8.294 | 5.4464
101 | 4039.722 | 8.3039 | 5.4349 |  |  | 101 | 4039.722 | 8.3039 | 5.4401 |  |  | 101 | 4039.722 | 8.3039 | 5.4419
102 | 4079.72 | 8.3138 | 5.4304 |  |  | 102 | 4079.72 | 8.3138 | 5.4356 |  |  | 102 | 4079.72 | 8.3138 | 5.4374
103 | 4119.717 | 8.3235 | 5.4259 |  |  | 103 | 4119.717 | 8.3235 | 5.4312 |  |  | 103 | 4119.717 | 8.3235 | 5.433
104 | 4159.714 | 8.3332 | 5.4216 |  |  | 104 | 4159.714 | 8.3332 | 5.4268 |  |  | 104 | 4159.714 | 8.3332 | 5.4286
105 | 4199.711 | 8.3428 | 5.4172 |  |  | 105 | 4199.711 | 8.3428 | 5.4224 |  |  | 105 | 4199.711 | 8.3428 | 5.4242
106 | 4239.709 | 8.3522 | 5.4129 |  |  | 106 | 4239.709 | 8.3522 | 5.4182 |  |  | 106 | 4239.709 | 8.3522 | 5.42
107 | 4279.706 | 8.3616 | 5.4087 |  |  | 107 | 4279.706 | 8.3616 | 5.4139 |  |  | 107 | 4279.706 | 8.3616 | 5.4157
108 | 4319.703 | 8.3709 | 5.4045 |  |  | 108 | 4319.703 | 8.3709 | 5.4097 |  |  | 108 | 4319.703 | 8.3709 | 5.4116
109 | 4359.7 | 8.3802 | 5.4004 |  |  | 109 | 4359.7 | 8.3802 | 5.4056 |  |  | 109 | 4359.7 | 8.3802 | 5.4074
110 | 4399.698 | 8.3893 | 5.3963 |  |  | 110 | 4399.698 | 8.3893 | 5.4015 |  |  | 110 | 4399.698 | 8.3893 | 5.4033
111 | 4439.695 | 8.3983 | 5.3923 |  |  | 111 | 4439.695 | 8.3983 | 5.3975 |  |  | 111 | 4439.695 | 8.3983 | 5.3993
112 | 4479.692 | 8.4073 | 5.3883 |  |  | 112 | 4479.692 | 8.4073 | 5.3935 |  |  | 112 | 4479.692 | 8.4073 | 5.3953
113 | 4519.689 | 8.4162 | 5.3844 |  |  | 113 | 4519.689 | 8.4162 | 5.3896 |  |  | 113 | 4519.689 | 8.4162 | 5.3914
114 | 4559.687 | 8.425 | 5.3805 |  |  | 114 | 4559.687 | 8.425 | 5.3857 |  |  | 114 | 4559.687 | 8.425 | 5.3875
115 | 4599.684 | 8.4337 | 5.3767 |  |  | 115 | 4599.684 | 8.4337 | 5.3818 |  |  | 115 | 4599.684 | 8.4337 | 5.3837
116 | 4639.681 | 8.4424 | 5.3729 |  |  | 116 | 4639.681 | 8.4424 | 5.378 |  |  | 116 | 4639.681 | 8.4424 | 5.3799
117 | 4679.678 | 8.451 | 5.3691 |  |  | 117 | 4679.678 | 8.451 | 5.3743 |  |  | 117 | 4679.678 | 8.451 | 5.3761
118 | 4719.676 | 8.4595 | 5.3654 |  |  | 118 | 4719.676 | 8.4595 | 5.3705 |  |  | 118 | 4719.676 | 8.4595 | 5.3724
119 | 4759.673 | 8.4679 | 5.3617 |  |  | 119 | 4759.673 | 8.4679 | 5.3669 |  |  | 119 | 4759.673 | 8.4679 | 5.3687
120 | 4799.67 | 8.4763 | 5.3581 |  |  | 120 | 4799.67 | 8.4763 | 5.3632 |  |  | 120 | 4799.67 | 8.4763 | 5.3651
121 | 4839.667 | 8.4846 | 5.3545 |  |  | 121 | 4839.667 | 8.4846 | 5.3596 |  |  | 121 | 4839.667 | 8.4846 | 5.3615
122 | 4879.665 | 8.4928 | 5.351 |  |  | 122 | 4879.665 | 8.4928 | 5.3561 |  |  | 122 | 4879.665 | 8.4928 | 5.3579
123 | 4919.662 | 8.501 | 5.3475 |  |  | 123 | 4919.662 | 8.501 | 5.3526 |  |  | 123 | 4919.662 | 8.501 | 5.3544
124 | 4959.659 | 8.5091 | 5.344 |  |  | 124 | 4959.659 | 8.5091 | 5.3491 |  |  | 124 | 4959.659 | 8.5091 | 5.351
125 | 4999.656 | 8.5171 | 5.3406 |  |  | 125 | 4999.656 | 8.5171 | 5.3457 |  |  | 125 | 4999.656 | 8.5171 | 5.3475
126 | 5039.654 | 8.5251 | 5.3372 |  |  | 126 | 5039.654 | 8.5251 | 5.3423 |  |  | 126 | 5039.654 | 8.5251 | 5.3441
127 | 5079.651 | 8.533 | 5.3338 |  |  | 127 | 5079.651 | 8.533 | 5.3389 |  |  | 127 | 5079.651 | 8.533 | 5.3408
128 | 5119.648 | 8.5408 | 5.3305 |  |  | 128 | 5119.648 | 8.5408 | 5.3356 |  |  | 128 | 5119.648 | 8.5408 | 5.3375
129 | 5159.645 | 8.5486 | 5.3272 |  |  | 129 | 5159.645 | 8.5486 | 5.3323 |  |  | 129 | 5159.645 | 8.5486 | 5.3342
130 | 5199.643 | 8.5563 | 5.324 |  |  | 130 | 5199.643 | 8.5563 | 5.3291 |  |  | 130 | 5199.643 | 8.5563 | 5.3309
131 | 5239.64 | 8.564 | 5.3208 |  |  | 131 | 5239.64 | 8.564 | 5.3258 |  |  | 131 | 5239.64 | 8.564 | 5.3277
132 | 5279.637 | 8.5716 | 5.3176 |  |  | 132 | 5279.637 | 8.5716 | 5.3227 |  |  | 132 | 5279.637 | 8.5716 | 5.3245
133 | 5319.634 | 8.5792 | 5.3144 |  |  | 133 | 5319.634 | 8.5792 | 5.3195 |  |  | 133 | 5319.634 | 8.5792 | 5.3214
134 | 5359.632 | 8.5867 | 5.3113 |  |  | 134 | 5359.632 | 8.5867 | 5.3164 |  |  | 134 | 5359.632 | 8.5867 | 5.3183
135 | 5399.629 | 8.5941 | 5.3083 |  |  | 135 | 5399.629 | 8.5941 | 5.3133 |  |  | 135 | 5399.629 | 8.5941 | 5.3152
136 | 5439.626 | 8.6015 | 5.3052 |  |  | 136 | 5439.626 | 8.6015 | 5.3103 |  |  | 136 | 5439.626 | 8.6015 | 5.3121
137 | 5479.623 | 8.6088 | 5.3022 |  |  | 137 | 5479.623 | 8.6088 | 5.3073 |  |  | 137 | 5479.623 | 8.6088 | 5.3091
138 | 5519.621 | 8.6161 | 5.2992 |  |  | 138 | 5519.621 | 8.6161 | 5.3043 |  |  | 138 | 5519.621 | 8.6161 | 5.3062
139 | 5559.618 | 8.6233 | 5.2963 |  |  | 139 | 5559.618 | 8.6233 | 5.3013 |  |  | 139 | 5559.618 | 8.6233 | 5.3032
140 | 5599.615 | 8.6305 | 5.2934 |  |  | 140 | 5599.615 | 8.6305 | 5.2984 |  |  | 140 | 5599.615 | 8.6305 | 5.3003
141 | 5639.612 | 8.6376 | 5.2905 |  |  | 141 | 5639.612 | 8.6376 | 5.2955 |  |  | 141 | 5639.612 | 8.6376 | 5.2974
142 | 5679.61 | 8.6446 | 5.2876 |  |  | 142 | 5679.61 | 8.6446 | 5.2926 |  |  | 142 | 5679.61 | 8.6446 | 5.2945
143 | 5719.607 | 8.6517 | 5.2848 |  |  | 143 | 5719.607 | 8.6517 | 5.2898 |  |  | 143 | 5719.607 | 8.6517 | 5.2917
144 | 5759.604 | 8.6586 | 5.282 |  |  | 144 | 5759.604 | 8.6586 | 5.287 |  |  | 144 | 5759.604 | 8.6586 | 5.2889
145 | 5799.601 | 8.6655 | 5.2792 |  |  | 145 | 5799.601 | 8.6655 | 5.2842 |  |  | 145 | 5799.601 | 8.6655 | 5.2861
146 | 5839.599 | 8.6724 | 5.2765 |  |  | 146 | 5839.599 | 8.6724 | 5.2815 |  |  | 146 | 5839.599 | 8.6724 | 5.2834
147 | 5879.596 | 8.6792 | 5.2738 |  |  | 147 | 5879.596 | 8.6792 | 5.2788 |  |  | 147 | 5879.596 | 8.6792 | 5.2807
148 | 5919.593 | 8.686 | 5.2711 |  |  | 148 | 5919.593 | 8.686 | 5.2761 |  |  | 148 | 5919.593 | 8.686 | 5.278
149 | 5959.59 | 8.6928 | 5.2685 |  |  | 149 | 5959.59 | 8.6928 | 5.2734 |  |  | 149 | 5959.59 | 8.6928 | 5.2753
150 | 5999.588 | 8.6994 | 5.2658 |  |  | 150 | 5999.588 | 8.6994 | 5.2708 |  |  | 150 | 5999.588 | 8.6994 | 5.2727
151 | 6039.585 | 8.7061 | 5.2632 |  |  | 151 | 6039.585 | 8.7061 | 5.2682 |  |  | 151 | 6039.585 | 8.7061 | 5.2701
152 | 6079.582 | 8.7127 | 5.2607 |  |  | 152 | 6079.582 | 8.7127 | 5.2656 |  |  | 152 | 6079.582 | 8.7127 | 5.2675
153 | 6119.579 | 8.7192 | 5.2581 |  |  | 153 | 6119.579 | 8.7192 | 5.2631 |  |  | 153 | 6119.579 | 8.7192 | 5.265
154 | 6159.577 | 8.7258 | 5.2556 |  |  | 154 | 6159.577 | 8.7258 | 5.2605 |  |  | 154 | 6159.577 | 8.7258 | 5.2624
155 | 6199.574 | 8.7322 | 5.2531 |  |  | 155 | 6199.574 | 8.7322 | 5.258 |  |  | 155 | 6199.574 | 8.7322 | 5.2599
156 | 6239.571 | 8.7387 | 5.2506 |  |  | 156 | 6239.571 | 8.7387 | 5.2556 |  |  | 156 | 6239.571 | 8.7387 | 5.2575
157 | 6279.568 | 8.7451 | 5.2482 |  |  | 157 | 6279.568 | 8.7451 | 5.2531 |  |  | 157 | 6279.568 | 8.7451 | 5.255
158 | 6319.566 | 8.7514 | 5.2458 |  |  | 158 | 6319.566 | 8.7514 | 5.2507 |  |  | 158 | 6319.566 | 8.7514 | 5.2526
159 | 6359.563 | 8.7577 | 5.2434 |  |  | 159 | 6359.563 | 8.7577 | 5.2483 |  |  | 159 | 6359.563 | 8.7577 | 5.2502
160 | 6399.56 | 8.764 | 5.241 |  |  | 160 | 6399.56 | 8.764 | 5.2459 |  |  | 160 | 6399.56 | 8.764 | 5.2478
161 | 6439.557 | 8.7702 | 5.2386 |  |  | 161 | 6439.557 | 8.7702 | 5.2435 |  |  | 161 | 6439.557 | 8.7702 | 5.2454
162 | 6479.555 | 8.7764 | 5.2363 |  |  | 162 | 6479.555 | 8.7764 | 5.2412 |  |  | 162 | 6479.555 | 8.7764 | 5.2431
163 | 6519.552 | 8.7826 | 5.234 |  |  | 163 | 6519.552 | 8.7826 | 5.2389 |  |  | 163 | 6519.552 | 8.7826 | 5.2408
164 | 6559.549 | 8.7887 | 5.2317 |  |  | 164 | 6559.549 | 8.7887 | 5.2366 |  |  | 164 | 6559.549 | 8.7887 | 5.2385
165 | 6599.546 | 8.7948 | 5.2295 |  |  | 165 | 6599.546 | 8.7948 | 5.2343 |  |  | 165 | 6599.546 | 8.7948 | 5.2362
166 | 6639.544 | 8.8008 | 5.2272 |  |  | 166 | 6639.544 | 8.8008 | 5.2321 |  |  | 166 | 6639.544 | 8.8008 | 5.234
167 | 6679.541 | 8.8068 | 5.225 |  |  | 167 | 6679.541 | 8.8068 | 5.2299 |  |  | 167 | 6679.541 | 8.8068 | 5.2318
168 | 6719.538 | 8.8128 | 5.2228 |  |  | 168 | 6719.538 | 8.8128 | 5.2277 |  |  | 168 | 6719.538 | 8.8128 | 5.2296
169 | 6759.535 | 8.8187 | 5.2206 |  |  | 169 | 6759.535 | 8.8187 | 5.2255 |  |  | 169 | 6759.535 | 8.8187 | 5.2274
170 | 6799.533 | 8.8246 | 5.2185 |  |  | 170 | 6799.533 | 8.8246 | 5.2233 |  |  | 170 | 6799.533 | 8.8246 | 5.2252
171 | 6839.53 | 8.8305 | 5.2164 |  |  | 171 | 6839.53 | 8.8305 | 5.2212 |  |  | 171 | 6839.53 | 8.8305 | 5.2231
172 | 6879.527 | 8.8363 | 5.2142 |  |  | 172 | 6879.527 | 8.8363 | 5.2191 |  |  | 172 | 6879.527 | 8.8363 | 5.221
173 | 6919.524 | 8.8421 | 5.2122 |  |  | 173 | 6919.524 | 8.8421 | 5.217 |  |  | 173 | 6919.524 | 8.8421 | 5.2189
174 | 6959.522 | 8.8479 | 5.2101 |  |  | 174 | 6959.522 | 8.8479 | 5.2149 |  |  | 174 | 6959.522 | 8.8479 | 5.2168
175 | 6999.519 | 8.8536 | 5.208 |  |  | 175 | 6999.519 | 8.8536 | 5.2129 |  |  | 175 | 6999.519 | 8.8536 | 5.2148
176 | 7039.516 | 8.8593 | 5.206 |  |  | 176 | 7039.516 | 8.8593 | 5.2108 |  |  | 176 | 7039.516 | 8.8593 | 5.2127
177 | 7079.513 | 8.865 | 5.204 |  |  | 177 | 7079.513 | 8.865 | 5.2088 |  |  | 177 | 7079.513 | 8.865 | 5.2107
178 | 7119.511 | 8.8706 | 5.202 |  |  | 178 | 7119.511 | 8.8706 | 5.2068 |  |  | 178 | 7119.511 | 8.8706 | 5.2087
179 | 7159.508 | 8.8762 | 5.2 |  |  | 179 | 7159.508 | 8.8762 | 5.2048 |  |  | 179 | 7159.508 | 8.8762 | 5.2067
180 | 7199.505 | 8.8818 | 5.1981 |  |  | 180 | 7199.505 | 8.8818 | 5.2029 |  |  | 180 | 7199.505 | 8.8818 | 5.2048
181 | 7239.502 | 8.8873 | 5.1961 |  |  | 181 | 7239.502 | 8.8873 | 5.2009 |  |  | 181 | 7239.502 | 8.8873 | 5.2028
182 | 7279.5 | 8.8928 | 5.1942 |  |  | 182 | 7279.5 | 8.8928 | 5.199 |  |  | 182 | 7279.5 | 8.8928 | 5.2009
183 | 7319.497 | 8.8983 | 5.1923 |  |  | 183 | 7319.497 | 8.8983 | 5.1971 |  |  | 183 | 7319.497 | 8.8983 | 5.199
184 | 7359.494 | 8.9037 | 5.1904 |  |  | 184 | 7359.494 | 8.9037 | 5.1952 |  |  | 184 | 7359.494 | 8.9037 | 5.1971
185 | 7399.491 | 8.9092 | 5.1886 |  |  | 185 | 7399.491 | 8.9092 | 5.1933 |  |  | 185 | 7399.491 | 8.9092 | 5.1952
186 | 7439.489 | 8.9146 | 5.1867 |  |  | 186 | 7439.489 | 8.9146 | 5.1915 |  |  | 186 | 7439.489 | 8.9146 | 5.1934
187 | 7479.486 | 8.9199 | 5.1849 |  |  | 187 | 7479.486 | 8.9199 | 5.1897 |  |  | 187 | 7479.486 | 8.9199 | 5.1915
188 | 7519.483 | 8.9253 | 5.1831 |  |  | 188 | 7519.483 | 8.9253 | 5.1878 |  |  | 188 | 7519.483 | 8.9253 | 5.1897
189 | 7559.48 | 8.9306 | 5.1813 |  |  | 189 | 7559.48 | 8.9306 | 5.186 |  |  | 189 | 7559.48 | 8.9306 | 5.1879
190 | 7599.478 | 8.9358 | 5.1795 |  |  | 190 | 7599.478 | 8.9358 | 5.1843 |  |  | 190 | 7599.478 | 8.9358 | 5.1861
191 | 7639.475 | 8.9411 | 5.1778 |  |  | 191 | 7639.475 | 8.9411 | 5.1825 |  |  | 191 | 7639.475 | 8.9411 | 5.1844
192 | 7679.472 | 8.9463 | 5.176 |  |  | 192 | 7679.472 | 8.9463 | 5.1807 |  |  | 192 | 7679.472 | 8.9463 | 5.1826
193 | 7719.469 | 8.9515 | 5.1743 |  |  | 193 | 7719.469 | 8.9515 | 5.179 |  |  | 193 | 7719.469 | 8.9515 | 5.1809
194 | 7759.467 | 8.9567 | 5.1726 |  |  | 194 | 7759.467 | 8.9567 | 5.1773 |  |  | 194 | 7759.467 | 8.9567 | 5.1792
195 | 7799.464 | 8.9618 | 5.1709 |  |  | 195 | 7799.464 | 8.9618 | 5.1756 |  |  | 195 | 7799.464 | 8.9618 | 5.1775
196 | 7839.461 | 8.9669 | 5.1692 |  |  | 196 | 7839.461 | 8.9669 | 5.1739 |  |  | 196 | 7839.461 | 8.9669 | 5.1758
197 | 7879.458 | 8.972 | 5.1675 |  |  | 197 | 7879.458 | 8.972 | 5.1722 |  |  | 197 | 7879.458 | 8.972 | 5.1741
198 | 7919.456 | 8.9771 | 5.1659 |  |  | 198 | 7919.456 | 8.9771 | 5.1706 |  |  | 198 | 7919.456 | 8.9771 | 5.1724
199 | 7959.453 | 8.9821 | 5.1642 |  |  | 199 | 7959.453 | 8.9821 | 5.1689 |  |  | 199 | 7959.453 | 8.9821 | 5.1708
200 | 7999.45 | 8.9871 | 5.1626 |  |  | 200 | 7999.45 | 8.9871 | 5.1673 |  |  | 200 | 7999.45 | 8.9871 | 5.1692
201 | 8039.447 | 8.9921 | 5.161 |  |  | 201 | 8039.447 | 8.9921 | 5.1657 |  |  | 201 | 8039.447 | 8.9921 | 5.1675
202 | 8079.445 | 8.9971 | 5.1594 |  |  | 202 | 8079.445 | 8.9971 | 5.1641 |  |  | 202 | 8079.445 | 8.9971 | 5.166
203 | 8119.442 | 9.002 | 5.1578 |  |  | 203 | 8119.442 | 9.002 | 5.1625 |  |  | 203 | 8119.442 | 9.002 | 5.1644
204 | 8159.439 | 9.0069 | 5.1563 |  |  | 204 | 8159.439 | 9.0069 | 5.1609 |  |  | 204 | 8159.439 | 9.0069 | 5.1628
205 | 8199.436 | 9.0118 | 5.1547 |  |  | 205 | 8199.436 | 9.0118 | 5.1594 |  |  | 205 | 8199.436 | 9.0118 | 5.1612
206 | 8239.434 | 9.0167 | 5.1532 |  |  | 206 | 8239.434 | 9.0167 | 5.1578 |  |  | 206 | 8239.434 | 9.0167 | 5.1597
207 | 8279.431 | 9.0215 | 5.1517 |  |  | 207 | 8279.431 | 9.0215 | 5.1563 |  |  | 207 | 8279.431 | 9.0215 | 5.1582
208 | 8319.428 | 9.0263 | 5.1502 |  |  | 208 | 8319.428 | 9.0263 | 5.1548 |  |  | 208 | 8319.428 | 9.0263 | 5.1567
209 | 8359.425 | 9.0311 | 5.1487 |  |  | 209 | 8359.425 | 9.0311 | 5.1533 |  |  | 209 | 8359.425 | 9.0311 | 5.1552
210 | 8399.423 | 9.0359 | 5.1472 |  |  | 210 | 8399.423 | 9.0359 | 5.1518 |  |  | 210 | 8399.423 | 9.0359 | 5.1537
211 | 8439.42 | 9.0407 | 5.1457 |  |  | 211 | 8439.42 | 9.0407 | 5.1503 |  |  | 211 | 8439.42 | 9.0407 | 5.1522
212 | 8479.417 | 9.0454 | 5.1443 |  |  | 212 | 8479.417 | 9.0454 | 5.1489 |  |  | 212 | 8479.417 | 9.0454 | 5.1507
213 | 8519.414 | 9.0501 | 5.1428 |  |  | 213 | 8519.414 | 9.0501 | 5.1474 |  |  | 213 | 8519.414 | 9.0501 | 5.1493
214 | 8559.412 | 9.0548 | 5.1414 |  |  | 214 | 8559.412 | 9.0548 | 5.146 |  |  | 214 | 8559.412 | 9.0548 | 5.1478
215 | 8599.409 | 9.0594 | 5.14 |  |  | 215 | 8599.409 | 9.0594 | 5.1446 |  |  | 215 | 8599.409 | 9.0594 | 5.1464
216 | 8639.406 | 9.0641 | 5.1386 |  |  | 216 | 8639.406 | 9.0641 | 5.1432 |  |  | 216 | 8639.406 | 9.0641 | 5.145
217 | 8679.403 | 9.0687 | 5.1372 |  |  | 217 | 8679.403 | 9.0687 | 5.1418 |  |  | 217 | 8679.403 | 9.0687 | 5.1436
218 | 8719.401 | 9.0733 | 5.1358 |  |  | 218 | 8719.401 | 9.0733 | 5.1404 |  |  | 218 | 8719.401 | 9.0733 | 5.1422
219 | 8759.398 | 9.0779 | 5.1345 |  |  | 219 | 8759.398 | 9.0779 | 5.139 |  |  | 219 | 8759.398 | 9.0779 | 5.1409
220 | 8799.395 | 9.0824 | 5.1331 |  |  | 220 | 8799.395 | 9.0824 | 5.1377 |  |  | 220 | 8799.395 | 9.0824 | 5.1395
221 | 8839.392 | 9.087 | 5.1318 |  |  | 221 | 8839.392 | 9.087 | 5.1363 |  |  | 221 | 8839.392 | 9.087 | 5.1381
222 | 8879.39 | 9.0915 | 5.1304 |  |  | 222 | 8879.39 | 9.0915 | 5.135 |  |  | 222 | 8879.39 | 9.0915 | 5.1368
223 | 8919.387 | 9.096 | 5.1291 |  |  | 223 | 8919.387 | 9.096 | 5.1337 |  |  | 223 | 8919.387 | 9.096 | 5.1355
224 | 8959.384 | 9.1005 | 5.1278 |  |  | 224 | 8959.384 | 9.1005 | 5.1323 |  |  | 224 | 8959.384 | 9.1005 | 5.1342
225 | 8999.381 | 9.1049 | 5.1265 |  |  | 225 | 8999.381 | 9.1049 | 5.131 |  |  | 225 | 8999.381 | 9.1049 | 5.1329
226 | 9039.379 | 9.1093 | 5.1252 |  |  | 226 | 9039.379 | 9.1093 | 5.1298 |  |  | 226 | 9039.379 | 9.1093 | 5.1316
227 | 9079.376 | 9.1138 | 5.124 |  |  | 227 | 9079.376 | 9.1138 | 5.1285 |  |  | 227 | 9079.376 | 9.1138 | 5.1303
228 | 9119.373 | 9.1182 | 5.1227 |  |  | 228 | 9119.373 | 9.1182 | 5.1272 |  |  | 228 | 9119.373 | 9.1182 | 5.129
229 | 9159.37 | 9.1225 | 5.1215 |  |  | 229 | 9159.37 | 9.1225 | 5.126 |  |  | 229 | 9159.37 | 9.1225 | 5.1278
230 | 9199.368 | 9.1269 | 5.1202 |  |  | 230 | 9199.368 | 9.1269 | 5.1247 |  |  | 230 | 9199.368 | 9.1269 | 5.1265
231 | 9239.365 | 9.1312 | 5.119 |  |  | 231 | 9239.365 | 9.1312 | 5.1235 |  |  | 231 | 9239.365 | 9.1312 | 5.1253
232 | 9279.362 | 9.1355 | 5.1178 |  |  | 232 | 9279.362 | 9.1355 | 5.1223 |  |  | 232 | 9279.362 | 9.1355 | 5.124
233 | 9319.359 | 9.1398 | 5.1166 |  |  | 233 | 9319.359 | 9.1398 | 5.121 |  |  | 233 | 9319.359 | 9.1398 | 5.1228
234 | 9359.357 | 9.1441 | 5.1154 |  |  | 234 | 9359.357 | 9.1441 | 5.1198 |  |  | 234 | 9359.357 | 9.1441 | 5.1216
235 | 9399.354 | 9.1484 | 5.1142 |  |  | 235 | 9399.354 | 9.1484 | 5.1186 |  |  | 235 | 9399.354 | 9.1484 | 5.1204
236 | 9439.351 | 9.1526 | 5.113 |  |  | 236 | 9439.351 | 9.1526 | 5.1175 |  |  | 236 | 9439.351 | 9.1526 | 5.1192
237 | 9479.348 | 9.1569 | 5.1118 |  |  | 237 | 9479.348 | 9.1569 | 5.1163 |  |  | 237 | 9479.348 | 9.1569 | 5.1181
238 | 9519.346 | 9.1611 | 5.1107 |  |  | 238 | 9519.346 | 9.1611 | 5.1151 |  |  | 238 | 9519.346 | 9.1611 | 5.1169
239 | 9559.343 | 9.1653 | 5.1095 |  |  | 239 | 9559.343 | 9.1653 | 5.114 |  |  | 239 | 9559.343 | 9.1653 | 5.1158
240 | 9599.34 | 9.1694 | 5.1084 |  |  | 240 | 9599.34 | 9.1694 | 5.1128 |  |  | 240 | 9599.34 | 9.1694 | 5.1146
241 | 9639.337 | 9.1736 | 5.1073 |  |  | 241 | 9639.337 | 9.1736 | 5.1117 |  |  | 241 | 9639.337 | 9.1736 | 5.1135
242 | 9679.335 | 9.1777 | 5.1062 |  |  | 242 | 9679.335 | 9.1777 | 5.1106 |  |  | 242 | 9679.335 | 9.1777 | 5.1123
243 | 9719.332 | 9.1819 | 5.1051 |  |  | 243 | 9719.332 | 9.1819 | 5.1095 |  |  | 243 | 9719.332 | 9.1819 | 5.1112
244 | 9759.329 | 9.186 | 5.104 |  |  | 244 | 9759.329 | 9.186 | 5.1084 |  |  | 244 | 9759.329 | 9.186 | 5.1101
245 | 9799.326 | 9.1901 | 5.1029 |  |  | 245 | 9799.326 | 9.1901 | 5.1073 |  |  | 245 | 9799.326 | 9.1901 | 5.109
246 | 9839.324 | 9.1941 | 5.1018 |  |  | 246 | 9839.324 | 9.1941 | 5.1062 |  |  | 246 | 9839.324 | 9.1941 | 5.1079
247 | 9879.321 | 9.1982 | 5.1007 |  |  | 247 | 9879.321 | 9.1982 | 5.1051 |  |  | 247 | 9879.321 | 9.1982 | 5.1069
248 | 9919.318 | 9.2022 | 5.0997 |  |  | 248 | 9919.318 | 9.2022 | 5.104 |  |  | 248 | 9919.318 | 9.2022 | 5.1058
249 | 9959.315 | 9.2063 | 5.0986 |  |  | 249 | 9959.315 | 9.2063 | 5.103 |  |  | 249 | 9959.315 | 9.2063 | 5.1047
250 | 9999.313 | 9.2103 | 5.0976 |  |  | 250 | 9999.313 | 9.2103 | 5.1019 |  |  | 250 | 9999.313 | 9.2103 | 5.1037
251 | 10039.31 | 9.2143 | 5.0965 |  |  | 251 | 10039.31 | 9.2143 | 5.1009 |  |  | 251 | 10039.31 | 9.2143 | 5.1026
252 | 10079.307 | 9.2182 | 5.0955 |  |  | 252 | 10079.307 | 9.2182 | 5.0999 |  |  | 252 | 10079.307 | 9.2182 | 5.1016
253 | 10119.304 | 9.2222 | 5.0945 |  |  | 253 | 10119.304 | 9.2222 | 5.0988 |  |  | 253 | 10119.304 | 9.2222 | 5.1006
254 | 10159.302 | 9.2261 | 5.0935 |  |  | 254 | 10159.302 | 9.2261 | 5.0978 |  |  | 254 | 10159.302 | 9.2261 | 5.0996
255 | 10199.299 | 9.2301 | 5.0925 |  |  | 255 | 10199.299 | 9.2301 | 5.0968 |  |  | 255 | 10199.299 | 9.2301 | 5.0986
256 | 10239.296 | 9.234 | 5.0915 |  |  | 256 | 10239.296 | 9.234 | 5.0958 |  |  | 256 | 10239.296 | 9.234 | 5.0976
257 | 10279.293 | 9.2379 | 5.0905 |  |  | 257 | 10279.293 | 9.2379 | 5.0948 |  |  | 257 | 10279.293 | 9.2379 | 5.0966
258 | 10319.291 | 9.2418 | 5.0895 |  |  | 258 | 10319.291 | 9.2418 | 5.0939 |  |  | 258 | 10319.291 | 9.2418 | 5.0956
259 | 10359.288 | 9.2456 | 5.0886 |  |  | 259 | 10359.288 | 9.2456 | 5.0929 |  |  | 259 | 10359.288 | 9.2456 | 5.0946
260 | 10399.285 | 9.2495 | 5.0876 |  |  | 260 | 10399.285 | 9.2495 | 5.0919 |  |  | 260 | 10399.285 | 9.2495 | 5.0936
261 | 10439.282 | 9.2533 | 5.0867 |  |  | 261 | 10439.282 | 9.2533 | 5.091 |  |  | 261 | 10439.282 | 9.2533 | 5.0927
262 | 10479.28 | 9.2572 | 5.0857 |  |  | 262 | 10479.28 | 9.2572 | 5.09 |  |  | 262 | 10479.28 | 9.2572 | 5.0917
263 | 10519.277 | 9.261 | 5.0848 |  |  | 263 | 10519.277 | 9.261 | 5.0891 |  |  | 263 | 10519.277 | 9.261 | 5.0908
264 | 10559.274 | 9.2648 | 5.0839 |  |  | 264 | 10559.274 | 9.2648 | 5.0882 |  |  | 264 | 10559.274 | 9.2648 | 5.0898
265 | 10599.271 | 9.2685 | 5.0829 |  |  | 265 | 10599.271 | 9.2685 | 5.0872 |  |  | 265 | 10599.271 | 9.2685 | 5.0889
266 | 10639.269 | 9.2723 | 5.082 |  |  | 266 | 10639.269 | 9.2723 | 5.0863 |  |  | 266 | 10639.269 | 9.2723 | 5.088
267 | 10679.266 | 9.2761 | 5.0811 |  |  | 267 | 10679.266 | 9.2761 | 5.0854 |  |  | 267 | 10679.266 | 9.2761 | 5.0871
268 | 10719.263 | 9.2798 | 5.0802 |  |  | 268 | 10719.263 | 9.2798 | 5.0845 |  |  | 268 | 10719.263 | 9.2798 | 5.0862
269 | 10759.26 | 9.2835 | 5.0793 |  |  | 269 | 10759.26 | 9.2835 | 5.0836 |  |  | 269 | 10759.26 | 9.2835 | 5.0853
270 | 10799.258 | 9.2872 | 5.0785 |  |  | 270 | 10799.258 | 9.2872 | 5.0827 |  |  | 270 | 10799.258 | 9.2872 | 5.0844
271 | 10839.255 | 9.2909 | 5.0776 |  |  | 271 | 10839.255 | 9.2909 | 5.0818 |  |  | 271 | 10839.255 | 9.2909 | 5.0835
272 | 10879.252 | 9.2946 | 5.0767 |  |  | 272 | 10879.252 | 9.2946 | 5.081 |  |  | 272 | 10879.252 | 9.2946 | 5.0826
273 | 10919.249 | 9.2983 | 5.0759 |  |  | 273 | 10919.249 | 9.2983 | 5.0801 |  |  | 273 | 10919.249 | 9.2983 | 5.0818
274 | 10959.247 | 9.3019 | 5.075 |  |  | 274 | 10959.247 | 9.3019 | 5.0792 |  |  | 274 | 10959.247 | 9.3019 | 5.0809
275 | 10999.244 | 9.3056 | 5.0742 |  |  | 275 | 10999.244 | 9.3056 | 5.0784 |  |  | 275 | 10999.244 | 9.3056 | 5.08
276 | 11039.241 | 9.3092 | 5.0733 |  |  | 276 | 11039.241 | 9.3092 | 5.0775 |  |  | 276 | 11039.241 | 9.3092 | 5.0792
277 | 11079.238 | 9.3128 | 5.0725 |  |  | 277 | 11079.238 | 9.3128 | 5.0767 |  |  | 277 | 11079.238 | 9.3128 | 5.0784
278 | 11119.236 | 9.3164 | 5.0717 |  |  | 278 | 11119.236 | 9.3164 | 5.0759 |  |  | 278 | 11119.236 | 9.3164 | 5.0775
279 | 11159.233 | 9.32 | 5.0709 |  |  | 279 | 11159.233 | 9.32 | 5.0751 |  |  | 279 | 11159.233 | 9.32 | 5.0767
280 | 11199.23 | 9.3236 | 5.0701 |  |  | 280 | 11199.23 | 9.3236 | 5.0742 |  |  | 280 | 11199.23 | 9.3236 | 5.0759
281 | 11239.227 | 9.3272 | 5.0692 |  |  | 281 | 11239.227 | 9.3272 | 5.0734 |  |  | 281 | 11239.227 | 9.3272 | 5.0751
282 | 11279.225 | 9.3307 | 5.0684 |  |  | 282 | 11279.225 | 9.3307 | 5.0726 |  |  | 282 | 11279.225 | 9.3307 | 5.0743
283 | 11319.222 | 9.3343 | 5.0677 |  |  | 283 | 11319.222 | 9.3343 | 5.0718 |  |  | 283 | 11319.222 | 9.3343 | 5.0735
284 | 11359.219 | 9.3378 | 5.0669 |  |  | 284 | 11359.219 | 9.3378 | 5.071 |  |  | 284 | 11359.219 | 9.3378 | 5.0727
285 | 11399.216 | 9.3413 | 5.0661 |  |  | 285 | 11399.216 | 9.3413 | 5.0702 |  |  | 285 | 11399.216 | 9.3413 | 5.0719
286 | 11439.214 | 9.3448 | 5.0653 |  |  | 286 | 11439.214 | 9.3448 | 5.0695 |  |  | 286 | 11439.214 | 9.3448 | 5.0711
287 | 11479.211 | 9.3483 | 5.0646 |  |  | 287 | 11479.211 | 9.3483 | 5.0687 |  |  | 287 | 11479.211 | 9.3483 | 5.0703
288 | 11519.208 | 9.3518 | 5.0638 |  |  | 288 | 11519.208 | 9.3518 | 5.0679 |  |  | 288 | 11519.208 | 9.3518 | 5.0695
289 | 11559.205 | 9.3552 | 5.063 |  |  | 289 | 11559.205 | 9.3552 | 5.0672 |  |  | 289 | 11559.205 | 9.3552 | 5.0688
290 | 11599.203 | 9.3587 | 5.0623 |  |  | 290 | 11599.203 | 9.3587 | 5.0664 |  |  | 290 | 11599.203 | 9.3587 | 5.068
291 | 11639.2 | 9.3621 | 5.0616 |  |  | 291 | 11639.2 | 9.3621 | 5.0657 |  |  | 291 | 11639.2 | 9.3621 | 5.0673
292 | 11679.197 | 9.3656 | 5.0608 |  |  | 292 | 11679.197 | 9.3656 | 5.0649 |  |  | 292 | 11679.197 | 9.3656 | 5.0665
293 | 11719.194 | 9.369 | 5.0601 |  |  | 293 | 11719.194 | 9.369 | 5.0642 |  |  | 293 | 11719.194 | 9.369 | 5.0658
294 | 11759.192 | 9.3724 | 5.0594 |  |  | 294 | 11759.192 | 9.3724 | 5.0635 |  |  | 294 | 11759.192 | 9.3724 | 5.0651
295 | 11799.189 | 9.3758 | 5.0587 |  |  | 295 | 11799.189 | 9.3758 | 5.0627 |  |  | 295 | 11799.189 | 9.3758 | 5.0643
296 | 11839.186 | 9.3792 | 5.0579 |  |  | 296 | 11839.186 | 9.3792 | 5.062 |  |  | 296 | 11839.186 | 9.3792 | 5.0636
297 | 11879.183 | 9.3825 | 5.0572 |  |  | 297 | 11879.183 | 9.3825 | 5.0613 |  |  | 297 | 11879.183 | 9.3825 | 5.0629
298 | 11919.181 | 9.3859 | 5.0565 |  |  | 298 | 11919.181 | 9.3859 | 5.0606 |  |  | 298 | 11919.181 | 9.3859 | 5.0622
299 | 11959.178 | 9.3893 | 5.0558 |  |  | 299 | 11959.178 | 9.3893 | 5.0599 |  |  | 299 | 11959.178 | 9.3893 | 5.0615
300 | 11999.175 | 9.3926 | 5.0551 |  |  | 300 | 11999.175 | 9.3926 | 5.0592 |  |  | 300 | 11999.175 | 9.3926 | 5.0608
301 | 12039.172 | 9.3959 | 5.0545 |  |  | 301 | 12039.172 | 9.3959 | 5.0585 |  |  | 301 | 12039.172 | 9.3959 | 5.0601
302 | 12079.17 | 9.3992 | 5.0538 |  |  | 302 | 12079.17 | 9.3992 | 5.0578 |  |  | 302 | 12079.17 | 9.3992 | 5.0594
303 | 12119.167 | 9.4025 | 5.0531 |  |  | 303 | 12119.167 | 9.4025 | 5.0571 |  |  | 303 | 12119.167 | 9.4025 | 5.0587
304 | 12159.164 | 9.4058 | 5.0524 |  |  | 304 | 12159.164 | 9.4058 | 5.0565 |  |  | 304 | 12159.164 | 9.4058 | 5.058
305 | 12199.161 | 9.4091 | 5.0518 |  |  | 305 | 12199.161 | 9.4091 | 5.0558 |  |  | 305 | 12199.161 | 9.4091 | 5.0574
306 | 12239.159 | 9.4124 | 5.0511 |  |  | 306 | 12239.159 | 9.4124 | 5.0551 |  |  | 306 | 12239.159 | 9.4124 | 5.0567
307 | 12279.156 | 9.4157 | 5.0505 |  |  | 307 | 12279.156 | 9.4157 | 5.0545 |  |  | 307 | 12279.156 | 9.4157 | 5.056
308 | 12319.153 | 9.4189 | 5.0498 |  |  | 308 | 12319.153 | 9.4189 | 5.0538 |  |  | 308 | 12319.153 | 9.4189 | 5.0554
309 | 12359.15 | 9.4222 | 5.0492 |  |  | 309 | 12359.15 | 9.4222 | 5.0532 |  |  | 309 | 12359.15 | 9.4222 | 5.0547
310 | 12399.148 | 9.4254 | 5.0486 |  |  | 310 | 12399.148 | 9.4254 | 5.0525 |  |  | 310 | 12399.148 | 9.4254 | 5.0541
311 | 12439.145 | 9.4286 | 5.0479 |  |  | 311 | 12439.145 | 9.4286 | 5.0519 |  |  | 311 | 12439.145 | 9.4286 | 5.0534
312 | 12479.142 | 9.4318 | 5.0473 |  |  | 312 | 12479.142 | 9.4318 | 5.0513 |  |  | 312 | 12479.142 | 9.4318 | 5.0528
313 | 12519.139 | 9.435 | 5.0467 |  |  | 313 | 12519.139 | 9.435 | 5.0506 |  |  | 313 | 12519.139 | 9.435 | 5.0522
314 | 12559.137 | 9.4382 | 5.0461 |  |  | 314 | 12559.137 | 9.4382 | 5.05 |  |  | 314 | 12559.137 | 9.4382 | 5.0515
315 | 12599.134 | 9.4414 | 5.0454 |  |  | 315 | 12599.134 | 9.4414 | 5.0494 |  |  | 315 | 12599.134 | 9.4414 | 5.0509
316 | 12639.131 | 9.4446 | 5.0448 |  |  | 316 | 12639.131 | 9.4446 | 5.0488 |  |  | 316 | 12639.131 | 9.4446 | 5.0503
317 | 12679.128 | 9.4477 | 5.0442 |  |  | 317 | 12679.128 | 9.4477 | 5.0482 |  |  | 317 | 12679.128 | 9.4477 | 5.0497
318 | 12719.126 | 9.4509 | 5.0436 |  |  | 318 | 12719.126 | 9.4509 | 5.0476 |  |  | 318 | 12719.126 | 9.4509 | 5.0491
319 | 12759.123 | 9.454 | 5.043 |  |  | 319 | 12759.123 | 9.454 | 5.047 |  |  | 319 | 12759.123 | 9.454 | 5.0485
320 | 12799.12 | 9.4571 | 5.0425 |  |  | 320 | 12799.12 | 9.4571 | 5.0464 |  |  | 320 | 12799.12 | 9.4571 | 5.0479
321 | 12839.117 | 9.4603 | 5.0419 |  |  | 321 | 12839.117 | 9.4603 | 5.0458 |  |  | 321 | 12839.117 | 9.4603 | 5.0473
322 | 12879.115 | 9.4634 | 5.0413 |  |  | 322 | 12879.115 | 9.4634 | 5.0452 |  |  | 322 | 12879.115 | 9.4634 | 5.0467
323 | 12919.112 | 9.4665 | 5.0407 |  |  | 323 | 12919.112 | 9.4665 | 5.0446 |  |  | 323 | 12919.112 | 9.4665 | 5.0461
324 | 12959.109 | 9.4696 | 5.0402 |  |  | 324 | 12959.109 | 9.4696 | 5.044 |  |  | 324 | 12959.109 | 9.4696 | 5.0455
325 | 12999.106 | 9.4726 | 5.0396 |  |  | 325 | 12999.106 | 9.4726 | 5.0435 |  |  | 325 | 12999.106 | 9.4726 | 5.045
326 | 13039.104 | 9.4757 | 5.039 |  |  | 326 | 13039.104 | 9.4757 | 5.0429 |  |  | 326 | 13039.104 | 9.4757 | 5.0444
327 | 13079.101 | 9.4788 | 5.0385 |  |  | 327 | 13079.101 | 9.4788 | 5.0423 |  |  | 327 | 13079.101 | 9.4788 | 5.0438
328 | 13119.098 | 9.4818 | 5.0379 |  |  | 328 | 13119.098 | 9.4818 | 5.0418 |  |  | 328 | 13119.098 | 9.4818 | 5.0433
329 | 13159.095 | 9.4849 | 5.0374 |  |  | 329 | 13159.095 | 9.4849 | 5.0412 |  |  | 329 | 13159.095 | 9.4849 | 5.0427
330 | 13199.093 | 9.4879 | 5.0368 |  |  | 330 | 13199.093 | 9.4879 | 5.0407 |  |  | 330 | 13199.093 | 9.4879 | 5.0422
331 | 13239.09 | 9.4909 | 5.0363 |  |  | 331 | 13239.09 | 9.4909 | 5.0401 |  |  | 331 | 13239.09 | 9.4909 | 5.0416
332 | 13279.087 | 9.4939 | 5.0357 |  |  | 332 | 13279.087 | 9.4939 | 5.0396 |  |  | 332 | 13279.087 | 9.4939 | 5.0411
333 | 13319.084 | 9.497 | 5.0352 |  |  | 333 | 13319.084 | 9.497 | 5.039 |  |  | 333 | 13319.084 | 9.497 | 5.0405
334 | 13359.082 | 9.5 | 5.0347 |  |  | 334 | 13359.082 | 9.5 | 5.0385 |  |  | 334 | 13359.082 | 9.5 | 5.04
335 | 13399.079 | 9.5029 | 5.0342 |  |  | 335 | 13399.079 | 9.5029 | 5.038 |  |  | 335 | 13399.079 | 9.5029 | 5.0394
336 | 13439.076 | 9.5059 | 5.0336 |  |  | 336 | 13439.076 | 9.5059 | 5.0374 |  |  | 336 | 13439.076 | 9.5059 | 5.0389
337 | 13479.073 | 9.5089 | 5.0331 |  |  | 337 | 13479.073 | 9.5089 | 5.0369 |  |  | 337 | 13479.073 | 9.5089 | 5.0384
338 | 13519.071 | 9.5119 | 5.0326 |  |  | 338 | 13519.071 | 9.5119 | 5.0364 |  |  | 338 | 13519.071 | 9.5119 | 5.0379
339 | 13559.068 | 9.5148 | 5.0321 |  |  | 339 | 13559.068 | 9.5148 | 5.0359 |  |  | 339 | 13559.068 | 9.5148 | 5.0374
340 | 13599.065 | 9.5178 | 5.0316 |  |  | 340 | 13599.065 | 9.5178 | 5.0354 |  |  | 340 | 13599.065 | 9.5178 | 5.0368
341 | 13639.062 | 9.5207 | 5.0311 |  |  | 341 | 13639.062 | 9.5207 | 5.0349 |  |  | 341 | 13639.062 | 9.5207 | 5.0363
342 | 13679.06 | 9.5236 | 5.0306 |  |  | 342 | 13679.06 | 9.5236 | 5.0344 |  |  | 342 | 13679.06 | 9.5236 | 5.0358
343 | 13719.057 | 9.5265 | 5.0301 |  |  | 343 | 13719.057 | 9.5265 | 5.0339 |  |  | 343 | 13719.057 | 9.5265 | 5.0353
344 | 13759.054 | 9.5295 | 5.0296 |  |  | 344 | 13759.054 | 9.5295 | 5.0334 |  |  | 344 | 13759.054 | 9.5295 | 5.0348
345 | 13799.051 | 9.5324 | 5.0292 |  |  | 345 | 13799.051 | 9.5324 | 5.0329 |  |  | 345 | 13799.051 | 9.5324 | 5.0343
346 | 13839.049 | 9.5352 | 5.0287 |  |  | 346 | 13839.049 | 9.5352 | 5.0324 |  |  | 346 | 13839.049 | 9.5352 | 5.0338
347 | 13879.046 | 9.5381 | 5.0282 |  |  | 347 | 13879.046 | 9.5381 | 5.0319 |  |  | 347 | 13879.046 | 9.5381 | 5.0334
348 | 13919.043 | 9.541 | 5.0277 |  |  | 348 | 13919.043 | 9.541 | 5.0314 |  |  | 348 | 13919.043 | 9.541 | 5.0329
349 | 13959.04 | 9.5439 | 5.0272 |  |  | 349 | 13959.04 | 9.5439 | 5.0309 |  |  | 349 | 13959.04 | 9.5439 | 5.0324
350 | 13999.038 | 9.5467 | 5.0268 |  |  | 350 | 13999.038 | 9.5467 | 5.0305 |  |  | 350 | 13999.038 | 9.5467 | 5.0319
351 | 14039.035 | 9.5496 | 5.0263 |  |  | 351 | 14039.035 | 9.5496 | 5.03 |  |  | 351 | 14039.035 | 9.5496 | 5.0314
352 | 14079.032 | 9.5524 | 5.0259 |  |  | 352 | 14079.032 | 9.5524 | 5.0295 |  |  | 352 | 14079.032 | 9.5524 | 5.031
353 | 14119.029 | 9.5553 | 5.0254 |  |  | 353 | 14119.029 | 9.5553 | 5.029 |  |  | 353 | 14119.029 | 9.5553 | 5.0305
354 | 14159.027 | 9.5581 | 5.025 |  |  | 354 | 14159.027 | 9.5581 | 5.0286 |  |  | 354 | 14159.027 | 9.5581 | 5.03
355 | 14199.024 | 9.5609 | 5.0245 |  |  | 355 | 14199.024 | 9.5609 | 5.0281 |  |  | 355 | 14199.024 | 9.5609 | 5.0296
356 | 14239.021 | 9.5637 | 5.0241 |  |  | 356 | 14239.021 | 9.5637 | 5.0277 |  |  | 356 | 14239.021 | 9.5637 | 5.0291
357 | 14279.018 | 9.5665 | 5.0236 |  |  | 357 | 14279.018 | 9.5665 | 5.0272 |  |  | 357 | 14279.018 | 9.5665 | 5.0287
358 | 14319.016 | 9.5693 | 5.0232 |  |  | 358 | 14319.016 | 9.5693 | 5.0268 |  |  | 358 | 14319.016 | 9.5693 | 5.0282
359 | 14359.013 | 9.5721 | 5.0227 |  |  | 359 | 14359.013 | 9.5721 | 5.0263 |  |  | 359 | 14359.013 | 9.5721 | 5.0278
360 | 14399.01 | 9.5749 | 5.0223 |  |  | 360 | 14399.01 | 9.5749 | 5.0259 |  |  | 360 | 14399.01 | 9.5749 | 5.0273
361 | 14439.007 | 9.5777 | 5.0219 |  |  | 361 | 14439.007 | 9.5777 | 5.0255 |  |  | 361 | 14439.007 | 9.5777 | 5.0269
362 | 14479.005 | 9.5805 | 5.0214 |  |  | 362 | 14479.005 | 9.5805 | 5.025 |  |  | 362 | 14479.005 | 9.5805 | 5.0265
363 | 14519.002 | 9.5832 | 5.021 |  |  | 363 | 14519.002 | 9.5832 | 5.0246 |  |  | 363 | 14519.002 | 9.5832 | 5.026
364 | 14558.999 | 9.586 | 5.0206 |  |  | 364 | 14558.999 | 9.586 | 5.0242 |  |  | 364 | 14558.999 | 9.586 | 5.0256
365 | 14598.996 | 9.5887 | 5.0202 |  |  | 365 | 14598.996 | 9.5887 | 5.0237 |  |  | 365 | 14598.996 | 9.5887 | 5.0252
366 | 14638.994 | 9.5914 | 5.0198 |  |  | 366 | 14638.994 | 9.5914 | 5.0233 |  |  | 366 | 14638.994 | 9.5914 | 5.0247
367 | 14678.991 | 9.5942 | 5.0194 |  |  | 367 | 14678.991 | 9.5942 | 5.0229 |  |  | 367 | 14678.991 | 9.5942 | 5.0243
368 | 14718.988 | 9.5969 | 5.019 |  |  | 368 | 14718.988 | 9.5969 | 5.0225 |  |  | 368 | 14718.988 | 9.5969 | 5.0239
369 | 14758.985 | 9.5996 | 5.0186 |  |  | 369 | 14758.985 | 9.5996 | 5.0221 |  |  | 369 | 14758.985 | 9.5996 | 5.0235
370 | 14798.983 | 9.6023 | 5.0182 |  |  | 370 | 14798.983 | 9.6023 | 5.0217 |  |  | 370 | 14798.983 | 9.6023 | 5.0231
371 | 14838.98 | 9.605 | 5.0178 |  |  | 371 | 14838.98 | 9.605 | 5.0212 |  |  | 371 | 14838.98 | 9.605 | 5.0227
372 | 14878.977 | 9.6077 | 5.0174 |  |  | 372 | 14878.977 | 9.6077 | 5.0208 |  |  | 372 | 14878.977 | 9.6077 | 5.0223
373 | 14918.974 | 9.6104 | 5.017 |  |  | 373 | 14918.974 | 9.6104 | 5.0204 |  |  | 373 | 14918.974 | 9.6104 | 5.0219
374 | 14958.972 | 9.6131 | 5.0166 |  |  | 374 | 14958.972 | 9.6131 | 5.02 |  |  | 374 | 14958.972 | 9.6131 | 5.0215
375 | 14998.969 | 9.6157 | 5.0162 |  |  | 375 | 14998.969 | 9.6157 | 5.0196 |  |  | 375 | 14998.969 | 9.6157 | 5.0211
376 | 15038.966 | 9.6184 | 5.0158 |  |  | 376 | 15038.966 | 9.6184 | 5.0192 |  |  | 376 | 15038.966 | 9.6184 | 5.0207
377 | 15078.963 | 9.6211 | 5.0154 |  |  | 377 | 15078.963 | 9.6211 | 5.0189 |  |  | 377 | 15078.963 | 9.6211 | 5.0203
378 | 15118.961 | 9.6237 | 5.015 |  |  | 378 | 15118.961 | 9.6237 | 5.0185 |  |  | 378 | 15118.961 | 9.6237 | 5.0199
379 | 15158.958 | 9.6263 | 5.0147 |  |  | 379 | 15158.958 | 9.6263 | 5.0181 |  |  | 379 | 15158.958 | 9.6263 | 5.0195
380 | 15198.955 | 9.629 | 5.0143 |  |  | 380 | 15198.955 | 9.629 | 5.0177 |  |  | 380 | 15198.955 | 9.629 | 5.0191
381 | 15238.952 | 9.6316 | 5.0139 |  |  | 381 | 15238.952 | 9.6316 | 5.0173 |  |  | 381 | 15238.952 | 9.6316 | 5.0187
382 | 15278.95 | 9.6342 | 5.0136 |  |  | 382 | 15278.95 | 9.6342 | 5.0169 |  |  | 382 | 15278.95 | 9.6342 | 5.0184
383 | 15318.947 | 9.6368 | 5.0132 |  |  | 383 | 15318.947 | 9.6368 | 5.0166 |  |  | 383 | 15318.947 | 9.6368 | 5.018
384 | 15358.944 | 9.6395 | 5.0128 |  |  | 384 | 15358.944 | 9.6395 | 5.0162 |  |  | 384 | 15358.944 | 9.6395 | 5.0176
385 | 15398.941 | 9.6421 | 5.0125 |  |  | 385 | 15398.941 | 9.6421 | 5.0158 |  |  | 385 | 15398.941 | 9.6421 | 5.0172
386 | 15438.939 | 9.6446 | 5.0121 |  |  | 386 | 15438.939 | 9.6446 | 5.0155 |  |  | 386 | 15438.939 | 9.6446 | 5.0169
387 | 15478.936 | 9.6472 | 5.0118 |  |  | 387 | 15478.936 | 9.6472 | 5.0151 |  |  | 387 | 15478.936 | 9.6472 | 5.0165
388 | 15518.933 | 9.6498 | 5.0114 |  |  | 388 | 15518.933 | 9.6498 | 5.0147 |  |  | 388 | 15518.933 | 9.6498 | 5.0161
389 | 15558.93 | 9.6524 | 5.011 |  |  | 389 | 15558.93 | 9.6524 | 5.0144 |  |  | 389 | 15558.93 | 9.6524 | 5.0158
390 | 15598.928 | 9.655 | 5.0107 |  |  | 390 | 15598.928 | 9.655 | 5.014 |  |  | 390 | 15598.928 | 9.655 | 5.0154
391 | 15638.925 | 9.6575 | 5.0104 |  |  | 391 | 15638.925 | 9.6575 | 5.0137 |  |  | 391 | 15638.925 | 9.6575 | 5.0151
392 | 15678.922 | 9.6601 | 5.01 |  |  | 392 | 15678.922 | 9.6601 | 5.0133 |  |  | 392 | 15678.922 | 9.6601 | 5.0147
393 | 15718.919 | 9.6626 | 5.0097 |  |  | 393 | 15718.919 | 9.6626 | 5.013 |  |  | 393 | 15718.919 | 9.6626 | 5.0144
394 | 15758.917 | 9.6652 | 5.0093 |  |  | 394 | 15758.917 | 9.6652 | 5.0126 |  |  | 394 | 15758.917 | 9.6652 | 5.014
395 | 15798.914 | 9.6677 | 5.009 |  |  | 395 | 15798.914 | 9.6677 | 5.0123 |  |  | 395 | 15798.914 | 9.6677 | 5.0137
396 | 15838.911 | 9.6702 | 5.0087 |  |  | 396 | 15838.911 | 9.6702 | 5.0119 |  |  | 396 | 15838.911 | 9.6702 | 5.0133
397 | 15878.908 | 9.6727 | 5.0083 |  |  | 397 | 15878.908 | 9.6727 | 5.0116 |  |  | 397 | 15878.908 | 9.6727 | 5.013
398 | 15918.906 | 9.6753 | 5.008 |  |  | 398 | 15918.906 | 9.6753 | 5.0113 |  |  | 398 | 15918.906 | 9.6753 | 5.0127
399 | 15958.903 | 9.6778 | 5.0077 |  |  | 399 | 15958.903 | 9.6778 | 5.0109 |  |  | 399 | 15958.903 | 9.6778 | 5.0123
400 | 15998.9 | 9.6803 | 5.0074 |  |  | 400 | 15998.9 | 9.6803 | 5.0106 |  |  | 400 | 15998.9 | 9.6803 | 5.012
401 | 16038.897 | 9.6828 | 5.007 |  |  | 401 | 16038.897 | 9.6828 | 5.0103 |  |  | 401 | 16038.897 | 9.6828 | 5.0117
402 | 16078.895 | 9.6853 | 5.0067 |  |  | 402 | 16078.895 | 9.6853 | 5.0099 |  |  | 402 | 16078.895 | 9.6853 | 5.0113
403 | 16118.892 | 9.6877 | 5.0064 |  |  | 403 | 16118.892 | 9.6877 | 5.0096 |  |  | 403 | 16118.892 | 9.6877 | 5.011
404 | 16158.889 | 9.6902 | 5.0061 |  |  | 404 | 16158.889 | 9.6902 | 5.0093 |  |  | 404 | 16158.889 | 9.6902 | 5.0107
405 | 16198.886 | 9.6927 | 5.0058 |  |  | 405 | 16198.886 | 9.6927 | 5.009 |  |  | 405 | 16198.886 | 9.6927 | 5.0104
406 | 16238.884 | 9.6952 | 5.0055 |  |  | 406 | 16238.884 | 9.6952 | 5.0086 |  |  | 406 | 16238.884 | 9.6952 | 5.01
407 | 16278.881 | 9.6976 | 5.0052 |  |  | 407 | 16278.881 | 9.6976 | 5.0083 |  |  | 407 | 16278.881 | 9.6976 | 5.0097
408 | 16318.878 | 9.7001 | 5.0049 |  |  | 408 | 16318.878 | 9.7001 | 5.008 |  |  | 408 | 16318.878 | 9.7001 | 5.0094
409 | 16358.875 | 9.7025 | 5.0045 |  |  | 409 | 16358.875 | 9.7025 | 5.0077 |  |  | 409 | 16358.875 | 9.7025 | 5.0091
410 | 16398.873 | 9.705 | 5.0042 |  |  | 410 | 16398.873 | 9.705 | 5.0074 |  |  | 410 | 16398.873 | 9.705 | 5.0088
411 | 16438.87 | 9.7074 | 5.0039 |  |  | 411 | 16438.87 | 9.7074 | 5.0071 |  |  | 411 | 16438.87 | 9.7074 | 5.0085
412 | 16478.867 | 9.7098 | 5.0037 |  |  | 412 | 16478.867 | 9.7098 | 5.0068 |  |  | 412 | 16478.867 | 9.7098 | 5.0082
413 | 16518.864 | 9.7123 | 5.0034 |  |  | 413 | 16518.864 | 9.7123 | 5.0064 |  |  | 413 | 16518.864 | 9.7123 | 5.0078
414 | 16558.862 | 9.7147 | 5.0031 |  |  | 414 | 16558.862 | 9.7147 | 5.0061 |  |  | 414 | 16558.862 | 9.7147 | 5.0075
415 | 16598.859 | 9.7171 | 5.0028 |  |  | 415 | 16598.859 | 9.7171 | 5.0058 |  |  | 415 | 16598.859 | 9.7171 | 5.0072
416 | 16638.856 | 9.7195 | 5.0025 |  |  | 416 | 16638.856 | 9.7195 | 5.0055 |  |  | 416 | 16638.856 | 9.7195 | 5.0069
417 | 16678.853 | 9.7219 | 5.0022 |  |  | 417 | 16678.853 | 9.7219 | 5.0052 |  |  | 417 | 16678.853 | 9.7219 | 5.0066
418 | 16718.851 | 9.7243 | 5.0019 |  |  | 418 | 16718.851 | 9.7243 | 5.0049 |  |  | 418 | 16718.851 | 9.7243 | 5.0064
419 | 16758.848 | 9.7267 | 5.0016 |  |  | 419 | 16758.848 | 9.7267 | 5.0047 |  |  | 419 | 16758.848 | 9.7267 | 5.0061
420 | 16798.845 | 9.7291 | 5.0013 |  |  | 420 | 16798.845 | 9.7291 | 5.0044 |  |  | 420 | 16798.845 | 9.7291 | 5.0058
421 | 16838.842 | 9.7314 | 5.0011 |  |  | 421 | 16838.842 | 9.7314 | 5.0041 |  |  | 421 | 16838.842 | 9.7314 | 5.0055
422 | 16878.84 | 9.7338 | 5.0008 |  |  | 422 | 16878.84 | 9.7338 | 5.0038 |  |  | 422 | 16878.84 | 9.7338 | 5.0052
423 | 16918.837 | 9.7362 | 5.0005 |  |  | 423 | 16918.837 | 9.7362 | 5.0035 |  |  | 423 | 16918.837 | 9.7362 | 5.0049
424 | 16958.834 | 9.7385 | 5.0002 |  |  | 424 | 16958.834 | 9.7385 | 5.0032 |  |  | 424 | 16958.834 | 9.7385 | 5.0046
425 | 16998.831 | 9.7409 | 5 |  |  | 425 | 16998.831 | 9.7409 | 5.0029 |  |  | 425 | 16998.831 | 9.7409 | 5.0043
426 | 17038.829 | 9.7433 | 4.9997 |  |  | 426 | 17038.829 | 9.7433 | 5.0027 |  |  | 426 | 17038.829 | 9.7433 | 5.0041
427 | 17078.826 | 9.7456 | 4.9994 |  |  | 427 | 17078.826 | 9.7456 | 5.0024 |  |  | 427 | 17078.826 | 9.7456 | 5.0038
428 | 17118.823 | 9.7479 | 4.9992 |  |  | 428 | 17118.823 | 9.7479 | 5.0021 |  |  | 428 | 17118.823 | 9.7479 | 5.0035
429 | 17158.82 | 9.7503 | 4.9989 |  |  | 429 | 17158.82 | 9.7503 | 5.0018 |  |  | 429 | 17158.82 | 9.7503 | 5.0032
430 | 17198.818 | 9.7526 | 4.9986 |  |  | 430 | 17198.818 | 9.7526 | 5.0015 |  |  | 430 | 17198.818 | 9.7526 | 5.003
431 | 17238.815 | 9.7549 | 4.9984 |  |  | 431 | 17238.815 | 9.7549 | 5.0013 |  |  | 431 | 17238.815 | 9.7549 | 5.0027
432 | 17278.812 | 9.7572 | 4.9981 |  |  | 432 | 17278.812 | 9.7572 | 5.001 |  |  | 432 | 17278.812 | 9.7572 | 5.0024
433 | 17318.809 | 9.7595 | 4.9979 |  |  | 433 | 17318.809 | 9.7595 | 5.0007 |  |  | 433 | 17318.809 | 9.7595 | 5.0021
434 | 17358.807 | 9.7619 | 4.9976 |  |  | 434 | 17358.807 | 9.7619 | 5.0005 |  |  | 434 | 17358.807 | 9.7619 | 5.0019
435 | 17398.804 | 9.7642 | 4.9973 |  |  | 435 | 17398.804 | 9.7642 | 5.0002 |  |  | 435 | 17398.804 | 9.7642 | 5.0016
436 | 17438.801 | 9.7665 | 4.9971 |  |  | 436 | 17438.801 | 9.7665 | 4.9999 |  |  | 436 | 17438.801 | 9.7665 | 5.0014
437 | 17478.798 | 9.7687 | 4.9968 |  |  | 437 | 17478.798 | 9.7687 | 4.9997 |  |  | 437 | 17478.798 | 9.7687 | 5.0011
438 | 17518.796 | 9.771 | 4.9966 |  |  | 438 | 17518.796 | 9.771 | 4.9994 |  |  | 438 | 17518.796 | 9.771 | 5.0008
439 | 17558.793 | 9.7733 | 4.9963 |  |  | 439 | 17558.793 | 9.7733 | 4.9992 |  |  | 439 | 17558.793 | 9.7733 | 5.0006
440 | 17598.79 | 9.7756 | 4.9961 |  |  | 440 | 17598.79 | 9.7756 | 4.9989 |  |  | 440 | 17598.79 | 9.7756 | 5.0003
441 | 17638.787 | 9.7779 | 4.9958 |  |  | 441 | 17638.787 | 9.7779 | 4.9986 |  |  | 441 | 17638.787 | 9.7779 | 5.0001
442 | 17678.785 | 9.7801 | 4.9956 |  |  | 442 | 17678.785 | 9.7801 | 4.9984 |  |  | 442 | 17678.785 | 9.7801 | 4.9998
443 | 17718.782 | 9.7824 | 4.9954 |  |  | 443 | 17718.782 | 9.7824 | 4.9981 |  |  | 443 | 17718.782 | 9.7824 | 4.9996
444 | 17758.779 | 9.7846 | 4.9951 |  |  | 444 | 17758.779 | 9.7846 | 4.9979 |  |  | 444 | 17758.779 | 9.7846 | 4.9993
445 | 17798.776 | 9.7869 | 4.9949 |  |  | 445 | 17798.776 | 9.7869 | 4.9976 |  |  | 445 | 17798.776 | 9.7869 | 4.9991
446 | 17838.774 | 9.7891 | 4.9946 |  |  | 446 | 17838.774 | 9.7891 | 4.9974 |  |  | 446 | 17838.774 | 9.7891 | 4.9988
447 | 17878.771 | 9.7914 | 4.9944 |  |  | 447 | 17878.771 | 9.7914 | 4.9971 |  |  | 447 | 17878.771 | 9.7914 | 4.9986
448 | 17918.768 | 9.7936 | 4.9942 |  |  | 448 | 17918.768 | 9.7936 | 4.9969 |  |  | 448 | 17918.768 | 9.7936 | 4.9983
449 | 17958.765 | 9.7958 | 4.9939 |  |  | 449 | 17958.765 | 9.7958 | 4.9967 |  |  | 449 | 17958.765 | 9.7958 | 4.9981
450 | 17998.763 | 9.7981 | 4.9937 |  |  | 450 | 17998.763 | 9.7981 | 4.9964 |  |  | 450 | 17998.763 | 9.7981 | 4.9978
451 | 18038.76 | 9.8003 | 4.9935 |  |  | 451 | 18038.76 | 9.8003 | 4.9962 |  |  | 451 | 18038.76 | 9.8003 | 4.9976
452 | 18078.757 | 9.8025 | 4.9933 |  |  | 452 | 18078.757 | 9.8025 | 4.9959 |  |  | 452 | 18078.757 | 9.8025 | 4.9974
453 | 18118.755 | 9.8047 | 4.993 |  |  | 453 | 18118.755 | 9.8047 | 4.9957 |  |  | 453 | 18118.755 | 9.8047 | 4.9971
454 | 18158.752 | 9.8069 | 4.9928 |  |  | 454 | 18158.752 | 9.8069 | 4.9955 |  |  | 454 | 18158.752 | 9.8069 | 4.9969
455 | 18198.749 | 9.8091 | 4.9926 |  |  | 455 | 18198.749 | 9.8091 | 4.9952 |  |  | 455 | 18198.749 | 9.8091 | 4.9967
456 | 18238.746 | 9.8113 | 4.9924 |  |  | 456 | 18238.746 | 9.8113 | 4.995 |  |  | 456 | 18238.746 | 9.8113 | 4.9964
457 | 18278.744 | 9.8135 | 4.9921 |  |  | 457 | 18278.744 | 9.8135 | 4.9948 |  |  | 457 | 18278.744 | 9.8135 | 4.9962
458 | 18318.741 | 9.8157 | 4.9919 |  |  | 458 | 18318.741 | 9.8157 | 4.9945 |  |  | 458 | 18318.741 | 9.8157 | 4.996
459 | 18358.738 | 9.8179 | 4.9917 |  |  | 459 | 18358.738 | 9.8179 | 4.9943 |  |  | 459 | 18358.738 | 9.8179 | 4.9957
460 | 18398.735 | 9.82 | 4.9915 |  |  | 460 | 18398.735 | 9.82 | 4.9941 |  |  | 460 | 18398.735 | 9.82 | 4.9955
461 | 18438.733 | 9.8222 | 4.9913 |  |  | 461 | 18438.733 | 9.8222 | 4.9938 |  |  | 461 | 18438.733 | 9.8222 | 4.9953
462 | 18478.73 | 9.8244 | 4.9911 |  |  | 462 | 18478.73 | 9.8244 | 4.9936 |  |  | 462 | 18478.73 | 9.8244 | 4.9951
463 | 18518.727 | 9.8265 | 4.9908 |  |  | 463 | 18518.727 | 9.8265 | 4.9934 |  |  | 463 | 18518.727 | 9.8265 | 4.9948
464 | 18558.724 | 9.8287 | 4.9906 |  |  | 464 | 18558.724 | 9.8287 | 4.9932 |  |  | 464 | 18558.724 | 9.8287 | 4.9946
465 | 18598.722 | 9.8308 | 4.9904 |  |  | 465 | 18598.722 | 9.8308 | 4.9929 |  |  | 465 | 18598.722 | 9.8308 | 4.9944
466 | 18638.719 | 9.833 | 4.9902 |  |  | 466 | 18638.719 | 9.833 | 4.9927 |  |  | 466 | 18638.719 | 9.833 | 4.9942
467 | 18678.716 | 9.8351 | 4.99 |  |  | 467 | 18678.716 | 9.8351 | 4.9925 |  |  | 467 | 18678.716 | 9.8351 | 4.994
468 | 18718.713 | 9.8373 | 4.9898 |  |  | 468 | 18718.713 | 9.8373 | 4.9923 |  |  | 468 | 18718.713 | 9.8373 | 4.9937
469 | 18758.711 | 9.8394 | 4.9896 |  |  | 469 | 18758.711 | 9.8394 | 4.9921 |  |  | 469 | 18758.711 | 9.8394 | 4.9935
470 | 18798.708 | 9.8415 | 4.9894 |  |  | 470 | 18798.708 | 9.8415 | 4.9919 |  |  | 470 | 18798.708 | 9.8415 | 4.9933
471 | 18838.705 | 9.8437 | 4.9892 |  |  | 471 | 18838.705 | 9.8437 | 4.9916 |  |  | 471 | 18838.705 | 9.8437 | 4.9931
472 | 18878.702 | 9.8458 | 4.989 |  |  | 472 | 18878.702 | 9.8458 | 4.9914 |  |  | 472 | 18878.702 | 9.8458 | 4.9929
473 | 18918.7 | 9.8479 | 4.9888 |  |  | 473 | 18918.7 | 9.8479 | 4.9912 |  |  | 473 | 18918.7 | 9.8479 | 4.9927
474 | 18958.697 | 9.85 | 4.9886 |  |  | 474 | 18958.697 | 9.85 | 4.991 |  |  | 474 | 18958.697 | 9.85 | 4.9925
475 | 18998.694 | 9.8521 | 4.9884 |  |  | 475 | 18998.694 | 9.8521 | 4.9908 |  |  | 475 | 18998.694 | 9.8521 | 4.9923
476 | 19038.691 | 9.8542 | 4.9882 |  |  | 476 | 19038.691 | 9.8542 | 4.9906 |  |  | 476 | 19038.691 | 9.8542 | 4.9921
477 | 19078.689 | 9.8563 | 4.988 |  |  | 477 | 19078.689 | 9.8563 | 4.9904 |  |  | 477 | 19078.689 | 9.8563 | 4.9919
478 | 19118.686 | 9.8584 | 4.9878 |  |  | 478 | 19118.686 | 9.8584 | 4.9902 |  |  | 478 | 19118.686 | 9.8584 | 4.9917
479 | 19158.683 | 9.8605 | 4.9876 |  |  | 479 | 19158.683 | 9.8605 | 4.99 |  |  | 479 | 19158.683 | 9.8605 | 4.9915
480 | 19198.68 | 9.8626 | 4.9874 |  |  | 480 | 19198.68 | 9.8626 | 4.9898 |  |  | 480 | 19198.68 | 9.8626 | 4.9913
481 | 19238.678 | 9.8647 | 4.9872 |  |  | 481 | 19238.678 | 9.8647 | 4.9896 |  |  | 481 | 19238.678 | 9.8647 | 4.9911
482 | 19278.675 | 9.8668 | 4.987 |  |  | 482 | 19278.675 | 9.8668 | 4.9894 |  |  | 482 | 19278.675 | 9.8668 | 4.9909
483 | 19318.672 | 9.8688 | 4.9869 |  |  | 483 | 19318.672 | 9.8688 | 4.9892 |  |  | 483 | 19318.672 | 9.8688 | 4.9907
484 | 19358.669 | 9.8709 | 4.9867 |  |  | 484 | 19358.669 | 9.8709 | 4.989 |  |  | 484 | 19358.669 | 9.8709 | 4.9905
485 | 19398.667 | 9.873 | 4.9865 |  |  | 485 | 19398.667 | 9.873 | 4.9888 |  |  | 485 | 19398.667 | 9.873 | 4.9903
486 | 19438.664 | 9.875 | 4.9863 |  |  | 486 | 19438.664 | 9.875 | 4.9886 |  |  | 486 | 19438.664 | 9.875 | 4.9901
487 | 19478.661 | 9.8771 | 4.9861 |  |  | 487 | 19478.661 | 9.8771 | 4.9884 |  |  | 487 | 19478.661 | 9.8771 | 4.9899
488 | 19518.658 | 9.8791 | 4.9859 |  |  | 488 | 19518.658 | 9.8791 | 4.9882 |  |  | 488 | 19518.658 | 9.8791 | 4.9897
489 | 19558.656 | 9.8812 | 4.9857 |  |  | 489 | 19558.656 | 9.8812 | 4.988 |  |  | 489 | 19558.656 | 9.8812 | 4.9895
490 | 19598.653 | 9.8832 | 4.9856 |  |  | 490 | 19598.653 | 9.8832 | 4.9878 |  |  | 490 | 19598.653 | 9.8832 | 4.9893
491 | 19638.65 | 9.8853 | 4.9854 |  |  | 491 | 19638.65 | 9.8853 | 4.9876 |  |  | 491 | 19638.65 | 9.8853 | 4.9891
492 | 19678.647 | 9.8873 | 4.9852 |  |  | 492 | 19678.647 | 9.8873 | 4.9874 |  |  | 492 | 19678.647 | 9.8873 | 4.9889
493 | 19718.645 | 9.8893 | 4.985 |  |  | 493 | 19718.645 | 9.8893 | 4.9872 |  |  | 493 | 19718.645 | 9.8893 | 4.9887
494 | 19758.642 | 9.8913 | 4.9849 |  |  | 494 | 19758.642 | 9.8913 | 4.987 |  |  | 494 | 19758.642 | 9.8913 | 4.9885
495 | 19798.639 | 9.8934 | 4.9847 |  |  | 495 | 19798.639 | 9.8934 | 4.9868 |  |  | 495 | 19798.639 | 9.8934 | 4.9884
496 | 19838.636 | 9.8954 | 4.9845 |  |  | 496 | 19838.636 | 9.8954 | 4.9866 |  |  | 496 | 19838.636 | 9.8954 | 4.9882
497 | 19878.634 | 9.8974 | 4.9843 |  |  | 497 | 19878.634 | 9.8974 | 4.9865 |  |  | 497 | 19878.634 | 9.8974 | 4.988
498 | 19918.631 | 9.8994 | 4.9842 |  |  | 498 | 19918.631 | 9.8994 | 4.9863 |  |  | 498 | 19918.631 | 9.8994 | 4.9878
499 | 19958.628 | 9.9014 | 4.984 |  |  | 499 | 19958.628 | 9.9014 | 4.9861 |  |  | 499 | 19958.628 | 9.9014 | 4.9876
500 | 19998.625 | 9.9034 | 4.9838 |  |  | 500 | 19998.625 | 9.9034 | 4.9859 |  |  | 500 | 19998.625 | 9.9034 | 4.9875
501 | 20038.623 | 9.9054 | 4.9836 |  |  | 501 | 20038.623 | 9.9054 | 4.9857 |  |  | 501 | 20038.623 | 9.9054 | 4.9873
502 | 20078.62 | 9.9074 | 4.9835 |  |  | 502 | 20078.62 | 9.9074 | 4.9855 |  |  | 502 | 20078.62 | 9.9074 | 4.9871
503 | 20118.617 | 9.9094 | 4.9833 |  |  | 503 | 20118.617 | 9.9094 | 4.9854 |  |  | 503 | 20118.617 | 9.9094 | 4.9869
504 | 20158.614 | 9.9114 | 4.9831 |  |  | 504 | 20158.614 | 9.9114 | 4.9852 |  |  | 504 | 20158.614 | 9.9114 | 4.9867
505 | 20198.612 | 9.9134 | 4.983 |  |  | 505 | 20198.612 | 9.9134 | 4.985 |  |  | 505 | 20198.612 | 9.9134 | 4.9866
506 | 20238.609 | 9.9153 | 4.9828 |  |  | 506 | 20238.609 | 9.9153 | 4.9848 |  |  | 506 | 20238.609 | 9.9153 | 4.9864
507 | 20278.606 | 9.9173 | 4.9826 |  |  | 507 | 20278.606 | 9.9173 | 4.9846 |  |  | 507 | 20278.606 | 9.9173 | 4.9862
508 | 20318.603 | 9.9193 | 4.9825 |  |  | 508 | 20318.603 | 9.9193 | 4.9845 |  |  | 508 | 20318.603 | 9.9193 | 4.986
509 | 20358.601 | 9.9213 | 4.9823 |  |  | 509 | 20358.601 | 9.9213 | 4.9843 |  |  | 509 | 20358.601 | 9.9213 | 4.9859
510 | 20398.598 | 9.9232 | 4.9822 |  |  | 510 | 20398.598 | 9.9232 | 4.9841 |  |  | 510 | 20398.598 | 9.9232 | 4.9857
511 | 20438.595 | 9.9252 | 4.982 |  |  | 511 | 20438.595 | 9.9252 | 4.9839 |  |  | 511 | 20438.595 | 9.9252 | 4.9855
512 | 20478.592 | 9.9271 | 4.9818 |  |  | 512 | 20478.592 | 9.9271 | 4.9838 |  |  | 512 | 20478.592 | 9.9271 | 4.9854
513 | 20518.59 | 9.9291 | 4.9817 |  |  | 513 | 20518.59 | 9.9291 | 4.9836 |  |  | 513 | 20518.59 | 9.9291 | 4.9852
514 | 20558.587 | 9.931 | 4.9815 |  |  | 514 | 20558.587 | 9.931 | 4.9834 |  |  | 514 | 20558.587 | 9.931 | 4.985
515 | 20598.584 | 9.933 | 4.9814 |  |  | 515 | 20598.584 | 9.933 | 4.9833 |  |  | 515 | 20598.584 | 9.933 | 4.9849
516 | 20638.581 | 9.9349 | 4.9812 |  |  | 516 | 20638.581 | 9.9349 | 4.9831 |  |  | 516 | 20638.581 | 9.9349 | 4.9847
517 | 20678.579 | 9.9369 | 4.981 |  |  | 517 | 20678.579 | 9.9369 | 4.9829 |  |  | 517 | 20678.579 | 9.9369 | 4.9845
518 | 20718.576 | 9.9388 | 4.9809 |  |  | 518 | 20718.576 | 9.9388 | 4.9828 |  |  | 518 | 20718.576 | 9.9388 | 4.9844
519 | 20758.573 | 9.9407 | 4.9807 |  |  | 519 | 20758.573 | 9.9407 | 4.9826 |  |  | 519 | 20758.573 | 9.9407 | 4.9842
520 | 20798.57 | 9.9426 | 4.9806 |  |  | 520 | 20798.57 | 9.9426 | 4.9824 |  |  | 520 | 20798.57 | 9.9426 | 4.984
521 | 20838.568 | 9.9446 | 4.9804 |  |  | 521 | 20838.568 | 9.9446 | 4.9823 |  |  | 521 | 20838.568 | 9.9446 | 4.9839
522 | 20878.565 | 9.9465 | 4.9803 |  |  | 522 | 20878.565 | 9.9465 | 4.9821 |  |  | 522 | 20878.565 | 9.9465 | 4.9837
523 | 20918.562 | 9.9484 | 4.9801 |  |  | 523 | 20918.562 | 9.9484 | 4.9819 |  |  | 523 | 20918.562 | 9.9484 | 4.9836
524 | 20958.559 | 9.9503 | 4.98 |  |  | 524 | 20958.559 | 9.9503 | 4.9818 |  |  | 524 | 20958.559 | 9.9503 | 4.9834
525 | 20998.557 | 9.9522 | 4.9798 |  |  | 525 | 20998.557 | 9.9522 | 4.9816 |  |  | 525 | 20998.557 | 9.9522 | 4.9832
526 | 21038.554 | 9.9541 | 4.9797 |  |  | 526 | 21038.554 | 9.9541 | 4.9814 |  |  | 526 | 21038.554 | 9.9541 | 4.9831
527 | 21078.551 | 9.956 | 4.9795 |  |  | 527 | 21078.551 | 9.956 | 4.9813 |  |  | 527 | 21078.551 | 9.956 | 4.9829
528 | 21118.548 | 9.9579 | 4.9794 |  |  | 528 | 21118.548 | 9.9579 | 4.9811 |  |  | 528 | 21118.548 | 9.9579 | 4.9828
529 | 21158.546 | 9.9598 | 4.9792 |  |  | 529 | 21158.546 | 9.9598 | 4.981 |  |  | 529 | 21158.546 | 9.9598 | 4.9826
530 | 21198.543 | 9.9617 | 4.9791 |  |  | 530 | 21198.543 | 9.9617 | 4.9808 |  |  | 530 | 21198.543 | 9.9617 | 4.9825
531 | 21238.54 | 9.9636 | 4.979 |  |  | 531 | 21238.54 | 9.9636 | 4.9806 |  |  | 531 | 21238.54 | 9.9636 | 4.9823
532 | 21278.537 | 9.9655 | 4.9788 |  |  | 532 | 21278.537 | 9.9655 | 4.9805 |  |  | 532 | 21278.537 | 9.9655 | 4.9821
533 | 21318.535 | 9.9673 | 4.9787 |  |  | 533 | 21318.535 | 9.9673 | 4.9803 |  |  | 533 | 21318.535 | 9.9673 | 4.982
534 | 21358.532 | 9.9692 | 4.9785 |  |  | 534 | 21358.532 | 9.9692 | 4.9802 |  |  | 534 | 21358.532 | 9.9692 | 4.9818
535 | 21398.529 | 9.9711 | 4.9784 |  |  | 535 | 21398.529 | 9.9711 | 4.98 |  |  | 535 | 21398.529 | 9.9711 | 4.9817
536 | 21438.526 | 9.9729 | 4.9782 |  |  | 536 | 21438.526 | 9.9729 | 4.9799 |  |  | 536 | 21438.526 | 9.9729 | 4.9815
537 | 21478.524 | 9.9748 | 4.9781 |  |  | 537 | 21478.524 | 9.9748 | 4.9797 |  |  | 537 | 21478.524 | 9.9748 | 4.9814
538 | 21518.521 | 9.9767 | 4.978 |  |  | 538 | 21518.521 | 9.9767 | 4.9795 |  |  | 538 | 21518.521 | 9.9767 | 4.9812
539 | 21558.518 | 9.9785 | 4.9778 |  |  | 539 | 21558.518 | 9.9785 | 4.9794 |  |  | 539 | 21558.518 | 9.9785 | 4.9811
540 | 21598.515 | 9.9804 | 4.9777 |  |  | 540 | 21598.515 | 9.9804 | 4.9792 |  |  | 540 | 21598.515 | 9.9804 | 4.9809
541 | 21638.513 | 9.9822 | 4.9775 |  |  | 541 | 21638.513 | 9.9822 | 4.9791 |  |  | 541 | 21638.513 | 9.9822 | 4.9808
542 | 21678.51 | 9.9841 | 4.9774 |  |  | 542 | 21678.51 | 9.9841 | 4.9789 |  |  | 542 | 21678.51 | 9.9841 | 4.9806
543 | 21718.507 | 9.9859 | 4.9773 |  |  | 543 | 21718.507 | 9.9859 | 4.9788 |  |  | 543 | 21718.507 | 9.9859 | 4.9805
544 | 21758.504 | 9.9878 | 4.9771 |  |  | 544 | 21758.504 | 9.9878 | 4.9786 |  |  | 544 | 21758.504 | 9.9878 | 4.9804
545 | 21798.502 | 9.9896 | 4.977 |  |  | 545 | 21798.502 | 9.9896 | 4.9785 |  |  | 545 | 21798.502 | 9.9896 | 4.9802
546 | 21838.499 | 9.9914 | 4.9769 |  |  | 546 | 21838.499 | 9.9914 | 4.9783 |  |  | 546 | 21838.499 | 9.9914 | 4.9801
547 | 21878.496 | 9.9933 | 4.9767 |  |  | 547 | 21878.496 | 9.9933 | 4.9782 |  |  | 547 | 21878.496 | 9.9933 | 4.9799
548 | 21918.493 | 9.9951 | 4.9766 |  |  | 548 | 21918.493 | 9.9951 | 4.978 |  |  | 548 | 21918.493 | 9.9951 | 4.9798
549 | 21958.491 | 9.9969 | 4.9765 |  |  | 549 | 21958.491 | 9.9969 | 4.9779 |  |  | 549 | 21958.491 | 9.9969 | 4.9796
550 | 21998.488 | 9.9987 | 4.9763 |  |  | 550 | 21998.488 | 9.9987 | 4.9777 |  |  | 550 | 21998.488 | 9.9987 | 4.9795
551 | 22038.485 | 10.0005 | 4.9762 |  |  | 551 | 22038.485 | 10.0005 | 4.9776 |  |  | 551 | 22038.485 | 10.0005 | 4.9794
552 | 22078.482 | 10.0024 | 4.9761 |  |  | 552 | 22078.482 | 10.0024 | 4.9775 |  |  | 552 | 22078.482 | 10.0024 | 4.9792
553 | 22118.48 | 10.0042 | 4.9759 |  |  | 553 | 22118.48 | 10.0042 | 4.9773 |  |  | 553 | 22118.48 | 10.0042 | 4.9791
554 | 22158.477 | 10.006 | 4.9758 |  |  | 554 | 22158.477 | 10.006 | 4.9772 |  |  | 554 | 22158.477 | 10.006 | 4.9789
555 | 22198.474 | 10.0078 | 4.9757 |  |  | 555 | 22198.474 | 10.0078 | 4.977 |  |  | 555 | 22198.474 | 10.0078 | 4.9788
556 | 22238.471 | 10.0096 | 4.9755 |  |  | 556 | 22238.471 | 10.0096 | 4.9769 |  |  | 556 | 22238.471 | 10.0096 | 4.9787
557 | 22278.469 | 10.0114 | 4.9754 |  |  | 557 | 22278.469 | 10.0114 | 4.9767 |  |  | 557 | 22278.469 | 10.0114 | 4.9785
558 | 22318.466 | 10.0132 | 4.9753 |  |  | 558 | 22318.466 | 10.0132 | 4.9766 |  |  | 558 | 22318.466 | 10.0132 | 4.9784
559 | 22358.463 | 10.015 | 4.9752 |  |  | 559 | 22358.463 | 10.015 | 4.9765 |  |  | 559 | 22358.463 | 10.015 | 4.9783
560 | 22398.46 | 10.0167 | 4.975 |  |  | 560 | 22398.46 | 10.0167 | 4.9763 |  |  | 560 | 22398.46 | 10.0167 | 4.9781
561 | 22438.458 | 10.0185 | 4.9749 |  |  | 561 | 22438.458 | 10.0185 | 4.9762 |  |  | 561 | 22438.458 | 10.0185 | 4.978
562 | 22478.455 | 10.0203 | 4.9748 |  |  | 562 | 22478.455 | 10.0203 | 4.976 |  |  | 562 | 22478.455 | 10.0203 | 4.9778
563 | 22518.452 | 10.0221 | 4.9747 |  |  | 563 | 22518.452 | 10.0221 | 4.9759 |  |  | 563 | 22518.452 | 10.0221 | 4.9777
564 | 22558.449 | 10.0239 | 4.9745 |  |  | 564 | 22558.449 | 10.0239 | 4.9758 |  |  | 564 | 22558.449 | 10.0239 | 4.9776
565 | 22598.447 | 10.0256 | 4.9744 |  |  | 565 | 22598.447 | 10.0256 | 4.9756 |  |  | 565 | 22598.447 | 10.0256 | 4.9774
566 | 22638.444 | 10.0274 | 4.9743 |  |  | 566 | 22638.444 | 10.0274 | 4.9755 |  |  | 566 | 22638.444 | 10.0274 | 4.9773
567 | 22678.441 | 10.0292 | 4.9742 |  |  | 567 | 22678.441 | 10.0292 | 4.9753 |  |  | 567 | 22678.441 | 10.0292 | 4.9772
568 | 22718.438 | 10.0309 | 4.974 |  |  | 568 | 22718.438 | 10.0309 | 4.9752 |  |  | 568 | 22718.438 | 10.0309 | 4.9771
569 | 22758.436 | 10.0327 | 4.9739 |  |  | 569 | 22758.436 | 10.0327 | 4.9751 |  |  | 569 | 22758.436 | 10.0327 | 4.9769
570 | 22798.433 | 10.0344 | 4.9738 |  |  | 570 | 22798.433 | 10.0344 | 4.9749 |  |  | 570 | 22798.433 | 10.0344 | 4.9768
571 | 22838.43 | 10.0362 | 4.9737 |  |  | 571 | 22838.43 | 10.0362 | 4.9748 |  |  | 571 | 22838.43 | 10.0362 | 4.9767
572 | 22878.427 | 10.0379 | 4.9736 |  |  | 572 | 22878.427 | 10.0379 | 4.9747 |  |  | 572 | 22878.427 | 10.0379 | 4.9765
573 | 22918.425 | 10.0397 | 4.9734 |  |  | 573 | 22918.425 | 10.0397 | 4.9745 |  |  | 573 | 22918.425 | 10.0397 | 4.9764
574 | 22958.422 | 10.0414 | 4.9733 |  |  | 574 | 22958.422 | 10.0414 | 4.9744 |  |  | 574 | 22958.422 | 10.0414 | 4.9763
575 | 22998.419 | 10.0432 | 4.9732 |  |  | 575 | 22998.419 | 10.0432 | 4.9743 |  |  | 575 | 22998.419 | 10.0432 | 4.9761
576 | 23038.416 | 10.0449 | 4.9731 |  |  | 576 | 23038.416 | 10.0449 | 4.9741 |  |  | 576 | 23038.416 | 10.0449 | 4.976
577 | 23078.414 | 10.0467 | 4.973 |  |  | 577 | 23078.414 | 10.0467 | 4.974 |  |  | 577 | 23078.414 | 10.0467 | 4.9759
578 | 23118.411 | 10.0484 | 4.9729 |  |  | 578 | 23118.411 | 10.0484 | 4.9739 |  |  | 578 | 23118.411 | 10.0484 | 4.9758
579 | 23158.408 | 10.0501 | 4.9727 |  |  | 579 | 23158.408 | 10.0501 | 4.9737 |  |  | 579 | 23158.408 | 10.0501 | 4.9756
580 | 23198.405 | 10.0518 | 4.9726 |  |  | 580 | 23198.405 | 10.0518 | 4.9736 |  |  | 580 | 23198.405 | 10.0518 | 4.9755
581 | 23238.403 | 10.0536 | 4.9725 |  |  | 581 | 23238.403 | 10.0536 | 4.9735 |  |  | 581 | 23238.403 | 10.0536 | 4.9754
582 | 23278.4 | 10.0553 | 4.9724 |  |  | 582 | 23278.4 | 10.0553 | 4.9733 |  |  | 582 | 23278.4 | 10.0553 | 4.9753
583 | 23318.397 | 10.057 | 4.9723 |  |  | 583 | 23318.397 | 10.057 | 4.9732 |  |  | 583 | 23318.397 | 10.057 | 4.9751
584 | 23358.394 | 10.0587 | 4.9722 |  |  | 584 | 23358.394 | 10.0587 | 4.9731 |  |  | 584 | 23358.394 | 10.0587 | 4.975
585 | 23398.392 | 10.0604 | 4.972 |  |  | 585 | 23398.392 | 10.0604 | 4.973 |  |  | 585 | 23398.392 | 10.0604 | 4.9749
586 | 23438.389 | 10.0621 | 4.9719 |  |  | 586 | 23438.389 | 10.0621 | 4.9728 |  |  | 586 | 23438.389 | 10.0621 | 4.9748
587 | 23478.386 | 10.0638 | 4.9718 |  |  | 587 | 23478.386 | 10.0638 | 4.9727 |  |  | 587 | 23478.386 | 10.0638 | 4.9747
588 | 23518.383 | 10.0655 | 4.9717 |  |  | 588 | 23518.383 | 10.0655 | 4.9726 |  |  | 588 | 23518.383 | 10.0655 | 4.9745
589 | 23558.381 | 10.0672 | 4.9716 |  |  | 589 | 23558.381 | 10.0672 | 4.9724 |  |  | 589 | 23558.381 | 10.0672 | 4.9744
590 | 23598.378 | 10.0689 | 4.9715 |  |  | 590 | 23598.378 | 10.0689 | 4.9723 |  |  | 590 | 23598.378 | 10.0689 | 4.9743
591 | 23638.375 | 10.0706 | 4.9714 |  |  | 591 | 23638.375 | 10.0706 | 4.9722 |  |  | 591 | 23638.375 | 10.0706 | 4.9742
592 | 23678.372 | 10.0723 | 4.9713 |  |  | 592 | 23678.372 | 10.0723 | 4.9721 |  |  | 592 | 23678.372 | 10.0723 | 4.974
593 | 23718.37 | 10.074 | 4.9712 |  |  | 593 | 23718.37 | 10.074 | 4.9719 |  |  | 593 | 23718.37 | 10.074 | 4.9739
594 | 23758.367 | 10.0757 | 4.971 |  |  | 594 | 23758.367 | 10.0757 | 4.9718 |  |  | 594 | 23758.367 | 10.0757 | 4.9738
595 | 23798.364 | 10.0774 | 4.9709 |  |  | 595 | 23798.364 | 10.0774 | 4.9717 |  |  | 595 | 23798.364 | 10.0774 | 4.9737
596 | 23838.361 | 10.0791 | 4.9708 |  |  | 596 | 23838.361 | 10.0791 | 4.9716 |  |  | 596 | 23838.361 | 10.0791 | 4.9736
597 | 23878.359 | 10.0807 | 4.9707 |  |  | 597 | 23878.359 | 10.0807 | 4.9714 |  |  | 597 | 23878.359 | 10.0807 | 4.9735
598 | 23918.356 | 10.0824 | 4.9706 |  |  | 598 | 23918.356 | 10.0824 | 4.9713 |  |  | 598 | 23918.356 | 10.0824 | 4.9733
599 | 23958.353 | 10.0841 | 4.9705 |  |  | 599 | 23958.353 | 10.0841 | 4.9712 |  |  | 599 | 23958.353 | 10.0841 | 4.9732
600 | 23998.35 | 10.0857 | 4.9704 |  |  | 600 | 23998.35 | 10.0857 | 4.9711 |  |  | 600 | 23998.35 | 10.0857 | 4.9731
601 | 24038.348 | 10.0874 | 4.9703 |  |  | 601 | 24038.348 | 10.0874 | 4.9709 |  |  | 601 | 24038.348 | 10.0874 | 4.973
602 | 24078.345 | 10.0891 | 4.9702 |  |  | 602 | 24078.345 | 10.0891 | 4.9708 |  |  | 602 | 24078.345 | 10.0891 | 4.9729
603 | 24118.342 | 10.0907 | 4.9701 |  |  | 603 | 24118.342 | 10.0907 | 4.9707 |  |  | 603 | 24118.342 | 10.0907 | 4.9728
604 | 24158.339 | 10.0924 | 4.97 |  |  | 604 | 24158.339 | 10.0924 | 4.9706 |  |  | 604 | 24158.339 | 10.0924 | 4.9726
605 | 24198.337 | 10.094 | 4.9699 |  |  | 605 | 24198.337 | 10.094 | 4.9705 |  |  | 605 | 24198.337 | 10.094 | 4.9725
606 | 24238.334 | 10.0957 | 4.9697 |  |  | 606 | 24238.334 | 10.0957 | 4.9703 |  |  | 606 | 24238.334 | 10.0957 | 4.9724
607 | 24278.331 | 10.0973 | 4.9696 |  |  | 607 | 24278.331 | 10.0973 | 4.9702 |  |  | 607 | 24278.331 | 10.0973 | 4.9723
608 | 24318.328 | 10.099 | 4.9695 |  |  | 608 | 24318.328 | 10.099 | 4.9701 |  |  | 608 | 24318.328 | 10.099 | 4.9722
609 | 24358.326 | 10.1006 | 4.9694 |  |  | 609 | 24358.326 | 10.1006 | 4.97 |  |  | 609 | 24358.326 | 10.1006 | 4.9721
610 | 24398.323 | 10.1023 | 4.9693 |  |  | 610 | 24398.323 | 10.1023 | 4.9699 |  |  | 610 | 24398.323 | 10.1023 | 4.972
611 | 24438.32 | 10.1039 | 4.9692 |  |  | 611 | 24438.32 | 10.1039 | 4.9697 |  |  | 611 | 24438.32 | 10.1039 | 4.9718
612 | 24478.317 | 10.1055 | 4.9691 |  |  | 612 | 24478.317 | 10.1055 | 4.9696 |  |  | 612 | 24478.317 | 10.1055 | 4.9717
613 | 24518.315 | 10.1072 | 4.969 |  |  | 613 | 24518.315 | 10.1072 | 4.9695 |  |  | 613 | 24518.315 | 10.1072 | 4.9716
614 | 24558.312 | 10.1088 | 4.9689 |  |  | 614 | 24558.312 | 10.1088 | 4.9694 |  |  | 614 | 24558.312 | 10.1088 | 4.9715
615 | 24598.309 | 10.1104 | 4.9688 |  |  | 615 | 24598.309 | 10.1104 | 4.9693 |  |  | 615 | 24598.309 | 10.1104 | 4.9714
616 | 24638.306 | 10.1121 | 4.9687 |  |  | 616 | 24638.306 | 10.1121 | 4.9691 |  |  | 616 | 24638.306 | 10.1121 | 4.9713
617 | 24678.304 | 10.1137 | 4.9686 |  |  | 617 | 24678.304 | 10.1137 | 4.969 |  |  | 617 | 24678.304 | 10.1137 | 4.9712
618 | 24718.301 | 10.1153 | 4.9685 |  |  | 618 | 24718.301 | 10.1153 | 4.9689 |  |  | 618 | 24718.301 | 10.1153 | 4.9711
619 | 24758.298 | 10.1169 | 4.9684 |  |  | 619 | 24758.298 | 10.1169 | 4.9688 |  |  | 619 | 24758.298 | 10.1169 | 4.971
620 | 24798.295 | 10.1185 | 4.9683 |  |  | 620 | 24798.295 | 10.1185 | 4.9687 |  |  | 620 | 24798.295 | 10.1185 | 4.9708
621 | 24838.293 | 10.1201 | 4.9682 |  |  | 621 | 24838.293 | 10.1201 | 4.9686 |  |  | 621 | 24838.293 | 10.1201 | 4.9707
622 | 24878.29 | 10.1218 | 4.9681 |  |  | 622 | 24878.29 | 10.1218 | 4.9684 |  |  | 622 | 24878.29 | 10.1218 | 4.9706
623 | 24918.287 | 10.1234 | 4.968 |  |  | 623 | 24918.287 | 10.1234 | 4.9683 |  |  | 623 | 24918.287 | 10.1234 | 4.9705
624 | 24958.284 | 10.125 | 4.9679 |  |  | 624 | 24958.284 | 10.125 | 4.9682 |  |  | 624 | 24958.284 | 10.125 | 4.9704
625 | 24998.282 | 10.1266 | 4.9678 |  |  | 625 | 24998.282 | 10.1266 | 4.9681 |  |  | 625 | 24998.282 | 10.1266 | 4.9703
626 | 25038.279 | 10.1282 | 4.9677 |  |  | 626 | 25038.279 | 10.1282 | 4.968 |  |  | 626 | 25038.279 | 10.1282 | 4.9702
627 | 25078.276 | 10.1298 | 4.9676 |  |  | 627 | 25078.276 | 10.1298 | 4.9679 |  |  | 627 | 25078.276 | 10.1298 | 4.9701
628 | 25118.273 | 10.1314 | 4.9675 |  |  | 628 | 25118.273 | 10.1314 | 4.9677 |  |  | 628 | 25118.273 | 10.1314 | 4.97
629 | 25158.271 | 10.1329 | 4.9674 |  |  | 629 | 25158.271 | 10.1329 | 4.9676 |  |  | 629 | 25158.271 | 10.1329 | 4.9699
630 | 25198.268 | 10.1345 | 4.9673 |  |  | 630 | 25198.268 | 10.1345 | 4.9675 |  |  | 630 | 25198.268 | 10.1345 | 4.9698
631 | 25238.265 | 10.1361 | 4.9672 |  |  | 631 | 25238.265 | 10.1361 | 4.9674 |  |  | 631 | 25238.265 | 10.1361 | 4.9697
632 | 25278.262 | 10.1377 | 4.9671 |  |  | 632 | 25278.262 | 10.1377 | 4.9673 |  |  | 632 | 25278.262 | 10.1377 | 4.9695
633 | 25318.26 | 10.1393 | 4.967 |  |  | 633 | 25318.26 | 10.1393 | 4.9672 |  |  | 633 | 25318.26 | 10.1393 | 4.9694
634 | 25358.257 | 10.1409 | 4.9669 |  |  | 634 | 25358.257 | 10.1409 | 4.9671 |  |  | 634 | 25358.257 | 10.1409 | 4.9693
635 | 25398.254 | 10.1424 | 4.9668 |  |  | 635 | 25398.254 | 10.1424 | 4.9669 |  |  | 635 | 25398.254 | 10.1424 | 4.9692
636 | 25438.251 | 10.144 | 4.9667 |  |  | 636 | 25438.251 | 10.144 | 4.9668 |  |  | 636 | 25438.251 | 10.144 | 4.9691
637 | 25478.249 | 10.1456 | 4.9666 |  |  | 637 | 25478.249 | 10.1456 | 4.9667 |  |  | 637 | 25478.249 | 10.1456 | 4.969
638 | 25518.246 | 10.1471 | 4.9665 |  |  | 638 | 25518.246 | 10.1471 | 4.9666 |  |  | 638 | 25518.246 | 10.1471 | 4.9689
639 | 25558.243 | 10.1487 | 4.9664 |  |  | 639 | 25558.243 | 10.1487 | 4.9665 |  |  | 639 | 25558.243 | 10.1487 | 4.9688
640 | 25598.24 | 10.1503 | 4.9663 |  |  | 640 | 25598.24 | 10.1503 | 4.9664 |  |  | 640 | 25598.24 | 10.1503 | 4.9687
641 | 25638.238 | 10.1518 | 4.9663 |  |  | 641 | 25638.238 | 10.1518 | 4.9663 |  |  | 641 | 25638.238 | 10.1518 | 4.9686
642 | 25678.235 | 10.1534 | 4.9662 |  |  | 642 | 25678.235 | 10.1534 | 4.9662 |  |  | 642 | 25678.235 | 10.1534 | 4.9685
643 | 25718.232 | 10.155 | 4.9661 |  |  | 643 | 25718.232 | 10.155 | 4.9661 |  |  | 643 | 25718.232 | 10.155 | 4.9684
644 | 25758.229 | 10.1565 | 4.966 |  |  | 644 | 25758.229 | 10.1565 | 4.9659 |  |  | 644 | 25758.229 | 10.1565 | 4.9683
645 | 25798.227 | 10.1581 | 4.9659 |  |  | 645 | 25798.227 | 10.1581 | 4.9658 |  |  | 645 | 25798.227 | 10.1581 | 4.9682
646 | 25838.224 | 10.1596 | 4.9658 |  |  | 646 | 25838.224 | 10.1596 | 4.9657 |  |  | 646 | 25838.224 | 10.1596 | 4.9681
647 | 25878.221 | 10.1612 | 4.9657 |  |  | 647 | 25878.221 | 10.1612 | 4.9656 |  |  | 647 | 25878.221 | 10.1612 | 4.968
648 | 25918.218 | 10.1627 | 4.9656 |  |  | 648 | 25918.218 | 10.1627 | 4.9655 |  |  | 648 | 25918.218 | 10.1627 | 4.9679
649 | 25958.216 | 10.1642 | 4.9655 |  |  | 649 | 25958.216 | 10.1642 | 4.9654 |  |  | 649 | 25958.216 | 10.1642 | 4.9678
650 | 25998.213 | 10.1658 | 4.9654 |  |  | 650 | 25998.213 | 10.1658 | 4.9653 |  |  | 650 | 25998.213 | 10.1658 | 4.9677
651 | 26038.21 | 10.1673 | 4.9653 |  |  | 651 | 26038.21 | 10.1673 | 4.9652 |  |  | 651 | 26038.21 | 10.1673 | 4.9676
652 | 26078.207 | 10.1689 | 4.9652 |  |  | 652 | 26078.207 | 10.1689 | 4.9651 |  |  | 652 | 26078.207 | 10.1689 | 4.9675
653 | 26118.205 | 10.1704 | 4.9651 |  |  | 653 | 26118.205 | 10.1704 | 4.9649 |  |  | 653 | 26118.205 | 10.1704 | 4.9674
654 | 26158.202 | 10.1719 | 4.965 |  |  | 654 | 26158.202 | 10.1719 | 4.9648 |  |  | 654 | 26158.202 | 10.1719 | 4.9673
655 | 26198.199 | 10.1734 | 4.9649 |  |  | 655 | 26198.199 | 10.1734 | 4.9647 |  |  | 655 | 26198.199 | 10.1734 | 4.9672
656 | 26238.196 | 10.175 | 4.9648 |  |  | 656 | 26238.196 | 10.175 | 4.9646 |  |  | 656 | 26238.196 | 10.175 | 4.9671
657 | 26278.194 | 10.1765 | 4.9648 |  |  | 657 | 26278.194 | 10.1765 | 4.9645 |  |  | 657 | 26278.194 | 10.1765 | 4.967
658 | 26318.191 | 10.178 | 4.9647 |  |  | 658 | 26318.191 | 10.178 | 4.9644 |  |  | 658 | 26318.191 | 10.178 | 4.9669
659 | 26358.188 | 10.1795 | 4.9646 |  |  | 659 | 26358.188 | 10.1795 | 4.9643 |  |  | 659 | 26358.188 | 10.1795 | 4.9668
660 | 26398.185 | 10.1811 | 4.9645 |  |  | 660 | 26398.185 | 10.1811 | 4.9642 |  |  | 660 | 26398.185 | 10.1811 | 4.9667
661 | 26438.183 | 10.1826 | 4.9644 |  |  | 661 | 26438.183 | 10.1826 | 4.9641 |  |  | 661 | 26438.183 | 10.1826 | 4.9666
662 | 26478.18 | 10.1841 | 4.9643 |  |  | 662 | 26478.18 | 10.1841 | 4.964 |  |  | 662 | 26478.18 | 10.1841 | 4.9665
663 | 26518.177 | 10.1856 | 4.9642 |  |  | 663 | 26518.177 | 10.1856 | 4.9639 |  |  | 663 | 26518.177 | 10.1856 | 4.9664
664 | 26558.174 | 10.1871 | 4.9641 |  |  | 664 | 26558.174 | 10.1871 | 4.9638 |  |  | 664 | 26558.174 | 10.1871 | 4.9663
665 | 26598.172 | 10.1886 | 4.964 |  |  | 665 | 26598.172 | 10.1886 | 4.9637 |  |  | 665 | 26598.172 | 10.1886 | 4.9662
666 | 26638.169 | 10.1901 | 4.9639 |  |  | 666 | 26638.169 | 10.1901 | 4.9635 |  |  | 666 | 26638.169 | 10.1901 | 4.9661
667 | 26678.166 | 10.1916 | 4.9639 |  |  | 667 | 26678.166 | 10.1916 | 4.9634 |  |  | 667 | 26678.166 | 10.1916 | 4.966
668 | 26718.163 | 10.1931 | 4.9638 |  |  | 668 | 26718.163 | 10.1931 | 4.9633 |  |  | 668 | 26718.163 | 10.1931 | 4.9659
669 | 26758.161 | 10.1946 | 4.9637 |  |  | 669 | 26758.161 | 10.1946 | 4.9632 |  |  | 669 | 26758.161 | 10.1946 | 4.9658
670 | 26798.158 | 10.1961 | 4.9636 |  |  | 670 | 26798.158 | 10.1961 | 4.9631 |  |  | 670 | 26798.158 | 10.1961 | 4.9657
671 | 26838.155 | 10.1976 | 4.9635 |  |  | 671 | 26838.155 | 10.1976 | 4.963 |  |  | 671 | 26838.155 | 10.1976 | 4.9656
672 | 26878.152 | 10.1991 | 4.9634 |  |  | 672 | 26878.152 | 10.1991 | 4.9629 |  |  | 672 | 26878.152 | 10.1991 | 4.9655
673 | 26918.15 | 10.2006 | 4.9633 |  |  | 673 | 26918.15 | 10.2006 | 4.9628 |  |  | 673 | 26918.15 | 10.2006 | 4.9654
674 | 26958.147 | 10.202 | 4.9632 |  |  | 674 | 26958.147 | 10.202 | 4.9627 |  |  | 674 | 26958.147 | 10.202 | 4.9653
675 | 26998.144 | 10.2035 | 4.9631 |  |  | 675 | 26998.144 | 10.2035 | 4.9626 |  |  | 675 | 26998.144 | 10.2035 | 4.9652
676 | 27038.141 | 10.205 | 4.9631 |  |  | 676 | 27038.141 | 10.205 | 4.9625 |  |  | 676 | 27038.141 | 10.205 | 4.9651
677 | 27078.139 | 10.2065 | 4.963 |  |  | 677 | 27078.139 | 10.2065 | 4.9624 |  |  | 677 | 27078.139 | 10.2065 | 4.965
678 | 27118.136 | 10.208 | 4.9629 |  |  | 678 | 27118.136 | 10.208 | 4.9623 |  |  | 678 | 27118.136 | 10.208 | 4.9649
679 | 27158.133 | 10.2094 | 4.9628 |  |  | 679 | 27158.133 | 10.2094 | 4.9622 |  |  | 679 | 27158.133 | 10.2094 | 4.9648
680 | 27198.13 | 10.2109 | 4.9627 |  |  | 680 | 27198.13 | 10.2109 | 4.9621 |  |  | 680 | 27198.13 | 10.2109 | 4.9647
681 | 27238.128 | 10.2124 | 4.9626 |  |  | 681 | 27238.128 | 10.2124 | 4.962 |  |  | 681 | 27238.128 | 10.2124 | 4.9646
682 | 27278.125 | 10.2138 | 4.9625 |  |  | 682 | 27278.125 | 10.2138 | 4.9619 |  |  | 682 | 27278.125 | 10.2138 | 4.9645
683 | 27318.122 | 10.2153 | 4.9624 |  |  | 683 | 27318.122 | 10.2153 | 4.9618 |  |  | 683 | 27318.122 | 10.2153 | 4.9644
684 | 27358.119 | 10.2168 | 4.9624 |  |  | 684 | 27358.119 | 10.2168 | 4.9617 |  |  | 684 | 27358.119 | 10.2168 | 4.9643
685 | 27398.117 | 10.2182 | 4.9623 |  |  | 685 | 27398.117 | 10.2182 | 4.9615 |  |  | 685 | 27398.117 | 10.2182 | 4.9642
686 | 27438.114 | 10.2197 | 4.9622 |  |  | 686 | 27438.114 | 10.2197 | 4.9614 |  |  | 686 | 27438.114 | 10.2197 | 4.9642
687 | 27478.111 | 10.2211 | 4.9621 |  |  | 687 | 27478.111 | 10.2211 | 4.9613 |  |  | 687 | 27478.111 | 10.2211 | 4.9641
688 | 27518.108 | 10.2226 | 4.962 |  |  | 688 | 27518.108 | 10.2226 | 4.9612 |  |  | 688 | 27518.108 | 10.2226 | 4.964
689 | 27558.106 | 10.2241 | 4.9619 |  |  | 689 | 27558.106 | 10.2241 | 4.9611 |  |  | 689 | 27558.106 | 10.2241 | 4.9639
690 | 27598.103 | 10.2255 | 4.9618 |  |  | 690 | 27598.103 | 10.2255 | 4.961 |  |  | 690 | 27598.103 | 10.2255 | 4.9638
691 | 27638.1 | 10.227 | 4.9618 |  |  | 691 | 27638.1 | 10.227 | 4.9609 |  |  | 691 | 27638.1 | 10.227 | 4.9637
692 | 27678.097 | 10.2284 | 4.9617 |  |  | 692 | 27678.097 | 10.2284 | 4.9608 |  |  | 692 | 27678.097 | 10.2284 | 4.9636
693 | 27718.095 | 10.2298 | 4.9616 |  |  | 693 | 27718.095 | 10.2298 | 4.9607 |  |  | 693 | 27718.095 | 10.2298 | 4.9635
694 | 27758.092 | 10.2313 | 4.9615 |  |  | 694 | 27758.092 | 10.2313 | 4.9606 |  |  | 694 | 27758.092 | 10.2313 | 4.9634
695 | 27798.089 | 10.2327 | 4.9614 |  |  | 695 | 27798.089 | 10.2327 | 4.9605 |  |  | 695 | 27798.089 | 10.2327 | 4.9633
696 | 27838.086 | 10.2342 | 4.9613 |  |  | 696 | 27838.086 | 10.2342 | 4.9604 |  |  | 696 | 27838.086 | 10.2342 | 4.9632
697 | 27878.084 | 10.2356 | 4.9612 |  |  | 697 | 27878.084 | 10.2356 | 4.9603 |  |  | 697 | 27878.084 | 10.2356 | 4.9631
698 | 27918.081 | 10.237 | 4.9612 |  |  | 698 | 27918.081 | 10.237 | 4.9602 |  |  | 698 | 27918.081 | 10.237 | 4.963
699 | 27958.078 | 10.2385 | 4.9611 |  |  | 699 | 27958.078 | 10.2385 | 4.9601 |  |  | 699 | 27958.078 | 10.2385 | 4.9629
700 | 27998.075 | 10.2399 | 4.961 |  |  | 700 | 27998.075 | 10.2399 | 4.96 |  |  | 700 | 27998.075 | 10.2399 | 4.9628
701 | 28038.073 | 10.2413 | 4.9609 |  |  | 701 | 28038.073 | 10.2413 | 4.9599 |  |  | 701 | 28038.073 | 10.2413 | 4.9627
702 | 28078.07 | 10.2427 | 4.9608 |  |  | 702 | 28078.07 | 10.2427 | 4.9598 |  |  | 702 | 28078.07 | 10.2427 | 4.9627
703 | 28118.067 | 10.2442 | 4.9607 |  |  | 703 | 28118.067 | 10.2442 | 4.9597 |  |  | 703 | 28118.067 | 10.2442 | 4.9626
704 | 28158.064 | 10.2456 | 4.9606 |  |  | 704 | 28158.064 | 10.2456 | 4.9596 |  |  | 704 | 28158.064 | 10.2456 | 4.9625
705 | 28198.062 | 10.247 | 4.9606 |  |  | 705 | 28198.062 | 10.247 | 4.9595 |  |  | 705 | 28198.062 | 10.247 | 4.9624
706 | 28238.059 | 10.2484 | 4.9605 |  |  | 706 | 28238.059 | 10.2484 | 4.9594 |  |  | 706 | 28238.059 | 10.2484 | 4.9623
707 | 28278.056 | 10.2498 | 4.9604 |  |  | 707 | 28278.056 | 10.2498 | 4.9593 |  |  | 707 | 28278.056 | 10.2498 | 4.9622
708 | 28318.053 | 10.2513 | 4.9603 |  |  | 708 | 28318.053 | 10.2513 | 4.9592 |  |  | 708 | 28318.053 | 10.2513 | 4.9621
709 | 28358.051 | 10.2527 | 4.9602 |  |  | 709 | 28358.051 | 10.2527 | 4.9591 |  |  | 709 | 28358.051 | 10.2527 | 4.962
710 | 28398.048 | 10.2541 | 4.9601 |  |  | 710 | 28398.048 | 10.2541 | 4.959 |  |  | 710 | 28398.048 | 10.2541 | 4.9619
711 | 28438.045 | 10.2555 | 4.9601 |  |  | 711 | 28438.045 | 10.2555 | 4.9589 |  |  | 711 | 28438.045 | 10.2555 | 4.9618
712 | 28478.042 | 10.2569 | 4.96 |  |  | 712 | 28478.042 | 10.2569 | 4.9588 |  |  | 712 | 28478.042 | 10.2569 | 4.9617
713 | 28518.04 | 10.2583 | 4.9599 |  |  | 713 | 28518.04 | 10.2583 | 4.9587 |  |  | 713 | 28518.04 | 10.2583 | 4.9616
714 | 28558.037 | 10.2597 | 4.9598 |  |  | 714 | 28558.037 | 10.2597 | 4.9586 |  |  | 714 | 28558.037 | 10.2597 | 4.9615
715 | 28598.034 | 10.2611 | 4.9597 |  |  | 715 | 28598.034 | 10.2611 | 4.9585 |  |  | 715 | 28598.034 | 10.2611 | 4.9615
716 | 28638.031 | 10.2625 | 4.9596 |  |  | 716 | 28638.031 | 10.2625 | 4.9584 |  |  | 716 | 28638.031 | 10.2625 | 4.9614
717 | 28678.029 | 10.2639 | 4.9596 |  |  | 717 | 28678.029 | 10.2639 | 4.9583 |  |  | 717 | 28678.029 | 10.2639 | 4.9613
718 | 28718.026 | 10.2653 | 4.9595 |  |  | 718 | 28718.026 | 10.2653 | 4.9582 |  |  | 718 | 28718.026 | 10.2653 | 4.9612
719 | 28758.023 | 10.2667 | 4.9594 |  |  | 719 | 28758.023 | 10.2667 | 4.9581 |  |  | 719 | 28758.023 | 10.2667 | 4.9611
720 | 28798.02 | 10.2681 | 4.9593 |  |  | 720 | 28798.02 | 10.2681 | 4.958 |  |  | 720 | 28798.02 | 10.2681 | 4.961
721 | 28838.018 | 10.2694 | 4.9592 |  |  | 721 | 28838.018 | 10.2694 | 4.9579 |  |  | 721 | 28838.018 | 10.2694 | 4.9609
722 | 28878.015 | 10.2708 | 4.9591 |  |  | 722 | 28878.015 | 10.2708 | 4.9578 |  |  | 722 | 28878.015 | 10.2708 | 4.9608
723 | 28918.012 | 10.2722 | 4.9591 |  |  | 723 | 28918.012 | 10.2722 | 4.9577 |  |  | 723 | 28918.012 | 10.2722 | 4.9607
724 | 28958.009 | 10.2736 | 4.959 |  |  | 724 | 28958.009 | 10.2736 | 4.9576 |  |  | 724 | 28958.009 | 10.2736 | 4.9606
725 | 28998.007 | 10.275 | 4.9589 |  |  | 725 | 28998.007 | 10.275 | 4.9575 |  |  | 725 | 28998.007 | 10.275 | 4.9605
726 | 29038.004 | 10.2764 | 4.9588 |  |  | 726 | 29038.004 | 10.2764 | 4.9574 |  |  | 726 | 29038.004 | 10.2764 | 4.9605
727 | 29078.001 | 10.2777 | 4.9587 |  |  | 727 | 29078.001 | 10.2777 | 4.9573 |  |  | 727 | 29078.001 | 10.2777 | 4.9604
728 | 29117.998 | 10.2791 | 4.9587 |  |  | 728 | 29117.998 | 10.2791 | 4.9572 |  |  | 728 | 29117.998 | 10.2791 | 4.9603
729 | 29157.996 | 10.2805 | 4.9586 |  |  | 729 | 29157.996 | 10.2805 | 4.9571 |  |  | 729 | 29157.996 | 10.2805 | 4.9602
730 | 29197.993 | 10.2819 | 4.9585 |  |  | 730 | 29197.993 | 10.2819 | 4.957 |  |  | 730 | 29197.993 | 10.2819 | 4.9601
731 | 29237.99 | 10.2832 | 4.9584 |  |  | 731 | 29237.99 | 10.2832 | 4.9569 |  |  | 731 | 29237.99 | 10.2832 | 4.96
732 | 29277.987 | 10.2846 | 4.9583 |  |  | 732 | 29277.987 | 10.2846 | 4.9568 |  |  | 732 | 29277.987 | 10.2846 | 4.9599
733 | 29317.985 | 10.286 | 4.9582 |  |  | 733 | 29317.985 | 10.286 | 4.9567 |  |  | 733 | 29317.985 | 10.286 | 4.9598
734 | 29357.982 | 10.2873 | 4.9582 |  |  | 734 | 29357.982 | 10.2873 | 4.9566 |  |  | 734 | 29357.982 | 10.2873 | 4.9597
735 | 29397.979 | 10.2887 | 4.9581 |  |  | 735 | 29397.979 | 10.2887 | 4.9565 |  |  | 735 | 29397.979 | 10.2887 | 4.9596
736 | 29437.976 | 10.29 | 4.958 |  |  | 736 | 29437.976 | 10.29 | 4.9564 |  |  | 736 | 29437.976 | 10.29 | 4.9596
737 | 29477.974 | 10.2914 | 4.9579 |  |  | 737 | 29477.974 | 10.2914 | 4.9563 |  |  | 737 | 29477.974 | 10.2914 | 4.9595
738 | 29517.971 | 10.2928 | 4.9578 |  |  | 738 | 29517.971 | 10.2928 | 4.9562 |  |  | 738 | 29517.971 | 10.2928 | 4.9594
739 | 29557.968 | 10.2941 | 4.9578 |  |  | 739 | 29557.968 | 10.2941 | 4.9561 |  |  | 739 | 29557.968 | 10.2941 | 4.9593
740 | 29597.965 | 10.2955 | 4.9577 |  |  | 740 | 29597.965 | 10.2955 | 4.956 |  |  | 740 | 29597.965 | 10.2955 | 4.9592
741 | 29637.963 | 10.2968 | 4.9576 |  |  | 741 | 29637.963 | 10.2968 | 4.9559 |  |  | 741 | 29637.963 | 10.2968 | 4.9591
742 | 29677.96 | 10.2982 | 4.9575 |  |  | 742 | 29677.96 | 10.2982 | 4.9558 |  |  | 742 | 29677.96 | 10.2982 | 4.959
743 | 29717.957 | 10.2995 | 4.9574 |  |  | 743 | 29717.957 | 10.2995 | 4.9557 |  |  | 743 | 29717.957 | 10.2995 | 4.9589
744 | 29757.954 | 10.3009 | 4.9573 |  |  | 744 | 29757.954 | 10.3009 | 4.9556 |  |  | 744 | 29757.954 | 10.3009 | 4.9588
745 | 29797.952 | 10.3022 | 4.9573 |  |  | 745 | 29797.952 | 10.3022 | 4.9555 |  |  | 745 | 29797.952 | 10.3022 | 4.9588
746 | 29837.949 | 10.3035 | 4.9572 |  |  | 746 | 29837.949 | 10.3035 | 4.9554 |  |  | 746 | 29837.949 | 10.3035 | 4.9587
747 | 29877.946 | 10.3049 | 4.9571 |  |  | 747 | 29877.946 | 10.3049 | 4.9553 |  |  | 747 | 29877.946 | 10.3049 | 4.9586
748 | 29917.943 | 10.3062 | 4.957 |  |  | 748 | 29917.943 | 10.3062 | 4.9552 |  |  | 748 | 29917.943 | 10.3062 | 4.9585
749 | 29957.941 | 10.3075 | 4.9569 |  |  | 749 | 29957.941 | 10.3075 | 4.9551 |  |  | 749 | 29957.941 | 10.3075 | 4.9584
750 | 29997.938 | 10.3089 | 4.9569 |  |  | 750 | 29997.938 | 10.3089 | 4.955 |  |  | 750 | 29997.938 | 10.3089 | 4.9583
751 | 30037.935 | 10.3102 | 4.9568 |  |  | 751 | 30037.935 | 10.3102 | 4.9549 |  |  | 751 | 30037.935 | 10.3102 | 4.9582
752 | 30077.932 | 10.3115 | 4.9567 |  |  | 752 | 30077.932 | 10.3115 | 4.9548 |  |  | 752 | 30077.932 | 10.3115 | 4.9581
753 | 30117.93 | 10.3129 | 4.9566 |  |  | 753 | 30117.93 | 10.3129 | 4.9547 |  |  | 753 | 30117.93 | 10.3129 | 4.958
754 | 30157.927 | 10.3142 | 4.9565 |  |  | 754 | 30157.927 | 10.3142 | 4.9546 |  |  | 754 | 30157.927 | 10.3142 | 4.958
755 | 30197.924 | 10.3155 | 4.9565 |  |  | 755 | 30197.924 | 10.3155 | 4.9545 |  |  | 755 | 30197.924 | 10.3155 | 4.9579
756 | 30237.921 | 10.3169 | 4.9564 |  |  | 756 | 30237.921 | 10.3169 | 4.9544 |  |  | 756 | 30237.921 | 10.3169 | 4.9578
757 | 30277.919 | 10.3182 | 4.9563 |  |  | 757 | 30277.919 | 10.3182 | 4.9543 |  |  | 757 | 30277.919 | 10.3182 | 4.9577
758 | 30317.916 | 10.3195 | 4.9562 |  |  | 758 | 30317.916 | 10.3195 | 4.9542 |  |  | 758 | 30317.916 | 10.3195 | 4.9576
759 | 30357.913 | 10.3208 | 4.9561 |  |  | 759 | 30357.913 | 10.3208 | 4.9541 |  |  | 759 | 30357.913 | 10.3208 | 4.9575
760 | 30397.91 | 10.3221 | 4.9561 |  |  | 760 | 30397.91 | 10.3221 | 4.954 |  |  | 760 | 30397.91 | 10.3221 | 4.9574
761 | 30437.908 | 10.3234 | 4.956 |  |  | 761 | 30437.908 | 10.3234 | 4.9539 |  |  | 761 | 30437.908 | 10.3234 | 4.9573
762 | 30477.905 | 10.3248 | 4.9559 |  |  | 762 | 30477.905 | 10.3248 | 4.9538 |  |  | 762 | 30477.905 | 10.3248 | 4.9573
763 | 30517.902 | 10.3261 | 4.9558 |  |  | 763 | 30517.902 | 10.3261 | 4.9537 |  |  | 763 | 30517.902 | 10.3261 | 4.9572
764 | 30557.899 | 10.3274 | 4.9557 |  |  | 764 | 30557.899 | 10.3274 | 4.9536 |  |  | 764 | 30557.899 | 10.3274 | 4.9571
765 | 30597.897 | 10.3287 | 4.9557 |  |  | 765 | 30597.897 | 10.3287 | 4.9535 |  |  | 765 | 30597.897 | 10.3287 | 4.957
766 | 30637.894 | 10.33 | 4.9556 |  |  | 766 | 30637.894 | 10.33 | 4.9534 |  |  | 766 | 30637.894 | 10.33 | 4.9569
767 | 30677.891 | 10.3313 | 4.9555 |  |  | 767 | 30677.891 | 10.3313 | 4.9533 |  |  | 767 | 30677.891 | 10.3313 | 4.9568
768 | 30717.888 | 10.3326 | 4.9554 |  |  | 768 | 30717.888 | 10.3326 | 4.9532 |  |  | 768 | 30717.888 | 10.3326 | 4.9567
769 | 30757.886 | 10.3339 | 4.9553 |  |  | 769 | 30757.886 | 10.3339 | 4.9531 |  |  | 769 | 30757.886 | 10.3339 | 4.9566
770 | 30797.883 | 10.3352 | 4.9553 |  |  | 770 | 30797.883 | 10.3352 | 4.953 |  |  | 770 | 30797.883 | 10.3352 | 4.9565
771 | 30837.88 | 10.3365 | 4.9552 |  |  | 771 | 30837.88 | 10.3365 | 4.9529 |  |  | 771 | 30837.88 | 10.3365 | 4.9565
772 | 30877.877 | 10.3378 | 4.9551 |  |  | 772 | 30877.877 | 10.3378 | 4.9528 |  |  | 772 | 30877.877 | 10.3378 | 4.9564
773 | 30917.875 | 10.3391 | 4.955 |  |  | 773 | 30917.875 | 10.3391 | 4.9527 |  |  | 773 | 30917.875 | 10.3391 | 4.9563
774 | 30957.872 | 10.3404 | 4.9549 |  |  | 774 | 30957.872 | 10.3404 | 4.9526 |  |  | 774 | 30957.872 | 10.3404 | 4.9562
775 | 30997.869 | 10.3417 | 4.9549 |  |  | 775 | 30997.869 | 10.3417 | 4.9525 |  |  | 775 | 30997.869 | 10.3417 | 4.9561
776 | 31037.866 | 10.343 | 4.9548 |  |  | 776 | 31037.866 | 10.343 | 4.9524 |  |  | 776 | 31037.866 | 10.343 | 4.956
777 | 31077.864 | 10.3443 | 4.9547 |  |  | 777 | 31077.864 | 10.3443 | 4.9523 |  |  | 777 | 31077.864 | 10.3443 | 4.9559
778 | 31117.861 | 10.3455 | 4.9546 |  |  | 778 | 31117.861 | 10.3455 | 4.9522 |  |  | 778 | 31117.861 | 10.3455 | 4.9558
779 | 31157.858 | 10.3468 | 4.9545 |  |  | 779 | 31157.858 | 10.3468 | 4.9521 |  |  | 779 | 31157.858 | 10.3468 | 4.9558
780 | 31197.855 | 10.3481 | 4.9545 |  |  | 780 | 31197.855 | 10.3481 | 4.952 |  |  | 780 | 31197.855 | 10.3481 | 4.9557
781 | 31237.853 | 10.3494 | 4.9544 |  |  | 781 | 31237.853 | 10.3494 | 4.9519 |  |  | 781 | 31237.853 | 10.3494 | 4.9556
782 | 31277.85 | 10.3507 | 4.9543 |  |  | 782 | 31277.85 | 10.3507 | 4.9518 |  |  | 782 | 31277.85 | 10.3507 | 4.9555
783 | 31317.847 | 10.3519 | 4.9542 |  |  | 783 | 31317.847 | 10.3519 | 4.9517 |  |  | 783 | 31317.847 | 10.3519 | 4.9554
784 | 31357.844 | 10.3532 | 4.9541 |  |  | 784 | 31357.844 | 10.3532 | 4.9516 |  |  | 784 | 31357.844 | 10.3532 | 4.9553
785 | 31397.842 | 10.3545 | 4.9541 |  |  | 785 | 31397.842 | 10.3545 | 4.9515 |  |  | 785 | 31397.842 | 10.3545 | 4.9552
786 | 31437.839 | 10.3558 | 4.954 |  |  | 786 | 31437.839 | 10.3558 | 4.9514 |  |  | 786 | 31437.839 | 10.3558 | 4.9551
787 | 31477.836 | 10.357 | 4.9539 |  |  | 787 | 31477.836 | 10.357 | 4.9513 |  |  | 787 | 31477.836 | 10.357 | 4.9551
788 | 31517.833 | 10.3583 | 4.9538 |  |  | 788 | 31517.833 | 10.3583 | 4.9512 |  |  | 788 | 31517.833 | 10.3583 | 4.955
789 | 31557.831 | 10.3596 | 4.9537 |  |  | 789 | 31557.831 | 10.3596 | 4.9511 |  |  | 789 | 31557.831 | 10.3596 | 4.9549
790 | 31597.828 | 10.3608 | 4.9537 |  |  | 790 | 31597.828 | 10.3608 | 4.951 |  |  | 790 | 31597.828 | 10.3608 | 4.9548
791 | 31637.825 | 10.3621 | 4.9536 |  |  | 791 | 31637.825 | 10.3621 | 4.9509 |  |  | 791 | 31637.825 | 10.3621 | 4.9547
792 | 31677.822 | 10.3634 | 4.9535 |  |  | 792 | 31677.822 | 10.3634 | 4.9508 |  |  | 792 | 31677.822 | 10.3634 | 4.9546
793 | 31717.82 | 10.3646 | 4.9534 |  |  | 793 | 31717.82 | 10.3646 | 4.9507 |  |  | 793 | 31717.82 | 10.3646 | 4.9545
794 | 31757.817 | 10.3659 | 4.9533 |  |  | 794 | 31757.817 | 10.3659 | 4.9506 |  |  | 794 | 31757.817 | 10.3659 | 4.9544
795 | 31797.814 | 10.3672 | 4.9533 |  |  | 795 | 31797.814 | 10.3672 | 4.9505 |  |  | 795 | 31797.814 | 10.3672 | 4.9544
796 | 31837.811 | 10.3684 | 4.9532 |  |  | 796 | 31837.811 | 10.3684 | 4.9504 |  |  | 796 | 31837.811 | 10.3684 | 4.9543
797 | 31877.809 | 10.3697 | 4.9531 |  |  | 797 | 31877.809 | 10.3697 | 4.9503 |  |  | 797 | 31877.809 | 10.3697 | 4.9542
798 | 31917.806 | 10.3709 | 4.953 |  |  | 798 | 31917.806 | 10.3709 | 4.9502 |  |  | 798 | 31917.806 | 10.3709 | 4.9541
799 | 31957.803 | 10.3722 | 4.9529 |  |  | 799 | 31957.803 | 10.3722 | 4.9501 |  |  | 799 | 31957.803 | 10.3722 | 4.954
800 | 31997.8 | 10.3734 | 4.9529 |  |  | 800 | 31997.8 | 10.3734 | 4.95 |  |  | 800 | 31997.8 | 10.3734 | 4.9539
801 | 32037.798 | 10.3747 | 4.9528 |  |  | 801 | 32037.798 | 10.3747 | 4.9499 |  |  | 801 | 32037.798 | 10.3747 | 4.9538
802 | 32077.795 | 10.3759 | 4.9527 |  |  | 802 | 32077.795 | 10.3759 | 4.9498 |  |  | 802 | 32077.795 | 10.3759 | 4.9537
803 | 32117.792 | 10.3772 | 4.9526 |  |  | 803 | 32117.792 | 10.3772 | 4.9497 |  |  | 803 | 32117.792 | 10.3772 | 4.9537
804 | 32157.789 | 10.3784 | 4.9525 |  |  | 804 | 32157.789 | 10.3784 | 4.9496 |  |  | 804 | 32157.789 | 10.3784 | 4.9536
805 | 32197.787 | 10.3797 | 4.9525 |  |  | 805 | 32197.787 | 10.3797 | 4.9495 |  |  | 805 | 32197.787 | 10.3797 | 4.9535
806 | 32237.784 | 10.3809 | 4.9524 |  |  | 806 | 32237.784 | 10.3809 | 4.9494 |  |  | 806 | 32237.784 | 10.3809 | 4.9534
807 | 32277.781 | 10.3821 | 4.9523 |  |  | 807 | 32277.781 | 10.3821 | 4.9493 |  |  | 807 | 32277.781 | 10.3821 | 4.9533
808 | 32317.778 | 10.3834 | 4.9522 |  |  | 808 | 32317.778 | 10.3834 | 4.9492 |  |  | 808 | 32317.778 | 10.3834 | 4.9532
809 | 32357.776 | 10.3846 | 4.9521 |  |  | 809 | 32357.776 | 10.3846 | 4.9491 |  |  | 809 | 32357.776 | 10.3846 | 4.9531
810 | 32397.773 | 10.3858 | 4.9521 |  |  | 810 | 32397.773 | 10.3858 | 4.949 |  |  | 810 | 32397.773 | 10.3858 | 4.953
811 | 32437.77 | 10.3871 | 4.952 |  |  | 811 | 32437.77 | 10.3871 | 4.9489 |  |  | 811 | 32437.77 | 10.3871 | 4.953
812 | 32477.767 | 10.3883 | 4.9519 |  |  | 812 | 32477.767 | 10.3883 | 4.9488 |  |  | 812 | 32477.767 | 10.3883 | 4.9529
813 | 32517.765 | 10.3895 | 4.9518 |  |  | 813 | 32517.765 | 10.3895 | 4.9487 |  |  | 813 | 32517.765 | 10.3895 | 4.9528
814 | 32557.762 | 10.3908 | 4.9517 |  |  | 814 | 32557.762 | 10.3908 | 4.9486 |  |  | 814 | 32557.762 | 10.3908 | 4.9527
815 | 32597.759 | 10.392 | 4.9517 |  |  | 815 | 32597.759 | 10.392 | 4.9485 |  |  | 815 | 32597.759 | 10.392 | 4.9526
816 | 32637.756 | 10.3932 | 4.9516 |  |  | 816 | 32637.756 | 10.3932 | 4.9484 |  |  | 816 | 32637.756 | 10.3932 | 4.9525
817 | 32677.754 | 10.3944 | 4.9515 |  |  | 817 | 32677.754 | 10.3944 | 4.9483 |  |  | 817 | 32677.754 | 10.3944 | 4.9524
818 | 32717.751 | 10.3957 | 4.9514 |  |  | 818 | 32717.751 | 10.3957 | 4.9483 |  |  | 818 | 32717.751 | 10.3957 | 4.9524
819 | 32757.748 | 10.3969 | 4.9514 |  |  | 819 | 32757.748 | 10.3969 | 4.9482 |  |  | 819 | 32757.748 | 10.3969 | 4.9523
820 | 32797.745 | 10.3981 | 4.9513 |  |  | 820 | 32797.745 | 10.3981 | 4.9481 |  |  | 820 | 32797.745 | 10.3981 | 4.9522
821 | 32837.743 | 10.3993 | 4.9512 |  |  | 821 | 32837.743 | 10.3993 | 4.948 |  |  | 821 | 32837.743 | 10.3993 | 4.9521
822 | 32877.74 | 10.4006 | 4.9511 |  |  | 822 | 32877.74 | 10.4006 | 4.9479 |  |  | 822 | 32877.74 | 10.4006 | 4.952
823 | 32917.737 | 10.4018 | 4.951 |  |  | 823 | 32917.737 | 10.4018 | 4.9478 |  |  | 823 | 32917.737 | 10.4018 | 4.9519
824 | 32957.734 | 10.403 | 4.951 |  |  | 824 | 32957.734 | 10.403 | 4.9477 |  |  | 824 | 32957.734 | 10.403 | 4.9518
825 | 32997.732 | 10.4042 | 4.9509 |  |  | 825 | 32997.732 | 10.4042 | 4.9476 |  |  | 825 | 32997.732 | 10.4042 | 4.9517
826 | 33037.729 | 10.4054 | 4.9508 |  |  | 826 | 33037.729 | 10.4054 | 4.9475 |  |  | 826 | 33037.729 | 10.4054 | 4.9517
827 | 33077.726 | 10.4066 | 4.9507 |  |  | 827 | 33077.726 | 10.4066 | 4.9474 |  |  | 827 | 33077.726 | 10.4066 | 4.9516
828 | 33117.723 | 10.4078 | 4.9506 |  |  | 828 | 33117.723 | 10.4078 | 4.9473 |  |  | 828 | 33117.723 | 10.4078 | 4.9515
829 | 33157.721 | 10.409 | 4.9506 |  |  | 829 | 33157.721 | 10.409 | 4.9472 |  |  | 829 | 33157.721 | 10.409 | 4.9514
830 | 33197.718 | 10.4102 | 4.9505 |  |  | 830 | 33197.718 | 10.4102 | 4.9471 |  |  | 830 | 33197.718 | 10.4102 | 4.9513
831 | 33237.715 | 10.4114 | 4.9504 |  |  | 831 | 33237.715 | 10.4114 | 4.947 |  |  | 831 | 33237.715 | 10.4114 | 4.9512
832 | 33277.712 | 10.4126 | 4.9503 |  |  | 832 | 33277.712 | 10.4126 | 4.9469 |  |  | 832 | 33277.712 | 10.4126 | 4.9511
833 | 33317.71 | 10.4138 | 4.9502 |  |  | 833 | 33317.71 | 10.4138 | 4.9468 |  |  | 833 | 33317.71 | 10.4138 | 4.951
834 | 33357.707 | 10.415 | 4.9502 |  |  | 834 | 33357.707 | 10.415 | 4.9467 |  |  | 834 | 33357.707 | 10.415 | 4.951
835 | 33397.704 | 10.4162 | 4.9501 |  |  | 835 | 33397.704 | 10.4162 | 4.9466 |  |  | 835 | 33397.704 | 10.4162 | 4.9509
836 | 33437.701 | 10.4174 | 4.95 |  |  | 836 | 33437.701 | 10.4174 | 4.9465 |  |  | 836 | 33437.701 | 10.4174 | 4.9508
837 | 33477.699 | 10.4186 | 4.9499 |  |  | 837 | 33477.699 | 10.4186 | 4.9464 |  |  | 837 | 33477.699 | 10.4186 | 4.9507
838 | 33517.696 | 10.4198 | 4.9498 |  |  | 838 | 33517.696 | 10.4198 | 4.9463 |  |  | 838 | 33517.696 | 10.4198 | 4.9506
839 | 33557.693 | 10.421 | 4.9498 |  |  | 839 | 33557.693 | 10.421 | 4.9462 |  |  | 839 | 33557.693 | 10.421 | 4.9505
840 | 33597.69 | 10.4222 | 4.9497 |  |  | 840 | 33597.69 | 10.4222 | 4.9461 |  |  | 840 | 33597.69 | 10.4222 | 4.9504
841 | 33637.688 | 10.4234 | 4.9496 |  |  | 841 | 33637.688 | 10.4234 | 4.946 |  |  | 841 | 33637.688 | 10.4234 | 4.9503
842 | 33677.685 | 10.4246 | 4.9495 |  |  | 842 | 33677.685 | 10.4246 | 4.9459 |  |  | 842 | 33677.685 | 10.4246 | 4.9503
843 | 33717.682 | 10.4258 | 4.9494 |  |  | 843 | 33717.682 | 10.4258 | 4.9458 |  |  | 843 | 33717.682 | 10.4258 | 4.9502
844 | 33757.679 | 10.427 | 4.9494 |  |  | 844 | 33757.679 | 10.427 | 4.9457 |  |  | 844 | 33757.679 | 10.427 | 4.9501
845 | 33797.677 | 10.4281 | 4.9493 |  |  | 845 | 33797.677 | 10.4281 | 4.9456 |  |  | 845 | 33797.677 | 10.4281 | 4.95
846 | 33837.674 | 10.4293 | 4.9492 |  |  | 846 | 33837.674 | 10.4293 | 4.9455 |  |  | 846 | 33837.674 | 10.4293 | 4.9499
847 | 33877.671 | 10.4305 | 4.9491 |  |  | 847 | 33877.671 | 10.4305 | 4.9454 |  |  | 847 | 33877.671 | 10.4305 | 4.9498
848 | 33917.668 | 10.4317 | 4.949 |  |  | 848 | 33917.668 | 10.4317 | 4.9453 |  |  | 848 | 33917.668 | 10.4317 | 4.9497
849 | 33957.666 | 10.4329 | 4.949 |  |  | 849 | 33957.666 | 10.4329 | 4.9452 |  |  | 849 | 33957.666 | 10.4329 | 4.9496
850 | 33997.663 | 10.434 | 4.9489 |  |  | 850 | 33997.663 | 10.434 | 4.9451 |  |  | 850 | 33997.663 | 10.434 | 4.9496
851 | 34037.66 | 10.4352 | 4.9488 |  |  | 851 | 34037.66 | 10.4352 | 4.945 |  |  | 851 | 34037.66 | 10.4352 | 4.9495
852 | 34077.657 | 10.4364 | 4.9487 |  |  | 852 | 34077.657 | 10.4364 | 4.9449 |  |  | 852 | 34077.657 | 10.4364 | 4.9494
853 | 34117.655 | 10.4376 | 4.9486 |  |  | 853 | 34117.655 | 10.4376 | 4.9448 |  |  | 853 | 34117.655 | 10.4376 | 4.9493
854 | 34157.652 | 10.4387 | 4.9486 |  |  | 854 | 34157.652 | 10.4387 | 4.9447 |  |  | 854 | 34157.652 | 10.4387 | 4.9492
855 | 34197.649 | 10.4399 | 4.9485 |  |  | 855 | 34197.649 | 10.4399 | 4.9446 |  |  | 855 | 34197.649 | 10.4399 | 4.9491
856 | 34237.646 | 10.4411 | 4.9484 |  |  | 856 | 34237.646 | 10.4411 | 4.9445 |  |  | 856 | 34237.646 | 10.4411 | 4.949
857 | 34277.644 | 10.4422 | 4.9483 |  |  | 857 | 34277.644 | 10.4422 | 4.9444 |  |  | 857 | 34277.644 | 10.4422 | 4.9489
858 | 34317.641 | 10.4434 | 4.9482 |  |  | 858 | 34317.641 | 10.4434 | 4.9443 |  |  | 858 | 34317.641 | 10.4434 | 4.9489
859 | 34357.638 | 10.4446 | 4.9482 |  |  | 859 | 34357.638 | 10.4446 | 4.9442 |  |  | 859 | 34357.638 | 10.4446 | 4.9488
860 | 34397.635 | 10.4457 | 4.9481 |  |  | 860 | 34397.635 | 10.4457 | 4.9441 |  |  | 860 | 34397.635 | 10.4457 | 4.9487
861 | 34437.633 | 10.4469 | 4.948 |  |  | 861 | 34437.633 | 10.4469 | 4.944 |  |  | 861 | 34437.633 | 10.4469 | 4.9486
862 | 34477.63 | 10.4481 | 4.9479 |  |  | 862 | 34477.63 | 10.4481 | 4.9439 |  |  | 862 | 34477.63 | 10.4481 | 4.9485
863 | 34517.627 | 10.4492 | 4.9478 |  |  | 863 | 34517.627 | 10.4492 | 4.9438 |  |  | 863 | 34517.627 | 10.4492 | 4.9484
864 | 34557.624 | 10.4504 | 4.9478 |  |  | 864 | 34557.624 | 10.4504 | 4.9437 |  |  | 864 | 34557.624 | 10.4504 | 4.9483
865 | 34597.622 | 10.4515 | 4.9477 |  |  | 865 | 34597.622 | 10.4515 | 4.9436 |  |  | 865 | 34597.622 | 10.4515 | 4.9482
866 | 34637.619 | 10.4527 | 4.9476 |  |  | 866 | 34637.619 | 10.4527 | 4.9435 |  |  | 866 | 34637.619 | 10.4527 | 4.9482
867 | 34677.616 | 10.4538 | 4.9475 |  |  | 867 | 34677.616 | 10.4538 | 4.9434 |  |  | 867 | 34677.616 | 10.4538 | 4.9481
868 | 34717.613 | 10.455 | 4.9474 |  |  | 868 | 34717.613 | 10.455 | 4.9433 |  |  | 868 | 34717.613 | 10.455 | 4.948
869 | 34757.611 | 10.4562 | 4.9473 |  |  | 869 | 34757.611 | 10.4562 | 4.9432 |  |  | 869 | 34757.611 | 10.4562 | 4.9479
870 | 34797.608 | 10.4573 | 4.9473 |  |  | 870 | 34797.608 | 10.4573 | 4.9431 |  |  | 870 | 34797.608 | 10.4573 | 4.9478
871 | 34837.605 | 10.4585 | 4.9472 |  |  | 871 | 34837.605 | 10.4585 | 4.943 |  |  | 871 | 34837.605 | 10.4585 | 4.9477
872 | 34877.602 | 10.4596 | 4.9471 |  |  | 872 | 34877.602 | 10.4596 | 4.9429 |  |  | 872 | 34877.602 | 10.4596 | 4.9476
873 | 34917.6 | 10.4607 | 4.947 |  |  | 873 | 34917.6 | 10.4607 | 4.9428 |  |  | 873 | 34917.6 | 10.4607 | 4.9475
874 | 34957.597 | 10.4619 | 4.9469 |  |  | 874 | 34957.597 | 10.4619 | 4.9427 |  |  | 874 | 34957.597 | 10.4619 | 4.9474
875 | 34997.594 | 10.463 | 4.9469 |  |  | 875 | 34997.594 | 10.463 | 4.9425 |  |  | 875 | 34997.594 | 10.463 | 4.9474
876 | 35037.591 | 10.4642 | 4.9468 |  |  | 876 | 35037.591 | 10.4642 | 4.9424 |  |  | 876 | 35037.591 | 10.4642 | 4.9473
877 | 35077.589 | 10.4653 | 4.9467 |  |  | 877 | 35077.589 | 10.4653 | 4.9423 |  |  | 877 | 35077.589 | 10.4653 | 4.9472
878 | 35117.586 | 10.4665 | 4.9466 |  |  | 878 | 35117.586 | 10.4665 | 4.9422 |  |  | 878 | 35117.586 | 10.4665 | 4.9471
879 | 35157.583 | 10.4676 | 4.9465 |  |  | 879 | 35157.583 | 10.4676 | 4.9421 |  |  | 879 | 35157.583 | 10.4676 | 4.947
880 | 35197.58 | 10.4687 | 4.9465 |  |  | 880 | 35197.58 | 10.4687 | 4.942 |  |  | 880 | 35197.58 | 10.4687 | 4.9469
881 | 35237.578 | 10.4699 | 4.9464 |  |  | 881 | 35237.578 | 10.4699 | 4.9419 |  |  | 881 | 35237.578 | 10.4699 | 4.9468
882 | 35277.575 | 10.471 | 4.9463 |  |  | 882 | 35277.575 | 10.471 | 4.9418 |  |  | 882 | 35277.575 | 10.471 | 4.9467
883 | 35317.572 | 10.4721 | 4.9462 |  |  | 883 | 35317.572 | 10.4721 | 4.9417 |  |  | 883 | 35317.572 | 10.4721 | 4.9467
884 | 35357.569 | 10.4733 | 4.9461 |  |  | 884 | 35357.569 | 10.4733 | 4.9416 |  |  | 884 | 35357.569 | 10.4733 | 4.9466
885 | 35397.567 | 10.4744 | 4.9461 |  |  | 885 | 35397.567 | 10.4744 | 4.9415 |  |  | 885 | 35397.567 | 10.4744 | 4.9465
886 | 35437.564 | 10.4755 | 4.946 |  |  | 886 | 35437.564 | 10.4755 | 4.9414 |  |  | 886 | 35437.564 | 10.4755 | 4.9464
887 | 35477.561 | 10.4767 | 4.9459 |  |  | 887 | 35477.561 | 10.4767 | 4.9413 |  |  | 887 | 35477.561 | 10.4767 | 4.9463
888 | 35517.558 | 10.4778 | 4.9458 |  |  | 888 | 35517.558 | 10.4778 | 4.9412 |  |  | 888 | 35517.558 | 10.4778 | 4.9462
889 | 35557.556 | 10.4789 | 4.9457 |  |  | 889 | 35557.556 | 10.4789 | 4.9411 |  |  | 889 | 35557.556 | 10.4789 | 4.9461
890 | 35597.553 | 10.48 | 4.9456 |  |  | 890 | 35597.553 | 10.48 | 4.941 |  |  | 890 | 35597.553 | 10.48 | 4.946
891 | 35637.55 | 10.4812 | 4.9456 |  |  | 891 | 35637.55 | 10.4812 | 4.9409 |  |  | 891 | 35637.55 | 10.4812 | 4.9459
892 | 35677.547 | 10.4823 | 4.9455 |  |  | 892 | 35677.547 | 10.4823 | 4.9408 |  |  | 892 | 35677.547 | 10.4823 | 4.9459
893 | 35717.545 | 10.4834 | 4.9454 |  |  | 893 | 35717.545 | 10.4834 | 4.9407 |  |  | 893 | 35717.545 | 10.4834 | 4.9458
894 | 35757.542 | 10.4845 | 4.9453 |  |  | 894 | 35757.542 | 10.4845 | 4.9406 |  |  | 894 | 35757.542 | 10.4845 | 4.9457
895 | 35797.539 | 10.4856 | 4.9452 |  |  | 895 | 35797.539 | 10.4856 | 4.9405 |  |  | 895 | 35797.539 | 10.4856 | 4.9456
896 | 35837.536 | 10.4868 | 4.9452 |  |  | 896 | 35837.536 | 10.4868 | 4.9404 |  |  | 896 | 35837.536 | 10.4868 | 4.9455
897 | 35877.534 | 10.4879 | 4.9451 |  |  | 897 | 35877.534 | 10.4879 | 4.9403 |  |  | 897 | 35877.534 | 10.4879 | 4.9454
898 | 35917.531 | 10.489 | 4.945 |  |  | 898 | 35917.531 | 10.489 | 4.9402 |  |  | 898 | 35917.531 | 10.489 | 4.9453
899 | 35957.528 | 10.4901 | 4.9449 |  |  | 899 | 35957.528 | 10.4901 | 4.9401 |  |  | 899 | 35957.528 | 10.4901 | 4.9452
900 | 35997.525 | 10.4912 | 4.9448 |  |  | 900 | 35997.525 | 10.4912 | 4.94 |  |  | 900 | 35997.525 | 10.4912 | 4.9451
901 | 36037.523 | 10.4923 | 4.9447 |  |  | 901 | 36037.523 | 10.4923 | 4.9399 |  |  | 901 | 36037.523 | 10.4923 | 4.9451
902 | 36077.52 | 10.4934 | 4.9447 |  |  | 902 | 36077.52 | 10.4934 | 4.9398 |  |  | 902 | 36077.52 | 10.4934 | 4.945
903 | 36117.517 | 10.4945 | 4.9446 |  |  | 903 | 36117.517 | 10.4945 | 4.9397 |  |  | 903 | 36117.517 | 10.4945 | 4.9449
904 | 36157.515 | 10.4956 | 4.9445 |  |  | 904 | 36157.515 | 10.4956 | 4.9396 |  |  | 904 | 36157.515 | 10.4956 | 4.9448
905 | 36197.512 | 10.4967 | 4.9444 |  |  | 905 | 36197.512 | 10.4967 | 4.9395 |  |  | 905 | 36197.512 | 10.4967 | 4.9447
906 | 36237.509 | 10.4978 | 4.9443 |  |  | 906 | 36237.509 | 10.4978 | 4.9394 |  |  | 906 | 36237.509 | 10.4978 | 4.9446
907 | 36277.506 | 10.499 | 4.9443 |  |  | 907 | 36277.506 | 10.499 | 4.9393 |  |  | 907 | 36277.506 | 10.499 | 4.9445
908 | 36317.504 | 10.5001 | 4.9442 |  |  | 908 | 36317.504 | 10.5001 | 4.9392 |  |  | 908 | 36317.504 | 10.5001 | 4.9444
909 | 36357.501 | 10.5012 | 4.9441 |  |  | 909 | 36357.501 | 10.5012 | 4.9391 |  |  | 909 | 36357.501 | 10.5012 | 4.9443
910 | 36397.498 | 10.5023 | 4.944 |  |  | 910 | 36397.498 | 10.5023 | 4.939 |  |  | 910 | 36397.498 | 10.5023 | 4.9443
911 | 36437.495 | 10.5034 | 4.9439 |  |  | 911 | 36437.495 | 10.5034 | 4.9389 |  |  | 911 | 36437.495 | 10.5034 | 4.9442
912 | 36477.493 | 10.5045 | 4.9438 |  |  | 912 | 36477.493 | 10.5045 | 4.9388 |  |  | 912 | 36477.493 | 10.5045 | 4.9441
913 | 36517.49 | 10.5055 | 4.9438 |  |  | 913 | 36517.49 | 10.5055 | 4.9387 |  |  | 913 | 36517.49 | 10.5055 | 4.944
914 | 36557.487 | 10.5066 | 4.9437 |  |  | 914 | 36557.487 | 10.5066 | 4.9386 |  |  | 914 | 36557.487 | 10.5066 | 4.9439
915 | 36597.484 | 10.5077 | 4.9436 |  |  | 915 | 36597.484 | 10.5077 | 4.9385 |  |  | 915 | 36597.484 | 10.5077 | 4.9438
916 | 36637.482 | 10.5088 | 4.9435 |  |  | 916 | 36637.482 | 10.5088 | 4.9384 |  |  | 916 | 36637.482 | 10.5088 | 4.9437
917 | 36677.479 | 10.5099 | 4.9434 |  |  | 917 | 36677.479 | 10.5099 | 4.9383 |  |  | 917 | 36677.479 | 10.5099 | 4.9436
918 | 36717.476 | 10.511 | 4.9433 |  |  | 918 | 36717.476 | 10.511 | 4.9382 |  |  | 918 | 36717.476 | 10.511 | 4.9435
919 | 36757.473 | 10.5121 | 4.9433 |  |  | 919 | 36757.473 | 10.5121 | 4.9381 |  |  | 919 | 36757.473 | 10.5121 | 4.9434
920 | 36797.471 | 10.5132 | 4.9432 |  |  | 920 | 36797.471 | 10.5132 | 4.9379 |  |  | 920 | 36797.471 | 10.5132 | 4.9434
921 | 36837.468 | 10.5143 | 4.9431 |  |  | 921 | 36837.468 | 10.5143 | 4.9378 |  |  | 921 | 36837.468 | 10.5143 | 4.9433
922 | 36877.465 | 10.5154 | 4.943 |  |  | 922 | 36877.465 | 10.5154 | 4.9377 |  |  | 922 | 36877.465 | 10.5154 | 4.9432
923 | 36917.462 | 10.5164 | 4.9429 |  |  | 923 | 36917.462 | 10.5164 | 4.9376 |  |  | 923 | 36917.462 | 10.5164 | 4.9431
924 | 36957.46 | 10.5175 | 4.9429 |  |  | 924 | 36957.46 | 10.5175 | 4.9375 |  |  | 924 | 36957.46 | 10.5175 | 4.943
925 | 36997.457 | 10.5186 | 4.9428 |  |  | 925 | 36997.457 | 10.5186 | 4.9374 |  |  | 925 | 36997.457 | 10.5186 | 4.9429
926 | 37037.454 | 10.5197 | 4.9427 |  |  | 926 | 37037.454 | 10.5197 | 4.9373 |  |  | 926 | 37037.454 | 10.5197 | 4.9428
927 | 37077.451 | 10.5208 | 4.9426 |  |  | 927 | 37077.451 | 10.5208 | 4.9372 |  |  | 927 | 37077.451 | 10.5208 | 4.9427
928 | 37117.449 | 10.5218 | 4.9425 |  |  | 928 | 37117.449 | 10.5218 | 4.9371 |  |  | 928 | 37117.449 | 10.5218 | 4.9426
929 | 37157.446 | 10.5229 | 4.9424 |  |  | 929 | 37157.446 | 10.5229 | 4.937 |  |  | 929 | 37157.446 | 10.5229 | 4.9425
930 | 37197.443 | 10.524 | 4.9424 |  |  | 930 | 37197.443 | 10.524 | 4.9369 |  |  | 930 | 37197.443 | 10.524 | 4.9425
931 | 37237.44 | 10.5251 | 4.9423 |  |  | 931 | 37237.44 | 10.5251 | 4.9368 |  |  | 931 | 37237.44 | 10.5251 | 4.9424
932 | 37277.438 | 10.5261 | 4.9422 |  |  | 932 | 37277.438 | 10.5261 | 4.9367 |  |  | 932 | 37277.438 | 10.5261 | 4.9423
933 | 37317.435 | 10.5272 | 4.9421 |  |  | 933 | 37317.435 | 10.5272 | 4.9366 |  |  | 933 | 37317.435 | 10.5272 | 4.9422
934 | 37357.432 | 10.5283 | 4.942 |  |  | 934 | 37357.432 | 10.5283 | 4.9365 |  |  | 934 | 37357.432 | 10.5283 | 4.9421
935 | 37397.429 | 10.5294 | 4.9419 |  |  | 935 | 37397.429 | 10.5294 | 4.9364 |  |  | 935 | 37397.429 | 10.5294 | 4.942
936 | 37437.427 | 10.5304 | 4.9419 |  |  | 936 | 37437.427 | 10.5304 | 4.9363 |  |  | 936 | 37437.427 | 10.5304 | 4.9419
937 | 37477.424 | 10.5315 | 4.9418 |  |  | 937 | 37477.424 | 10.5315 | 4.9362 |  |  | 937 | 37477.424 | 10.5315 | 4.9418
938 | 37517.421 | 10.5326 | 4.9417 |  |  | 938 | 37517.421 | 10.5326 | 4.9361 |  |  | 938 | 37517.421 | 10.5326 | 4.9417
939 | 37557.418 | 10.5336 | 4.9416 |  |  | 939 | 37557.418 | 10.5336 | 4.936 |  |  | 939 | 37557.418 | 10.5336 | 4.9416
940 | 37597.416 | 10.5347 | 4.9415 |  |  | 940 | 37597.416 | 10.5347 | 4.9359 |  |  | 940 | 37597.416 | 10.5347 | 4.9416
941 | 37637.413 | 10.5358 | 4.9414 |  |  | 941 | 37637.413 | 10.5358 | 4.9358 |  |  | 941 | 37637.413 | 10.5358 | 4.9415
942 | 37677.41 | 10.5368 | 4.9414 |  |  | 942 | 37677.41 | 10.5368 | 4.9357 |  |  | 942 | 37677.41 | 10.5368 | 4.9414
943 | 37717.407 | 10.5379 | 4.9413 |  |  | 943 | 37717.407 | 10.5379 | 4.9356 |  |  | 943 | 37717.407 | 10.5379 | 4.9413
944 | 37757.405 | 10.5389 | 4.9412 |  |  | 944 | 37757.405 | 10.5389 | 4.9354 |  |  | 944 | 37757.405 | 10.5389 | 4.9412
945 | 37797.402 | 10.54 | 4.9411 |  |  | 945 | 37797.402 | 10.54 | 4.9353 |  |  | 945 | 37797.402 | 10.54 | 4.9411
946 | 37837.399 | 10.5411 | 4.941 |  |  | 946 | 37837.399 | 10.5411 | 4.9352 |  |  | 946 | 37837.399 | 10.5411 | 4.941
947 | 37877.396 | 10.5421 | 4.9409 |  |  | 947 | 37877.396 | 10.5421 | 4.9351 |  |  | 947 | 37877.396 | 10.5421 | 4.9409
948 | 37917.394 | 10.5432 | 4.9408 |  |  | 948 | 37917.394 | 10.5432 | 4.935 |  |  | 948 | 37917.394 | 10.5432 | 4.9408
949 | 37957.391 | 10.5442 | 4.9408 |  |  | 949 | 37957.391 | 10.5442 | 4.9349 |  |  | 949 | 37957.391 | 10.5442 | 4.9407
950 | 37997.388 | 10.5453 | 4.9407 |  |  | 950 | 37997.388 | 10.5453 | 4.9348 |  |  | 950 | 37997.388 | 10.5453 | 4.9406
951 | 38037.385 | 10.5463 | 4.9406 |  |  | 951 | 38037.385 | 10.5463 | 4.9347 |  |  | 951 | 38037.385 | 10.5463 | 4.9406
952 | 38077.383 | 10.5474 | 4.9405 |  |  | 952 | 38077.383 | 10.5474 | 4.9346 |  |  | 952 | 38077.383 | 10.5474 | 4.9405
953 | 38117.38 | 10.5484 | 4.9404 |  |  | 953 | 38117.38 | 10.5484 | 4.9345 |  |  | 953 | 38117.38 | 10.5484 | 4.9404
954 | 38157.377 | 10.5495 | 4.9403 |  |  | 954 | 38157.377 | 10.5495 | 4.9344 |  |  | 954 | 38157.377 | 10.5495 | 4.9403
955 | 38197.374 | 10.5505 | 4.9403 |  |  | 955 | 38197.374 | 10.5505 | 4.9343 |  |  | 955 | 38197.374 | 10.5505 | 4.9402
956 | 38237.372 | 10.5516 | 4.9402 |  |  | 956 | 38237.372 | 10.5516 | 4.9342 |  |  | 956 | 38237.372 | 10.5516 | 4.9401
957 | 38277.369 | 10.5526 | 4.9401 |  |  | 957 | 38277.369 | 10.5526 | 4.9341 |  |  | 957 | 38277.369 | 10.5526 | 4.94
958 | 38317.366 | 10.5537 | 4.94 |  |  | 958 | 38317.366 | 10.5537 | 4.934 |  |  | 958 | 38317.366 | 10.5537 | 4.9399
959 | 38357.363 | 10.5547 | 4.9399 |  |  | 959 | 38357.363 | 10.5547 | 4.9339 |  |  | 959 | 38357.363 | 10.5547 | 4.9398
960 | 38397.361 | 10.5557 | 4.9398 |  |  | 960 | 38397.361 | 10.5557 | 4.9338 |  |  | 960 | 38397.361 | 10.5557 | 4.9397
961 | 38437.358 | 10.5568 | 4.9398 |  |  | 961 | 38437.358 | 10.5568 | 4.9337 |  |  | 961 | 38437.358 | 10.5568 | 4.9396
962 | 38477.355 | 10.5578 | 4.9397 |  |  | 962 | 38477.355 | 10.5578 | 4.9336 |  |  | 962 | 38477.355 | 10.5578 | 4.9395
963 | 38517.352 | 10.5589 | 4.9396 |  |  | 963 | 38517.352 | 10.5589 | 4.9334 |  |  | 963 | 38517.352 | 10.5589 | 4.9395
964 | 38557.35 | 10.5599 | 4.9395 |  |  | 964 | 38557.35 | 10.5599 | 4.9333 |  |  | 964 | 38557.35 | 10.5599 | 4.9394
965 | 38597.347 | 10.5609 | 4.9394 |  |  | 965 | 38597.347 | 10.5609 | 4.9332 |  |  | 965 | 38597.347 | 10.5609 | 4.9393
966 | 38637.344 | 10.562 | 4.9393 |  |  | 966 | 38637.344 | 10.562 | 4.9331 |  |  | 966 | 38637.344 | 10.562 | 4.9392
967 | 38677.341 | 10.563 | 4.9392 |  |  | 967 | 38677.341 | 10.563 | 4.933 |  |  | 967 | 38677.341 | 10.563 | 4.9391
968 | 38717.339 | 10.564 | 4.9392 |  |  | 968 | 38717.339 | 10.564 | 4.9329 |  |  | 968 | 38717.339 | 10.564 | 4.939
969 | 38757.336 | 10.5651 | 4.9391 |  |  | 969 | 38757.336 | 10.5651 | 4.9328 |  |  | 969 | 38757.336 | 10.5651 | 4.9389
970 | 38797.333 | 10.5661 | 4.939 |  |  | 970 | 38797.333 | 10.5661 | 4.9327 |  |  | 970 | 38797.333 | 10.5661 | 4.9388
971 | 38837.33 | 10.5671 | 4.9389 |  |  | 971 | 38837.33 | 10.5671 | 4.9326 |  |  | 971 | 38837.33 | 10.5671 | 4.9387
972 | 38877.328 | 10.5682 | 4.9388 |  |  | 972 | 38877.328 | 10.5682 | 4.9325 |  |  | 972 | 38877.328 | 10.5682 | 4.9386
973 | 38917.325 | 10.5692 | 4.9387 |  |  | 973 | 38917.325 | 10.5692 | 4.9324 |  |  | 973 | 38917.325 | 10.5692 | 4.9385
974 | 38957.322 | 10.5702 | 4.9386 |  |  | 974 | 38957.322 | 10.5702 | 4.9323 |  |  | 974 | 38957.322 | 10.5702 | 4.9384
975 | 38997.319 | 10.5712 | 4.9386 |  |  | 975 | 38997.319 | 10.5712 | 4.9322 |  |  | 975 | 38997.319 | 10.5712 | 4.9384
976 | 39037.317 | 10.5723 | 4.9385 |  |  | 976 | 39037.317 | 10.5723 | 4.9321 |  |  | 976 | 39037.317 | 10.5723 | 4.9383
977 | 39077.314 | 10.5733 | 4.9384 |  |  | 977 | 39077.314 | 10.5733 | 4.932 |  |  | 977 | 39077.314 | 10.5733 | 4.9382
978 | 39117.311 | 10.5743 | 4.9383 |  |  | 978 | 39117.311 | 10.5743 | 4.9319 |  |  | 978 | 39117.311 | 10.5743 | 4.9381
979 | 39157.308 | 10.5753 | 4.9382 |  |  | 979 | 39157.308 | 10.5753 | 4.9317 |  |  | 979 | 39157.308 | 10.5753 | 4.938
980 | 39197.306 | 10.5764 | 4.9381 |  |  | 980 | 39197.306 | 10.5764 | 4.9316 |  |  | 980 | 39197.306 | 10.5764 | 4.9379
981 | 39237.303 | 10.5774 | 4.938 |  |  | 981 | 39237.303 | 10.5774 | 4.9315 |  |  | 981 | 39237.303 | 10.5774 | 4.9378
982 | 39277.3 | 10.5784 | 4.938 |  |  | 982 | 39277.3 | 10.5784 | 4.9314 |  |  | 982 | 39277.3 | 10.5784 | 4.9377
983 | 39317.297 | 10.5794 | 4.9379 |  |  | 983 | 39317.297 | 10.5794 | 4.9313 |  |  | 983 | 39317.297 | 10.5794 | 4.9376
984 | 39357.295 | 10.5804 | 4.9378 |  |  | 984 | 39357.295 | 10.5804 | 4.9312 |  |  | 984 | 39357.295 | 10.5804 | 4.9375
985 | 39397.292 | 10.5815 | 4.9377 |  |  | 985 | 39397.292 | 10.5815 | 4.9311 |  |  | 985 | 39397.292 | 10.5815 | 4.9374
986 | 39437.289 | 10.5825 | 4.9376 |  |  | 986 | 39437.289 | 10.5825 | 4.931 |  |  | 986 | 39437.289 | 10.5825 | 4.9373
987 | 39477.286 | 10.5835 | 4.9375 |  |  | 987 | 39477.286 | 10.5835 | 4.9309 |  |  | 987 | 39477.286 | 10.5835 | 4.9372
988 | 39517.284 | 10.5845 | 4.9374 |  |  | 988 | 39517.284 | 10.5845 | 4.9308 |  |  | 988 | 39517.284 | 10.5845 | 4.9372
989 | 39557.281 | 10.5855 | 4.9374 |  |  | 989 | 39557.281 | 10.5855 | 4.9307 |  |  | 989 | 39557.281 | 10.5855 | 4.9371
990 | 39597.278 | 10.5865 | 4.9373 |  |  | 990 | 39597.278 | 10.5865 | 4.9306 |  |  | 990 | 39597.278 | 10.5865 | 4.937

## Run6
MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR
File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run6\1 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run6\2 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run6\3 01 001.RAW
Cursor Location: |  |  | (178, 193) |  |  | Cursor Location: |  |  | (179, 193) |  |  | Cursor Location: |  |  | (177, 193)
Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15)
Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev
1 | 39.997 | 3.6888 | 7.7384 | 5.1237 | 0.32501 | 1 | 39.997 | 3.6888 | 7.7445 | 5.1214 | 0.32719 | 1 | 39.997 | 3.6888 | 7.7354 | 5.1257 | 0.32531
2 | 79.995 | 4.382 | 7.403 |  |  | 2 | 79.995 | 4.382 | 7.4107 |  |  | 2 | 79.995 | 4.382 | 7.4043
3 | 119.992 | 4.7874 | 7.208 |  |  | 3 | 119.992 | 4.7874 | 7.2145 |  |  | 3 | 119.992 | 4.7874 | 7.211
4 | 159.989 | 5.0751 | 7.0679 |  |  | 4 | 159.989 | 5.0751 | 7.0731 |  |  | 4 | 159.989 | 5.0751 | 7.0714
5 | 199.986 | 5.2982 | 6.9579 |  |  | 5 | 199.986 | 5.2982 | 6.9622 |  |  | 5 | 199.986 | 5.2982 | 6.9615
6 | 239.984 | 5.4806 | 6.8672 |  |  | 6 | 239.984 | 5.4806 | 6.8707 |  |  | 6 | 239.984 | 5.4806 | 6.8707
7 | 279.981 | 5.6347 | 6.79 |  |  | 7 | 279.981 | 5.6347 | 6.793 |  |  | 7 | 279.981 | 5.6347 | 6.7933
8 | 319.978 | 5.7683 | 6.7228 |  |  | 8 | 319.978 | 5.7683 | 6.7253 |  |  | 8 | 319.978 | 5.7683 | 6.7259
9 | 359.975 | 5.886 | 6.6633 |  |  | 9 | 359.975 | 5.886 | 6.6655 |  |  | 9 | 359.975 | 5.886 | 6.6661
10 | 399.973 | 5.9914 | 6.6099 |  |  | 10 | 399.973 | 5.9914 | 6.6118 |  |  | 10 | 399.973 | 5.9914 | 6.6125
11 | 439.97 | 6.0867 | 6.5615 |  |  | 11 | 439.97 | 6.0867 | 6.5632 |  |  | 11 | 439.97 | 6.0867 | 6.5639
12 | 479.967 | 6.1737 | 6.5172 |  |  | 12 | 479.967 | 6.1737 | 6.5188 |  |  | 12 | 479.967 | 6.1737 | 6.5194
13 | 519.964 | 6.2538 | 6.4764 |  |  | 13 | 519.964 | 6.2538 | 6.4779 |  |  | 13 | 519.964 | 6.2538 | 6.4785
14 | 559.962 | 6.3279 | 6.4386 |  |  | 14 | 559.962 | 6.3279 | 6.44 |  |  | 14 | 559.962 | 6.3279 | 6.4405
15 | 599.959 | 6.3969 | 6.4034 |  |  | 15 | 599.959 | 6.3969 | 6.4048 |  |  | 15 | 599.959 | 6.3969 | 6.4052
16 | 639.956 | 6.4614 | 6.3704 |  |  | 16 | 639.956 | 6.4614 | 6.3718 |  |  | 16 | 639.956 | 6.4614 | 6.3721
17 | 679.953 | 6.522 | 6.3394 |  |  | 17 | 679.953 | 6.522 | 6.3408 |  |  | 17 | 679.953 | 6.522 | 6.341
18 | 719.951 | 6.5792 | 6.3102 |  |  | 18 | 719.951 | 6.5792 | 6.3116 |  |  | 18 | 719.951 | 6.5792 | 6.3117
19 | 759.948 | 6.6332 | 6.2826 |  |  | 19 | 759.948 | 6.6332 | 6.2839 |  |  | 19 | 759.948 | 6.6332 | 6.284
20 | 799.945 | 6.6845 | 6.2563 |  |  | 20 | 799.945 | 6.6845 | 6.2577 |  |  | 20 | 799.945 | 6.6845 | 6.2577
21 | 839.942 | 6.7333 | 6.2314 |  |  | 21 | 839.942 | 6.7333 | 6.2328 |  |  | 21 | 839.942 | 6.7333 | 6.2327
22 | 879.94 | 6.7799 | 6.2076 |  |  | 22 | 879.94 | 6.7799 | 6.209 |  |  | 22 | 879.94 | 6.7799 | 6.2088
23 | 919.937 | 6.8243 | 6.1849 |  |  | 23 | 919.937 | 6.8243 | 6.1863 |  |  | 23 | 919.937 | 6.8243 | 6.1861
24 | 959.934 | 6.8669 | 6.1631 |  |  | 24 | 959.934 | 6.8669 | 6.1646 |  |  | 24 | 959.934 | 6.8669 | 6.1643
25 | 999.931 | 6.9077 | 6.1422 |  |  | 25 | 999.931 | 6.9077 | 6.1438 |  |  | 25 | 999.931 | 6.9077 | 6.1434
26 | 1039.929 | 6.9469 | 6.1222 |  |  | 26 | 1039.929 | 6.9469 | 6.1238 |  |  | 26 | 1039.929 | 6.9469 | 6.1233
27 | 1079.926 | 6.9846 | 6.1029 |  |  | 27 | 1079.926 | 6.9846 | 6.1045 |  |  | 27 | 1079.926 | 6.9846 | 6.104
28 | 1119.923 | 7.021 | 6.0843 |  |  | 28 | 1119.923 | 7.021 | 6.086 |  |  | 28 | 1119.923 | 7.021 | 6.0855
29 | 1159.92 | 7.0561 | 6.0664 |  |  | 29 | 1159.92 | 7.0561 | 6.0681 |  |  | 29 | 1159.92 | 7.0561 | 6.0675
30 | 1199.918 | 7.09 | 6.0491 |  |  | 30 | 1199.918 | 7.09 | 6.0508 |  |  | 30 | 1199.918 | 7.09 | 6.0502
31 | 1239.915 | 7.1228 | 6.0323 |  |  | 31 | 1239.915 | 7.1228 | 6.0341 |  |  | 31 | 1239.915 | 7.1228 | 6.0335
32 | 1279.912 | 7.1545 | 6.0161 |  |  | 32 | 1279.912 | 7.1545 | 6.018 |  |  | 32 | 1279.912 | 7.1545 | 6.0173
33 | 1319.909 | 7.1853 | 6.0004 |  |  | 33 | 1319.909 | 7.1853 | 6.0023 |  |  | 33 | 1319.909 | 7.1853 | 6.0016
34 | 1359.907 | 7.2152 | 5.9852 |  |  | 34 | 1359.907 | 7.2152 | 5.9871 |  |  | 34 | 1359.907 | 7.2152 | 5.9864
35 | 1399.904 | 7.2442 | 5.9704 |  |  | 35 | 1399.904 | 7.2442 | 5.9724 |  |  | 35 | 1399.904 | 7.2442 | 5.9717
36 | 1439.901 | 7.2723 | 5.9561 |  |  | 36 | 1439.901 | 7.2723 | 5.9581 |  |  | 36 | 1439.901 | 7.2723 | 5.9574
37 | 1479.898 | 7.2997 | 5.9421 |  |  | 37 | 1479.898 | 7.2997 | 5.9442 |  |  | 37 | 1479.898 | 7.2997 | 5.9435
38 | 1519.896 | 7.3264 | 5.9286 |  |  | 38 | 1519.896 | 7.3264 | 5.9307 |  |  | 38 | 1519.896 | 7.3264 | 5.9299
39 | 1559.893 | 7.3524 | 5.9154 |  |  | 39 | 1559.893 | 7.3524 | 5.9175 |  |  | 39 | 1559.893 | 7.3524 | 5.9167
40 | 1599.89 | 7.3777 | 5.9025 |  |  | 40 | 1599.89 | 7.3777 | 5.9047 |  |  | 40 | 1599.89 | 7.3777 | 5.9039
41 | 1639.887 | 7.4024 | 5.89 |  |  | 41 | 1639.887 | 7.4024 | 5.8922 |  |  | 41 | 1639.887 | 7.4024 | 5.8914
42 | 1679.885 | 7.4265 | 5.8778 |  |  | 42 | 1679.885 | 7.4265 | 5.88 |  |  | 42 | 1679.885 | 7.4265 | 5.8792
43 | 1719.882 | 7.45 | 5.8658 |  |  | 43 | 1719.882 | 7.45 | 5.8681 |  |  | 43 | 1719.882 | 7.45 | 5.8673
44 | 1759.879 | 7.473 | 5.8542 |  |  | 44 | 1759.879 | 7.473 | 5.8565 |  |  | 44 | 1759.879 | 7.473 | 5.8557
45 | 1799.876 | 7.4955 | 5.8428 |  |  | 45 | 1799.876 | 7.4955 | 5.8452 |  |  | 45 | 1799.876 | 7.4955 | 5.8444
46 | 1839.874 | 7.5175 | 5.8317 |  |  | 46 | 1839.874 | 7.5175 | 5.8341 |  |  | 46 | 1839.874 | 7.5175 | 5.8333
47 | 1879.871 | 7.539 | 5.8208 |  |  | 47 | 1879.871 | 7.539 | 5.8233 |  |  | 47 | 1879.871 | 7.539 | 5.8225
48 | 1919.868 | 7.56 | 5.8102 |  |  | 48 | 1919.868 | 7.56 | 5.8127 |  |  | 48 | 1919.868 | 7.56 | 5.8119
49 | 1959.865 | 7.5806 | 5.7998 |  |  | 49 | 1959.865 | 7.5806 | 5.8023 |  |  | 49 | 1959.865 | 7.5806 | 5.8016
50 | 1999.863 | 7.6008 | 5.7897 |  |  | 50 | 1999.863 | 7.6008 | 5.7922 |  |  | 50 | 1999.863 | 7.6008 | 5.7914
51 | 2039.86 | 7.6206 | 5.7797 |  |  | 51 | 2039.86 | 7.6206 | 5.7822 |  |  | 51 | 2039.86 | 7.6206 | 5.7815
52 | 2079.857 | 7.6401 | 5.7699 |  |  | 52 | 2079.857 | 7.6401 | 5.7725 |  |  | 52 | 2079.857 | 7.6401 | 5.7718
53 | 2119.854 | 7.6591 | 5.7604 |  |  | 53 | 2119.854 | 7.6591 | 5.763 |  |  | 53 | 2119.854 | 7.6591 | 5.7623
54 | 2159.852 | 7.6778 | 5.751 |  |  | 54 | 2159.852 | 7.6778 | 5.7536 |  |  | 54 | 2159.852 | 7.6778 | 5.7529
55 | 2199.849 | 7.6961 | 5.7418 |  |  | 55 | 2199.849 | 7.6961 | 5.7444 |  |  | 55 | 2199.849 | 7.6961 | 5.7438
56 | 2239.846 | 7.7142 | 5.7328 |  |  | 56 | 2239.846 | 7.7142 | 5.7355 |  |  | 56 | 2239.846 | 7.7142 | 5.7348
57 | 2279.843 | 7.7319 | 5.724 |  |  | 57 | 2279.843 | 7.7319 | 5.7266 |  |  | 57 | 2279.843 | 7.7319 | 5.726
58 | 2319.841 | 7.7493 | 5.7153 |  |  | 58 | 2319.841 | 7.7493 | 5.718 |  |  | 58 | 2319.841 | 7.7493 | 5.7174
59 | 2359.838 | 7.7663 | 5.7068 |  |  | 59 | 2359.838 | 7.7663 | 5.7095 |  |  | 59 | 2359.838 | 7.7663 | 5.7089
60 | 2399.835 | 7.7832 | 5.6985 |  |  | 60 | 2399.835 | 7.7832 | 5.7012 |  |  | 60 | 2399.835 | 7.7832 | 5.7006
61 | 2439.832 | 7.7997 | 5.6902 |  |  | 61 | 2439.832 | 7.7997 | 5.693 |  |  | 61 | 2439.832 | 7.7997 | 5.6924
62 | 2479.83 | 7.8159 | 5.6822 |  |  | 62 | 2479.83 | 7.8159 | 5.6849 |  |  | 62 | 2479.83 | 7.8159 | 5.6844
63 | 2519.827 | 7.8319 | 5.6743 |  |  | 63 | 2519.827 | 7.8319 | 5.677 |  |  | 63 | 2519.827 | 7.8319 | 5.6765
64 | 2559.824 | 7.8477 | 5.6665 |  |  | 64 | 2559.824 | 7.8477 | 5.6693 |  |  | 64 | 2559.824 | 7.8477 | 5.6688
65 | 2599.821 | 7.8632 | 5.6588 |  |  | 65 | 2599.821 | 7.8632 | 5.6616 |  |  | 65 | 2599.821 | 7.8632 | 5.6611
66 | 2639.819 | 7.8785 | 5.6513 |  |  | 66 | 2639.819 | 7.8785 | 5.6541 |  |  | 66 | 2639.819 | 7.8785 | 5.6537
67 | 2679.816 | 7.8935 | 5.6439 |  |  | 67 | 2679.816 | 7.8935 | 5.6467 |  |  | 67 | 2679.816 | 7.8935 | 5.6463
68 | 2719.813 | 7.9083 | 5.6366 |  |  | 68 | 2719.813 | 7.9083 | 5.6395 |  |  | 68 | 2719.813 | 7.9083 | 5.639
69 | 2759.81 | 7.9229 | 5.6295 |  |  | 69 | 2759.81 | 7.9229 | 5.6323 |  |  | 69 | 2759.81 | 7.9229 | 5.6319
70 | 2799.808 | 7.9373 | 5.6225 |  |  | 70 | 2799.808 | 7.9373 | 5.6253 |  |  | 70 | 2799.808 | 7.9373 | 5.6249
71 | 2839.805 | 7.9515 | 5.6155 |  |  | 71 | 2839.805 | 7.9515 | 5.6184 |  |  | 71 | 2839.805 | 7.9515 | 5.618
72 | 2879.802 | 7.9655 | 5.6087 |  |  | 72 | 2879.802 | 7.9655 | 5.6116 |  |  | 72 | 2879.802 | 7.9655 | 5.6112
73 | 2919.799 | 7.9793 | 5.602 |  |  | 73 | 2919.799 | 7.9793 | 5.6049 |  |  | 73 | 2919.799 | 7.9793 | 5.6046
74 | 2959.797 | 7.9929 | 5.5954 |  |  | 74 | 2959.797 | 7.9929 | 5.5983 |  |  | 74 | 2959.797 | 7.9929 | 5.598
75 | 2999.794 | 8.0063 | 5.5889 |  |  | 75 | 2999.794 | 8.0063 | 5.5918 |  |  | 75 | 2999.794 | 8.0063 | 5.5915
76 | 3039.791 | 8.0195 | 5.5825 |  |  | 76 | 3039.791 | 8.0195 | 5.5853 |  |  | 76 | 3039.791 | 8.0195 | 5.5851
77 | 3079.788 | 8.0326 | 5.5762 |  |  | 77 | 3079.788 | 8.0326 | 5.579 |  |  | 77 | 3079.788 | 8.0326 | 5.5788
78 | 3119.786 | 8.0455 | 5.5699 |  |  | 78 | 3119.786 | 8.0455 | 5.5728 |  |  | 78 | 3119.786 | 8.0455 | 5.5727
79 | 3159.783 | 8.0583 | 5.5638 |  |  | 79 | 3159.783 | 8.0583 | 5.5667 |  |  | 79 | 3159.783 | 8.0583 | 5.5666
80 | 3199.78 | 8.0708 | 5.5578 |  |  | 80 | 3199.78 | 8.0708 | 5.5607 |  |  | 80 | 3199.78 | 8.0708 | 5.5606
81 | 3239.777 | 8.0833 | 5.5518 |  |  | 81 | 3239.777 | 8.0833 | 5.5547 |  |  | 81 | 3239.777 | 8.0833 | 5.5546
82 | 3279.775 | 8.0955 | 5.546 |  |  | 82 | 3279.775 | 8.0955 | 5.5489 |  |  | 82 | 3279.775 | 8.0955 | 5.5488
83 | 3319.772 | 8.1077 | 5.5402 |  |  | 83 | 3319.772 | 8.1077 | 5.5431 |  |  | 83 | 3319.772 | 8.1077 | 5.543
84 | 3359.769 | 8.1196 | 5.5345 |  |  | 84 | 3359.769 | 8.1196 | 5.5374 |  |  | 84 | 3359.769 | 8.1196 | 5.5374
85 | 3399.766 | 8.1315 | 5.5289 |  |  | 85 | 3399.766 | 8.1315 | 5.5318 |  |  | 85 | 3399.766 | 8.1315 | 5.5318
86 | 3439.764 | 8.1432 | 5.5233 |  |  | 86 | 3439.764 | 8.1432 | 5.5262 |  |  | 86 | 3439.764 | 8.1432 | 5.5263
87 | 3479.761 | 8.1547 | 5.5179 |  |  | 87 | 3479.761 | 8.1547 | 5.5208 |  |  | 87 | 3479.761 | 8.1547 | 5.5208
88 | 3519.758 | 8.1661 | 5.5125 |  |  | 88 | 3519.758 | 8.1661 | 5.5154 |  |  | 88 | 3519.758 | 8.1661 | 5.5155
89 | 3559.755 | 8.1774 | 5.5072 |  |  | 89 | 3559.755 | 8.1774 | 5.5101 |  |  | 89 | 3559.755 | 8.1774 | 5.5102
90 | 3599.753 | 8.1886 | 5.5019 |  |  | 90 | 3599.753 | 8.1886 | 5.5048 |  |  | 90 | 3599.753 | 8.1886 | 5.505
91 | 3639.75 | 8.1997 | 5.4968 |  |  | 91 | 3639.75 | 8.1997 | 5.4997 |  |  | 91 | 3639.75 | 8.1997 | 5.4998
92 | 3679.747 | 8.2106 | 5.4917 |  |  | 92 | 3679.747 | 8.2106 | 5.4946 |  |  | 92 | 3679.747 | 8.2106 | 5.4947
93 | 3719.744 | 8.2214 | 5.4866 |  |  | 93 | 3719.744 | 8.2214 | 5.4895 |  |  | 93 | 3719.744 | 8.2214 | 5.4897
94 | 3759.742 | 8.2321 | 5.4817 |  |  | 94 | 3759.742 | 8.2321 | 5.4845 |  |  | 94 | 3759.742 | 8.2321 | 5.4848
95 | 3799.739 | 8.2427 | 5.4768 |  |  | 95 | 3799.739 | 8.2427 | 5.4796 |  |  | 95 | 3799.739 | 8.2427 | 5.4799
96 | 3839.736 | 8.2532 | 5.4719 |  |  | 96 | 3839.736 | 8.2532 | 5.4748 |  |  | 96 | 3839.736 | 8.2532 | 5.4751
97 | 3879.733 | 8.2635 | 5.4671 |  |  | 97 | 3879.733 | 8.2635 | 5.47 |  |  | 97 | 3879.733 | 8.2635 | 5.4703
98 | 3919.731 | 8.2738 | 5.4624 |  |  | 98 | 3919.731 | 8.2738 | 5.4653 |  |  | 98 | 3919.731 | 8.2738 | 5.4656
99 | 3959.728 | 8.2839 | 5.4578 |  |  | 99 | 3959.728 | 8.2839 | 5.4606 |  |  | 99 | 3959.728 | 8.2839 | 5.461
100 | 3999.725 | 8.294 | 5.4532 |  |  | 100 | 3999.725 | 8.294 | 5.456 |  |  | 100 | 3999.725 | 8.294 | 5.4564
101 | 4039.722 | 8.3039 | 5.4486 |  |  | 101 | 4039.722 | 8.3039 | 5.4515 |  |  | 101 | 4039.722 | 8.3039 | 5.4519
102 | 4079.72 | 8.3138 | 5.4442 |  |  | 102 | 4079.72 | 8.3138 | 5.447 |  |  | 102 | 4079.72 | 8.3138 | 5.4474
103 | 4119.717 | 8.3235 | 5.4397 |  |  | 103 | 4119.717 | 8.3235 | 5.4426 |  |  | 103 | 4119.717 | 8.3235 | 5.443
104 | 4159.714 | 8.3332 | 5.4354 |  |  | 104 | 4159.714 | 8.3332 | 5.4382 |  |  | 104 | 4159.714 | 8.3332 | 5.4386
105 | 4199.711 | 8.3428 | 5.431 |  |  | 105 | 4199.711 | 8.3428 | 5.4338 |  |  | 105 | 4199.711 | 8.3428 | 5.4343
106 | 4239.709 | 8.3522 | 5.4268 |  |  | 106 | 4239.709 | 8.3522 | 5.4296 |  |  | 106 | 4239.709 | 8.3522 | 5.4301
107 | 4279.706 | 8.3616 | 5.4226 |  |  | 107 | 4279.706 | 8.3616 | 5.4254 |  |  | 107 | 4279.706 | 8.3616 | 5.4259
108 | 4319.703 | 8.3709 | 5.4184 |  |  | 108 | 4319.703 | 8.3709 | 5.4212 |  |  | 108 | 4319.703 | 8.3709 | 5.4217
109 | 4359.7 | 8.3802 | 5.4143 |  |  | 109 | 4359.7 | 8.3802 | 5.4171 |  |  | 109 | 4359.7 | 8.3802 | 5.4176
110 | 4399.698 | 8.3893 | 5.4102 |  |  | 110 | 4399.698 | 8.3893 | 5.413 |  |  | 110 | 4399.698 | 8.3893 | 5.4136
111 | 4439.695 | 8.3983 | 5.4062 |  |  | 111 | 4439.695 | 8.3983 | 5.409 |  |  | 111 | 4439.695 | 8.3983 | 5.4096
112 | 4479.692 | 8.4073 | 5.4022 |  |  | 112 | 4479.692 | 8.4073 | 5.405 |  |  | 112 | 4479.692 | 8.4073 | 5.4056
113 | 4519.689 | 8.4162 | 5.3983 |  |  | 113 | 4519.689 | 8.4162 | 5.4011 |  |  | 113 | 4519.689 | 8.4162 | 5.4017
114 | 4559.687 | 8.425 | 5.3944 |  |  | 114 | 4559.687 | 8.425 | 5.3972 |  |  | 114 | 4559.687 | 8.425 | 5.3979
115 | 4599.684 | 8.4337 | 5.3906 |  |  | 115 | 4599.684 | 8.4337 | 5.3933 |  |  | 115 | 4599.684 | 8.4337 | 5.3941
116 | 4639.681 | 8.4424 | 5.3868 |  |  | 116 | 4639.681 | 8.4424 | 5.3895 |  |  | 116 | 4639.681 | 8.4424 | 5.3903
117 | 4679.678 | 8.451 | 5.3831 |  |  | 117 | 4679.678 | 8.451 | 5.3858 |  |  | 117 | 4679.678 | 8.451 | 5.3866
118 | 4719.676 | 8.4595 | 5.3794 |  |  | 118 | 4719.676 | 8.4595 | 5.3821 |  |  | 118 | 4719.676 | 8.4595 | 5.3829
119 | 4759.673 | 8.4679 | 5.3757 |  |  | 119 | 4759.673 | 8.4679 | 5.3784 |  |  | 119 | 4759.673 | 8.4679 | 5.3792
120 | 4799.67 | 8.4763 | 5.3721 |  |  | 120 | 4799.67 | 8.4763 | 5.3748 |  |  | 120 | 4799.67 | 8.4763 | 5.3756
121 | 4839.667 | 8.4846 | 5.3685 |  |  | 121 | 4839.667 | 8.4846 | 5.3712 |  |  | 121 | 4839.667 | 8.4846 | 5.3721
122 | 4879.665 | 8.4928 | 5.365 |  |  | 122 | 4879.665 | 8.4928 | 5.3677 |  |  | 122 | 4879.665 | 8.4928 | 5.3685
123 | 4919.662 | 8.501 | 5.3615 |  |  | 123 | 4919.662 | 8.501 | 5.3642 |  |  | 123 | 4919.662 | 8.501 | 5.3651
124 | 4959.659 | 8.5091 | 5.3581 |  |  | 124 | 4959.659 | 8.5091 | 5.3607 |  |  | 124 | 4959.659 | 8.5091 | 5.3616
125 | 4999.656 | 8.5171 | 5.3547 |  |  | 125 | 4999.656 | 8.5171 | 5.3573 |  |  | 125 | 4999.656 | 8.5171 | 5.3582
126 | 5039.654 | 8.5251 | 5.3513 |  |  | 126 | 5039.654 | 8.5251 | 5.3539 |  |  | 126 | 5039.654 | 8.5251 | 5.3548
127 | 5079.651 | 8.533 | 5.3479 |  |  | 127 | 5079.651 | 8.533 | 5.3505 |  |  | 127 | 5079.651 | 8.533 | 5.3515
128 | 5119.648 | 8.5408 | 5.3446 |  |  | 128 | 5119.648 | 8.5408 | 5.3472 |  |  | 128 | 5119.648 | 8.5408 | 5.3482
129 | 5159.645 | 8.5486 | 5.3414 |  |  | 129 | 5159.645 | 8.5486 | 5.3439 |  |  | 129 | 5159.645 | 8.5486 | 5.345
130 | 5199.643 | 8.5563 | 5.3381 |  |  | 130 | 5199.643 | 8.5563 | 5.3407 |  |  | 130 | 5199.643 | 8.5563 | 5.3417
131 | 5239.64 | 8.564 | 5.3349 |  |  | 131 | 5239.64 | 8.564 | 5.3375 |  |  | 131 | 5239.64 | 8.564 | 5.3386
132 | 5279.637 | 8.5716 | 5.3318 |  |  | 132 | 5279.637 | 8.5716 | 5.3343 |  |  | 132 | 5279.637 | 8.5716 | 5.3354
133 | 5319.634 | 8.5792 | 5.3287 |  |  | 133 | 5319.634 | 8.5792 | 5.3312 |  |  | 133 | 5319.634 | 8.5792 | 5.3323
134 | 5359.632 | 8.5867 | 5.3256 |  |  | 134 | 5359.632 | 8.5867 | 5.3281 |  |  | 134 | 5359.632 | 8.5867 | 5.3292
135 | 5399.629 | 8.5941 | 5.3225 |  |  | 135 | 5399.629 | 8.5941 | 5.325 |  |  | 135 | 5399.629 | 8.5941 | 5.3262
136 | 5439.626 | 8.6015 | 5.3195 |  |  | 136 | 5439.626 | 8.6015 | 5.3219 |  |  | 136 | 5439.626 | 8.6015 | 5.3231
137 | 5479.623 | 8.6088 | 5.3165 |  |  | 137 | 5479.623 | 8.6088 | 5.3189 |  |  | 137 | 5479.623 | 8.6088 | 5.3202
138 | 5519.621 | 8.6161 | 5.3135 |  |  | 138 | 5519.621 | 8.6161 | 5.316 |  |  | 138 | 5519.621 | 8.6161 | 5.3172
139 | 5559.618 | 8.6233 | 5.3106 |  |  | 139 | 5559.618 | 8.6233 | 5.313 |  |  | 139 | 5559.618 | 8.6233 | 5.3143
140 | 5599.615 | 8.6305 | 5.3077 |  |  | 140 | 5599.615 | 8.6305 | 5.3101 |  |  | 140 | 5599.615 | 8.6305 | 5.3114
141 | 5639.612 | 8.6376 | 5.3048 |  |  | 141 | 5639.612 | 8.6376 | 5.3072 |  |  | 141 | 5639.612 | 8.6376 | 5.3085
142 | 5679.61 | 8.6446 | 5.302 |  |  | 142 | 5679.61 | 8.6446 | 5.3044 |  |  | 142 | 5679.61 | 8.6446 | 5.3057
143 | 5719.607 | 8.6517 | 5.2992 |  |  | 143 | 5719.607 | 8.6517 | 5.3015 |  |  | 143 | 5719.607 | 8.6517 | 5.3029
144 | 5759.604 | 8.6586 | 5.2964 |  |  | 144 | 5759.604 | 8.6586 | 5.2988 |  |  | 144 | 5759.604 | 8.6586 | 5.3001
145 | 5799.601 | 8.6655 | 5.2937 |  |  | 145 | 5799.601 | 8.6655 | 5.296 |  |  | 145 | 5799.601 | 8.6655 | 5.2974
146 | 5839.599 | 8.6724 | 5.2909 |  |  | 146 | 5839.599 | 8.6724 | 5.2933 |  |  | 146 | 5839.599 | 8.6724 | 5.2946
147 | 5879.596 | 8.6792 | 5.2882 |  |  | 147 | 5879.596 | 8.6792 | 5.2905 |  |  | 147 | 5879.596 | 8.6792 | 5.292
148 | 5919.593 | 8.686 | 5.2856 |  |  | 148 | 5919.593 | 8.686 | 5.2879 |  |  | 148 | 5919.593 | 8.686 | 5.2893
149 | 5959.59 | 8.6928 | 5.2829 |  |  | 149 | 5959.59 | 8.6928 | 5.2852 |  |  | 149 | 5959.59 | 8.6928 | 5.2867
150 | 5999.588 | 8.6994 | 5.2803 |  |  | 150 | 5999.588 | 8.6994 | 5.2826 |  |  | 150 | 5999.588 | 8.6994 | 5.2841
151 | 6039.585 | 8.7061 | 5.2777 |  |  | 151 | 6039.585 | 8.7061 | 5.28 |  |  | 151 | 6039.585 | 8.7061 | 5.2815
152 | 6079.582 | 8.7127 | 5.2752 |  |  | 152 | 6079.582 | 8.7127 | 5.2774 |  |  | 152 | 6079.582 | 8.7127 | 5.2789
153 | 6119.579 | 8.7192 | 5.2727 |  |  | 153 | 6119.579 | 8.7192 | 5.2749 |  |  | 153 | 6119.579 | 8.7192 | 5.2764
154 | 6159.577 | 8.7258 | 5.2702 |  |  | 154 | 6159.577 | 8.7258 | 5.2723 |  |  | 154 | 6159.577 | 8.7258 | 5.2739
155 | 6199.574 | 8.7322 | 5.2677 |  |  | 155 | 6199.574 | 8.7322 | 5.2698 |  |  | 155 | 6199.574 | 8.7322 | 5.2714
156 | 6239.571 | 8.7387 | 5.2652 |  |  | 156 | 6239.571 | 8.7387 | 5.2674 |  |  | 156 | 6239.571 | 8.7387 | 5.269
157 | 6279.568 | 8.7451 | 5.2628 |  |  | 157 | 6279.568 | 8.7451 | 5.2649 |  |  | 157 | 6279.568 | 8.7451 | 5.2665
158 | 6319.566 | 8.7514 | 5.2604 |  |  | 158 | 6319.566 | 8.7514 | 5.2625 |  |  | 158 | 6319.566 | 8.7514 | 5.2641
159 | 6359.563 | 8.7577 | 5.258 |  |  | 159 | 6359.563 | 8.7577 | 5.2601 |  |  | 159 | 6359.563 | 8.7577 | 5.2617
160 | 6399.56 | 8.764 | 5.2556 |  |  | 160 | 6399.56 | 8.764 | 5.2577 |  |  | 160 | 6399.56 | 8.764 | 5.2594
161 | 6439.557 | 8.7702 | 5.2533 |  |  | 161 | 6439.557 | 8.7702 | 5.2554 |  |  | 161 | 6439.557 | 8.7702 | 5.2571
162 | 6479.555 | 8.7764 | 5.251 |  |  | 162 | 6479.555 | 8.7764 | 5.253 |  |  | 162 | 6479.555 | 8.7764 | 5.2547
163 | 6519.552 | 8.7826 | 5.2487 |  |  | 163 | 6519.552 | 8.7826 | 5.2507 |  |  | 163 | 6519.552 | 8.7826 | 5.2525
164 | 6559.549 | 8.7887 | 5.2464 |  |  | 164 | 6559.549 | 8.7887 | 5.2484 |  |  | 164 | 6559.549 | 8.7887 | 5.2502
165 | 6599.546 | 8.7948 | 5.2442 |  |  | 165 | 6599.546 | 8.7948 | 5.2462 |  |  | 165 | 6599.546 | 8.7948 | 5.2479
166 | 6639.544 | 8.8008 | 5.242 |  |  | 166 | 6639.544 | 8.8008 | 5.2439 |  |  | 166 | 6639.544 | 8.8008 | 5.2457
167 | 6679.541 | 8.8068 | 5.2398 |  |  | 167 | 6679.541 | 8.8068 | 5.2417 |  |  | 167 | 6679.541 | 8.8068 | 5.2435
168 | 6719.538 | 8.8128 | 5.2376 |  |  | 168 | 6719.538 | 8.8128 | 5.2395 |  |  | 168 | 6719.538 | 8.8128 | 5.2413
169 | 6759.535 | 8.8187 | 5.2354 |  |  | 169 | 6759.535 | 8.8187 | 5.2374 |  |  | 169 | 6759.535 | 8.8187 | 5.2392
170 | 6799.533 | 8.8246 | 5.2333 |  |  | 170 | 6799.533 | 8.8246 | 5.2352 |  |  | 170 | 6799.533 | 8.8246 | 5.2371
171 | 6839.53 | 8.8305 | 5.2312 |  |  | 171 | 6839.53 | 8.8305 | 5.2331 |  |  | 171 | 6839.53 | 8.8305 | 5.2349
172 | 6879.527 | 8.8363 | 5.2291 |  |  | 172 | 6879.527 | 8.8363 | 5.231 |  |  | 172 | 6879.527 | 8.8363 | 5.2328
173 | 6919.524 | 8.8421 | 5.227 |  |  | 173 | 6919.524 | 8.8421 | 5.2289 |  |  | 173 | 6919.524 | 8.8421 | 5.2308
174 | 6959.522 | 8.8479 | 5.2249 |  |  | 174 | 6959.522 | 8.8479 | 5.2268 |  |  | 174 | 6959.522 | 8.8479 | 5.2287
175 | 6999.519 | 8.8536 | 5.2229 |  |  | 175 | 6999.519 | 8.8536 | 5.2247 |  |  | 175 | 6999.519 | 8.8536 | 5.2267
176 | 7039.516 | 8.8593 | 5.2209 |  |  | 176 | 7039.516 | 8.8593 | 5.2227 |  |  | 176 | 7039.516 | 8.8593 | 5.2247
177 | 7079.513 | 8.865 | 5.2189 |  |  | 177 | 7079.513 | 8.865 | 5.2207 |  |  | 177 | 7079.513 | 8.865 | 5.2227
178 | 7119.511 | 8.8706 | 5.2169 |  |  | 178 | 7119.511 | 8.8706 | 5.2187 |  |  | 178 | 7119.511 | 8.8706 | 5.2207
179 | 7159.508 | 8.8762 | 5.2149 |  |  | 179 | 7159.508 | 8.8762 | 5.2167 |  |  | 179 | 7159.508 | 8.8762 | 5.2187
180 | 7199.505 | 8.8818 | 5.213 |  |  | 180 | 7199.505 | 8.8818 | 5.2148 |  |  | 180 | 7199.505 | 8.8818 | 5.2168
181 | 7239.502 | 8.8873 | 5.2111 |  |  | 181 | 7239.502 | 8.8873 | 5.2128 |  |  | 181 | 7239.502 | 8.8873 | 5.2149
182 | 7279.5 | 8.8928 | 5.2092 |  |  | 182 | 7279.5 | 8.8928 | 5.2109 |  |  | 182 | 7279.5 | 8.8928 | 5.213
183 | 7319.497 | 8.8983 | 5.2073 |  |  | 183 | 7319.497 | 8.8983 | 5.209 |  |  | 183 | 7319.497 | 8.8983 | 5.2111
184 | 7359.494 | 8.9037 | 5.2054 |  |  | 184 | 7359.494 | 8.9037 | 5.2071 |  |  | 184 | 7359.494 | 8.9037 | 5.2092
185 | 7399.491 | 8.9092 | 5.2036 |  |  | 185 | 7399.491 | 8.9092 | 5.2052 |  |  | 185 | 7399.491 | 8.9092 | 5.2073
186 | 7439.489 | 8.9146 | 5.2017 |  |  | 186 | 7439.489 | 8.9146 | 5.2034 |  |  | 186 | 7439.489 | 8.9146 | 5.2055
187 | 7479.486 | 8.9199 | 5.1999 |  |  | 187 | 7479.486 | 8.9199 | 5.2015 |  |  | 187 | 7479.486 | 8.9199 | 5.2037
188 | 7519.483 | 8.9253 | 5.1981 |  |  | 188 | 7519.483 | 8.9253 | 5.1997 |  |  | 188 | 7519.483 | 8.9253 | 5.2019
189 | 7559.48 | 8.9306 | 5.1963 |  |  | 189 | 7559.48 | 8.9306 | 5.1979 |  |  | 189 | 7559.48 | 8.9306 | 5.2001
190 | 7599.478 | 8.9358 | 5.1946 |  |  | 190 | 7599.478 | 8.9358 | 5.1961 |  |  | 190 | 7599.478 | 8.9358 | 5.1983
191 | 7639.475 | 8.9411 | 5.1928 |  |  | 191 | 7639.475 | 8.9411 | 5.1944 |  |  | 191 | 7639.475 | 8.9411 | 5.1966
192 | 7679.472 | 8.9463 | 5.1911 |  |  | 192 | 7679.472 | 8.9463 | 5.1926 |  |  | 192 | 7679.472 | 8.9463 | 5.1949
193 | 7719.469 | 8.9515 | 5.1894 |  |  | 193 | 7719.469 | 8.9515 | 5.1909 |  |  | 193 | 7719.469 | 8.9515 | 5.1931
194 | 7759.467 | 8.9567 | 5.1877 |  |  | 194 | 7759.467 | 8.9567 | 5.1892 |  |  | 194 | 7759.467 | 8.9567 | 5.1914
195 | 7799.464 | 8.9618 | 5.186 |  |  | 195 | 7799.464 | 8.9618 | 5.1875 |  |  | 195 | 7799.464 | 8.9618 | 5.1897
196 | 7839.461 | 8.9669 | 5.1843 |  |  | 196 | 7839.461 | 8.9669 | 5.1858 |  |  | 196 | 7839.461 | 8.9669 | 5.1881
197 | 7879.458 | 8.972 | 5.1827 |  |  | 197 | 7879.458 | 8.972 | 5.1841 |  |  | 197 | 7879.458 | 8.972 | 5.1864
198 | 7919.456 | 8.9771 | 5.181 |  |  | 198 | 7919.456 | 8.9771 | 5.1824 |  |  | 198 | 7919.456 | 8.9771 | 5.1848
199 | 7959.453 | 8.9821 | 5.1794 |  |  | 199 | 7959.453 | 8.9821 | 5.1808 |  |  | 199 | 7959.453 | 8.9821 | 5.1831
200 | 7999.45 | 8.9871 | 5.1778 |  |  | 200 | 7999.45 | 8.9871 | 5.1792 |  |  | 200 | 7999.45 | 8.9871 | 5.1815
201 | 8039.447 | 8.9921 | 5.1762 |  |  | 201 | 8039.447 | 8.9921 | 5.1776 |  |  | 201 | 8039.447 | 8.9921 | 5.1799
202 | 8079.445 | 8.9971 | 5.1746 |  |  | 202 | 8079.445 | 8.9971 | 5.176 |  |  | 202 | 8079.445 | 8.9971 | 5.1784
203 | 8119.442 | 9.002 | 5.173 |  |  | 203 | 8119.442 | 9.002 | 5.1744 |  |  | 203 | 8119.442 | 9.002 | 5.1768
204 | 8159.439 | 9.0069 | 5.1715 |  |  | 204 | 8159.439 | 9.0069 | 5.1728 |  |  | 204 | 8159.439 | 9.0069 | 5.1752
205 | 8199.436 | 9.0118 | 5.1699 |  |  | 205 | 8199.436 | 9.0118 | 5.1712 |  |  | 205 | 8199.436 | 9.0118 | 5.1737
206 | 8239.434 | 9.0167 | 5.1684 |  |  | 206 | 8239.434 | 9.0167 | 5.1697 |  |  | 206 | 8239.434 | 9.0167 | 5.1722
207 | 8279.431 | 9.0215 | 5.1669 |  |  | 207 | 8279.431 | 9.0215 | 5.1682 |  |  | 207 | 8279.431 | 9.0215 | 5.1706
208 | 8319.428 | 9.0263 | 5.1654 |  |  | 208 | 8319.428 | 9.0263 | 5.1667 |  |  | 208 | 8319.428 | 9.0263 | 5.1691
209 | 8359.425 | 9.0311 | 5.1639 |  |  | 209 | 8359.425 | 9.0311 | 5.1652 |  |  | 209 | 8359.425 | 9.0311 | 5.1677
210 | 8399.423 | 9.0359 | 5.1625 |  |  | 210 | 8399.423 | 9.0359 | 5.1637 |  |  | 210 | 8399.423 | 9.0359 | 5.1662
211 | 8439.42 | 9.0407 | 5.161 |  |  | 211 | 8439.42 | 9.0407 | 5.1622 |  |  | 211 | 8439.42 | 9.0407 | 5.1647
212 | 8479.417 | 9.0454 | 5.1595 |  |  | 212 | 8479.417 | 9.0454 | 5.1607 |  |  | 212 | 8479.417 | 9.0454 | 5.1633
213 | 8519.414 | 9.0501 | 5.1581 |  |  | 213 | 8519.414 | 9.0501 | 5.1593 |  |  | 213 | 8519.414 | 9.0501 | 5.1618
214 | 8559.412 | 9.0548 | 5.1567 |  |  | 214 | 8559.412 | 9.0548 | 5.1578 |  |  | 214 | 8559.412 | 9.0548 | 5.1604
215 | 8599.409 | 9.0594 | 5.1553 |  |  | 215 | 8599.409 | 9.0594 | 5.1564 |  |  | 215 | 8599.409 | 9.0594 | 5.159
216 | 8639.406 | 9.0641 | 5.1539 |  |  | 216 | 8639.406 | 9.0641 | 5.155 |  |  | 216 | 8639.406 | 9.0641 | 5.1576
217 | 8679.403 | 9.0687 | 5.1525 |  |  | 217 | 8679.403 | 9.0687 | 5.1536 |  |  | 217 | 8679.403 | 9.0687 | 5.1562
218 | 8719.401 | 9.0733 | 5.1511 |  |  | 218 | 8719.401 | 9.0733 | 5.1522 |  |  | 218 | 8719.401 | 9.0733 | 5.1549
219 | 8759.398 | 9.0779 | 5.1498 |  |  | 219 | 8759.398 | 9.0779 | 5.1508 |  |  | 219 | 8759.398 | 9.0779 | 5.1535
220 | 8799.395 | 9.0824 | 5.1484 |  |  | 220 | 8799.395 | 9.0824 | 5.1495 |  |  | 220 | 8799.395 | 9.0824 | 5.1521
221 | 8839.392 | 9.087 | 5.1471 |  |  | 221 | 8839.392 | 9.087 | 5.1481 |  |  | 221 | 8839.392 | 9.087 | 5.1508
222 | 8879.39 | 9.0915 | 5.1458 |  |  | 222 | 8879.39 | 9.0915 | 5.1468 |  |  | 222 | 8879.39 | 9.0915 | 5.1495
223 | 8919.387 | 9.096 | 5.1445 |  |  | 223 | 8919.387 | 9.096 | 5.1455 |  |  | 223 | 8919.387 | 9.096 | 5.1482
224 | 8959.384 | 9.1005 | 5.1432 |  |  | 224 | 8959.384 | 9.1005 | 5.1441 |  |  | 224 | 8959.384 | 9.1005 | 5.1469
225 | 8999.381 | 9.1049 | 5.1419 |  |  | 225 | 8999.381 | 9.1049 | 5.1428 |  |  | 225 | 8999.381 | 9.1049 | 5.1456
226 | 9039.379 | 9.1093 | 5.1406 |  |  | 226 | 9039.379 | 9.1093 | 5.1415 |  |  | 226 | 9039.379 | 9.1093 | 5.1443
227 | 9079.376 | 9.1138 | 5.1393 |  |  | 227 | 9079.376 | 9.1138 | 5.1403 |  |  | 227 | 9079.376 | 9.1138 | 5.143
228 | 9119.373 | 9.1182 | 5.1381 |  |  | 228 | 9119.373 | 9.1182 | 5.139 |  |  | 228 | 9119.373 | 9.1182 | 5.1418
229 | 9159.37 | 9.1225 | 5.1368 |  |  | 229 | 9159.37 | 9.1225 | 5.1377 |  |  | 229 | 9159.37 | 9.1225 | 5.1405
230 | 9199.368 | 9.1269 | 5.1356 |  |  | 230 | 9199.368 | 9.1269 | 5.1365 |  |  | 230 | 9199.368 | 9.1269 | 5.1393
231 | 9239.365 | 9.1312 | 5.1344 |  |  | 231 | 9239.365 | 9.1312 | 5.1352 |  |  | 231 | 9239.365 | 9.1312 | 5.1381
232 | 9279.362 | 9.1355 | 5.1332 |  |  | 232 | 9279.362 | 9.1355 | 5.134 |  |  | 232 | 9279.362 | 9.1355 | 5.1368
233 | 9319.359 | 9.1398 | 5.132 |  |  | 233 | 9319.359 | 9.1398 | 5.1328 |  |  | 233 | 9319.359 | 9.1398 | 5.1356
234 | 9359.357 | 9.1441 | 5.1308 |  |  | 234 | 9359.357 | 9.1441 | 5.1316 |  |  | 234 | 9359.357 | 9.1441 | 5.1344
235 | 9399.354 | 9.1484 | 5.1296 |  |  | 235 | 9399.354 | 9.1484 | 5.1304 |  |  | 235 | 9399.354 | 9.1484 | 5.1333
236 | 9439.351 | 9.1526 | 5.1284 |  |  | 236 | 9439.351 | 9.1526 | 5.1292 |  |  | 236 | 9439.351 | 9.1526 | 5.1321
237 | 9479.348 | 9.1569 | 5.1273 |  |  | 237 | 9479.348 | 9.1569 | 5.128 |  |  | 237 | 9479.348 | 9.1569 | 5.1309
238 | 9519.346 | 9.1611 | 5.1261 |  |  | 238 | 9519.346 | 9.1611 | 5.1268 |  |  | 238 | 9519.346 | 9.1611 | 5.1298
239 | 9559.343 | 9.1653 | 5.125 |  |  | 239 | 9559.343 | 9.1653 | 5.1257 |  |  | 239 | 9559.343 | 9.1653 | 5.1286
240 | 9599.34 | 9.1694 | 5.1238 |  |  | 240 | 9599.34 | 9.1694 | 5.1245 |  |  | 240 | 9599.34 | 9.1694 | 5.1275
241 | 9639.337 | 9.1736 | 5.1227 |  |  | 241 | 9639.337 | 9.1736 | 5.1234 |  |  | 241 | 9639.337 | 9.1736 | 5.1264
242 | 9679.335 | 9.1777 | 5.1216 |  |  | 242 | 9679.335 | 9.1777 | 5.1223 |  |  | 242 | 9679.335 | 9.1777 | 5.1252
243 | 9719.332 | 9.1819 | 5.1205 |  |  | 243 | 9719.332 | 9.1819 | 5.1211 |  |  | 243 | 9719.332 | 9.1819 | 5.1241
244 | 9759.329 | 9.186 | 5.1194 |  |  | 244 | 9759.329 | 9.186 | 5.12 |  |  | 244 | 9759.329 | 9.186 | 5.123
245 | 9799.326 | 9.1901 | 5.1183 |  |  | 245 | 9799.326 | 9.1901 | 5.1189 |  |  | 245 | 9799.326 | 9.1901 | 5.1219
246 | 9839.324 | 9.1941 | 5.1173 |  |  | 246 | 9839.324 | 9.1941 | 5.1178 |  |  | 246 | 9839.324 | 9.1941 | 5.1209
247 | 9879.321 | 9.1982 | 5.1162 |  |  | 247 | 9879.321 | 9.1982 | 5.1167 |  |  | 247 | 9879.321 | 9.1982 | 5.1198
248 | 9919.318 | 9.2022 | 5.1151 |  |  | 248 | 9919.318 | 9.2022 | 5.1157 |  |  | 248 | 9919.318 | 9.2022 | 5.1187
249 | 9959.315 | 9.2063 | 5.1141 |  |  | 249 | 9959.315 | 9.2063 | 5.1146 |  |  | 249 | 9959.315 | 9.2063 | 5.1177
250 | 9999.313 | 9.2103 | 5.113 |  |  | 250 | 9999.313 | 9.2103 | 5.1136 |  |  | 250 | 9999.313 | 9.2103 | 5.1166
251 | 10039.31 | 9.2143 | 5.112 |  |  | 251 | 10039.31 | 9.2143 | 5.1125 |  |  | 251 | 10039.31 | 9.2143 | 5.1156
252 | 10079.307 | 9.2182 | 5.111 |  |  | 252 | 10079.307 | 9.2182 | 5.1115 |  |  | 252 | 10079.307 | 9.2182 | 5.1146
253 | 10119.304 | 9.2222 | 5.11 |  |  | 253 | 10119.304 | 9.2222 | 5.1104 |  |  | 253 | 10119.304 | 9.2222 | 5.1136
254 | 10159.302 | 9.2261 | 5.109 |  |  | 254 | 10159.302 | 9.2261 | 5.1094 |  |  | 254 | 10159.302 | 9.2261 | 5.1126
255 | 10199.299 | 9.2301 | 5.108 |  |  | 255 | 10199.299 | 9.2301 | 5.1084 |  |  | 255 | 10199.299 | 9.2301 | 5.1116
256 | 10239.296 | 9.234 | 5.107 |  |  | 256 | 10239.296 | 9.234 | 5.1074 |  |  | 256 | 10239.296 | 9.234 | 5.1106
257 | 10279.293 | 9.2379 | 5.106 |  |  | 257 | 10279.293 | 9.2379 | 5.1064 |  |  | 257 | 10279.293 | 9.2379 | 5.1096
258 | 10319.291 | 9.2418 | 5.105 |  |  | 258 | 10319.291 | 9.2418 | 5.1054 |  |  | 258 | 10319.291 | 9.2418 | 5.1086
259 | 10359.288 | 9.2456 | 5.1041 |  |  | 259 | 10359.288 | 9.2456 | 5.1044 |  |  | 259 | 10359.288 | 9.2456 | 5.1076
260 | 10399.285 | 9.2495 | 5.1031 |  |  | 260 | 10399.285 | 9.2495 | 5.1034 |  |  | 260 | 10399.285 | 9.2495 | 5.1067
261 | 10439.282 | 9.2533 | 5.1022 |  |  | 261 | 10439.282 | 9.2533 | 5.1025 |  |  | 261 | 10439.282 | 9.2533 | 5.1057
262 | 10479.28 | 9.2572 | 5.1012 |  |  | 262 | 10479.28 | 9.2572 | 5.1015 |  |  | 262 | 10479.28 | 9.2572 | 5.1048
263 | 10519.277 | 9.261 | 5.1003 |  |  | 263 | 10519.277 | 9.261 | 5.1006 |  |  | 263 | 10519.277 | 9.261 | 5.1038
264 | 10559.274 | 9.2648 | 5.0994 |  |  | 264 | 10559.274 | 9.2648 | 5.0996 |  |  | 264 | 10559.274 | 9.2648 | 5.1029
265 | 10599.271 | 9.2685 | 5.0984 |  |  | 265 | 10599.271 | 9.2685 | 5.0987 |  |  | 265 | 10599.271 | 9.2685 | 5.102
266 | 10639.269 | 9.2723 | 5.0975 |  |  | 266 | 10639.269 | 9.2723 | 5.0978 |  |  | 266 | 10639.269 | 9.2723 | 5.1011
267 | 10679.266 | 9.2761 | 5.0966 |  |  | 267 | 10679.266 | 9.2761 | 5.0968 |  |  | 267 | 10679.266 | 9.2761 | 5.1002
268 | 10719.263 | 9.2798 | 5.0957 |  |  | 268 | 10719.263 | 9.2798 | 5.0959 |  |  | 268 | 10719.263 | 9.2798 | 5.0993
269 | 10759.26 | 9.2835 | 5.0948 |  |  | 269 | 10759.26 | 9.2835 | 5.095 |  |  | 269 | 10759.26 | 9.2835 | 5.0984
270 | 10799.258 | 9.2872 | 5.094 |  |  | 270 | 10799.258 | 9.2872 | 5.0941 |  |  | 270 | 10799.258 | 9.2872 | 5.0975
271 | 10839.255 | 9.2909 | 5.0931 |  |  | 271 | 10839.255 | 9.2909 | 5.0932 |  |  | 271 | 10839.255 | 9.2909 | 5.0966
272 | 10879.252 | 9.2946 | 5.0922 |  |  | 272 | 10879.252 | 9.2946 | 5.0924 |  |  | 272 | 10879.252 | 9.2946 | 5.0957
273 | 10919.249 | 9.2983 | 5.0914 |  |  | 273 | 10919.249 | 9.2983 | 5.0915 |  |  | 273 | 10919.249 | 9.2983 | 5.0949
274 | 10959.247 | 9.3019 | 5.0905 |  |  | 274 | 10959.247 | 9.3019 | 5.0906 |  |  | 274 | 10959.247 | 9.3019 | 5.094
275 | 10999.244 | 9.3056 | 5.0897 |  |  | 275 | 10999.244 | 9.3056 | 5.0898 |  |  | 275 | 10999.244 | 9.3056 | 5.0932
276 | 11039.241 | 9.3092 | 5.0888 |  |  | 276 | 11039.241 | 9.3092 | 5.0889 |  |  | 276 | 11039.241 | 9.3092 | 5.0923
277 | 11079.238 | 9.3128 | 5.088 |  |  | 277 | 11079.238 | 9.3128 | 5.088 |  |  | 277 | 11079.238 | 9.3128 | 5.0915
278 | 11119.236 | 9.3164 | 5.0872 |  |  | 278 | 11119.236 | 9.3164 | 5.0872 |  |  | 278 | 11119.236 | 9.3164 | 5.0907
279 | 11159.233 | 9.32 | 5.0864 |  |  | 279 | 11159.233 | 9.32 | 5.0864 |  |  | 279 | 11159.233 | 9.32 | 5.0898
280 | 11199.23 | 9.3236 | 5.0855 |  |  | 280 | 11199.23 | 9.3236 | 5.0855 |  |  | 280 | 11199.23 | 9.3236 | 5.089
281 | 11239.227 | 9.3272 | 5.0847 |  |  | 281 | 11239.227 | 9.3272 | 5.0847 |  |  | 281 | 11239.227 | 9.3272 | 5.0882
282 | 11279.225 | 9.3307 | 5.0839 |  |  | 282 | 11279.225 | 9.3307 | 5.0839 |  |  | 282 | 11279.225 | 9.3307 | 5.0874
283 | 11319.222 | 9.3343 | 5.0831 |  |  | 283 | 11319.222 | 9.3343 | 5.0831 |  |  | 283 | 11319.222 | 9.3343 | 5.0866
284 | 11359.219 | 9.3378 | 5.0824 |  |  | 284 | 11359.219 | 9.3378 | 5.0823 |  |  | 284 | 11359.219 | 9.3378 | 5.0858
285 | 11399.216 | 9.3413 | 5.0816 |  |  | 285 | 11399.216 | 9.3413 | 5.0815 |  |  | 285 | 11399.216 | 9.3413 | 5.085
286 | 11439.214 | 9.3448 | 5.0808 |  |  | 286 | 11439.214 | 9.3448 | 5.0807 |  |  | 286 | 11439.214 | 9.3448 | 5.0842
287 | 11479.211 | 9.3483 | 5.08 |  |  | 287 | 11479.211 | 9.3483 | 5.0799 |  |  | 287 | 11479.211 | 9.3483 | 5.0835
288 | 11519.208 | 9.3518 | 5.0793 |  |  | 288 | 11519.208 | 9.3518 | 5.0791 |  |  | 288 | 11519.208 | 9.3518 | 5.0827
289 | 11559.205 | 9.3552 | 5.0785 |  |  | 289 | 11559.205 | 9.3552 | 5.0784 |  |  | 289 | 11559.205 | 9.3552 | 5.0819
290 | 11599.203 | 9.3587 | 5.0778 |  |  | 290 | 11599.203 | 9.3587 | 5.0776 |  |  | 290 | 11599.203 | 9.3587 | 5.0812
291 | 11639.2 | 9.3621 | 5.077 |  |  | 291 | 11639.2 | 9.3621 | 5.0768 |  |  | 291 | 11639.2 | 9.3621 | 5.0804
292 | 11679.197 | 9.3656 | 5.0763 |  |  | 292 | 11679.197 | 9.3656 | 5.0761 |  |  | 292 | 11679.197 | 9.3656 | 5.0797
293 | 11719.194 | 9.369 | 5.0755 |  |  | 293 | 11719.194 | 9.369 | 5.0753 |  |  | 293 | 11719.194 | 9.369 | 5.079
294 | 11759.192 | 9.3724 | 5.0748 |  |  | 294 | 11759.192 | 9.3724 | 5.0746 |  |  | 294 | 11759.192 | 9.3724 | 5.0782
295 | 11799.189 | 9.3758 | 5.0741 |  |  | 295 | 11799.189 | 9.3758 | 5.0739 |  |  | 295 | 11799.189 | 9.3758 | 5.0775
296 | 11839.186 | 9.3792 | 5.0734 |  |  | 296 | 11839.186 | 9.3792 | 5.0731 |  |  | 296 | 11839.186 | 9.3792 | 5.0768
297 | 11879.183 | 9.3825 | 5.0727 |  |  | 297 | 11879.183 | 9.3825 | 5.0724 |  |  | 297 | 11879.183 | 9.3825 | 5.0761
298 | 11919.181 | 9.3859 | 5.072 |  |  | 298 | 11919.181 | 9.3859 | 5.0717 |  |  | 298 | 11919.181 | 9.3859 | 5.0754
299 | 11959.178 | 9.3893 | 5.0713 |  |  | 299 | 11959.178 | 9.3893 | 5.071 |  |  | 299 | 11959.178 | 9.3893 | 5.0747
300 | 11999.175 | 9.3926 | 5.0706 |  |  | 300 | 11999.175 | 9.3926 | 5.0703 |  |  | 300 | 11999.175 | 9.3926 | 5.074
301 | 12039.172 | 9.3959 | 5.0699 |  |  | 301 | 12039.172 | 9.3959 | 5.0696 |  |  | 301 | 12039.172 | 9.3959 | 5.0733
302 | 12079.17 | 9.3992 | 5.0692 |  |  | 302 | 12079.17 | 9.3992 | 5.0689 |  |  | 302 | 12079.17 | 9.3992 | 5.0726
303 | 12119.167 | 9.4025 | 5.0685 |  |  | 303 | 12119.167 | 9.4025 | 5.0682 |  |  | 303 | 12119.167 | 9.4025 | 5.0719
304 | 12159.164 | 9.4058 | 5.0679 |  |  | 304 | 12159.164 | 9.4058 | 5.0675 |  |  | 304 | 12159.164 | 9.4058 | 5.0712
305 | 12199.161 | 9.4091 | 5.0672 |  |  | 305 | 12199.161 | 9.4091 | 5.0668 |  |  | 305 | 12199.161 | 9.4091 | 5.0705
306 | 12239.159 | 9.4124 | 5.0665 |  |  | 306 | 12239.159 | 9.4124 | 5.0661 |  |  | 306 | 12239.159 | 9.4124 | 5.0699
307 | 12279.156 | 9.4157 | 5.0659 |  |  | 307 | 12279.156 | 9.4157 | 5.0654 |  |  | 307 | 12279.156 | 9.4157 | 5.0692
308 | 12319.153 | 9.4189 | 5.0652 |  |  | 308 | 12319.153 | 9.4189 | 5.0648 |  |  | 308 | 12319.153 | 9.4189 | 5.0686
309 | 12359.15 | 9.4222 | 5.0646 |  |  | 309 | 12359.15 | 9.4222 | 5.0641 |  |  | 309 | 12359.15 | 9.4222 | 5.0679
310 | 12399.148 | 9.4254 | 5.0639 |  |  | 310 | 12399.148 | 9.4254 | 5.0635 |  |  | 310 | 12399.148 | 9.4254 | 5.0673
311 | 12439.145 | 9.4286 | 5.0633 |  |  | 311 | 12439.145 | 9.4286 | 5.0628 |  |  | 311 | 12439.145 | 9.4286 | 5.0666
312 | 12479.142 | 9.4318 | 5.0627 |  |  | 312 | 12479.142 | 9.4318 | 5.0622 |  |  | 312 | 12479.142 | 9.4318 | 5.066
313 | 12519.139 | 9.435 | 5.062 |  |  | 313 | 12519.139 | 9.435 | 5.0615 |  |  | 313 | 12519.139 | 9.435 | 5.0654
314 | 12559.137 | 9.4382 | 5.0614 |  |  | 314 | 12559.137 | 9.4382 | 5.0609 |  |  | 314 | 12559.137 | 9.4382 | 5.0647
315 | 12599.134 | 9.4414 | 5.0608 |  |  | 315 | 12599.134 | 9.4414 | 5.0602 |  |  | 315 | 12599.134 | 9.4414 | 5.0641
316 | 12639.131 | 9.4446 | 5.0602 |  |  | 316 | 12639.131 | 9.4446 | 5.0596 |  |  | 316 | 12639.131 | 9.4446 | 5.0635
317 | 12679.128 | 9.4477 | 5.0596 |  |  | 317 | 12679.128 | 9.4477 | 5.059 |  |  | 317 | 12679.128 | 9.4477 | 5.0629
318 | 12719.126 | 9.4509 | 5.059 |  |  | 318 | 12719.126 | 9.4509 | 5.0584 |  |  | 318 | 12719.126 | 9.4509 | 5.0623
319 | 12759.123 | 9.454 | 5.0584 |  |  | 319 | 12759.123 | 9.454 | 5.0578 |  |  | 319 | 12759.123 | 9.454 | 5.0617
320 | 12799.12 | 9.4571 | 5.0578 |  |  | 320 | 12799.12 | 9.4571 | 5.0571 |  |  | 320 | 12799.12 | 9.4571 | 5.0611
321 | 12839.117 | 9.4603 | 5.0572 |  |  | 321 | 12839.117 | 9.4603 | 5.0565 |  |  | 321 | 12839.117 | 9.4603 | 5.0605
322 | 12879.115 | 9.4634 | 5.0566 |  |  | 322 | 12879.115 | 9.4634 | 5.0559 |  |  | 322 | 12879.115 | 9.4634 | 5.0599
323 | 12919.112 | 9.4665 | 5.056 |  |  | 323 | 12919.112 | 9.4665 | 5.0553 |  |  | 323 | 12919.112 | 9.4665 | 5.0593
324 | 12959.109 | 9.4696 | 5.0554 |  |  | 324 | 12959.109 | 9.4696 | 5.0547 |  |  | 324 | 12959.109 | 9.4696 | 5.0587
325 | 12999.106 | 9.4726 | 5.0549 |  |  | 325 | 12999.106 | 9.4726 | 5.0542 |  |  | 325 | 12999.106 | 9.4726 | 5.0581
326 | 13039.104 | 9.4757 | 5.0543 |  |  | 326 | 13039.104 | 9.4757 | 5.0536 |  |  | 326 | 13039.104 | 9.4757 | 5.0576
327 | 13079.101 | 9.4788 | 5.0537 |  |  | 327 | 13079.101 | 9.4788 | 5.053 |  |  | 327 | 13079.101 | 9.4788 | 5.057
328 | 13119.098 | 9.4818 | 5.0532 |  |  | 328 | 13119.098 | 9.4818 | 5.0524 |  |  | 328 | 13119.098 | 9.4818 | 5.0564
329 | 13159.095 | 9.4849 | 5.0526 |  |  | 329 | 13159.095 | 9.4849 | 5.0519 |  |  | 329 | 13159.095 | 9.4849 | 5.0559
330 | 13199.093 | 9.4879 | 5.0521 |  |  | 330 | 13199.093 | 9.4879 | 5.0513 |  |  | 330 | 13199.093 | 9.4879 | 5.0553
331 | 13239.09 | 9.4909 | 5.0515 |  |  | 331 | 13239.09 | 9.4909 | 5.0507 |  |  | 331 | 13239.09 | 9.4909 | 5.0548
332 | 13279.087 | 9.4939 | 5.051 |  |  | 332 | 13279.087 | 9.4939 | 5.0502 |  |  | 332 | 13279.087 | 9.4939 | 5.0542
333 | 13319.084 | 9.497 | 5.0504 |  |  | 333 | 13319.084 | 9.497 | 5.0496 |  |  | 333 | 13319.084 | 9.497 | 5.0537
334 | 13359.082 | 9.5 | 5.0499 |  |  | 334 | 13359.082 | 9.5 | 5.0491 |  |  | 334 | 13359.082 | 9.5 | 5.0531
335 | 13399.079 | 9.5029 | 5.0494 |  |  | 335 | 13399.079 | 9.5029 | 5.0485 |  |  | 335 | 13399.079 | 9.5029 | 5.0526
336 | 13439.076 | 9.5059 | 5.0489 |  |  | 336 | 13439.076 | 9.5059 | 5.048 |  |  | 336 | 13439.076 | 9.5059 | 5.0521
337 | 13479.073 | 9.5089 | 5.0483 |  |  | 337 | 13479.073 | 9.5089 | 5.0474 |  |  | 337 | 13479.073 | 9.5089 | 5.0515
338 | 13519.071 | 9.5119 | 5.0478 |  |  | 338 | 13519.071 | 9.5119 | 5.0469 |  |  | 338 | 13519.071 | 9.5119 | 5.051
339 | 13559.068 | 9.5148 | 5.0473 |  |  | 339 | 13559.068 | 9.5148 | 5.0464 |  |  | 339 | 13559.068 | 9.5148 | 5.0505
340 | 13599.065 | 9.5178 | 5.0468 |  |  | 340 | 13599.065 | 9.5178 | 5.0458 |  |  | 340 | 13599.065 | 9.5178 | 5.05
341 | 13639.062 | 9.5207 | 5.0463 |  |  | 341 | 13639.062 | 9.5207 | 5.0453 |  |  | 341 | 13639.062 | 9.5207 | 5.0495
342 | 13679.06 | 9.5236 | 5.0458 |  |  | 342 | 13679.06 | 9.5236 | 5.0448 |  |  | 342 | 13679.06 | 9.5236 | 5.0489
343 | 13719.057 | 9.5265 | 5.0453 |  |  | 343 | 13719.057 | 9.5265 | 5.0443 |  |  | 343 | 13719.057 | 9.5265 | 5.0484
344 | 13759.054 | 9.5295 | 5.0448 |  |  | 344 | 13759.054 | 9.5295 | 5.0438 |  |  | 344 | 13759.054 | 9.5295 | 5.0479
345 | 13799.051 | 9.5324 | 5.0443 |  |  | 345 | 13799.051 | 9.5324 | 5.0433 |  |  | 345 | 13799.051 | 9.5324 | 5.0474
346 | 13839.049 | 9.5352 | 5.0438 |  |  | 346 | 13839.049 | 9.5352 | 5.0428 |  |  | 346 | 13839.049 | 9.5352 | 5.0469
347 | 13879.046 | 9.5381 | 5.0433 |  |  | 347 | 13879.046 | 9.5381 | 5.0423 |  |  | 347 | 13879.046 | 9.5381 | 5.0465
348 | 13919.043 | 9.541 | 5.0428 |  |  | 348 | 13919.043 | 9.541 | 5.0418 |  |  | 348 | 13919.043 | 9.541 | 5.046
349 | 13959.04 | 9.5439 | 5.0423 |  |  | 349 | 13959.04 | 9.5439 | 5.0413 |  |  | 349 | 13959.04 | 9.5439 | 5.0455
350 | 13999.038 | 9.5467 | 5.0419 |  |  | 350 | 13999.038 | 9.5467 | 5.0408 |  |  | 350 | 13999.038 | 9.5467 | 5.045
351 | 14039.035 | 9.5496 | 5.0414 |  |  | 351 | 14039.035 | 9.5496 | 5.0403 |  |  | 351 | 14039.035 | 9.5496 | 5.0445
352 | 14079.032 | 9.5524 | 5.0409 |  |  | 352 | 14079.032 | 9.5524 | 5.0398 |  |  | 352 | 14079.032 | 9.5524 | 5.044
353 | 14119.029 | 9.5553 | 5.0405 |  |  | 353 | 14119.029 | 9.5553 | 5.0393 |  |  | 353 | 14119.029 | 9.5553 | 5.0436
354 | 14159.027 | 9.5581 | 5.04 |  |  | 354 | 14159.027 | 9.5581 | 5.0388 |  |  | 354 | 14159.027 | 9.5581 | 5.0431
355 | 14199.024 | 9.5609 | 5.0395 |  |  | 355 | 14199.024 | 9.5609 | 5.0384 |  |  | 355 | 14199.024 | 9.5609 | 5.0426
356 | 14239.021 | 9.5637 | 5.0391 |  |  | 356 | 14239.021 | 9.5637 | 5.0379 |  |  | 356 | 14239.021 | 9.5637 | 5.0422
357 | 14279.018 | 9.5665 | 5.0386 |  |  | 357 | 14279.018 | 9.5665 | 5.0374 |  |  | 357 | 14279.018 | 9.5665 | 5.0417
358 | 14319.016 | 9.5693 | 5.0382 |  |  | 358 | 14319.016 | 9.5693 | 5.037 |  |  | 358 | 14319.016 | 9.5693 | 5.0413
359 | 14359.013 | 9.5721 | 5.0377 |  |  | 359 | 14359.013 | 9.5721 | 5.0365 |  |  | 359 | 14359.013 | 9.5721 | 5.0408
360 | 14399.01 | 9.5749 | 5.0373 |  |  | 360 | 14399.01 | 9.5749 | 5.0361 |  |  | 360 | 14399.01 | 9.5749 | 5.0404
361 | 14439.007 | 9.5777 | 5.0368 |  |  | 361 | 14439.007 | 9.5777 | 5.0356 |  |  | 361 | 14439.007 | 9.5777 | 5.0399
362 | 14479.005 | 9.5805 | 5.0364 |  |  | 362 | 14479.005 | 9.5805 | 5.0351 |  |  | 362 | 14479.005 | 9.5805 | 5.0395
363 | 14519.002 | 9.5832 | 5.036 |  |  | 363 | 14519.002 | 9.5832 | 5.0347 |  |  | 363 | 14519.002 | 9.5832 | 5.039
364 | 14558.999 | 9.586 | 5.0355 |  |  | 364 | 14558.999 | 9.586 | 5.0343 |  |  | 364 | 14558.999 | 9.586 | 5.0386
365 | 14598.996 | 9.5887 | 5.0351 |  |  | 365 | 14598.996 | 9.5887 | 5.0338 |  |  | 365 | 14598.996 | 9.5887 | 5.0382
366 | 14638.994 | 9.5914 | 5.0347 |  |  | 366 | 14638.994 | 9.5914 | 5.0334 |  |  | 366 | 14638.994 | 9.5914 | 5.0377
367 | 14678.991 | 9.5942 | 5.0343 |  |  | 367 | 14678.991 | 9.5942 | 5.0329 |  |  | 367 | 14678.991 | 9.5942 | 5.0373
368 | 14718.988 | 9.5969 | 5.0339 |  |  | 368 | 14718.988 | 9.5969 | 5.0325 |  |  | 368 | 14718.988 | 9.5969 | 5.0369
369 | 14758.985 | 9.5996 | 5.0334 |  |  | 369 | 14758.985 | 9.5996 | 5.0321 |  |  | 369 | 14758.985 | 9.5996 | 5.0365
370 | 14798.983 | 9.6023 | 5.033 |  |  | 370 | 14798.983 | 9.6023 | 5.0316 |  |  | 370 | 14798.983 | 9.6023 | 5.0361
371 | 14838.98 | 9.605 | 5.0326 |  |  | 371 | 14838.98 | 9.605 | 5.0312 |  |  | 371 | 14838.98 | 9.605 | 5.0356
372 | 14878.977 | 9.6077 | 5.0322 |  |  | 372 | 14878.977 | 9.6077 | 5.0308 |  |  | 372 | 14878.977 | 9.6077 | 5.0352
373 | 14918.974 | 9.6104 | 5.0318 |  |  | 373 | 14918.974 | 9.6104 | 5.0304 |  |  | 373 | 14918.974 | 9.6104 | 5.0348
374 | 14958.972 | 9.6131 | 5.0314 |  |  | 374 | 14958.972 | 9.6131 | 5.03 |  |  | 374 | 14958.972 | 9.6131 | 5.0344
375 | 14998.969 | 9.6157 | 5.031 |  |  | 375 | 14998.969 | 9.6157 | 5.0296 |  |  | 375 | 14998.969 | 9.6157 | 5.034
376 | 15038.966 | 9.6184 | 5.0306 |  |  | 376 | 15038.966 | 9.6184 | 5.0291 |  |  | 376 | 15038.966 | 9.6184 | 5.0336
377 | 15078.963 | 9.6211 | 5.0302 |  |  | 377 | 15078.963 | 9.6211 | 5.0287 |  |  | 377 | 15078.963 | 9.6211 | 5.0332
378 | 15118.961 | 9.6237 | 5.0298 |  |  | 378 | 15118.961 | 9.6237 | 5.0283 |  |  | 378 | 15118.961 | 9.6237 | 5.0328
379 | 15158.958 | 9.6263 | 5.0294 |  |  | 379 | 15158.958 | 9.6263 | 5.0279 |  |  | 379 | 15158.958 | 9.6263 | 5.0324
380 | 15198.955 | 9.629 | 5.029 |  |  | 380 | 15198.955 | 9.629 | 5.0275 |  |  | 380 | 15198.955 | 9.629 | 5.032
381 | 15238.952 | 9.6316 | 5.0287 |  |  | 381 | 15238.952 | 9.6316 | 5.0271 |  |  | 381 | 15238.952 | 9.6316 | 5.0316
382 | 15278.95 | 9.6342 | 5.0283 |  |  | 382 | 15278.95 | 9.6342 | 5.0267 |  |  | 382 | 15278.95 | 9.6342 | 5.0313
383 | 15318.947 | 9.6368 | 5.0279 |  |  | 383 | 15318.947 | 9.6368 | 5.0264 |  |  | 383 | 15318.947 | 9.6368 | 5.0309
384 | 15358.944 | 9.6395 | 5.0275 |  |  | 384 | 15358.944 | 9.6395 | 5.026 |  |  | 384 | 15358.944 | 9.6395 | 5.0305
385 | 15398.941 | 9.6421 | 5.0272 |  |  | 385 | 15398.941 | 9.6421 | 5.0256 |  |  | 385 | 15398.941 | 9.6421 | 5.0301
386 | 15438.939 | 9.6446 | 5.0268 |  |  | 386 | 15438.939 | 9.6446 | 5.0252 |  |  | 386 | 15438.939 | 9.6446 | 5.0297
387 | 15478.936 | 9.6472 | 5.0264 |  |  | 387 | 15478.936 | 9.6472 | 5.0248 |  |  | 387 | 15478.936 | 9.6472 | 5.0294
388 | 15518.933 | 9.6498 | 5.0261 |  |  | 388 | 15518.933 | 9.6498 | 5.0244 |  |  | 388 | 15518.933 | 9.6498 | 5.029
389 | 15558.93 | 9.6524 | 5.0257 |  |  | 389 | 15558.93 | 9.6524 | 5.0241 |  |  | 389 | 15558.93 | 9.6524 | 5.0286
390 | 15598.928 | 9.655 | 5.0253 |  |  | 390 | 15598.928 | 9.655 | 5.0237 |  |  | 390 | 15598.928 | 9.655 | 5.0283
391 | 15638.925 | 9.6575 | 5.025 |  |  | 391 | 15638.925 | 9.6575 | 5.0233 |  |  | 391 | 15638.925 | 9.6575 | 5.0279
392 | 15678.922 | 9.6601 | 5.0246 |  |  | 392 | 15678.922 | 9.6601 | 5.0229 |  |  | 392 | 15678.922 | 9.6601 | 5.0275
393 | 15718.919 | 9.6626 | 5.0243 |  |  | 393 | 15718.919 | 9.6626 | 5.0226 |  |  | 393 | 15718.919 | 9.6626 | 5.0272
394 | 15758.917 | 9.6652 | 5.0239 |  |  | 394 | 15758.917 | 9.6652 | 5.0222 |  |  | 394 | 15758.917 | 9.6652 | 5.0268
395 | 15798.914 | 9.6677 | 5.0236 |  |  | 395 | 15798.914 | 9.6677 | 5.0219 |  |  | 395 | 15798.914 | 9.6677 | 5.0265
396 | 15838.911 | 9.6702 | 5.0232 |  |  | 396 | 15838.911 | 9.6702 | 5.0215 |  |  | 396 | 15838.911 | 9.6702 | 5.0261
397 | 15878.908 | 9.6727 | 5.0229 |  |  | 397 | 15878.908 | 9.6727 | 5.0211 |  |  | 397 | 15878.908 | 9.6727 | 5.0258
398 | 15918.906 | 9.6753 | 5.0225 |  |  | 398 | 15918.906 | 9.6753 | 5.0208 |  |  | 398 | 15918.906 | 9.6753 | 5.0254
399 | 15958.903 | 9.6778 | 5.0222 |  |  | 399 | 15958.903 | 9.6778 | 5.0204 |  |  | 399 | 15958.903 | 9.6778 | 5.0251
400 | 15998.9 | 9.6803 | 5.0219 |  |  | 400 | 15998.9 | 9.6803 | 5.0201 |  |  | 400 | 15998.9 | 9.6803 | 5.0247
401 | 16038.897 | 9.6828 | 5.0215 |  |  | 401 | 16038.897 | 9.6828 | 5.0197 |  |  | 401 | 16038.897 | 9.6828 | 5.0244
402 | 16078.895 | 9.6853 | 5.0212 |  |  | 402 | 16078.895 | 9.6853 | 5.0194 |  |  | 402 | 16078.895 | 9.6853 | 5.0241
403 | 16118.892 | 9.6877 | 5.0209 |  |  | 403 | 16118.892 | 9.6877 | 5.019 |  |  | 403 | 16118.892 | 9.6877 | 5.0237
404 | 16158.889 | 9.6902 | 5.0205 |  |  | 404 | 16158.889 | 9.6902 | 5.0187 |  |  | 404 | 16158.889 | 9.6902 | 5.0234
405 | 16198.886 | 9.6927 | 5.0202 |  |  | 405 | 16198.886 | 9.6927 | 5.0184 |  |  | 405 | 16198.886 | 9.6927 | 5.0231
406 | 16238.884 | 9.6952 | 5.0199 |  |  | 406 | 16238.884 | 9.6952 | 5.018 |  |  | 406 | 16238.884 | 9.6952 | 5.0227
407 | 16278.881 | 9.6976 | 5.0196 |  |  | 407 | 16278.881 | 9.6976 | 5.0177 |  |  | 407 | 16278.881 | 9.6976 | 5.0224
408 | 16318.878 | 9.7001 | 5.0192 |  |  | 408 | 16318.878 | 9.7001 | 5.0174 |  |  | 408 | 16318.878 | 9.7001 | 5.0221
409 | 16358.875 | 9.7025 | 5.0189 |  |  | 409 | 16358.875 | 9.7025 | 5.017 |  |  | 409 | 16358.875 | 9.7025 | 5.0218
410 | 16398.873 | 9.705 | 5.0186 |  |  | 410 | 16398.873 | 9.705 | 5.0167 |  |  | 410 | 16398.873 | 9.705 | 5.0214
411 | 16438.87 | 9.7074 | 5.0183 |  |  | 411 | 16438.87 | 9.7074 | 5.0164 |  |  | 411 | 16438.87 | 9.7074 | 5.0211
412 | 16478.867 | 9.7098 | 5.018 |  |  | 412 | 16478.867 | 9.7098 | 5.016 |  |  | 412 | 16478.867 | 9.7098 | 5.0208
413 | 16518.864 | 9.7123 | 5.0177 |  |  | 413 | 16518.864 | 9.7123 | 5.0157 |  |  | 413 | 16518.864 | 9.7123 | 5.0205
414 | 16558.862 | 9.7147 | 5.0174 |  |  | 414 | 16558.862 | 9.7147 | 5.0154 |  |  | 414 | 16558.862 | 9.7147 | 5.0202
415 | 16598.859 | 9.7171 | 5.0171 |  |  | 415 | 16598.859 | 9.7171 | 5.0151 |  |  | 415 | 16598.859 | 9.7171 | 5.0199
416 | 16638.856 | 9.7195 | 5.0168 |  |  | 416 | 16638.856 | 9.7195 | 5.0148 |  |  | 416 | 16638.856 | 9.7195 | 5.0196
417 | 16678.853 | 9.7219 | 5.0165 |  |  | 417 | 16678.853 | 9.7219 | 5.0145 |  |  | 417 | 16678.853 | 9.7219 | 5.0193
418 | 16718.851 | 9.7243 | 5.0162 |  |  | 418 | 16718.851 | 9.7243 | 5.0141 |  |  | 418 | 16718.851 | 9.7243 | 5.019
419 | 16758.848 | 9.7267 | 5.0159 |  |  | 419 | 16758.848 | 9.7267 | 5.0138 |  |  | 419 | 16758.848 | 9.7267 | 5.0187
420 | 16798.845 | 9.7291 | 5.0156 |  |  | 420 | 16798.845 | 9.7291 | 5.0135 |  |  | 420 | 16798.845 | 9.7291 | 5.0184
421 | 16838.842 | 9.7314 | 5.0153 |  |  | 421 | 16838.842 | 9.7314 | 5.0132 |  |  | 421 | 16838.842 | 9.7314 | 5.0181
422 | 16878.84 | 9.7338 | 5.015 |  |  | 422 | 16878.84 | 9.7338 | 5.0129 |  |  | 422 | 16878.84 | 9.7338 | 5.0178
423 | 16918.837 | 9.7362 | 5.0147 |  |  | 423 | 16918.837 | 9.7362 | 5.0126 |  |  | 423 | 16918.837 | 9.7362 | 5.0175
424 | 16958.834 | 9.7385 | 5.0144 |  |  | 424 | 16958.834 | 9.7385 | 5.0123 |  |  | 424 | 16958.834 | 9.7385 | 5.0172
425 | 16998.831 | 9.7409 | 5.0141 |  |  | 425 | 16998.831 | 9.7409 | 5.012 |  |  | 425 | 16998.831 | 9.7409 | 5.0169
426 | 17038.829 | 9.7433 | 5.0138 |  |  | 426 | 17038.829 | 9.7433 | 5.0117 |  |  | 426 | 17038.829 | 9.7433 | 5.0166
427 | 17078.826 | 9.7456 | 5.0135 |  |  | 427 | 17078.826 | 9.7456 | 5.0114 |  |  | 427 | 17078.826 | 9.7456 | 5.0163
428 | 17118.823 | 9.7479 | 5.0133 |  |  | 428 | 17118.823 | 9.7479 | 5.0111 |  |  | 428 | 17118.823 | 9.7479 | 5.016
429 | 17158.82 | 9.7503 | 5.013 |  |  | 429 | 17158.82 | 9.7503 | 5.0108 |  |  | 429 | 17158.82 | 9.7503 | 5.0157
430 | 17198.818 | 9.7526 | 5.0127 |  |  | 430 | 17198.818 | 9.7526 | 5.0105 |  |  | 430 | 17198.818 | 9.7526 | 5.0154
431 | 17238.815 | 9.7549 | 5.0124 |  |  | 431 | 17238.815 | 9.7549 | 5.0102 |  |  | 431 | 17238.815 | 9.7549 | 5.0152
432 | 17278.812 | 9.7572 | 5.0121 |  |  | 432 | 17278.812 | 9.7572 | 5.01 |  |  | 432 | 17278.812 | 9.7572 | 5.0149
433 | 17318.809 | 9.7595 | 5.0119 |  |  | 433 | 17318.809 | 9.7595 | 5.0097 |  |  | 433 | 17318.809 | 9.7595 | 5.0146
434 | 17358.807 | 9.7619 | 5.0116 |  |  | 434 | 17358.807 | 9.7619 | 5.0094 |  |  | 434 | 17358.807 | 9.7619 | 5.0143
435 | 17398.804 | 9.7642 | 5.0113 |  |  | 435 | 17398.804 | 9.7642 | 5.0091 |  |  | 435 | 17398.804 | 9.7642 | 5.014
436 | 17438.801 | 9.7665 | 5.0111 |  |  | 436 | 17438.801 | 9.7665 | 5.0088 |  |  | 436 | 17438.801 | 9.7665 | 5.0138
437 | 17478.798 | 9.7687 | 5.0108 |  |  | 437 | 17478.798 | 9.7687 | 5.0085 |  |  | 437 | 17478.798 | 9.7687 | 5.0135
438 | 17518.796 | 9.771 | 5.0105 |  |  | 438 | 17518.796 | 9.771 | 5.0083 |  |  | 438 | 17518.796 | 9.771 | 5.0132
439 | 17558.793 | 9.7733 | 5.0103 |  |  | 439 | 17558.793 | 9.7733 | 5.008 |  |  | 439 | 17558.793 | 9.7733 | 5.013
440 | 17598.79 | 9.7756 | 5.01 |  |  | 440 | 17598.79 | 9.7756 | 5.0077 |  |  | 440 | 17598.79 | 9.7756 | 5.0127
441 | 17638.787 | 9.7779 | 5.0097 |  |  | 441 | 17638.787 | 9.7779 | 5.0074 |  |  | 441 | 17638.787 | 9.7779 | 5.0124
442 | 17678.785 | 9.7801 | 5.0095 |  |  | 442 | 17678.785 | 9.7801 | 5.0072 |  |  | 442 | 17678.785 | 9.7801 | 5.0122
443 | 17718.782 | 9.7824 | 5.0092 |  |  | 443 | 17718.782 | 9.7824 | 5.0069 |  |  | 443 | 17718.782 | 9.7824 | 5.0119
444 | 17758.779 | 9.7846 | 5.009 |  |  | 444 | 17758.779 | 9.7846 | 5.0066 |  |  | 444 | 17758.779 | 9.7846 | 5.0116
445 | 17798.776 | 9.7869 | 5.0087 |  |  | 445 | 17798.776 | 9.7869 | 5.0064 |  |  | 445 | 17798.776 | 9.7869 | 5.0114
446 | 17838.774 | 9.7891 | 5.0085 |  |  | 446 | 17838.774 | 9.7891 | 5.0061 |  |  | 446 | 17838.774 | 9.7891 | 5.0111
447 | 17878.771 | 9.7914 | 5.0082 |  |  | 447 | 17878.771 | 9.7914 | 5.0058 |  |  | 447 | 17878.771 | 9.7914 | 5.0109
448 | 17918.768 | 9.7936 | 5.008 |  |  | 448 | 17918.768 | 9.7936 | 5.0056 |  |  | 448 | 17918.768 | 9.7936 | 5.0106
449 | 17958.765 | 9.7958 | 5.0077 |  |  | 449 | 17958.765 | 9.7958 | 5.0053 |  |  | 449 | 17958.765 | 9.7958 | 5.0104
450 | 17998.763 | 9.7981 | 5.0075 |  |  | 450 | 17998.763 | 9.7981 | 5.0051 |  |  | 450 | 17998.763 | 9.7981 | 5.0101
451 | 18038.76 | 9.8003 | 5.0072 |  |  | 451 | 18038.76 | 9.8003 | 5.0048 |  |  | 451 | 18038.76 | 9.8003 | 5.0099
452 | 18078.757 | 9.8025 | 5.007 |  |  | 452 | 18078.757 | 9.8025 | 5.0045 |  |  | 452 | 18078.757 | 9.8025 | 5.0096
453 | 18118.755 | 9.8047 | 5.0067 |  |  | 453 | 18118.755 | 9.8047 | 5.0043 |  |  | 453 | 18118.755 | 9.8047 | 5.0094
454 | 18158.752 | 9.8069 | 5.0065 |  |  | 454 | 18158.752 | 9.8069 | 5.004 |  |  | 454 | 18158.752 | 9.8069 | 5.0091
455 | 18198.749 | 9.8091 | 5.0063 |  |  | 455 | 18198.749 | 9.8091 | 5.0038 |  |  | 455 | 18198.749 | 9.8091 | 5.0089
456 | 18238.746 | 9.8113 | 5.006 |  |  | 456 | 18238.746 | 9.8113 | 5.0035 |  |  | 456 | 18238.746 | 9.8113 | 5.0086
457 | 18278.744 | 9.8135 | 5.0058 |  |  | 457 | 18278.744 | 9.8135 | 5.0033 |  |  | 457 | 18278.744 | 9.8135 | 5.0084
458 | 18318.741 | 9.8157 | 5.0055 |  |  | 458 | 18318.741 | 9.8157 | 5.003 |  |  | 458 | 18318.741 | 9.8157 | 5.0081
459 | 18358.738 | 9.8179 | 5.0053 |  |  | 459 | 18358.738 | 9.8179 | 5.0028 |  |  | 459 | 18358.738 | 9.8179 | 5.0079
460 | 18398.735 | 9.82 | 5.0051 |  |  | 460 | 18398.735 | 9.82 | 5.0026 |  |  | 460 | 18398.735 | 9.82 | 5.0077
461 | 18438.733 | 9.8222 | 5.0048 |  |  | 461 | 18438.733 | 9.8222 | 5.0023 |  |  | 461 | 18438.733 | 9.8222 | 5.0074
462 | 18478.73 | 9.8244 | 5.0046 |  |  | 462 | 18478.73 | 9.8244 | 5.0021 |  |  | 462 | 18478.73 | 9.8244 | 5.0072
463 | 18518.727 | 9.8265 | 5.0044 |  |  | 463 | 18518.727 | 9.8265 | 5.0018 |  |  | 463 | 18518.727 | 9.8265 | 5.007
464 | 18558.724 | 9.8287 | 5.0042 |  |  | 464 | 18558.724 | 9.8287 | 5.0016 |  |  | 464 | 18558.724 | 9.8287 | 5.0067
465 | 18598.722 | 9.8308 | 5.0039 |  |  | 465 | 18598.722 | 9.8308 | 5.0013 |  |  | 465 | 18598.722 | 9.8308 | 5.0065
466 | 18638.719 | 9.833 | 5.0037 |  |  | 466 | 18638.719 | 9.833 | 5.0011 |  |  | 466 | 18638.719 | 9.833 | 5.0063
467 | 18678.716 | 9.8351 | 5.0035 |  |  | 467 | 18678.716 | 9.8351 | 5.0009 |  |  | 467 | 18678.716 | 9.8351 | 5.006
468 | 18718.713 | 9.8373 | 5.0033 |  |  | 468 | 18718.713 | 9.8373 | 5.0006 |  |  | 468 | 18718.713 | 9.8373 | 5.0058
469 | 18758.711 | 9.8394 | 5.003 |  |  | 469 | 18758.711 | 9.8394 | 5.0004 |  |  | 469 | 18758.711 | 9.8394 | 5.0056
470 | 18798.708 | 9.8415 | 5.0028 |  |  | 470 | 18798.708 | 9.8415 | 5.0002 |  |  | 470 | 18798.708 | 9.8415 | 5.0054
471 | 18838.705 | 9.8437 | 5.0026 |  |  | 471 | 18838.705 | 9.8437 | 4.9999 |  |  | 471 | 18838.705 | 9.8437 | 5.0051
472 | 18878.702 | 9.8458 | 5.0024 |  |  | 472 | 18878.702 | 9.8458 | 4.9997 |  |  | 472 | 18878.702 | 9.8458 | 5.0049
473 | 18918.7 | 9.8479 | 5.0022 |  |  | 473 | 18918.7 | 9.8479 | 4.9995 |  |  | 473 | 18918.7 | 9.8479 | 5.0047
474 | 18958.697 | 9.85 | 5.0019 |  |  | 474 | 18958.697 | 9.85 | 4.9993 |  |  | 474 | 18958.697 | 9.85 | 5.0045
475 | 18998.694 | 9.8521 | 5.0017 |  |  | 475 | 18998.694 | 9.8521 | 4.999 |  |  | 475 | 18998.694 | 9.8521 | 5.0042
476 | 19038.691 | 9.8542 | 5.0015 |  |  | 476 | 19038.691 | 9.8542 | 4.9988 |  |  | 476 | 19038.691 | 9.8542 | 5.004
477 | 19078.689 | 9.8563 | 5.0013 |  |  | 477 | 19078.689 | 9.8563 | 4.9986 |  |  | 477 | 19078.689 | 9.8563 | 5.0038
478 | 19118.686 | 9.8584 | 5.0011 |  |  | 478 | 19118.686 | 9.8584 | 4.9984 |  |  | 478 | 19118.686 | 9.8584 | 5.0036
479 | 19158.683 | 9.8605 | 5.0009 |  |  | 479 | 19158.683 | 9.8605 | 4.9981 |  |  | 479 | 19158.683 | 9.8605 | 5.0034
480 | 19198.68 | 9.8626 | 5.0007 |  |  | 480 | 19198.68 | 9.8626 | 4.9979 |  |  | 480 | 19198.68 | 9.8626 | 5.0032
481 | 19238.678 | 9.8647 | 5.0005 |  |  | 481 | 19238.678 | 9.8647 | 4.9977 |  |  | 481 | 19238.678 | 9.8647 | 5.0029
482 | 19278.675 | 9.8668 | 5.0003 |  |  | 482 | 19278.675 | 9.8668 | 4.9975 |  |  | 482 | 19278.675 | 9.8668 | 5.0027
483 | 19318.672 | 9.8688 | 5.0001 |  |  | 483 | 19318.672 | 9.8688 | 4.9973 |  |  | 483 | 19318.672 | 9.8688 | 5.0025
484 | 19358.669 | 9.8709 | 4.9999 |  |  | 484 | 19358.669 | 9.8709 | 4.997 |  |  | 484 | 19358.669 | 9.8709 | 5.0023
485 | 19398.667 | 9.873 | 4.9996 |  |  | 485 | 19398.667 | 9.873 | 4.9968 |  |  | 485 | 19398.667 | 9.873 | 5.0021
486 | 19438.664 | 9.875 | 4.9994 |  |  | 486 | 19438.664 | 9.875 | 4.9966 |  |  | 486 | 19438.664 | 9.875 | 5.0019
487 | 19478.661 | 9.8771 | 4.9992 |  |  | 487 | 19478.661 | 9.8771 | 4.9964 |  |  | 487 | 19478.661 | 9.8771 | 5.0017
488 | 19518.658 | 9.8791 | 4.999 |  |  | 488 | 19518.658 | 9.8791 | 4.9962 |  |  | 488 | 19518.658 | 9.8791 | 5.0015
489 | 19558.656 | 9.8812 | 4.9988 |  |  | 489 | 19558.656 | 9.8812 | 4.996 |  |  | 489 | 19558.656 | 9.8812 | 5.0013
490 | 19598.653 | 9.8832 | 4.9986 |  |  | 490 | 19598.653 | 9.8832 | 4.9958 |  |  | 490 | 19598.653 | 9.8832 | 5.0011
491 | 19638.65 | 9.8853 | 4.9984 |  |  | 491 | 19638.65 | 9.8853 | 4.9956 |  |  | 491 | 19638.65 | 9.8853 | 5.0009
492 | 19678.647 | 9.8873 | 4.9983 |  |  | 492 | 19678.647 | 9.8873 | 4.9954 |  |  | 492 | 19678.647 | 9.8873 | 5.0007
493 | 19718.645 | 9.8893 | 4.9981 |  |  | 493 | 19718.645 | 9.8893 | 4.9952 |  |  | 493 | 19718.645 | 9.8893 | 5.0005
494 | 19758.642 | 9.8913 | 4.9979 |  |  | 494 | 19758.642 | 9.8913 | 4.9949 |  |  | 494 | 19758.642 | 9.8913 | 5.0003
495 | 19798.639 | 9.8934 | 4.9977 |  |  | 495 | 19798.639 | 9.8934 | 4.9947 |  |  | 495 | 19798.639 | 9.8934 | 5.0001
496 | 19838.636 | 9.8954 | 4.9975 |  |  | 496 | 19838.636 | 9.8954 | 4.9945 |  |  | 496 | 19838.636 | 9.8954 | 4.9999
497 | 19878.634 | 9.8974 | 4.9973 |  |  | 497 | 19878.634 | 9.8974 | 4.9943 |  |  | 497 | 19878.634 | 9.8974 | 4.9997
498 | 19918.631 | 9.8994 | 4.9971 |  |  | 498 | 19918.631 | 9.8994 | 4.9941 |  |  | 498 | 19918.631 | 9.8994 | 4.9995
499 | 19958.628 | 9.9014 | 4.9969 |  |  | 499 | 19958.628 | 9.9014 | 4.9939 |  |  | 499 | 19958.628 | 9.9014 | 4.9993
500 | 19998.625 | 9.9034 | 4.9967 |  |  | 500 | 19998.625 | 9.9034 | 4.9937 |  |  | 500 | 19998.625 | 9.9034 | 4.9991
501 | 20038.623 | 9.9054 | 4.9965 |  |  | 501 | 20038.623 | 9.9054 | 4.9935 |  |  | 501 | 20038.623 | 9.9054 | 4.9989
502 | 20078.62 | 9.9074 | 4.9963 |  |  | 502 | 20078.62 | 9.9074 | 4.9933 |  |  | 502 | 20078.62 | 9.9074 | 4.9987
503 | 20118.617 | 9.9094 | 4.9962 |  |  | 503 | 20118.617 | 9.9094 | 4.9931 |  |  | 503 | 20118.617 | 9.9094 | 4.9985
504 | 20158.614 | 9.9114 | 4.996 |  |  | 504 | 20158.614 | 9.9114 | 4.9929 |  |  | 504 | 20158.614 | 9.9114 | 4.9983
505 | 20198.612 | 9.9134 | 4.9958 |  |  | 505 | 20198.612 | 9.9134 | 4.9928 |  |  | 505 | 20198.612 | 9.9134 | 4.9981
506 | 20238.609 | 9.9153 | 4.9956 |  |  | 506 | 20238.609 | 9.9153 | 4.9926 |  |  | 506 | 20238.609 | 9.9153 | 4.998
507 | 20278.606 | 9.9173 | 4.9954 |  |  | 507 | 20278.606 | 9.9173 | 4.9924 |  |  | 507 | 20278.606 | 9.9173 | 4.9978
508 | 20318.603 | 9.9193 | 4.9952 |  |  | 508 | 20318.603 | 9.9193 | 4.9922 |  |  | 508 | 20318.603 | 9.9193 | 4.9976
509 | 20358.601 | 9.9213 | 4.9951 |  |  | 509 | 20358.601 | 9.9213 | 4.992 |  |  | 509 | 20358.601 | 9.9213 | 4.9974
510 | 20398.598 | 9.9232 | 4.9949 |  |  | 510 | 20398.598 | 9.9232 | 4.9918 |  |  | 510 | 20398.598 | 9.9232 | 4.9972
511 | 20438.595 | 9.9252 | 4.9947 |  |  | 511 | 20438.595 | 9.9252 | 4.9916 |  |  | 511 | 20438.595 | 9.9252 | 4.997
512 | 20478.592 | 9.9271 | 4.9945 |  |  | 512 | 20478.592 | 9.9271 | 4.9914 |  |  | 512 | 20478.592 | 9.9271 | 4.9968
513 | 20518.59 | 9.9291 | 4.9943 |  |  | 513 | 20518.59 | 9.9291 | 4.9912 |  |  | 513 | 20518.59 | 9.9291 | 4.9967
514 | 20558.587 | 9.931 | 4.9942 |  |  | 514 | 20558.587 | 9.931 | 4.991 |  |  | 514 | 20558.587 | 9.931 | 4.9965
515 | 20598.584 | 9.933 | 4.994 |  |  | 515 | 20598.584 | 9.933 | 4.9909 |  |  | 515 | 20598.584 | 9.933 | 4.9963
516 | 20638.581 | 9.9349 | 4.9938 |  |  | 516 | 20638.581 | 9.9349 | 4.9907 |  |  | 516 | 20638.581 | 9.9349 | 4.9961
517 | 20678.579 | 9.9369 | 4.9936 |  |  | 517 | 20678.579 | 9.9369 | 4.9905 |  |  | 517 | 20678.579 | 9.9369 | 4.9959
518 | 20718.576 | 9.9388 | 4.9935 |  |  | 518 | 20718.576 | 9.9388 | 4.9903 |  |  | 518 | 20718.576 | 9.9388 | 4.9958
519 | 20758.573 | 9.9407 | 4.9933 |  |  | 519 | 20758.573 | 9.9407 | 4.9901 |  |  | 519 | 20758.573 | 9.9407 | 4.9956
520 | 20798.57 | 9.9426 | 4.9931 |  |  | 520 | 20798.57 | 9.9426 | 4.9899 |  |  | 520 | 20798.57 | 9.9426 | 4.9954
521 | 20838.568 | 9.9446 | 4.993 |  |  | 521 | 20838.568 | 9.9446 | 4.9898 |  |  | 521 | 20838.568 | 9.9446 | 4.9952
522 | 20878.565 | 9.9465 | 4.9928 |  |  | 522 | 20878.565 | 9.9465 | 4.9896 |  |  | 522 | 20878.565 | 9.9465 | 4.9951
523 | 20918.562 | 9.9484 | 4.9926 |  |  | 523 | 20918.562 | 9.9484 | 4.9894 |  |  | 523 | 20918.562 | 9.9484 | 4.9949
524 | 20958.559 | 9.9503 | 4.9925 |  |  | 524 | 20958.559 | 9.9503 | 4.9892 |  |  | 524 | 20958.559 | 9.9503 | 4.9947
525 | 20998.557 | 9.9522 | 4.9923 |  |  | 525 | 20998.557 | 9.9522 | 4.989 |  |  | 525 | 20998.557 | 9.9522 | 4.9945
526 | 21038.554 | 9.9541 | 4.9921 |  |  | 526 | 21038.554 | 9.9541 | 4.9889 |  |  | 526 | 21038.554 | 9.9541 | 4.9944
527 | 21078.551 | 9.956 | 4.992 |  |  | 527 | 21078.551 | 9.956 | 4.9887 |  |  | 527 | 21078.551 | 9.956 | 4.9942
528 | 21118.548 | 9.9579 | 4.9918 |  |  | 528 | 21118.548 | 9.9579 | 4.9885 |  |  | 528 | 21118.548 | 9.9579 | 4.994
529 | 21158.546 | 9.9598 | 4.9916 |  |  | 529 | 21158.546 | 9.9598 | 4.9883 |  |  | 529 | 21158.546 | 9.9598 | 4.9939
530 | 21198.543 | 9.9617 | 4.9915 |  |  | 530 | 21198.543 | 9.9617 | 4.9882 |  |  | 530 | 21198.543 | 9.9617 | 4.9937
531 | 21238.54 | 9.9636 | 4.9913 |  |  | 531 | 21238.54 | 9.9636 | 4.988 |  |  | 531 | 21238.54 | 9.9636 | 4.9935
532 | 21278.537 | 9.9655 | 4.9911 |  |  | 532 | 21278.537 | 9.9655 | 4.9878 |  |  | 532 | 21278.537 | 9.9655 | 4.9933
533 | 21318.535 | 9.9673 | 4.991 |  |  | 533 | 21318.535 | 9.9673 | 4.9876 |  |  | 533 | 21318.535 | 9.9673 | 4.9932
534 | 21358.532 | 9.9692 | 4.9908 |  |  | 534 | 21358.532 | 9.9692 | 4.9875 |  |  | 534 | 21358.532 | 9.9692 | 4.993
535 | 21398.529 | 9.9711 | 4.9906 |  |  | 535 | 21398.529 | 9.9711 | 4.9873 |  |  | 535 | 21398.529 | 9.9711 | 4.9928
536 | 21438.526 | 9.9729 | 4.9905 |  |  | 536 | 21438.526 | 9.9729 | 4.9871 |  |  | 536 | 21438.526 | 9.9729 | 4.9927
537 | 21478.524 | 9.9748 | 4.9903 |  |  | 537 | 21478.524 | 9.9748 | 4.987 |  |  | 537 | 21478.524 | 9.9748 | 4.9925
538 | 21518.521 | 9.9767 | 4.9902 |  |  | 538 | 21518.521 | 9.9767 | 4.9868 |  |  | 538 | 21518.521 | 9.9767 | 4.9924
539 | 21558.518 | 9.9785 | 4.99 |  |  | 539 | 21558.518 | 9.9785 | 4.9866 |  |  | 539 | 21558.518 | 9.9785 | 4.9922
540 | 21598.515 | 9.9804 | 4.9899 |  |  | 540 | 21598.515 | 9.9804 | 4.9865 |  |  | 540 | 21598.515 | 9.9804 | 4.992
541 | 21638.513 | 9.9822 | 4.9897 |  |  | 541 | 21638.513 | 9.9822 | 4.9863 |  |  | 541 | 21638.513 | 9.9822 | 4.9919
542 | 21678.51 | 9.9841 | 4.9895 |  |  | 542 | 21678.51 | 9.9841 | 4.9861 |  |  | 542 | 21678.51 | 9.9841 | 4.9917
543 | 21718.507 | 9.9859 | 4.9894 |  |  | 543 | 21718.507 | 9.9859 | 4.986 |  |  | 543 | 21718.507 | 9.9859 | 4.9915
544 | 21758.504 | 9.9878 | 4.9892 |  |  | 544 | 21758.504 | 9.9878 | 4.9858 |  |  | 544 | 21758.504 | 9.9878 | 4.9914
545 | 21798.502 | 9.9896 | 4.9891 |  |  | 545 | 21798.502 | 9.9896 | 4.9856 |  |  | 545 | 21798.502 | 9.9896 | 4.9912
546 | 21838.499 | 9.9914 | 4.9889 |  |  | 546 | 21838.499 | 9.9914 | 4.9855 |  |  | 546 | 21838.499 | 9.9914 | 4.9911
547 | 21878.496 | 9.9933 | 4.9888 |  |  | 547 | 21878.496 | 9.9933 | 4.9853 |  |  | 547 | 21878.496 | 9.9933 | 4.9909
548 | 21918.493 | 9.9951 | 4.9886 |  |  | 548 | 21918.493 | 9.9951 | 4.9851 |  |  | 548 | 21918.493 | 9.9951 | 4.9908
549 | 21958.491 | 9.9969 | 4.9885 |  |  | 549 | 21958.491 | 9.9969 | 4.985 |  |  | 549 | 21958.491 | 9.9969 | 4.9906
550 | 21998.488 | 9.9987 | 4.9883 |  |  | 550 | 21998.488 | 9.9987 | 4.9848 |  |  | 550 | 21998.488 | 9.9987 | 4.9904
551 | 22038.485 | 10.0005 | 4.9882 |  |  | 551 | 22038.485 | 10.0005 | 4.9847 |  |  | 551 | 22038.485 | 10.0005 | 4.9903
552 | 22078.482 | 10.0024 | 4.988 |  |  | 552 | 22078.482 | 10.0024 | 4.9845 |  |  | 552 | 22078.482 | 10.0024 | 4.9901
553 | 22118.48 | 10.0042 | 4.9879 |  |  | 553 | 22118.48 | 10.0042 | 4.9843 |  |  | 553 | 22118.48 | 10.0042 | 4.99
554 | 22158.477 | 10.006 | 4.9877 |  |  | 554 | 22158.477 | 10.006 | 4.9842 |  |  | 554 | 22158.477 | 10.006 | 4.9898
555 | 22198.474 | 10.0078 | 4.9876 |  |  | 555 | 22198.474 | 10.0078 | 4.984 |  |  | 555 | 22198.474 | 10.0078 | 4.9897
556 | 22238.471 | 10.0096 | 4.9874 |  |  | 556 | 22238.471 | 10.0096 | 4.9839 |  |  | 556 | 22238.471 | 10.0096 | 4.9895
557 | 22278.469 | 10.0114 | 4.9873 |  |  | 557 | 22278.469 | 10.0114 | 4.9837 |  |  | 557 | 22278.469 | 10.0114 | 4.9894
558 | 22318.466 | 10.0132 | 4.9871 |  |  | 558 | 22318.466 | 10.0132 | 4.9836 |  |  | 558 | 22318.466 | 10.0132 | 4.9892
559 | 22358.463 | 10.015 | 4.987 |  |  | 559 | 22358.463 | 10.015 | 4.9834 |  |  | 559 | 22358.463 | 10.015 | 4.9891
560 | 22398.46 | 10.0167 | 4.9868 |  |  | 560 | 22398.46 | 10.0167 | 4.9832 |  |  | 560 | 22398.46 | 10.0167 | 4.9889
561 | 22438.458 | 10.0185 | 4.9867 |  |  | 561 | 22438.458 | 10.0185 | 4.9831 |  |  | 561 | 22438.458 | 10.0185 | 4.9888
562 | 22478.455 | 10.0203 | 4.9866 |  |  | 562 | 22478.455 | 10.0203 | 4.9829 |  |  | 562 | 22478.455 | 10.0203 | 4.9886
563 | 22518.452 | 10.0221 | 4.9864 |  |  | 563 | 22518.452 | 10.0221 | 4.9828 |  |  | 563 | 22518.452 | 10.0221 | 4.9885
564 | 22558.449 | 10.0239 | 4.9863 |  |  | 564 | 22558.449 | 10.0239 | 4.9826 |  |  | 564 | 22558.449 | 10.0239 | 4.9883
565 | 22598.447 | 10.0256 | 4.9861 |  |  | 565 | 22598.447 | 10.0256 | 4.9825 |  |  | 565 | 22598.447 | 10.0256 | 4.9882
566 | 22638.444 | 10.0274 | 4.986 |  |  | 566 | 22638.444 | 10.0274 | 4.9823 |  |  | 566 | 22638.444 | 10.0274 | 4.988
567 | 22678.441 | 10.0292 | 4.9858 |  |  | 567 | 22678.441 | 10.0292 | 4.9822 |  |  | 567 | 22678.441 | 10.0292 | 4.9879
568 | 22718.438 | 10.0309 | 4.9857 |  |  | 568 | 22718.438 | 10.0309 | 4.982 |  |  | 568 | 22718.438 | 10.0309 | 4.9877
569 | 22758.436 | 10.0327 | 4.9856 |  |  | 569 | 22758.436 | 10.0327 | 4.9819 |  |  | 569 | 22758.436 | 10.0327 | 4.9876
570 | 22798.433 | 10.0344 | 4.9854 |  |  | 570 | 22798.433 | 10.0344 | 4.9817 |  |  | 570 | 22798.433 | 10.0344 | 4.9874
571 | 22838.43 | 10.0362 | 4.9853 |  |  | 571 | 22838.43 | 10.0362 | 4.9816 |  |  | 571 | 22838.43 | 10.0362 | 4.9873
572 | 22878.427 | 10.0379 | 4.9851 |  |  | 572 | 22878.427 | 10.0379 | 4.9814 |  |  | 572 | 22878.427 | 10.0379 | 4.9871
573 | 22918.425 | 10.0397 | 4.985 |  |  | 573 | 22918.425 | 10.0397 | 4.9813 |  |  | 573 | 22918.425 | 10.0397 | 4.987
574 | 22958.422 | 10.0414 | 4.9849 |  |  | 574 | 22958.422 | 10.0414 | 4.9811 |  |  | 574 | 22958.422 | 10.0414 | 4.9869
575 | 22998.419 | 10.0432 | 4.9847 |  |  | 575 | 22998.419 | 10.0432 | 4.981 |  |  | 575 | 22998.419 | 10.0432 | 4.9867
576 | 23038.416 | 10.0449 | 4.9846 |  |  | 576 | 23038.416 | 10.0449 | 4.9808 |  |  | 576 | 23038.416 | 10.0449 | 4.9866
577 | 23078.414 | 10.0467 | 4.9845 |  |  | 577 | 23078.414 | 10.0467 | 4.9807 |  |  | 577 | 23078.414 | 10.0467 | 4.9864
578 | 23118.411 | 10.0484 | 4.9843 |  |  | 578 | 23118.411 | 10.0484 | 4.9805 |  |  | 578 | 23118.411 | 10.0484 | 4.9863
579 | 23158.408 | 10.0501 | 4.9842 |  |  | 579 | 23158.408 | 10.0501 | 4.9804 |  |  | 579 | 23158.408 | 10.0501 | 4.9861
580 | 23198.405 | 10.0518 | 4.984 |  |  | 580 | 23198.405 | 10.0518 | 4.9802 |  |  | 580 | 23198.405 | 10.0518 | 4.986
581 | 23238.403 | 10.0536 | 4.9839 |  |  | 581 | 23238.403 | 10.0536 | 4.9801 |  |  | 581 | 23238.403 | 10.0536 | 4.9859
582 | 23278.4 | 10.0553 | 4.9838 |  |  | 582 | 23278.4 | 10.0553 | 4.98 |  |  | 582 | 23278.4 | 10.0553 | 4.9857
583 | 23318.397 | 10.057 | 4.9836 |  |  | 583 | 23318.397 | 10.057 | 4.9798 |  |  | 583 | 23318.397 | 10.057 | 4.9856
584 | 23358.394 | 10.0587 | 4.9835 |  |  | 584 | 23358.394 | 10.0587 | 4.9797 |  |  | 584 | 23358.394 | 10.0587 | 4.9854
585 | 23398.392 | 10.0604 | 4.9834 |  |  | 585 | 23398.392 | 10.0604 | 4.9795 |  |  | 585 | 23398.392 | 10.0604 | 4.9853
586 | 23438.389 | 10.0621 | 4.9832 |  |  | 586 | 23438.389 | 10.0621 | 4.9794 |  |  | 586 | 23438.389 | 10.0621 | 4.9852
587 | 23478.386 | 10.0638 | 4.9831 |  |  | 587 | 23478.386 | 10.0638 | 4.9792 |  |  | 587 | 23478.386 | 10.0638 | 4.985
588 | 23518.383 | 10.0655 | 4.983 |  |  | 588 | 23518.383 | 10.0655 | 4.9791 |  |  | 588 | 23518.383 | 10.0655 | 4.9849
589 | 23558.381 | 10.0672 | 4.9828 |  |  | 589 | 23558.381 | 10.0672 | 4.979 |  |  | 589 | 23558.381 | 10.0672 | 4.9848
590 | 23598.378 | 10.0689 | 4.9827 |  |  | 590 | 23598.378 | 10.0689 | 4.9788 |  |  | 590 | 23598.378 | 10.0689 | 4.9846
591 | 23638.375 | 10.0706 | 4.9826 |  |  | 591 | 23638.375 | 10.0706 | 4.9787 |  |  | 591 | 23638.375 | 10.0706 | 4.9845
592 | 23678.372 | 10.0723 | 4.9825 |  |  | 592 | 23678.372 | 10.0723 | 4.9785 |  |  | 592 | 23678.372 | 10.0723 | 4.9843
593 | 23718.37 | 10.074 | 4.9823 |  |  | 593 | 23718.37 | 10.074 | 4.9784 |  |  | 593 | 23718.37 | 10.074 | 4.9842
594 | 23758.367 | 10.0757 | 4.9822 |  |  | 594 | 23758.367 | 10.0757 | 4.9783 |  |  | 594 | 23758.367 | 10.0757 | 4.9841
595 | 23798.364 | 10.0774 | 4.9821 |  |  | 595 | 23798.364 | 10.0774 | 4.9781 |  |  | 595 | 23798.364 | 10.0774 | 4.9839
596 | 23838.361 | 10.0791 | 4.9819 |  |  | 596 | 23838.361 | 10.0791 | 4.978 |  |  | 596 | 23838.361 | 10.0791 | 4.9838
597 | 23878.359 | 10.0807 | 4.9818 |  |  | 597 | 23878.359 | 10.0807 | 4.9779 |  |  | 597 | 23878.359 | 10.0807 | 4.9837
598 | 23918.356 | 10.0824 | 4.9817 |  |  | 598 | 23918.356 | 10.0824 | 4.9777 |  |  | 598 | 23918.356 | 10.0824 | 4.9835
599 | 23958.353 | 10.0841 | 4.9816 |  |  | 599 | 23958.353 | 10.0841 | 4.9776 |  |  | 599 | 23958.353 | 10.0841 | 4.9834
600 | 23998.35 | 10.0857 | 4.9814 |  |  | 600 | 23998.35 | 10.0857 | 4.9774 |  |  | 600 | 23998.35 | 10.0857 | 4.9833
601 | 24038.348 | 10.0874 | 4.9813 |  |  | 601 | 24038.348 | 10.0874 | 4.9773 |  |  | 601 | 24038.348 | 10.0874 | 4.9831
602 | 24078.345 | 10.0891 | 4.9812 |  |  | 602 | 24078.345 | 10.0891 | 4.9772 |  |  | 602 | 24078.345 | 10.0891 | 4.983
603 | 24118.342 | 10.0907 | 4.981 |  |  | 603 | 24118.342 | 10.0907 | 4.977 |  |  | 603 | 24118.342 | 10.0907 | 4.9829
604 | 24158.339 | 10.0924 | 4.9809 |  |  | 604 | 24158.339 | 10.0924 | 4.9769 |  |  | 604 | 24158.339 | 10.0924 | 4.9827
605 | 24198.337 | 10.094 | 4.9808 |  |  | 605 | 24198.337 | 10.094 | 4.9768 |  |  | 605 | 24198.337 | 10.094 | 4.9826
606 | 24238.334 | 10.0957 | 4.9807 |  |  | 606 | 24238.334 | 10.0957 | 4.9766 |  |  | 606 | 24238.334 | 10.0957 | 4.9825
607 | 24278.331 | 10.0973 | 4.9805 |  |  | 607 | 24278.331 | 10.0973 | 4.9765 |  |  | 607 | 24278.331 | 10.0973 | 4.9824
608 | 24318.328 | 10.099 | 4.9804 |  |  | 608 | 24318.328 | 10.099 | 4.9764 |  |  | 608 | 24318.328 | 10.099 | 4.9822
609 | 24358.326 | 10.1006 | 4.9803 |  |  | 609 | 24358.326 | 10.1006 | 4.9762 |  |  | 609 | 24358.326 | 10.1006 | 4.9821
610 | 24398.323 | 10.1023 | 4.9802 |  |  | 610 | 24398.323 | 10.1023 | 4.9761 |  |  | 610 | 24398.323 | 10.1023 | 4.982
611 | 24438.32 | 10.1039 | 4.9801 |  |  | 611 | 24438.32 | 10.1039 | 4.976 |  |  | 611 | 24438.32 | 10.1039 | 4.9818
612 | 24478.317 | 10.1055 | 4.9799 |  |  | 612 | 24478.317 | 10.1055 | 4.9758 |  |  | 612 | 24478.317 | 10.1055 | 4.9817
613 | 24518.315 | 10.1072 | 4.9798 |  |  | 613 | 24518.315 | 10.1072 | 4.9757 |  |  | 613 | 24518.315 | 10.1072 | 4.9816
614 | 24558.312 | 10.1088 | 4.9797 |  |  | 614 | 24558.312 | 10.1088 | 4.9756 |  |  | 614 | 24558.312 | 10.1088 | 4.9815
615 | 24598.309 | 10.1104 | 4.9796 |  |  | 615 | 24598.309 | 10.1104 | 4.9754 |  |  | 615 | 24598.309 | 10.1104 | 4.9813
616 | 24638.306 | 10.1121 | 4.9794 |  |  | 616 | 24638.306 | 10.1121 | 4.9753 |  |  | 616 | 24638.306 | 10.1121 | 4.9812
617 | 24678.304 | 10.1137 | 4.9793 |  |  | 617 | 24678.304 | 10.1137 | 4.9752 |  |  | 617 | 24678.304 | 10.1137 | 4.9811
618 | 24718.301 | 10.1153 | 4.9792 |  |  | 618 | 24718.301 | 10.1153 | 4.9751 |  |  | 618 | 24718.301 | 10.1153 | 4.9809
619 | 24758.298 | 10.1169 | 4.9791 |  |  | 619 | 24758.298 | 10.1169 | 4.9749 |  |  | 619 | 24758.298 | 10.1169 | 4.9808
620 | 24798.295 | 10.1185 | 4.979 |  |  | 620 | 24798.295 | 10.1185 | 4.9748 |  |  | 620 | 24798.295 | 10.1185 | 4.9807
621 | 24838.293 | 10.1201 | 4.9788 |  |  | 621 | 24838.293 | 10.1201 | 4.9747 |  |  | 621 | 24838.293 | 10.1201 | 4.9806
622 | 24878.29 | 10.1218 | 4.9787 |  |  | 622 | 24878.29 | 10.1218 | 4.9745 |  |  | 622 | 24878.29 | 10.1218 | 4.9804
623 | 24918.287 | 10.1234 | 4.9786 |  |  | 623 | 24918.287 | 10.1234 | 4.9744 |  |  | 623 | 24918.287 | 10.1234 | 4.9803
624 | 24958.284 | 10.125 | 4.9785 |  |  | 624 | 24958.284 | 10.125 | 4.9743 |  |  | 624 | 24958.284 | 10.125 | 4.9802
625 | 24998.282 | 10.1266 | 4.9784 |  |  | 625 | 24998.282 | 10.1266 | 4.9742 |  |  | 625 | 24998.282 | 10.1266 | 4.9801
626 | 25038.279 | 10.1282 | 4.9782 |  |  | 626 | 25038.279 | 10.1282 | 4.974 |  |  | 626 | 25038.279 | 10.1282 | 4.9799
627 | 25078.276 | 10.1298 | 4.9781 |  |  | 627 | 25078.276 | 10.1298 | 4.9739 |  |  | 627 | 25078.276 | 10.1298 | 4.9798
628 | 25118.273 | 10.1314 | 4.978 |  |  | 628 | 25118.273 | 10.1314 | 4.9738 |  |  | 628 | 25118.273 | 10.1314 | 4.9797
629 | 25158.271 | 10.1329 | 4.9779 |  |  | 629 | 25158.271 | 10.1329 | 4.9736 |  |  | 629 | 25158.271 | 10.1329 | 4.9796
630 | 25198.268 | 10.1345 | 4.9778 |  |  | 630 | 25198.268 | 10.1345 | 4.9735 |  |  | 630 | 25198.268 | 10.1345 | 4.9794
631 | 25238.265 | 10.1361 | 4.9777 |  |  | 631 | 25238.265 | 10.1361 | 4.9734 |  |  | 631 | 25238.265 | 10.1361 | 4.9793
632 | 25278.262 | 10.1377 | 4.9775 |  |  | 632 | 25278.262 | 10.1377 | 4.9733 |  |  | 632 | 25278.262 | 10.1377 | 4.9792
633 | 25318.26 | 10.1393 | 4.9774 |  |  | 633 | 25318.26 | 10.1393 | 4.9731 |  |  | 633 | 25318.26 | 10.1393 | 4.9791
634 | 25358.257 | 10.1409 | 4.9773 |  |  | 634 | 25358.257 | 10.1409 | 4.973 |  |  | 634 | 25358.257 | 10.1409 | 4.979
635 | 25398.254 | 10.1424 | 4.9772 |  |  | 635 | 25398.254 | 10.1424 | 4.9729 |  |  | 635 | 25398.254 | 10.1424 | 4.9788
636 | 25438.251 | 10.144 | 4.9771 |  |  | 636 | 25438.251 | 10.144 | 4.9728 |  |  | 636 | 25438.251 | 10.144 | 4.9787
637 | 25478.249 | 10.1456 | 4.977 |  |  | 637 | 25478.249 | 10.1456 | 4.9726 |  |  | 637 | 25478.249 | 10.1456 | 4.9786
638 | 25518.246 | 10.1471 | 4.9768 |  |  | 638 | 25518.246 | 10.1471 | 4.9725 |  |  | 638 | 25518.246 | 10.1471 | 4.9785
639 | 25558.243 | 10.1487 | 4.9767 |  |  | 639 | 25558.243 | 10.1487 | 4.9724 |  |  | 639 | 25558.243 | 10.1487 | 4.9783
640 | 25598.24 | 10.1503 | 4.9766 |  |  | 640 | 25598.24 | 10.1503 | 4.9723 |  |  | 640 | 25598.24 | 10.1503 | 4.9782
641 | 25638.238 | 10.1518 | 4.9765 |  |  | 641 | 25638.238 | 10.1518 | 4.9722 |  |  | 641 | 25638.238 | 10.1518 | 4.9781
642 | 25678.235 | 10.1534 | 4.9764 |  |  | 642 | 25678.235 | 10.1534 | 4.972 |  |  | 642 | 25678.235 | 10.1534 | 4.978
643 | 25718.232 | 10.155 | 4.9763 |  |  | 643 | 25718.232 | 10.155 | 4.9719 |  |  | 643 | 25718.232 | 10.155 | 4.9779
644 | 25758.229 | 10.1565 | 4.9762 |  |  | 644 | 25758.229 | 10.1565 | 4.9718 |  |  | 644 | 25758.229 | 10.1565 | 4.9777
645 | 25798.227 | 10.1581 | 4.976 |  |  | 645 | 25798.227 | 10.1581 | 4.9717 |  |  | 645 | 25798.227 | 10.1581 | 4.9776
646 | 25838.224 | 10.1596 | 4.9759 |  |  | 646 | 25838.224 | 10.1596 | 4.9715 |  |  | 646 | 25838.224 | 10.1596 | 4.9775
647 | 25878.221 | 10.1612 | 4.9758 |  |  | 647 | 25878.221 | 10.1612 | 4.9714 |  |  | 647 | 25878.221 | 10.1612 | 4.9774
648 | 25918.218 | 10.1627 | 4.9757 |  |  | 648 | 25918.218 | 10.1627 | 4.9713 |  |  | 648 | 25918.218 | 10.1627 | 4.9773
649 | 25958.216 | 10.1642 | 4.9756 |  |  | 649 | 25958.216 | 10.1642 | 4.9712 |  |  | 649 | 25958.216 | 10.1642 | 4.9772
650 | 25998.213 | 10.1658 | 4.9755 |  |  | 650 | 25998.213 | 10.1658 | 4.9711 |  |  | 650 | 25998.213 | 10.1658 | 4.977
651 | 26038.21 | 10.1673 | 4.9754 |  |  | 651 | 26038.21 | 10.1673 | 4.9709 |  |  | 651 | 26038.21 | 10.1673 | 4.9769
652 | 26078.207 | 10.1689 | 4.9753 |  |  | 652 | 26078.207 | 10.1689 | 4.9708 |  |  | 652 | 26078.207 | 10.1689 | 4.9768
653 | 26118.205 | 10.1704 | 4.9751 |  |  | 653 | 26118.205 | 10.1704 | 4.9707 |  |  | 653 | 26118.205 | 10.1704 | 4.9767
654 | 26158.202 | 10.1719 | 4.975 |  |  | 654 | 26158.202 | 10.1719 | 4.9706 |  |  | 654 | 26158.202 | 10.1719 | 4.9766
655 | 26198.199 | 10.1734 | 4.9749 |  |  | 655 | 26198.199 | 10.1734 | 4.9705 |  |  | 655 | 26198.199 | 10.1734 | 4.9764
656 | 26238.196 | 10.175 | 4.9748 |  |  | 656 | 26238.196 | 10.175 | 4.9703 |  |  | 656 | 26238.196 | 10.175 | 4.9763
657 | 26278.194 | 10.1765 | 4.9747 |  |  | 657 | 26278.194 | 10.1765 | 4.9702 |  |  | 657 | 26278.194 | 10.1765 | 4.9762
658 | 26318.191 | 10.178 | 4.9746 |  |  | 658 | 26318.191 | 10.178 | 4.9701 |  |  | 658 | 26318.191 | 10.178 | 4.9761
659 | 26358.188 | 10.1795 | 4.9745 |  |  | 659 | 26358.188 | 10.1795 | 4.97 |  |  | 659 | 26358.188 | 10.1795 | 4.976
660 | 26398.185 | 10.1811 | 4.9744 |  |  | 660 | 26398.185 | 10.1811 | 4.9699 |  |  | 660 | 26398.185 | 10.1811 | 4.9759
661 | 26438.183 | 10.1826 | 4.9743 |  |  | 661 | 26438.183 | 10.1826 | 4.9697 |  |  | 661 | 26438.183 | 10.1826 | 4.9757
662 | 26478.18 | 10.1841 | 4.9741 |  |  | 662 | 26478.18 | 10.1841 | 4.9696 |  |  | 662 | 26478.18 | 10.1841 | 4.9756
663 | 26518.177 | 10.1856 | 4.974 |  |  | 663 | 26518.177 | 10.1856 | 4.9695 |  |  | 663 | 26518.177 | 10.1856 | 4.9755
664 | 26558.174 | 10.1871 | 4.9739 |  |  | 664 | 26558.174 | 10.1871 | 4.9694 |  |  | 664 | 26558.174 | 10.1871 | 4.9754
665 | 26598.172 | 10.1886 | 4.9738 |  |  | 665 | 26598.172 | 10.1886 | 4.9693 |  |  | 665 | 26598.172 | 10.1886 | 4.9753
666 | 26638.169 | 10.1901 | 4.9737 |  |  | 666 | 26638.169 | 10.1901 | 4.9691 |  |  | 666 | 26638.169 | 10.1901 | 4.9752
667 | 26678.166 | 10.1916 | 4.9736 |  |  | 667 | 26678.166 | 10.1916 | 4.969 |  |  | 667 | 26678.166 | 10.1916 | 4.9751
668 | 26718.163 | 10.1931 | 4.9735 |  |  | 668 | 26718.163 | 10.1931 | 4.9689 |  |  | 668 | 26718.163 | 10.1931 | 4.9749
669 | 26758.161 | 10.1946 | 4.9734 |  |  | 669 | 26758.161 | 10.1946 | 4.9688 |  |  | 669 | 26758.161 | 10.1946 | 4.9748
670 | 26798.158 | 10.1961 | 4.9733 |  |  | 670 | 26798.158 | 10.1961 | 4.9687 |  |  | 670 | 26798.158 | 10.1961 | 4.9747
671 | 26838.155 | 10.1976 | 4.9732 |  |  | 671 | 26838.155 | 10.1976 | 4.9686 |  |  | 671 | 26838.155 | 10.1976 | 4.9746
672 | 26878.152 | 10.1991 | 4.9731 |  |  | 672 | 26878.152 | 10.1991 | 4.9684 |  |  | 672 | 26878.152 | 10.1991 | 4.9745
673 | 26918.15 | 10.2006 | 4.973 |  |  | 673 | 26918.15 | 10.2006 | 4.9683 |  |  | 673 | 26918.15 | 10.2006 | 4.9744
674 | 26958.147 | 10.202 | 4.9728 |  |  | 674 | 26958.147 | 10.202 | 4.9682 |  |  | 674 | 26958.147 | 10.202 | 4.9743
675 | 26998.144 | 10.2035 | 4.9727 |  |  | 675 | 26998.144 | 10.2035 | 4.9681 |  |  | 675 | 26998.144 | 10.2035 | 4.9741
676 | 27038.141 | 10.205 | 4.9726 |  |  | 676 | 27038.141 | 10.205 | 4.968 |  |  | 676 | 27038.141 | 10.205 | 4.974
677 | 27078.139 | 10.2065 | 4.9725 |  |  | 677 | 27078.139 | 10.2065 | 4.9679 |  |  | 677 | 27078.139 | 10.2065 | 4.9739
678 | 27118.136 | 10.208 | 4.9724 |  |  | 678 | 27118.136 | 10.208 | 4.9678 |  |  | 678 | 27118.136 | 10.208 | 4.9738
679 | 27158.133 | 10.2094 | 4.9723 |  |  | 679 | 27158.133 | 10.2094 | 4.9676 |  |  | 679 | 27158.133 | 10.2094 | 4.9737
680 | 27198.13 | 10.2109 | 4.9722 |  |  | 680 | 27198.13 | 10.2109 | 4.9675 |  |  | 680 | 27198.13 | 10.2109 | 4.9736
681 | 27238.128 | 10.2124 | 4.9721 |  |  | 681 | 27238.128 | 10.2124 | 4.9674 |  |  | 681 | 27238.128 | 10.2124 | 4.9735
682 | 27278.125 | 10.2138 | 4.972 |  |  | 682 | 27278.125 | 10.2138 | 4.9673 |  |  | 682 | 27278.125 | 10.2138 | 4.9733
683 | 27318.122 | 10.2153 | 4.9719 |  |  | 683 | 27318.122 | 10.2153 | 4.9672 |  |  | 683 | 27318.122 | 10.2153 | 4.9732
684 | 27358.119 | 10.2168 | 4.9718 |  |  | 684 | 27358.119 | 10.2168 | 4.9671 |  |  | 684 | 27358.119 | 10.2168 | 4.9731
685 | 27398.117 | 10.2182 | 4.9717 |  |  | 685 | 27398.117 | 10.2182 | 4.967 |  |  | 685 | 27398.117 | 10.2182 | 4.973
686 | 27438.114 | 10.2197 | 4.9716 |  |  | 686 | 27438.114 | 10.2197 | 4.9668 |  |  | 686 | 27438.114 | 10.2197 | 4.9729
687 | 27478.111 | 10.2211 | 4.9715 |  |  | 687 | 27478.111 | 10.2211 | 4.9667 |  |  | 687 | 27478.111 | 10.2211 | 4.9728
688 | 27518.108 | 10.2226 | 4.9714 |  |  | 688 | 27518.108 | 10.2226 | 4.9666 |  |  | 688 | 27518.108 | 10.2226 | 4.9727
689 | 27558.106 | 10.2241 | 4.9713 |  |  | 689 | 27558.106 | 10.2241 | 4.9665 |  |  | 689 | 27558.106 | 10.2241 | 4.9726
690 | 27598.103 | 10.2255 | 4.9711 |  |  | 690 | 27598.103 | 10.2255 | 4.9664 |  |  | 690 | 27598.103 | 10.2255 | 4.9725
691 | 27638.1 | 10.227 | 4.971 |  |  | 691 | 27638.1 | 10.227 | 4.9663 |  |  | 691 | 27638.1 | 10.227 | 4.9723
692 | 27678.097 | 10.2284 | 4.9709 |  |  | 692 | 27678.097 | 10.2284 | 4.9662 |  |  | 692 | 27678.097 | 10.2284 | 4.9722
693 | 27718.095 | 10.2298 | 4.9708 |  |  | 693 | 27718.095 | 10.2298 | 4.966 |  |  | 693 | 27718.095 | 10.2298 | 4.9721
694 | 27758.092 | 10.2313 | 4.9707 |  |  | 694 | 27758.092 | 10.2313 | 4.9659 |  |  | 694 | 27758.092 | 10.2313 | 4.972
695 | 27798.089 | 10.2327 | 4.9706 |  |  | 695 | 27798.089 | 10.2327 | 4.9658 |  |  | 695 | 27798.089 | 10.2327 | 4.9719
696 | 27838.086 | 10.2342 | 4.9705 |  |  | 696 | 27838.086 | 10.2342 | 4.9657 |  |  | 696 | 27838.086 | 10.2342 | 4.9718
697 | 27878.084 | 10.2356 | 4.9704 |  |  | 697 | 27878.084 | 10.2356 | 4.9656 |  |  | 697 | 27878.084 | 10.2356 | 4.9717
698 | 27918.081 | 10.237 | 4.9703 |  |  | 698 | 27918.081 | 10.237 | 4.9655 |  |  | 698 | 27918.081 | 10.237 | 4.9716
699 | 27958.078 | 10.2385 | 4.9702 |  |  | 699 | 27958.078 | 10.2385 | 4.9654 |  |  | 699 | 27958.078 | 10.2385 | 4.9715
700 | 27998.075 | 10.2399 | 4.9701 |  |  | 700 | 27998.075 | 10.2399 | 4.9653 |  |  | 700 | 27998.075 | 10.2399 | 4.9713
701 | 28038.073 | 10.2413 | 4.97 |  |  | 701 | 28038.073 | 10.2413 | 4.9651 |  |  | 701 | 28038.073 | 10.2413 | 4.9712
702 | 28078.07 | 10.2427 | 4.9699 |  |  | 702 | 28078.07 | 10.2427 | 4.965 |  |  | 702 | 28078.07 | 10.2427 | 4.9711
703 | 28118.067 | 10.2442 | 4.9698 |  |  | 703 | 28118.067 | 10.2442 | 4.9649 |  |  | 703 | 28118.067 | 10.2442 | 4.971
704 | 28158.064 | 10.2456 | 4.9697 |  |  | 704 | 28158.064 | 10.2456 | 4.9648 |  |  | 704 | 28158.064 | 10.2456 | 4.9709
705 | 28198.062 | 10.247 | 4.9696 |  |  | 705 | 28198.062 | 10.247 | 4.9647 |  |  | 705 | 28198.062 | 10.247 | 4.9708
706 | 28238.059 | 10.2484 | 4.9695 |  |  | 706 | 28238.059 | 10.2484 | 4.9646 |  |  | 706 | 28238.059 | 10.2484 | 4.9707
707 | 28278.056 | 10.2498 | 4.9694 |  |  | 707 | 28278.056 | 10.2498 | 4.9645 |  |  | 707 | 28278.056 | 10.2498 | 4.9706
708 | 28318.053 | 10.2513 | 4.9693 |  |  | 708 | 28318.053 | 10.2513 | 4.9644 |  |  | 708 | 28318.053 | 10.2513 | 4.9705
709 | 28358.051 | 10.2527 | 4.9692 |  |  | 709 | 28358.051 | 10.2527 | 4.9643 |  |  | 709 | 28358.051 | 10.2527 | 4.9704
710 | 28398.048 | 10.2541 | 4.9691 |  |  | 710 | 28398.048 | 10.2541 | 4.9641 |  |  | 710 | 28398.048 | 10.2541 | 4.9703
711 | 28438.045 | 10.2555 | 4.969 |  |  | 711 | 28438.045 | 10.2555 | 4.964 |  |  | 711 | 28438.045 | 10.2555 | 4.9701
712 | 28478.042 | 10.2569 | 4.9689 |  |  | 712 | 28478.042 | 10.2569 | 4.9639 |  |  | 712 | 28478.042 | 10.2569 | 4.97
713 | 28518.04 | 10.2583 | 4.9688 |  |  | 713 | 28518.04 | 10.2583 | 4.9638 |  |  | 713 | 28518.04 | 10.2583 | 4.9699
714 | 28558.037 | 10.2597 | 4.9687 |  |  | 714 | 28558.037 | 10.2597 | 4.9637 |  |  | 714 | 28558.037 | 10.2597 | 4.9698
715 | 28598.034 | 10.2611 | 4.9686 |  |  | 715 | 28598.034 | 10.2611 | 4.9636 |  |  | 715 | 28598.034 | 10.2611 | 4.9697
716 | 28638.031 | 10.2625 | 4.9685 |  |  | 716 | 28638.031 | 10.2625 | 4.9635 |  |  | 716 | 28638.031 | 10.2625 | 4.9696
717 | 28678.029 | 10.2639 | 4.9684 |  |  | 717 | 28678.029 | 10.2639 | 4.9634 |  |  | 717 | 28678.029 | 10.2639 | 4.9695
718 | 28718.026 | 10.2653 | 4.9683 |  |  | 718 | 28718.026 | 10.2653 | 4.9633 |  |  | 718 | 28718.026 | 10.2653 | 4.9694
719 | 28758.023 | 10.2667 | 4.9682 |  |  | 719 | 28758.023 | 10.2667 | 4.9632 |  |  | 719 | 28758.023 | 10.2667 | 4.9693
720 | 28798.02 | 10.2681 | 4.968 |  |  | 720 | 28798.02 | 10.2681 | 4.963 |  |  | 720 | 28798.02 | 10.2681 | 4.9692
721 | 28838.018 | 10.2694 | 4.9679 |  |  | 721 | 28838.018 | 10.2694 | 4.9629 |  |  | 721 | 28838.018 | 10.2694 | 4.9691
722 | 28878.015 | 10.2708 | 4.9678 |  |  | 722 | 28878.015 | 10.2708 | 4.9628 |  |  | 722 | 28878.015 | 10.2708 | 4.9689
723 | 28918.012 | 10.2722 | 4.9677 |  |  | 723 | 28918.012 | 10.2722 | 4.9627 |  |  | 723 | 28918.012 | 10.2722 | 4.9688
724 | 28958.009 | 10.2736 | 4.9676 |  |  | 724 | 28958.009 | 10.2736 | 4.9626 |  |  | 724 | 28958.009 | 10.2736 | 4.9687
725 | 28998.007 | 10.275 | 4.9675 |  |  | 725 | 28998.007 | 10.275 | 4.9625 |  |  | 725 | 28998.007 | 10.275 | 4.9686
726 | 29038.004 | 10.2764 | 4.9674 |  |  | 726 | 29038.004 | 10.2764 | 4.9624 |  |  | 726 | 29038.004 | 10.2764 | 4.9685
727 | 29078.001 | 10.2777 | 4.9673 |  |  | 727 | 29078.001 | 10.2777 | 4.9623 |  |  | 727 | 29078.001 | 10.2777 | 4.9684
728 | 29117.998 | 10.2791 | 4.9672 |  |  | 728 | 29117.998 | 10.2791 | 4.9622 |  |  | 728 | 29117.998 | 10.2791 | 4.9683
729 | 29157.996 | 10.2805 | 4.9671 |  |  | 729 | 29157.996 | 10.2805 | 4.9621 |  |  | 729 | 29157.996 | 10.2805 | 4.9682
730 | 29197.993 | 10.2819 | 4.967 |  |  | 730 | 29197.993 | 10.2819 | 4.962 |  |  | 730 | 29197.993 | 10.2819 | 4.9681
731 | 29237.99 | 10.2832 | 4.9669 |  |  | 731 | 29237.99 | 10.2832 | 4.9619 |  |  | 731 | 29237.99 | 10.2832 | 4.968
732 | 29277.987 | 10.2846 | 4.9668 |  |  | 732 | 29277.987 | 10.2846 | 4.9617 |  |  | 732 | 29277.987 | 10.2846 | 4.9679
733 | 29317.985 | 10.286 | 4.9667 |  |  | 733 | 29317.985 | 10.286 | 4.9616 |  |  | 733 | 29317.985 | 10.286 | 4.9678
734 | 29357.982 | 10.2873 | 4.9666 |  |  | 734 | 29357.982 | 10.2873 | 4.9615 |  |  | 734 | 29357.982 | 10.2873 | 4.9677
735 | 29397.979 | 10.2887 | 4.9665 |  |  | 735 | 29397.979 | 10.2887 | 4.9614 |  |  | 735 | 29397.979 | 10.2887 | 4.9676
736 | 29437.976 | 10.29 | 4.9664 |  |  | 736 | 29437.976 | 10.29 | 4.9613 |  |  | 736 | 29437.976 | 10.29 | 4.9674
737 | 29477.974 | 10.2914 | 4.9663 |  |  | 737 | 29477.974 | 10.2914 | 4.9612 |  |  | 737 | 29477.974 | 10.2914 | 4.9673
738 | 29517.971 | 10.2928 | 4.9662 |  |  | 738 | 29517.971 | 10.2928 | 4.9611 |  |  | 738 | 29517.971 | 10.2928 | 4.9672
739 | 29557.968 | 10.2941 | 4.9661 |  |  | 739 | 29557.968 | 10.2941 | 4.961 |  |  | 739 | 29557.968 | 10.2941 | 4.9671
740 | 29597.965 | 10.2955 | 4.966 |  |  | 740 | 29597.965 | 10.2955 | 4.9609 |  |  | 740 | 29597.965 | 10.2955 | 4.967
741 | 29637.963 | 10.2968 | 4.9659 |  |  | 741 | 29637.963 | 10.2968 | 4.9608 |  |  | 741 | 29637.963 | 10.2968 | 4.9669
742 | 29677.96 | 10.2982 | 4.9658 |  |  | 742 | 29677.96 | 10.2982 | 4.9607 |  |  | 742 | 29677.96 | 10.2982 | 4.9668
743 | 29717.957 | 10.2995 | 4.9657 |  |  | 743 | 29717.957 | 10.2995 | 4.9606 |  |  | 743 | 29717.957 | 10.2995 | 4.9667
744 | 29757.954 | 10.3009 | 4.9656 |  |  | 744 | 29757.954 | 10.3009 | 4.9604 |  |  | 744 | 29757.954 | 10.3009 | 4.9666
745 | 29797.952 | 10.3022 | 4.9655 |  |  | 745 | 29797.952 | 10.3022 | 4.9603 |  |  | 745 | 29797.952 | 10.3022 | 4.9665
746 | 29837.949 | 10.3035 | 4.9654 |  |  | 746 | 29837.949 | 10.3035 | 4.9602 |  |  | 746 | 29837.949 | 10.3035 | 4.9664
747 | 29877.946 | 10.3049 | 4.9653 |  |  | 747 | 29877.946 | 10.3049 | 4.9601 |  |  | 747 | 29877.946 | 10.3049 | 4.9663
748 | 29917.943 | 10.3062 | 4.9652 |  |  | 748 | 29917.943 | 10.3062 | 4.96 |  |  | 748 | 29917.943 | 10.3062 | 4.9662
749 | 29957.941 | 10.3075 | 4.9651 |  |  | 749 | 29957.941 | 10.3075 | 4.9599 |  |  | 749 | 29957.941 | 10.3075 | 4.9661
750 | 29997.938 | 10.3089 | 4.965 |  |  | 750 | 29997.938 | 10.3089 | 4.9598 |  |  | 750 | 29997.938 | 10.3089 | 4.966
751 | 30037.935 | 10.3102 | 4.9649 |  |  | 751 | 30037.935 | 10.3102 | 4.9597 |  |  | 751 | 30037.935 | 10.3102 | 4.9658
752 | 30077.932 | 10.3115 | 4.9648 |  |  | 752 | 30077.932 | 10.3115 | 4.9596 |  |  | 752 | 30077.932 | 10.3115 | 4.9657
753 | 30117.93 | 10.3129 | 4.9647 |  |  | 753 | 30117.93 | 10.3129 | 4.9595 |  |  | 753 | 30117.93 | 10.3129 | 4.9656
754 | 30157.927 | 10.3142 | 4.9646 |  |  | 754 | 30157.927 | 10.3142 | 4.9594 |  |  | 754 | 30157.927 | 10.3142 | 4.9655
755 | 30197.924 | 10.3155 | 4.9645 |  |  | 755 | 30197.924 | 10.3155 | 4.9593 |  |  | 755 | 30197.924 | 10.3155 | 4.9654
756 | 30237.921 | 10.3169 | 4.9644 |  |  | 756 | 30237.921 | 10.3169 | 4.9592 |  |  | 756 | 30237.921 | 10.3169 | 4.9653
757 | 30277.919 | 10.3182 | 4.9643 |  |  | 757 | 30277.919 | 10.3182 | 4.9591 |  |  | 757 | 30277.919 | 10.3182 | 4.9652
758 | 30317.916 | 10.3195 | 4.9642 |  |  | 758 | 30317.916 | 10.3195 | 4.959 |  |  | 758 | 30317.916 | 10.3195 | 4.9651
759 | 30357.913 | 10.3208 | 4.9641 |  |  | 759 | 30357.913 | 10.3208 | 4.9588 |  |  | 759 | 30357.913 | 10.3208 | 4.965
760 | 30397.91 | 10.3221 | 4.964 |  |  | 760 | 30397.91 | 10.3221 | 4.9587 |  |  | 760 | 30397.91 | 10.3221 | 4.9649
761 | 30437.908 | 10.3234 | 4.9639 |  |  | 761 | 30437.908 | 10.3234 | 4.9586 |  |  | 761 | 30437.908 | 10.3234 | 4.9648
762 | 30477.905 | 10.3248 | 4.9638 |  |  | 762 | 30477.905 | 10.3248 | 4.9585 |  |  | 762 | 30477.905 | 10.3248 | 4.9647
763 | 30517.902 | 10.3261 | 4.9637 |  |  | 763 | 30517.902 | 10.3261 | 4.9584 |  |  | 763 | 30517.902 | 10.3261 | 4.9646
764 | 30557.899 | 10.3274 | 4.9636 |  |  | 764 | 30557.899 | 10.3274 | 4.9583 |  |  | 764 | 30557.899 | 10.3274 | 4.9645
765 | 30597.897 | 10.3287 | 4.9635 |  |  | 765 | 30597.897 | 10.3287 | 4.9582 |  |  | 765 | 30597.897 | 10.3287 | 4.9644
766 | 30637.894 | 10.33 | 4.9635 |  |  | 766 | 30637.894 | 10.33 | 4.9581 |  |  | 766 | 30637.894 | 10.33 | 4.9643
767 | 30677.891 | 10.3313 | 4.9634 |  |  | 767 | 30677.891 | 10.3313 | 4.958 |  |  | 767 | 30677.891 | 10.3313 | 4.9642
768 | 30717.888 | 10.3326 | 4.9633 |  |  | 768 | 30717.888 | 10.3326 | 4.9579 |  |  | 768 | 30717.888 | 10.3326 | 4.9641
769 | 30757.886 | 10.3339 | 4.9632 |  |  | 769 | 30757.886 | 10.3339 | 4.9578 |  |  | 769 | 30757.886 | 10.3339 | 4.9639
770 | 30797.883 | 10.3352 | 4.9631 |  |  | 770 | 30797.883 | 10.3352 | 4.9577 |  |  | 770 | 30797.883 | 10.3352 | 4.9638
771 | 30837.88 | 10.3365 | 4.963 |  |  | 771 | 30837.88 | 10.3365 | 4.9576 |  |  | 771 | 30837.88 | 10.3365 | 4.9637
772 | 30877.877 | 10.3378 | 4.9629 |  |  | 772 | 30877.877 | 10.3378 | 4.9575 |  |  | 772 | 30877.877 | 10.3378 | 4.9636
773 | 30917.875 | 10.3391 | 4.9628 |  |  | 773 | 30917.875 | 10.3391 | 4.9574 |  |  | 773 | 30917.875 | 10.3391 | 4.9635
774 | 30957.872 | 10.3404 | 4.9627 |  |  | 774 | 30957.872 | 10.3404 | 4.9573 |  |  | 774 | 30957.872 | 10.3404 | 4.9634
775 | 30997.869 | 10.3417 | 4.9626 |  |  | 775 | 30997.869 | 10.3417 | 4.9571 |  |  | 775 | 30997.869 | 10.3417 | 4.9633
776 | 31037.866 | 10.343 | 4.9625 |  |  | 776 | 31037.866 | 10.343 | 4.957 |  |  | 776 | 31037.866 | 10.343 | 4.9632
777 | 31077.864 | 10.3443 | 4.9624 |  |  | 777 | 31077.864 | 10.3443 | 4.9569 |  |  | 777 | 31077.864 | 10.3443 | 4.9631
778 | 31117.861 | 10.3455 | 4.9623 |  |  | 778 | 31117.861 | 10.3455 | 4.9568 |  |  | 778 | 31117.861 | 10.3455 | 4.963
779 | 31157.858 | 10.3468 | 4.9622 |  |  | 779 | 31157.858 | 10.3468 | 4.9567 |  |  | 779 | 31157.858 | 10.3468 | 4.9629
780 | 31197.855 | 10.3481 | 4.9621 |  |  | 780 | 31197.855 | 10.3481 | 4.9566 |  |  | 780 | 31197.855 | 10.3481 | 4.9628
781 | 31237.853 | 10.3494 | 4.962 |  |  | 781 | 31237.853 | 10.3494 | 4.9565 |  |  | 781 | 31237.853 | 10.3494 | 4.9627
782 | 31277.85 | 10.3507 | 4.9619 |  |  | 782 | 31277.85 | 10.3507 | 4.9564 |  |  | 782 | 31277.85 | 10.3507 | 4.9626
783 | 31317.847 | 10.3519 | 4.9618 |  |  | 783 | 31317.847 | 10.3519 | 4.9563 |  |  | 783 | 31317.847 | 10.3519 | 4.9625
784 | 31357.844 | 10.3532 | 4.9617 |  |  | 784 | 31357.844 | 10.3532 | 4.9562 |  |  | 784 | 31357.844 | 10.3532 | 4.9624
785 | 31397.842 | 10.3545 | 4.9616 |  |  | 785 | 31397.842 | 10.3545 | 4.9561 |  |  | 785 | 31397.842 | 10.3545 | 4.9623
786 | 31437.839 | 10.3558 | 4.9615 |  |  | 786 | 31437.839 | 10.3558 | 4.956 |  |  | 786 | 31437.839 | 10.3558 | 4.9622
787 | 31477.836 | 10.357 | 4.9614 |  |  | 787 | 31477.836 | 10.357 | 4.9559 |  |  | 787 | 31477.836 | 10.357 | 4.9621
788 | 31517.833 | 10.3583 | 4.9613 |  |  | 788 | 31517.833 | 10.3583 | 4.9558 |  |  | 788 | 31517.833 | 10.3583 | 4.9619
789 | 31557.831 | 10.3596 | 4.9612 |  |  | 789 | 31557.831 | 10.3596 | 4.9557 |  |  | 789 | 31557.831 | 10.3596 | 4.9618
790 | 31597.828 | 10.3608 | 4.9611 |  |  | 790 | 31597.828 | 10.3608 | 4.9556 |  |  | 790 | 31597.828 | 10.3608 | 4.9617
791 | 31637.825 | 10.3621 | 4.961 |  |  | 791 | 31637.825 | 10.3621 | 4.9555 |  |  | 791 | 31637.825 | 10.3621 | 4.9616
792 | 31677.822 | 10.3634 | 4.9609 |  |  | 792 | 31677.822 | 10.3634 | 4.9554 |  |  | 792 | 31677.822 | 10.3634 | 4.9615
793 | 31717.82 | 10.3646 | 4.9608 |  |  | 793 | 31717.82 | 10.3646 | 4.9553 |  |  | 793 | 31717.82 | 10.3646 | 4.9614
794 | 31757.817 | 10.3659 | 4.9607 |  |  | 794 | 31757.817 | 10.3659 | 4.9551 |  |  | 794 | 31757.817 | 10.3659 | 4.9613
795 | 31797.814 | 10.3672 | 4.9606 |  |  | 795 | 31797.814 | 10.3672 | 4.955 |  |  | 795 | 31797.814 | 10.3672 | 4.9612
796 | 31837.811 | 10.3684 | 4.9605 |  |  | 796 | 31837.811 | 10.3684 | 4.9549 |  |  | 796 | 31837.811 | 10.3684 | 4.9611
797 | 31877.809 | 10.3697 | 4.9604 |  |  | 797 | 31877.809 | 10.3697 | 4.9548 |  |  | 797 | 31877.809 | 10.3697 | 4.961
798 | 31917.806 | 10.3709 | 4.9603 |  |  | 798 | 31917.806 | 10.3709 | 4.9547 |  |  | 798 | 31917.806 | 10.3709 | 4.9609
799 | 31957.803 | 10.3722 | 4.9602 |  |  | 799 | 31957.803 | 10.3722 | 4.9546 |  |  | 799 | 31957.803 | 10.3722 | 4.9608
800 | 31997.8 | 10.3734 | 4.9601 |  |  | 800 | 31997.8 | 10.3734 | 4.9545 |  |  | 800 | 31997.8 | 10.3734 | 4.9607
801 | 32037.798 | 10.3747 | 4.96 |  |  | 801 | 32037.798 | 10.3747 | 4.9544 |  |  | 801 | 32037.798 | 10.3747 | 4.9606
802 | 32077.795 | 10.3759 | 4.9599 |  |  | 802 | 32077.795 | 10.3759 | 4.9543 |  |  | 802 | 32077.795 | 10.3759 | 4.9605
803 | 32117.792 | 10.3772 | 4.9598 |  |  | 803 | 32117.792 | 10.3772 | 4.9542 |  |  | 803 | 32117.792 | 10.3772 | 4.9604
804 | 32157.789 | 10.3784 | 4.9597 |  |  | 804 | 32157.789 | 10.3784 | 4.9541 |  |  | 804 | 32157.789 | 10.3784 | 4.9603
805 | 32197.787 | 10.3797 | 4.9596 |  |  | 805 | 32197.787 | 10.3797 | 4.954 |  |  | 805 | 32197.787 | 10.3797 | 4.9602
806 | 32237.784 | 10.3809 | 4.9595 |  |  | 806 | 32237.784 | 10.3809 | 4.9539 |  |  | 806 | 32237.784 | 10.3809 | 4.9601
807 | 32277.781 | 10.3821 | 4.9594 |  |  | 807 | 32277.781 | 10.3821 | 4.9538 |  |  | 807 | 32277.781 | 10.3821 | 4.96
808 | 32317.778 | 10.3834 | 4.9593 |  |  | 808 | 32317.778 | 10.3834 | 4.9537 |  |  | 808 | 32317.778 | 10.3834 | 4.9599
809 | 32357.776 | 10.3846 | 4.9592 |  |  | 809 | 32357.776 | 10.3846 | 4.9536 |  |  | 809 | 32357.776 | 10.3846 | 4.9597
810 | 32397.773 | 10.3858 | 4.9591 |  |  | 810 | 32397.773 | 10.3858 | 4.9535 |  |  | 810 | 32397.773 | 10.3858 | 4.9596
811 | 32437.77 | 10.3871 | 4.959 |  |  | 811 | 32437.77 | 10.3871 | 4.9534 |  |  | 811 | 32437.77 | 10.3871 | 4.9595
812 | 32477.767 | 10.3883 | 4.9589 |  |  | 812 | 32477.767 | 10.3883 | 4.9533 |  |  | 812 | 32477.767 | 10.3883 | 4.9594
813 | 32517.765 | 10.3895 | 4.9588 |  |  | 813 | 32517.765 | 10.3895 | 4.9532 |  |  | 813 | 32517.765 | 10.3895 | 4.9593
814 | 32557.762 | 10.3908 | 4.9587 |  |  | 814 | 32557.762 | 10.3908 | 4.9531 |  |  | 814 | 32557.762 | 10.3908 | 4.9592
815 | 32597.759 | 10.392 | 4.9586 |  |  | 815 | 32597.759 | 10.392 | 4.9529 |  |  | 815 | 32597.759 | 10.392 | 4.9591
816 | 32637.756 | 10.3932 | 4.9586 |  |  | 816 | 32637.756 | 10.3932 | 4.9528 |  |  | 816 | 32637.756 | 10.3932 | 4.959
817 | 32677.754 | 10.3944 | 4.9585 |  |  | 817 | 32677.754 | 10.3944 | 4.9527 |  |  | 817 | 32677.754 | 10.3944 | 4.9589
818 | 32717.751 | 10.3957 | 4.9584 |  |  | 818 | 32717.751 | 10.3957 | 4.9526 |  |  | 818 | 32717.751 | 10.3957 | 4.9588
819 | 32757.748 | 10.3969 | 4.9583 |  |  | 819 | 32757.748 | 10.3969 | 4.9525 |  |  | 819 | 32757.748 | 10.3969 | 4.9587
820 | 32797.745 | 10.3981 | 4.9582 |  |  | 820 | 32797.745 | 10.3981 | 4.9524 |  |  | 820 | 32797.745 | 10.3981 | 4.9586
821 | 32837.743 | 10.3993 | 4.9581 |  |  | 821 | 32837.743 | 10.3993 | 4.9523 |  |  | 821 | 32837.743 | 10.3993 | 4.9585
822 | 32877.74 | 10.4006 | 4.958 |  |  | 822 | 32877.74 | 10.4006 | 4.9522 |  |  | 822 | 32877.74 | 10.4006 | 4.9584
823 | 32917.737 | 10.4018 | 4.9579 |  |  | 823 | 32917.737 | 10.4018 | 4.9521 |  |  | 823 | 32917.737 | 10.4018 | 4.9583
824 | 32957.734 | 10.403 | 4.9578 |  |  | 824 | 32957.734 | 10.403 | 4.952 |  |  | 824 | 32957.734 | 10.403 | 4.9582
825 | 32997.732 | 10.4042 | 4.9577 |  |  | 825 | 32997.732 | 10.4042 | 4.9519 |  |  | 825 | 32997.732 | 10.4042 | 4.9581
826 | 33037.729 | 10.4054 | 4.9576 |  |  | 826 | 33037.729 | 10.4054 | 4.9518 |  |  | 826 | 33037.729 | 10.4054 | 4.958
827 | 33077.726 | 10.4066 | 4.9575 |  |  | 827 | 33077.726 | 10.4066 | 4.9517 |  |  | 827 | 33077.726 | 10.4066 | 4.9579
828 | 33117.723 | 10.4078 | 4.9574 |  |  | 828 | 33117.723 | 10.4078 | 4.9516 |  |  | 828 | 33117.723 | 10.4078 | 4.9578
829 | 33157.721 | 10.409 | 4.9573 |  |  | 829 | 33157.721 | 10.409 | 4.9515 |  |  | 829 | 33157.721 | 10.409 | 4.9577
830 | 33197.718 | 10.4102 | 4.9572 |  |  | 830 | 33197.718 | 10.4102 | 4.9514 |  |  | 830 | 33197.718 | 10.4102 | 4.9575
831 | 33237.715 | 10.4114 | 4.9571 |  |  | 831 | 33237.715 | 10.4114 | 4.9513 |  |  | 831 | 33237.715 | 10.4114 | 4.9574
832 | 33277.712 | 10.4126 | 4.957 |  |  | 832 | 33277.712 | 10.4126 | 4.9512 |  |  | 832 | 33277.712 | 10.4126 | 4.9573
833 | 33317.71 | 10.4138 | 4.9569 |  |  | 833 | 33317.71 | 10.4138 | 4.9511 |  |  | 833 | 33317.71 | 10.4138 | 4.9572
834 | 33357.707 | 10.415 | 4.9568 |  |  | 834 | 33357.707 | 10.415 | 4.951 |  |  | 834 | 33357.707 | 10.415 | 4.9571
835 | 33397.704 | 10.4162 | 4.9567 |  |  | 835 | 33397.704 | 10.4162 | 4.9509 |  |  | 835 | 33397.704 | 10.4162 | 4.957
836 | 33437.701 | 10.4174 | 4.9566 |  |  | 836 | 33437.701 | 10.4174 | 4.9508 |  |  | 836 | 33437.701 | 10.4174 | 4.9569
837 | 33477.699 | 10.4186 | 4.9565 |  |  | 837 | 33477.699 | 10.4186 | 4.9507 |  |  | 837 | 33477.699 | 10.4186 | 4.9568
838 | 33517.696 | 10.4198 | 4.9564 |  |  | 838 | 33517.696 | 10.4198 | 4.9505 |  |  | 838 | 33517.696 | 10.4198 | 4.9567
839 | 33557.693 | 10.421 | 4.9563 |  |  | 839 | 33557.693 | 10.421 | 4.9504 |  |  | 839 | 33557.693 | 10.421 | 4.9566
840 | 33597.69 | 10.4222 | 4.9562 |  |  | 840 | 33597.69 | 10.4222 | 4.9503 |  |  | 840 | 33597.69 | 10.4222 | 4.9565
841 | 33637.688 | 10.4234 | 4.9561 |  |  | 841 | 33637.688 | 10.4234 | 4.9502 |  |  | 841 | 33637.688 | 10.4234 | 4.9564
842 | 33677.685 | 10.4246 | 4.956 |  |  | 842 | 33677.685 | 10.4246 | 4.9501 |  |  | 842 | 33677.685 | 10.4246 | 4.9563
843 | 33717.682 | 10.4258 | 4.9559 |  |  | 843 | 33717.682 | 10.4258 | 4.95 |  |  | 843 | 33717.682 | 10.4258 | 4.9562
844 | 33757.679 | 10.427 | 4.9558 |  |  | 844 | 33757.679 | 10.427 | 4.9499 |  |  | 844 | 33757.679 | 10.427 | 4.9561
845 | 33797.677 | 10.4281 | 4.9557 |  |  | 845 | 33797.677 | 10.4281 | 4.9498 |  |  | 845 | 33797.677 | 10.4281 | 4.956
846 | 33837.674 | 10.4293 | 4.9556 |  |  | 846 | 33837.674 | 10.4293 | 4.9497 |  |  | 846 | 33837.674 | 10.4293 | 4.9559
847 | 33877.671 | 10.4305 | 4.9555 |  |  | 847 | 33877.671 | 10.4305 | 4.9496 |  |  | 847 | 33877.671 | 10.4305 | 4.9558
848 | 33917.668 | 10.4317 | 4.9554 |  |  | 848 | 33917.668 | 10.4317 | 4.9495 |  |  | 848 | 33917.668 | 10.4317 | 4.9557
849 | 33957.666 | 10.4329 | 4.9553 |  |  | 849 | 33957.666 | 10.4329 | 4.9494 |  |  | 849 | 33957.666 | 10.4329 | 4.9556
850 | 33997.663 | 10.434 | 4.9552 |  |  | 850 | 33997.663 | 10.434 | 4.9493 |  |  | 850 | 33997.663 | 10.434 | 4.9554
851 | 34037.66 | 10.4352 | 4.9551 |  |  | 851 | 34037.66 | 10.4352 | 4.9492 |  |  | 851 | 34037.66 | 10.4352 | 4.9553
852 | 34077.657 | 10.4364 | 4.955 |  |  | 852 | 34077.657 | 10.4364 | 4.9491 |  |  | 852 | 34077.657 | 10.4364 | 4.9552
853 | 34117.655 | 10.4376 | 4.9549 |  |  | 853 | 34117.655 | 10.4376 | 4.949 |  |  | 853 | 34117.655 | 10.4376 | 4.9551
854 | 34157.652 | 10.4387 | 4.9548 |  |  | 854 | 34157.652 | 10.4387 | 4.9489 |  |  | 854 | 34157.652 | 10.4387 | 4.955
855 | 34197.649 | 10.4399 | 4.9547 |  |  | 855 | 34197.649 | 10.4399 | 4.9488 |  |  | 855 | 34197.649 | 10.4399 | 4.9549
856 | 34237.646 | 10.4411 | 4.9546 |  |  | 856 | 34237.646 | 10.4411 | 4.9487 |  |  | 856 | 34237.646 | 10.4411 | 4.9548
857 | 34277.644 | 10.4422 | 4.9545 |  |  | 857 | 34277.644 | 10.4422 | 4.9486 |  |  | 857 | 34277.644 | 10.4422 | 4.9547
858 | 34317.641 | 10.4434 | 4.9544 |  |  | 858 | 34317.641 | 10.4434 | 4.9485 |  |  | 858 | 34317.641 | 10.4434 | 4.9546
859 | 34357.638 | 10.4446 | 4.9543 |  |  | 859 | 34357.638 | 10.4446 | 4.9484 |  |  | 859 | 34357.638 | 10.4446 | 4.9545
860 | 34397.635 | 10.4457 | 4.9543 |  |  | 860 | 34397.635 | 10.4457 | 4.9482 |  |  | 860 | 34397.635 | 10.4457 | 4.9544
861 | 34437.633 | 10.4469 | 4.9542 |  |  | 861 | 34437.633 | 10.4469 | 4.9481 |  |  | 861 | 34437.633 | 10.4469 | 4.9543
862 | 34477.63 | 10.4481 | 4.9541 |  |  | 862 | 34477.63 | 10.4481 | 4.948 |  |  | 862 | 34477.63 | 10.4481 | 4.9542
863 | 34517.627 | 10.4492 | 4.954 |  |  | 863 | 34517.627 | 10.4492 | 4.9479 |  |  | 863 | 34517.627 | 10.4492 | 4.9541
864 | 34557.624 | 10.4504 | 4.9539 |  |  | 864 | 34557.624 | 10.4504 | 4.9478 |  |  | 864 | 34557.624 | 10.4504 | 4.954
865 | 34597.622 | 10.4515 | 4.9538 |  |  | 865 | 34597.622 | 10.4515 | 4.9477 |  |  | 865 | 34597.622 | 10.4515 | 4.9539
866 | 34637.619 | 10.4527 | 4.9537 |  |  | 866 | 34637.619 | 10.4527 | 4.9476 |  |  | 866 | 34637.619 | 10.4527 | 4.9538
867 | 34677.616 | 10.4538 | 4.9536 |  |  | 867 | 34677.616 | 10.4538 | 4.9475 |  |  | 867 | 34677.616 | 10.4538 | 4.9537
868 | 34717.613 | 10.455 | 4.9535 |  |  | 868 | 34717.613 | 10.455 | 4.9474 |  |  | 868 | 34717.613 | 10.455 | 4.9535
869 | 34757.611 | 10.4562 | 4.9534 |  |  | 869 | 34757.611 | 10.4562 | 4.9473 |  |  | 869 | 34757.611 | 10.4562 | 4.9534
870 | 34797.608 | 10.4573 | 4.9533 |  |  | 870 | 34797.608 | 10.4573 | 4.9472 |  |  | 870 | 34797.608 | 10.4573 | 4.9533
871 | 34837.605 | 10.4585 | 4.9532 |  |  | 871 | 34837.605 | 10.4585 | 4.9471 |  |  | 871 | 34837.605 | 10.4585 | 4.9532
872 | 34877.602 | 10.4596 | 4.9531 |  |  | 872 | 34877.602 | 10.4596 | 4.947 |  |  | 872 | 34877.602 | 10.4596 | 4.9531
873 | 34917.6 | 10.4607 | 4.953 |  |  | 873 | 34917.6 | 10.4607 | 4.9469 |  |  | 873 | 34917.6 | 10.4607 | 4.953
874 | 34957.597 | 10.4619 | 4.9529 |  |  | 874 | 34957.597 | 10.4619 | 4.9468 |  |  | 874 | 34957.597 | 10.4619 | 4.9529
875 | 34997.594 | 10.463 | 4.9528 |  |  | 875 | 34997.594 | 10.463 | 4.9467 |  |  | 875 | 34997.594 | 10.463 | 4.9528
876 | 35037.591 | 10.4642 | 4.9527 |  |  | 876 | 35037.591 | 10.4642 | 4.9466 |  |  | 876 | 35037.591 | 10.4642 | 4.9527
877 | 35077.589 | 10.4653 | 4.9526 |  |  | 877 | 35077.589 | 10.4653 | 4.9465 |  |  | 877 | 35077.589 | 10.4653 | 4.9526
878 | 35117.586 | 10.4665 | 4.9525 |  |  | 878 | 35117.586 | 10.4665 | 4.9464 |  |  | 878 | 35117.586 | 10.4665 | 4.9525
879 | 35157.583 | 10.4676 | 4.9524 |  |  | 879 | 35157.583 | 10.4676 | 4.9463 |  |  | 879 | 35157.583 | 10.4676 | 4.9524
880 | 35197.58 | 10.4687 | 4.9523 |  |  | 880 | 35197.58 | 10.4687 | 4.9462 |  |  | 880 | 35197.58 | 10.4687 | 4.9523
881 | 35237.578 | 10.4699 | 4.9522 |  |  | 881 | 35237.578 | 10.4699 | 4.946 |  |  | 881 | 35237.578 | 10.4699 | 4.9522
882 | 35277.575 | 10.471 | 4.9521 |  |  | 882 | 35277.575 | 10.471 | 4.9459 |  |  | 882 | 35277.575 | 10.471 | 4.9521
883 | 35317.572 | 10.4721 | 4.952 |  |  | 883 | 35317.572 | 10.4721 | 4.9458 |  |  | 883 | 35317.572 | 10.4721 | 4.952
884 | 35357.569 | 10.4733 | 4.9519 |  |  | 884 | 35357.569 | 10.4733 | 4.9457 |  |  | 884 | 35357.569 | 10.4733 | 4.9519
885 | 35397.567 | 10.4744 | 4.9518 |  |  | 885 | 35397.567 | 10.4744 | 4.9456 |  |  | 885 | 35397.567 | 10.4744 | 4.9517
886 | 35437.564 | 10.4755 | 4.9517 |  |  | 886 | 35437.564 | 10.4755 | 4.9455 |  |  | 886 | 35437.564 | 10.4755 | 4.9516
887 | 35477.561 | 10.4767 | 4.9516 |  |  | 887 | 35477.561 | 10.4767 | 4.9454 |  |  | 887 | 35477.561 | 10.4767 | 4.9515
888 | 35517.558 | 10.4778 | 4.9515 |  |  | 888 | 35517.558 | 10.4778 | 4.9453 |  |  | 888 | 35517.558 | 10.4778 | 4.9514
889 | 35557.556 | 10.4789 | 4.9514 |  |  | 889 | 35557.556 | 10.4789 | 4.9452 |  |  | 889 | 35557.556 | 10.4789 | 4.9513
890 | 35597.553 | 10.48 | 4.9513 |  |  | 890 | 35597.553 | 10.48 | 4.9451 |  |  | 890 | 35597.553 | 10.48 | 4.9512
891 | 35637.55 | 10.4812 | 4.9512 |  |  | 891 | 35637.55 | 10.4812 | 4.945 |  |  | 891 | 35637.55 | 10.4812 | 4.9511
892 | 35677.547 | 10.4823 | 4.9511 |  |  | 892 | 35677.547 | 10.4823 | 4.9449 |  |  | 892 | 35677.547 | 10.4823 | 4.951
893 | 35717.545 | 10.4834 | 4.951 |  |  | 893 | 35717.545 | 10.4834 | 4.9448 |  |  | 893 | 35717.545 | 10.4834 | 4.9509
894 | 35757.542 | 10.4845 | 4.9509 |  |  | 894 | 35757.542 | 10.4845 | 4.9447 |  |  | 894 | 35757.542 | 10.4845 | 4.9508
895 | 35797.539 | 10.4856 | 4.9508 |  |  | 895 | 35797.539 | 10.4856 | 4.9446 |  |  | 895 | 35797.539 | 10.4856 | 4.9507
896 | 35837.536 | 10.4868 | 4.9507 |  |  | 896 | 35837.536 | 10.4868 | 4.9445 |  |  | 896 | 35837.536 | 10.4868 | 4.9506
897 | 35877.534 | 10.4879 | 4.9506 |  |  | 897 | 35877.534 | 10.4879 | 4.9444 |  |  | 897 | 35877.534 | 10.4879 | 4.9505
898 | 35917.531 | 10.489 | 4.9505 |  |  | 898 | 35917.531 | 10.489 | 4.9443 |  |  | 898 | 35917.531 | 10.489 | 4.9504
899 | 35957.528 | 10.4901 | 4.9504 |  |  | 899 | 35957.528 | 10.4901 | 4.9442 |  |  | 899 | 35957.528 | 10.4901 | 4.9503
900 | 35997.525 | 10.4912 | 4.9503 |  |  | 900 | 35997.525 | 10.4912 | 4.9441 |  |  | 900 | 35997.525 | 10.4912 | 4.9502
901 | 36037.523 | 10.4923 | 4.9502 |  |  | 901 | 36037.523 | 10.4923 | 4.9439 |  |  | 901 | 36037.523 | 10.4923 | 4.95
902 | 36077.52 | 10.4934 | 4.9501 |  |  | 902 | 36077.52 | 10.4934 | 4.9438 |  |  | 902 | 36077.52 | 10.4934 | 4.9499
903 | 36117.517 | 10.4945 | 4.95 |  |  | 903 | 36117.517 | 10.4945 | 4.9437 |  |  | 903 | 36117.517 | 10.4945 | 4.9498
904 | 36157.515 | 10.4956 | 4.9499 |  |  | 904 | 36157.515 | 10.4956 | 4.9436 |  |  | 904 | 36157.515 | 10.4956 | 4.9497
905 | 36197.512 | 10.4967 | 4.9498 |  |  | 905 | 36197.512 | 10.4967 | 4.9435 |  |  | 905 | 36197.512 | 10.4967 | 4.9496
906 | 36237.509 | 10.4978 | 4.9497 |  |  | 906 | 36237.509 | 10.4978 | 4.9434 |  |  | 906 | 36237.509 | 10.4978 | 4.9495
907 | 36277.506 | 10.499 | 4.9496 |  |  | 907 | 36277.506 | 10.499 | 4.9433 |  |  | 907 | 36277.506 | 10.499 | 4.9494
908 | 36317.504 | 10.5001 | 4.9495 |  |  | 908 | 36317.504 | 10.5001 | 4.9432 |  |  | 908 | 36317.504 | 10.5001 | 4.9493
909 | 36357.501 | 10.5012 | 4.9494 |  |  | 909 | 36357.501 | 10.5012 | 4.9431 |  |  | 909 | 36357.501 | 10.5012 | 4.9492
910 | 36397.498 | 10.5023 | 4.9493 |  |  | 910 | 36397.498 | 10.5023 | 4.943 |  |  | 910 | 36397.498 | 10.5023 | 4.9491
911 | 36437.495 | 10.5034 | 4.9492 |  |  | 911 | 36437.495 | 10.5034 | 4.9429 |  |  | 911 | 36437.495 | 10.5034 | 4.949
912 | 36477.493 | 10.5045 | 4.9491 |  |  | 912 | 36477.493 | 10.5045 | 4.9428 |  |  | 912 | 36477.493 | 10.5045 | 4.9489
913 | 36517.49 | 10.5055 | 4.949 |  |  | 913 | 36517.49 | 10.5055 | 4.9427 |  |  | 913 | 36517.49 | 10.5055 | 4.9488
914 | 36557.487 | 10.5066 | 4.9489 |  |  | 914 | 36557.487 | 10.5066 | 4.9426 |  |  | 914 | 36557.487 | 10.5066 | 4.9487
915 | 36597.484 | 10.5077 | 4.9488 |  |  | 915 | 36597.484 | 10.5077 | 4.9425 |  |  | 915 | 36597.484 | 10.5077 | 4.9485
916 | 36637.482 | 10.5088 | 4.9487 |  |  | 916 | 36637.482 | 10.5088 | 4.9424 |  |  | 916 | 36637.482 | 10.5088 | 4.9484
917 | 36677.479 | 10.5099 | 4.9486 |  |  | 917 | 36677.479 | 10.5099 | 4.9423 |  |  | 917 | 36677.479 | 10.5099 | 4.9483
918 | 36717.476 | 10.511 | 4.9485 |  |  | 918 | 36717.476 | 10.511 | 4.9422 |  |  | 918 | 36717.476 | 10.511 | 4.9482
919 | 36757.473 | 10.5121 | 4.9484 |  |  | 919 | 36757.473 | 10.5121 | 4.942 |  |  | 919 | 36757.473 | 10.5121 | 4.9481
920 | 36797.471 | 10.5132 | 4.9483 |  |  | 920 | 36797.471 | 10.5132 | 4.9419 |  |  | 920 | 36797.471 | 10.5132 | 4.948
921 | 36837.468 | 10.5143 | 4.9482 |  |  | 921 | 36837.468 | 10.5143 | 4.9418 |  |  | 921 | 36837.468 | 10.5143 | 4.9479
922 | 36877.465 | 10.5154 | 4.9481 |  |  | 922 | 36877.465 | 10.5154 | 4.9417 |  |  | 922 | 36877.465 | 10.5154 | 4.9478
923 | 36917.462 | 10.5164 | 4.948 |  |  | 923 | 36917.462 | 10.5164 | 4.9416 |  |  | 923 | 36917.462 | 10.5164 | 4.9477
924 | 36957.46 | 10.5175 | 4.9479 |  |  | 924 | 36957.46 | 10.5175 | 4.9415 |  |  | 924 | 36957.46 | 10.5175 | 4.9476
925 | 36997.457 | 10.5186 | 4.9478 |  |  | 925 | 36997.457 | 10.5186 | 4.9414 |  |  | 925 | 36997.457 | 10.5186 | 4.9475
926 | 37037.454 | 10.5197 | 4.9477 |  |  | 926 | 37037.454 | 10.5197 | 4.9413 |  |  | 926 | 37037.454 | 10.5197 | 4.9474
927 | 37077.451 | 10.5208 | 4.9476 |  |  | 927 | 37077.451 | 10.5208 | 4.9412 |  |  | 927 | 37077.451 | 10.5208 | 4.9473
928 | 37117.449 | 10.5218 | 4.9475 |  |  | 928 | 37117.449 | 10.5218 | 4.9411 |  |  | 928 | 37117.449 | 10.5218 | 4.9472
929 | 37157.446 | 10.5229 | 4.9474 |  |  | 929 | 37157.446 | 10.5229 | 4.941 |  |  | 929 | 37157.446 | 10.5229 | 4.947
930 | 37197.443 | 10.524 | 4.9473 |  |  | 930 | 37197.443 | 10.524 | 4.9409 |  |  | 930 | 37197.443 | 10.524 | 4.9469
931 | 37237.44 | 10.5251 | 4.9472 |  |  | 931 | 37237.44 | 10.5251 | 4.9408 |  |  | 931 | 37237.44 | 10.5251 | 4.9468
932 | 37277.438 | 10.5261 | 4.9471 |  |  | 932 | 37277.438 | 10.5261 | 4.9407 |  |  | 932 | 37277.438 | 10.5261 | 4.9467
933 | 37317.435 | 10.5272 | 4.947 |  |  | 933 | 37317.435 | 10.5272 | 4.9406 |  |  | 933 | 37317.435 | 10.5272 | 4.9466
934 | 37357.432 | 10.5283 | 4.9469 |  |  | 934 | 37357.432 | 10.5283 | 4.9405 |  |  | 934 | 37357.432 | 10.5283 | 4.9465
935 | 37397.429 | 10.5294 | 4.9468 |  |  | 935 | 37397.429 | 10.5294 | 4.9404 |  |  | 935 | 37397.429 | 10.5294 | 4.9464
936 | 37437.427 | 10.5304 | 4.9467 |  |  | 936 | 37437.427 | 10.5304 | 4.9402 |  |  | 936 | 37437.427 | 10.5304 | 4.9463
937 | 37477.424 | 10.5315 | 4.9466 |  |  | 937 | 37477.424 | 10.5315 | 4.9401 |  |  | 937 | 37477.424 | 10.5315 | 4.9462
938 | 37517.421 | 10.5326 | 4.9465 |  |  | 938 | 37517.421 | 10.5326 | 4.94 |  |  | 938 | 37517.421 | 10.5326 | 4.9461
939 | 37557.418 | 10.5336 | 4.9464 |  |  | 939 | 37557.418 | 10.5336 | 4.9399 |  |  | 939 | 37557.418 | 10.5336 | 4.946
940 | 37597.416 | 10.5347 | 4.9463 |  |  | 940 | 37597.416 | 10.5347 | 4.9398 |  |  | 940 | 37597.416 | 10.5347 | 4.9459
941 | 37637.413 | 10.5358 | 4.9462 |  |  | 941 | 37637.413 | 10.5358 | 4.9397 |  |  | 941 | 37637.413 | 10.5358 | 4.9457
942 | 37677.41 | 10.5368 | 4.9461 |  |  | 942 | 37677.41 | 10.5368 | 4.9396 |  |  | 942 | 37677.41 | 10.5368 | 4.9456
943 | 37717.407 | 10.5379 | 4.946 |  |  | 943 | 37717.407 | 10.5379 | 4.9395 |  |  | 943 | 37717.407 | 10.5379 | 4.9455
944 | 37757.405 | 10.5389 | 4.9459 |  |  | 944 | 37757.405 | 10.5389 | 4.9394 |  |  | 944 | 37757.405 | 10.5389 | 4.9454
945 | 37797.402 | 10.54 | 4.9458 |  |  | 945 | 37797.402 | 10.54 | 4.9393 |  |  | 945 | 37797.402 | 10.54 | 4.9453
946 | 37837.399 | 10.5411 | 4.9457 |  |  | 946 | 37837.399 | 10.5411 | 4.9392 |  |  | 946 | 37837.399 | 10.5411 | 4.9452
947 | 37877.396 | 10.5421 | 4.9456 |  |  | 947 | 37877.396 | 10.5421 | 4.9391 |  |  | 947 | 37877.396 | 10.5421 | 4.9451
948 | 37917.394 | 10.5432 | 4.9455 |  |  | 948 | 37917.394 | 10.5432 | 4.939 |  |  | 948 | 37917.394 | 10.5432 | 4.945
949 | 37957.391 | 10.5442 | 4.9454 |  |  | 949 | 37957.391 | 10.5442 | 4.9389 |  |  | 949 | 37957.391 | 10.5442 | 4.9449
950 | 37997.388 | 10.5453 | 4.9453 |  |  | 950 | 37997.388 | 10.5453 | 4.9388 |  |  | 950 | 37997.388 | 10.5453 | 4.9448
951 | 38037.385 | 10.5463 | 4.9452 |  |  | 951 | 38037.385 | 10.5463 | 4.9386 |  |  | 951 | 38037.385 | 10.5463 | 4.9447
952 | 38077.383 | 10.5474 | 4.9451 |  |  | 952 | 38077.383 | 10.5474 | 4.9385 |  |  | 952 | 38077.383 | 10.5474 | 4.9446
953 | 38117.38 | 10.5484 | 4.945 |  |  | 953 | 38117.38 | 10.5484 | 4.9384 |  |  | 953 | 38117.38 | 10.5484 | 4.9444
954 | 38157.377 | 10.5495 | 4.9449 |  |  | 954 | 38157.377 | 10.5495 | 4.9383 |  |  | 954 | 38157.377 | 10.5495 | 4.9443
955 | 38197.374 | 10.5505 | 4.9448 |  |  | 955 | 38197.374 | 10.5505 | 4.9382 |  |  | 955 | 38197.374 | 10.5505 | 4.9442
956 | 38237.372 | 10.5516 | 4.9447 |  |  | 956 | 38237.372 | 10.5516 | 4.9381 |  |  | 956 | 38237.372 | 10.5516 | 4.9441
957 | 38277.369 | 10.5526 | 4.9446 |  |  | 957 | 38277.369 | 10.5526 | 4.938 |  |  | 957 | 38277.369 | 10.5526 | 4.944
958 | 38317.366 | 10.5537 | 4.9445 |  |  | 958 | 38317.366 | 10.5537 | 4.9379 |  |  | 958 | 38317.366 | 10.5537 | 4.9439
959 | 38357.363 | 10.5547 | 4.9444 |  |  | 959 | 38357.363 | 10.5547 | 4.9378 |  |  | 959 | 38357.363 | 10.5547 | 4.9438
960 | 38397.361 | 10.5557 | 4.9443 |  |  | 960 | 38397.361 | 10.5557 | 4.9377 |  |  | 960 | 38397.361 | 10.5557 | 4.9437
961 | 38437.358 | 10.5568 | 4.9442 |  |  | 961 | 38437.358 | 10.5568 | 4.9376 |  |  | 961 | 38437.358 | 10.5568 | 4.9436
962 | 38477.355 | 10.5578 | 4.9441 |  |  | 962 | 38477.355 | 10.5578 | 4.9375 |  |  | 962 | 38477.355 | 10.5578 | 4.9435
963 | 38517.352 | 10.5589 | 4.944 |  |  | 963 | 38517.352 | 10.5589 | 4.9374 |  |  | 963 | 38517.352 | 10.5589 | 4.9434
964 | 38557.35 | 10.5599 | 4.9439 |  |  | 964 | 38557.35 | 10.5599 | 4.9373 |  |  | 964 | 38557.35 | 10.5599 | 4.9432
965 | 38597.347 | 10.5609 | 4.9438 |  |  | 965 | 38597.347 | 10.5609 | 4.9372 |  |  | 965 | 38597.347 | 10.5609 | 4.9431
966 | 38637.344 | 10.562 | 4.9437 |  |  | 966 | 38637.344 | 10.562 | 4.937 |  |  | 966 | 38637.344 | 10.562 | 4.943
967 | 38677.341 | 10.563 | 4.9436 |  |  | 967 | 38677.341 | 10.563 | 4.9369 |  |  | 967 | 38677.341 | 10.563 | 4.9429
968 | 38717.339 | 10.564 | 4.9435 |  |  | 968 | 38717.339 | 10.564 | 4.9368 |  |  | 968 | 38717.339 | 10.564 | 4.9428
969 | 38757.336 | 10.5651 | 4.9434 |  |  | 969 | 38757.336 | 10.5651 | 4.9367 |  |  | 969 | 38757.336 | 10.5651 | 4.9427
970 | 38797.333 | 10.5661 | 4.9433 |  |  | 970 | 38797.333 | 10.5661 | 4.9366 |  |  | 970 | 38797.333 | 10.5661 | 4.9426
971 | 38837.33 | 10.5671 | 4.9432 |  |  | 971 | 38837.33 | 10.5671 | 4.9365 |  |  | 971 | 38837.33 | 10.5671 | 4.9425
972 | 38877.328 | 10.5682 | 4.9431 |  |  | 972 | 38877.328 | 10.5682 | 4.9364 |  |  | 972 | 38877.328 | 10.5682 | 4.9424
973 | 38917.325 | 10.5692 | 4.943 |  |  | 973 | 38917.325 | 10.5692 | 4.9363 |  |  | 973 | 38917.325 | 10.5692 | 4.9423
974 | 38957.322 | 10.5702 | 4.9429 |  |  | 974 | 38957.322 | 10.5702 | 4.9362 |  |  | 974 | 38957.322 | 10.5702 | 4.9422
975 | 38997.319 | 10.5712 | 4.9428 |  |  | 975 | 38997.319 | 10.5712 | 4.9361 |  |  | 975 | 38997.319 | 10.5712 | 4.942
976 | 39037.317 | 10.5723 | 4.9427 |  |  | 976 | 39037.317 | 10.5723 | 4.936 |  |  | 976 | 39037.317 | 10.5723 | 4.9419
977 | 39077.314 | 10.5733 | 4.9426 |  |  | 977 | 39077.314 | 10.5733 | 4.9359 |  |  | 977 | 39077.314 | 10.5733 | 4.9418
978 | 39117.311 | 10.5743 | 4.9425 |  |  | 978 | 39117.311 | 10.5743 | 4.9358 |  |  | 978 | 39117.311 | 10.5743 | 4.9417
979 | 39157.308 | 10.5753 | 4.9424 |  |  | 979 | 39157.308 | 10.5753 | 4.9356 |  |  | 979 | 39157.308 | 10.5753 | 4.9416
980 | 39197.306 | 10.5764 | 4.9423 |  |  | 980 | 39197.306 | 10.5764 | 4.9355 |  |  | 980 | 39197.306 | 10.5764 | 4.9415
981 | 39237.303 | 10.5774 | 4.9422 |  |  | 981 | 39237.303 | 10.5774 | 4.9354 |  |  | 981 | 39237.303 | 10.5774 | 4.9414
982 | 39277.3 | 10.5784 | 4.9421 |  |  | 982 | 39277.3 | 10.5784 | 4.9353 |  |  | 982 | 39277.3 | 10.5784 | 4.9413
983 | 39317.297 | 10.5794 | 4.942 |  |  | 983 | 39317.297 | 10.5794 | 4.9352 |  |  | 983 | 39317.297 | 10.5794 | 4.9412
984 | 39357.295 | 10.5804 | 4.9419 |  |  | 984 | 39357.295 | 10.5804 | 4.9351 |  |  | 984 | 39357.295 | 10.5804 | 4.9411
985 | 39397.292 | 10.5815 | 4.9418 |  |  | 985 | 39397.292 | 10.5815 | 4.935 |  |  | 985 | 39397.292 | 10.5815 | 4.9409
986 | 39437.289 | 10.5825 | 4.9417 |  |  | 986 | 39437.289 | 10.5825 | 4.9349 |  |  | 986 | 39437.289 | 10.5825 | 4.9408
987 | 39477.286 | 10.5835 | 4.9416 |  |  | 987 | 39477.286 | 10.5835 | 4.9348 |  |  | 987 | 39477.286 | 10.5835 | 4.9407
988 | 39517.284 | 10.5845 | 4.9415 |  |  | 988 | 39517.284 | 10.5845 | 4.9347 |  |  | 988 | 39517.284 | 10.5845 | 4.9406
989 | 39557.281 | 10.5855 | 4.9414 |  |  | 989 | 39557.281 | 10.5855 | 4.9346 |  |  | 989 | 39557.281 | 10.5855 | 4.9405
990 | 39597.278 | 10.5865 | 4.9413 |  |  | 990 | 39597.278 | 10.5865 | 4.9345 |  |  | 990 | 39597.278 | 10.5865 | 4.9404

## Run7
MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR
File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run7\1 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run7\2 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run7\3 01 001.RAW
Cursor Location: |  |  | (178, 194) |  |  | Cursor Location: |  |  | (178, 193) |  |  | Cursor Location: |  |  | (178, 193)
Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15)
Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev
1 | 39.997 | 3.6888 | 7.7292 | 5.1266 | 0.32159 | 1 | 39.997 | 3.6888 | 7.7399 | 5.1228 | 0.3257 | 1 | 39.997 | 3.6888 | 7.7354 | 5.1228 | 0.32602
2 | 79.995 | 4.382 | 7.3976 |  |  | 2 | 79.995 | 4.382 | 7.4044 |  |  | 2 | 79.995 | 4.382 | 7.4027
3 | 119.992 | 4.7874 | 7.2026 |  |  | 3 | 119.992 | 4.7874 | 7.2094 |  |  | 3 | 119.992 | 4.7874 | 7.2092
4 | 159.989 | 5.0751 | 7.062 |  |  | 4 | 159.989 | 5.0751 | 7.0692 |  |  | 4 | 159.989 | 5.0751 | 7.0698
5 | 199.986 | 5.2982 | 6.9516 |  |  | 5 | 199.986 | 5.2982 | 6.959 |  |  | 5 | 199.986 | 5.2982 | 6.9601
6 | 239.984 | 5.4806 | 6.8605 |  |  | 6 | 239.984 | 5.4806 | 6.8681 |  |  | 6 | 239.984 | 5.4806 | 6.8695
7 | 279.981 | 5.6347 | 6.783 |  |  | 7 | 279.981 | 5.6347 | 6.7908 |  |  | 7 | 279.981 | 5.6347 | 6.7923
8 | 319.978 | 5.7683 | 6.7156 |  |  | 8 | 319.978 | 5.7683 | 6.7234 |  |  | 8 | 319.978 | 5.7683 | 6.725
9 | 359.975 | 5.886 | 6.6559 |  |  | 9 | 359.975 | 5.886 | 6.6637 |  |  | 9 | 359.975 | 5.886 | 6.6654
10 | 399.973 | 5.9914 | 6.6024 |  |  | 10 | 399.973 | 5.9914 | 6.6102 |  |  | 10 | 399.973 | 5.9914 | 6.6119
11 | 439.97 | 6.0867 | 6.5539 |  |  | 11 | 439.97 | 6.0867 | 6.5617 |  |  | 11 | 439.97 | 6.0867 | 6.5633
12 | 479.967 | 6.1737 | 6.5096 |  |  | 12 | 479.967 | 6.1737 | 6.5173 |  |  | 12 | 479.967 | 6.1737 | 6.519
13 | 519.964 | 6.2538 | 6.4688 |  |  | 13 | 519.964 | 6.2538 | 6.4765 |  |  | 13 | 519.964 | 6.2538 | 6.4781
14 | 559.962 | 6.3279 | 6.431 |  |  | 14 | 559.962 | 6.3279 | 6.4386 |  |  | 14 | 559.962 | 6.3279 | 6.4402
15 | 599.959 | 6.3969 | 6.3958 |  |  | 15 | 599.959 | 6.3969 | 6.4033 |  |  | 15 | 599.959 | 6.3969 | 6.4049
16 | 639.956 | 6.4614 | 6.3628 |  |  | 16 | 639.956 | 6.4614 | 6.3703 |  |  | 16 | 639.956 | 6.4614 | 6.3718
17 | 679.953 | 6.522 | 6.3319 |  |  | 17 | 679.953 | 6.522 | 6.3393 |  |  | 17 | 679.953 | 6.522 | 6.3408
18 | 719.951 | 6.5792 | 6.3027 |  |  | 18 | 719.951 | 6.5792 | 6.3101 |  |  | 18 | 719.951 | 6.5792 | 6.3115
19 | 759.948 | 6.6332 | 6.2751 |  |  | 19 | 759.948 | 6.6332 | 6.2824 |  |  | 19 | 759.948 | 6.6332 | 6.2838
20 | 799.945 | 6.6845 | 6.2489 |  |  | 20 | 799.945 | 6.6845 | 6.2562 |  |  | 20 | 799.945 | 6.6845 | 6.2575
21 | 839.942 | 6.7333 | 6.224 |  |  | 21 | 839.942 | 6.7333 | 6.2312 |  |  | 21 | 839.942 | 6.7333 | 6.2325
22 | 879.94 | 6.7799 | 6.2003 |  |  | 22 | 879.94 | 6.7799 | 6.2075 |  |  | 22 | 879.94 | 6.7799 | 6.2087
23 | 919.937 | 6.8243 | 6.1776 |  |  | 23 | 919.937 | 6.8243 | 6.1847 |  |  | 23 | 919.937 | 6.8243 | 6.1859
24 | 959.934 | 6.8669 | 6.1559 |  |  | 24 | 959.934 | 6.8669 | 6.163 |  |  | 24 | 959.934 | 6.8669 | 6.1641
25 | 999.931 | 6.9077 | 6.1351 |  |  | 25 | 999.931 | 6.9077 | 6.1421 |  |  | 25 | 999.931 | 6.9077 | 6.1432
26 | 1039.929 | 6.9469 | 6.1151 |  |  | 26 | 1039.929 | 6.9469 | 6.1221 |  |  | 26 | 1039.929 | 6.9469 | 6.1232
27 | 1079.926 | 6.9846 | 6.0959 |  |  | 27 | 1079.926 | 6.9846 | 6.1028 |  |  | 27 | 1079.926 | 6.9846 | 6.1038
28 | 1119.923 | 7.021 | 6.0774 |  |  | 28 | 1119.923 | 7.021 | 6.0843 |  |  | 28 | 1119.923 | 7.021 | 6.0852
29 | 1159.92 | 7.0561 | 6.0595 |  |  | 29 | 1159.92 | 7.0561 | 6.0664 |  |  | 29 | 1159.92 | 7.0561 | 6.0673
30 | 1199.918 | 7.09 | 6.0423 |  |  | 30 | 1199.918 | 7.09 | 6.0491 |  |  | 30 | 1199.918 | 7.09 | 6.05
31 | 1239.915 | 7.1228 | 6.0256 |  |  | 31 | 1239.915 | 7.1228 | 6.0324 |  |  | 31 | 1239.915 | 7.1228 | 6.0332
32 | 1279.912 | 7.1545 | 6.0095 |  |  | 32 | 1279.912 | 7.1545 | 6.0162 |  |  | 32 | 1279.912 | 7.1545 | 6.017
33 | 1319.909 | 7.1853 | 5.9939 |  |  | 33 | 1319.909 | 7.1853 | 6.0005 |  |  | 33 | 1319.909 | 7.1853 | 6.0013
34 | 1359.907 | 7.2152 | 5.9787 |  |  | 34 | 1359.907 | 7.2152 | 5.9853 |  |  | 34 | 1359.907 | 7.2152 | 5.9861
35 | 1399.904 | 7.2442 | 5.964 |  |  | 35 | 1399.904 | 7.2442 | 5.9706 |  |  | 35 | 1399.904 | 7.2442 | 5.9713
36 | 1439.901 | 7.2723 | 5.9497 |  |  | 36 | 1439.901 | 7.2723 | 5.9563 |  |  | 36 | 1439.901 | 7.2723 | 5.957
37 | 1479.898 | 7.2997 | 5.9359 |  |  | 37 | 1479.898 | 7.2997 | 5.9424 |  |  | 37 | 1479.898 | 7.2997 | 5.943
38 | 1519.896 | 7.3264 | 5.9223 |  |  | 38 | 1519.896 | 7.3264 | 5.9288 |  |  | 38 | 1519.896 | 7.3264 | 5.9295
39 | 1559.893 | 7.3524 | 5.9092 |  |  | 39 | 1559.893 | 7.3524 | 5.9156 |  |  | 39 | 1559.893 | 7.3524 | 5.9162
40 | 1599.89 | 7.3777 | 5.8964 |  |  | 40 | 1599.89 | 7.3777 | 5.9028 |  |  | 40 | 1599.89 | 7.3777 | 5.9034
41 | 1639.887 | 7.4024 | 5.8839 |  |  | 41 | 1639.887 | 7.4024 | 5.8903 |  |  | 41 | 1639.887 | 7.4024 | 5.8909
42 | 1679.885 | 7.4265 | 5.8718 |  |  | 42 | 1679.885 | 7.4265 | 5.8781 |  |  | 42 | 1679.885 | 7.4265 | 5.8786
43 | 1719.882 | 7.45 | 5.8599 |  |  | 43 | 1719.882 | 7.45 | 5.8662 |  |  | 43 | 1719.882 | 7.45 | 5.8667
44 | 1759.879 | 7.473 | 5.8483 |  |  | 44 | 1759.879 | 7.473 | 5.8546 |  |  | 44 | 1759.879 | 7.473 | 5.8551
45 | 1799.876 | 7.4955 | 5.837 |  |  | 45 | 1799.876 | 7.4955 | 5.8433 |  |  | 45 | 1799.876 | 7.4955 | 5.8437
46 | 1839.874 | 7.5175 | 5.826 |  |  | 46 | 1839.874 | 7.5175 | 5.8322 |  |  | 46 | 1839.874 | 7.5175 | 5.8326
47 | 1879.871 | 7.539 | 5.8152 |  |  | 47 | 1879.871 | 7.539 | 5.8214 |  |  | 47 | 1879.871 | 7.539 | 5.8217
48 | 1919.868 | 7.56 | 5.8046 |  |  | 48 | 1919.868 | 7.56 | 5.8108 |  |  | 48 | 1919.868 | 7.56 | 5.8111
49 | 1959.865 | 7.5806 | 5.7943 |  |  | 49 | 1959.865 | 7.5806 | 5.8004 |  |  | 49 | 1959.865 | 7.5806 | 5.8007
50 | 1999.863 | 7.6008 | 5.7841 |  |  | 50 | 1999.863 | 7.6008 | 5.7903 |  |  | 50 | 1999.863 | 7.6008 | 5.7906
51 | 2039.86 | 7.6206 | 5.7742 |  |  | 51 | 2039.86 | 7.6206 | 5.7803 |  |  | 51 | 2039.86 | 7.6206 | 5.7806
52 | 2079.857 | 7.6401 | 5.7645 |  |  | 52 | 2079.857 | 7.6401 | 5.7706 |  |  | 52 | 2079.857 | 7.6401 | 5.7709
53 | 2119.854 | 7.6591 | 5.755 |  |  | 53 | 2119.854 | 7.6591 | 5.7611 |  |  | 53 | 2119.854 | 7.6591 | 5.7613
54 | 2159.852 | 7.6778 | 5.7457 |  |  | 54 | 2159.852 | 7.6778 | 5.7517 |  |  | 54 | 2159.852 | 7.6778 | 5.7519
55 | 2199.849 | 7.6961 | 5.7366 |  |  | 55 | 2199.849 | 7.6961 | 5.7426 |  |  | 55 | 2199.849 | 7.6961 | 5.7428
56 | 2239.846 | 7.7142 | 5.7276 |  |  | 56 | 2239.846 | 7.7142 | 5.7336 |  |  | 56 | 2239.846 | 7.7142 | 5.7338
57 | 2279.843 | 7.7319 | 5.7189 |  |  | 57 | 2279.843 | 7.7319 | 5.7248 |  |  | 57 | 2279.843 | 7.7319 | 5.7249
58 | 2319.841 | 7.7493 | 5.7102 |  |  | 58 | 2319.841 | 7.7493 | 5.7161 |  |  | 58 | 2319.841 | 7.7493 | 5.7163
59 | 2359.838 | 7.7663 | 5.7018 |  |  | 59 | 2359.838 | 7.7663 | 5.7076 |  |  | 59 | 2359.838 | 7.7663 | 5.7078
60 | 2399.835 | 7.7832 | 5.6935 |  |  | 60 | 2399.835 | 7.7832 | 5.6993 |  |  | 60 | 2399.835 | 7.7832 | 5.6994
61 | 2439.832 | 7.7997 | 5.6853 |  |  | 61 | 2439.832 | 7.7997 | 5.6911 |  |  | 61 | 2439.832 | 7.7997 | 5.6912
62 | 2479.83 | 7.8159 | 5.6773 |  |  | 62 | 2479.83 | 7.8159 | 5.6831 |  |  | 62 | 2479.83 | 7.8159 | 5.6831
63 | 2519.827 | 7.8319 | 5.6694 |  |  | 63 | 2519.827 | 7.8319 | 5.6752 |  |  | 63 | 2519.827 | 7.8319 | 5.6752
64 | 2559.824 | 7.8477 | 5.6617 |  |  | 64 | 2559.824 | 7.8477 | 5.6674 |  |  | 64 | 2559.824 | 7.8477 | 5.6675
65 | 2599.821 | 7.8632 | 5.6541 |  |  | 65 | 2599.821 | 7.8632 | 5.6598 |  |  | 65 | 2599.821 | 7.8632 | 5.6598
66 | 2639.819 | 7.8785 | 5.6466 |  |  | 66 | 2639.819 | 7.8785 | 5.6523 |  |  | 66 | 2639.819 | 7.8785 | 5.6523
67 | 2679.816 | 7.8935 | 5.6393 |  |  | 67 | 2679.816 | 7.8935 | 5.6449 |  |  | 67 | 2679.816 | 7.8935 | 5.6449
68 | 2719.813 | 7.9083 | 5.632 |  |  | 68 | 2719.813 | 7.9083 | 5.6376 |  |  | 68 | 2719.813 | 7.9083 | 5.6376
69 | 2759.81 | 7.9229 | 5.6249 |  |  | 69 | 2759.81 | 7.9229 | 5.6305 |  |  | 69 | 2759.81 | 7.9229 | 5.6305
70 | 2799.808 | 7.9373 | 5.6179 |  |  | 70 | 2799.808 | 7.9373 | 5.6235 |  |  | 70 | 2799.808 | 7.9373 | 5.6234
71 | 2839.805 | 7.9515 | 5.611 |  |  | 71 | 2839.805 | 7.9515 | 5.6166 |  |  | 71 | 2839.805 | 7.9515 | 5.6165
72 | 2879.802 | 7.9655 | 5.6043 |  |  | 72 | 2879.802 | 7.9655 | 5.6098 |  |  | 72 | 2879.802 | 7.9655 | 5.6097
73 | 2919.799 | 7.9793 | 5.5976 |  |  | 73 | 2919.799 | 7.9793 | 5.6031 |  |  | 73 | 2919.799 | 7.9793 | 5.603
74 | 2959.797 | 7.9929 | 5.591 |  |  | 74 | 2959.797 | 7.9929 | 5.5965 |  |  | 74 | 2959.797 | 7.9929 | 5.5964
75 | 2999.794 | 8.0063 | 5.5846 |  |  | 75 | 2999.794 | 8.0063 | 5.59 |  |  | 75 | 2999.794 | 8.0063 | 5.5899
76 | 3039.791 | 8.0195 | 5.5782 |  |  | 76 | 3039.791 | 8.0195 | 5.5836 |  |  | 76 | 3039.791 | 8.0195 | 5.5835
77 | 3079.788 | 8.0326 | 5.5719 |  |  | 77 | 3079.788 | 8.0326 | 5.5773 |  |  | 77 | 3079.788 | 8.0326 | 5.5772
78 | 3119.786 | 8.0455 | 5.5658 |  |  | 78 | 3119.786 | 8.0455 | 5.5711 |  |  | 78 | 3119.786 | 8.0455 | 5.571
79 | 3159.783 | 8.0583 | 5.5597 |  |  | 79 | 3159.783 | 8.0583 | 5.565 |  |  | 79 | 3159.783 | 8.0583 | 5.5648
80 | 3199.78 | 8.0708 | 5.5537 |  |  | 80 | 3199.78 | 8.0708 | 5.5589 |  |  | 80 | 3199.78 | 8.0708 | 5.5588
81 | 3239.777 | 8.0833 | 5.5478 |  |  | 81 | 3239.777 | 8.0833 | 5.553 |  |  | 81 | 3239.777 | 8.0833 | 5.5529
82 | 3279.775 | 8.0955 | 5.5419 |  |  | 82 | 3279.775 | 8.0955 | 5.5472 |  |  | 82 | 3279.775 | 8.0955 | 5.547
83 | 3319.772 | 8.1077 | 5.5362 |  |  | 83 | 3319.772 | 8.1077 | 5.5414 |  |  | 83 | 3319.772 | 8.1077 | 5.5412
84 | 3359.769 | 8.1196 | 5.5305 |  |  | 84 | 3359.769 | 8.1196 | 5.5357 |  |  | 84 | 3359.769 | 8.1196 | 5.5355
85 | 3399.766 | 8.1315 | 5.525 |  |  | 85 | 3399.766 | 8.1315 | 5.5301 |  |  | 85 | 3399.766 | 8.1315 | 5.5299
86 | 3439.764 | 8.1432 | 5.5195 |  |  | 86 | 3439.764 | 8.1432 | 5.5246 |  |  | 86 | 3439.764 | 8.1432 | 5.5244
87 | 3479.761 | 8.1547 | 5.514 |  |  | 87 | 3479.761 | 8.1547 | 5.5191 |  |  | 87 | 3479.761 | 8.1547 | 5.5189
88 | 3519.758 | 8.1661 | 5.5087 |  |  | 88 | 3519.758 | 8.1661 | 5.5137 |  |  | 88 | 3519.758 | 8.1661 | 5.5136
89 | 3559.755 | 8.1774 | 5.5034 |  |  | 89 | 3559.755 | 8.1774 | 5.5084 |  |  | 89 | 3559.755 | 8.1774 | 5.5082
90 | 3599.753 | 8.1886 | 5.4982 |  |  | 90 | 3599.753 | 8.1886 | 5.5032 |  |  | 90 | 3599.753 | 8.1886 | 5.503
91 | 3639.75 | 8.1997 | 5.4931 |  |  | 91 | 3639.75 | 8.1997 | 5.498 |  |  | 91 | 3639.75 | 8.1997 | 5.4978
92 | 3679.747 | 8.2106 | 5.488 |  |  | 92 | 3679.747 | 8.2106 | 5.4929 |  |  | 92 | 3679.747 | 8.2106 | 5.4927
93 | 3719.744 | 8.2214 | 5.483 |  |  | 93 | 3719.744 | 8.2214 | 5.4879 |  |  | 93 | 3719.744 | 8.2214 | 5.4877
94 | 3759.742 | 8.2321 | 5.4781 |  |  | 94 | 3759.742 | 8.2321 | 5.4829 |  |  | 94 | 3759.742 | 8.2321 | 5.4827
95 | 3799.739 | 8.2427 | 5.4732 |  |  | 95 | 3799.739 | 8.2427 | 5.478 |  |  | 95 | 3799.739 | 8.2427 | 5.4778
96 | 3839.736 | 8.2532 | 5.4684 |  |  | 96 | 3839.736 | 8.2532 | 5.4732 |  |  | 96 | 3839.736 | 8.2532 | 5.473
97 | 3879.733 | 8.2635 | 5.4636 |  |  | 97 | 3879.733 | 8.2635 | 5.4684 |  |  | 97 | 3879.733 | 8.2635 | 5.4682
98 | 3919.731 | 8.2738 | 5.459 |  |  | 98 | 3919.731 | 8.2738 | 5.4637 |  |  | 98 | 3919.731 | 8.2738 | 5.4635
99 | 3959.728 | 8.2839 | 5.4543 |  |  | 99 | 3959.728 | 8.2839 | 5.4591 |  |  | 99 | 3959.728 | 8.2839 | 5.4588
100 | 3999.725 | 8.294 | 5.4498 |  |  | 100 | 3999.725 | 8.294 | 5.4545 |  |  | 100 | 3999.725 | 8.294 | 5.4542
101 | 4039.722 | 8.3039 | 5.4453 |  |  | 101 | 4039.722 | 8.3039 | 5.4499 |  |  | 101 | 4039.722 | 8.3039 | 5.4497
102 | 4079.72 | 8.3138 | 5.4408 |  |  | 102 | 4079.72 | 8.3138 | 5.4454 |  |  | 102 | 4079.72 | 8.3138 | 5.4452
103 | 4119.717 | 8.3235 | 5.4364 |  |  | 103 | 4119.717 | 8.3235 | 5.441 |  |  | 103 | 4119.717 | 8.3235 | 5.4408
104 | 4159.714 | 8.3332 | 5.4321 |  |  | 104 | 4159.714 | 8.3332 | 5.4367 |  |  | 104 | 4159.714 | 8.3332 | 5.4364
105 | 4199.711 | 8.3428 | 5.4278 |  |  | 105 | 4199.711 | 8.3428 | 5.4323 |  |  | 105 | 4199.711 | 8.3428 | 5.4321
106 | 4239.709 | 8.3522 | 5.4236 |  |  | 106 | 4239.709 | 8.3522 | 5.4281 |  |  | 106 | 4239.709 | 8.3522 | 5.4279
107 | 4279.706 | 8.3616 | 5.4194 |  |  | 107 | 4279.706 | 8.3616 | 5.4239 |  |  | 107 | 4279.706 | 8.3616 | 5.4236
108 | 4319.703 | 8.3709 | 5.4152 |  |  | 108 | 4319.703 | 8.3709 | 5.4197 |  |  | 108 | 4319.703 | 8.3709 | 5.4195
109 | 4359.7 | 8.3802 | 5.4112 |  |  | 109 | 4359.7 | 8.3802 | 5.4156 |  |  | 109 | 4359.7 | 8.3802 | 5.4154
110 | 4399.698 | 8.3893 | 5.4071 |  |  | 110 | 4399.698 | 8.3893 | 5.4115 |  |  | 110 | 4399.698 | 8.3893 | 5.4113
111 | 4439.695 | 8.3983 | 5.4031 |  |  | 111 | 4439.695 | 8.3983 | 5.4075 |  |  | 111 | 4439.695 | 8.3983 | 5.4073
112 | 4479.692 | 8.4073 | 5.3992 |  |  | 112 | 4479.692 | 8.4073 | 5.4035 |  |  | 112 | 4479.692 | 8.4073 | 5.4033
113 | 4519.689 | 8.4162 | 5.3953 |  |  | 113 | 4519.689 | 8.4162 | 5.3996 |  |  | 113 | 4519.689 | 8.4162 | 5.3994
114 | 4559.687 | 8.425 | 5.3915 |  |  | 114 | 4559.687 | 8.425 | 5.3957 |  |  | 114 | 4559.687 | 8.425 | 5.3955
115 | 4599.684 | 8.4337 | 5.3877 |  |  | 115 | 4599.684 | 8.4337 | 5.3919 |  |  | 115 | 4599.684 | 8.4337 | 5.3917
116 | 4639.681 | 8.4424 | 5.3839 |  |  | 116 | 4639.681 | 8.4424 | 5.3881 |  |  | 116 | 4639.681 | 8.4424 | 5.3879
117 | 4679.678 | 8.451 | 5.3802 |  |  | 117 | 4679.678 | 8.451 | 5.3844 |  |  | 117 | 4679.678 | 8.451 | 5.3842
118 | 4719.676 | 8.4595 | 5.3765 |  |  | 118 | 4719.676 | 8.4595 | 5.3807 |  |  | 118 | 4719.676 | 8.4595 | 5.3805
119 | 4759.673 | 8.4679 | 5.3729 |  |  | 119 | 4759.673 | 8.4679 | 5.377 |  |  | 119 | 4759.673 | 8.4679 | 5.3768
120 | 4799.67 | 8.4763 | 5.3693 |  |  | 120 | 4799.67 | 8.4763 | 5.3734 |  |  | 120 | 4799.67 | 8.4763 | 5.3732
121 | 4839.667 | 8.4846 | 5.3658 |  |  | 121 | 4839.667 | 8.4846 | 5.3698 |  |  | 121 | 4839.667 | 8.4846 | 5.3696
122 | 4879.665 | 8.4928 | 5.3623 |  |  | 122 | 4879.665 | 8.4928 | 5.3663 |  |  | 122 | 4879.665 | 8.4928 | 5.3661
123 | 4919.662 | 8.501 | 5.3588 |  |  | 123 | 4919.662 | 8.501 | 5.3628 |  |  | 123 | 4919.662 | 8.501 | 5.3626
124 | 4959.659 | 8.5091 | 5.3554 |  |  | 124 | 4959.659 | 8.5091 | 5.3593 |  |  | 124 | 4959.659 | 8.5091 | 5.3592
125 | 4999.656 | 8.5171 | 5.352 |  |  | 125 | 4999.656 | 8.5171 | 5.3559 |  |  | 125 | 4999.656 | 8.5171 | 5.3557
126 | 5039.654 | 8.5251 | 5.3486 |  |  | 126 | 5039.654 | 8.5251 | 5.3525 |  |  | 126 | 5039.654 | 8.5251 | 5.3524
127 | 5079.651 | 8.533 | 5.3453 |  |  | 127 | 5079.651 | 8.533 | 5.3492 |  |  | 127 | 5079.651 | 8.533 | 5.349
128 | 5119.648 | 8.5408 | 5.342 |  |  | 128 | 5119.648 | 8.5408 | 5.3459 |  |  | 128 | 5119.648 | 8.5408 | 5.3457
129 | 5159.645 | 8.5486 | 5.3388 |  |  | 129 | 5159.645 | 8.5486 | 5.3426 |  |  | 129 | 5159.645 | 8.5486 | 5.3425
130 | 5199.643 | 8.5563 | 5.3356 |  |  | 130 | 5199.643 | 8.5563 | 5.3394 |  |  | 130 | 5199.643 | 8.5563 | 5.3392
131 | 5239.64 | 8.564 | 5.3324 |  |  | 131 | 5239.64 | 8.564 | 5.3362 |  |  | 131 | 5239.64 | 8.564 | 5.336
132 | 5279.637 | 8.5716 | 5.3293 |  |  | 132 | 5279.637 | 8.5716 | 5.333 |  |  | 132 | 5279.637 | 8.5716 | 5.3329
133 | 5319.634 | 8.5792 | 5.3262 |  |  | 133 | 5319.634 | 8.5792 | 5.3299 |  |  | 133 | 5319.634 | 8.5792 | 5.3298
134 | 5359.632 | 8.5867 | 5.3231 |  |  | 134 | 5359.632 | 8.5867 | 5.3268 |  |  | 134 | 5359.632 | 8.5867 | 5.3267
135 | 5399.629 | 8.5941 | 5.3201 |  |  | 135 | 5399.629 | 8.5941 | 5.3237 |  |  | 135 | 5399.629 | 8.5941 | 5.3236
136 | 5439.626 | 8.6015 | 5.3171 |  |  | 136 | 5439.626 | 8.6015 | 5.3207 |  |  | 136 | 5439.626 | 8.6015 | 5.3206
137 | 5479.623 | 8.6088 | 5.3141 |  |  | 137 | 5479.623 | 8.6088 | 5.3177 |  |  | 137 | 5479.623 | 8.6088 | 5.3176
138 | 5519.621 | 8.6161 | 5.3112 |  |  | 138 | 5519.621 | 8.6161 | 5.3147 |  |  | 138 | 5519.621 | 8.6161 | 5.3146
139 | 5559.618 | 8.6233 | 5.3083 |  |  | 139 | 5559.618 | 8.6233 | 5.3118 |  |  | 139 | 5559.618 | 8.6233 | 5.3117
140 | 5599.615 | 8.6305 | 5.3054 |  |  | 140 | 5599.615 | 8.6305 | 5.3089 |  |  | 140 | 5599.615 | 8.6305 | 5.3088
141 | 5639.612 | 8.6376 | 5.3026 |  |  | 141 | 5639.612 | 8.6376 | 5.306 |  |  | 141 | 5639.612 | 8.6376 | 5.3059
142 | 5679.61 | 8.6446 | 5.2997 |  |  | 142 | 5679.61 | 8.6446 | 5.3032 |  |  | 142 | 5679.61 | 8.6446 | 5.3031
143 | 5719.607 | 8.6517 | 5.297 |  |  | 143 | 5719.607 | 8.6517 | 5.3003 |  |  | 143 | 5719.607 | 8.6517 | 5.3003
144 | 5759.604 | 8.6586 | 5.2942 |  |  | 144 | 5759.604 | 8.6586 | 5.2976 |  |  | 144 | 5759.604 | 8.6586 | 5.2975
145 | 5799.601 | 8.6655 | 5.2915 |  |  | 145 | 5799.601 | 8.6655 | 5.2948 |  |  | 145 | 5799.601 | 8.6655 | 5.2948
146 | 5839.599 | 8.6724 | 5.2888 |  |  | 146 | 5839.599 | 8.6724 | 5.2921 |  |  | 146 | 5839.599 | 8.6724 | 5.292
147 | 5879.596 | 8.6792 | 5.2861 |  |  | 147 | 5879.596 | 8.6792 | 5.2894 |  |  | 147 | 5879.596 | 8.6792 | 5.2893
148 | 5919.593 | 8.686 | 5.2835 |  |  | 148 | 5919.593 | 8.686 | 5.2867 |  |  | 148 | 5919.593 | 8.686 | 5.2867
149 | 5959.59 | 8.6928 | 5.2808 |  |  | 149 | 5959.59 | 8.6928 | 5.2841 |  |  | 149 | 5959.59 | 8.6928 | 5.284
150 | 5999.588 | 8.6994 | 5.2783 |  |  | 150 | 5999.588 | 8.6994 | 5.2814 |  |  | 150 | 5999.588 | 8.6994 | 5.2814
151 | 6039.585 | 8.7061 | 5.2757 |  |  | 151 | 6039.585 | 8.7061 | 5.2788 |  |  | 151 | 6039.585 | 8.7061 | 5.2788
152 | 6079.582 | 8.7127 | 5.2732 |  |  | 152 | 6079.582 | 8.7127 | 5.2763 |  |  | 152 | 6079.582 | 8.7127 | 5.2763
153 | 6119.579 | 8.7192 | 5.2707 |  |  | 153 | 6119.579 | 8.7192 | 5.2737 |  |  | 153 | 6119.579 | 8.7192 | 5.2737
154 | 6159.577 | 8.7258 | 5.2682 |  |  | 154 | 6159.577 | 8.7258 | 5.2712 |  |  | 154 | 6159.577 | 8.7258 | 5.2712
155 | 6199.574 | 8.7322 | 5.2657 |  |  | 155 | 6199.574 | 8.7322 | 5.2687 |  |  | 155 | 6199.574 | 8.7322 | 5.2688
156 | 6239.571 | 8.7387 | 5.2633 |  |  | 156 | 6239.571 | 8.7387 | 5.2663 |  |  | 156 | 6239.571 | 8.7387 | 5.2663
157 | 6279.568 | 8.7451 | 5.2609 |  |  | 157 | 6279.568 | 8.7451 | 5.2638 |  |  | 157 | 6279.568 | 8.7451 | 5.2639
158 | 6319.566 | 8.7514 | 5.2585 |  |  | 158 | 6319.566 | 8.7514 | 5.2614 |  |  | 158 | 6319.566 | 8.7514 | 5.2615
159 | 6359.563 | 8.7577 | 5.2561 |  |  | 159 | 6359.563 | 8.7577 | 5.259 |  |  | 159 | 6359.563 | 8.7577 | 5.2591
160 | 6399.56 | 8.764 | 5.2538 |  |  | 160 | 6399.56 | 8.764 | 5.2566 |  |  | 160 | 6399.56 | 8.764 | 5.2567
161 | 6439.557 | 8.7702 | 5.2515 |  |  | 161 | 6439.557 | 8.7702 | 5.2543 |  |  | 161 | 6439.557 | 8.7702 | 5.2544
162 | 6479.555 | 8.7764 | 5.2492 |  |  | 162 | 6479.555 | 8.7764 | 5.252 |  |  | 162 | 6479.555 | 8.7764 | 5.2521
163 | 6519.552 | 8.7826 | 5.2469 |  |  | 163 | 6519.552 | 8.7826 | 5.2497 |  |  | 163 | 6519.552 | 8.7826 | 5.2498
164 | 6559.549 | 8.7887 | 5.2447 |  |  | 164 | 6559.549 | 8.7887 | 5.2474 |  |  | 164 | 6559.549 | 8.7887 | 5.2475
165 | 6599.546 | 8.7948 | 5.2424 |  |  | 165 | 6599.546 | 8.7948 | 5.2451 |  |  | 165 | 6599.546 | 8.7948 | 5.2453
166 | 6639.544 | 8.8008 | 5.2402 |  |  | 166 | 6639.544 | 8.8008 | 5.2429 |  |  | 166 | 6639.544 | 8.8008 | 5.243
167 | 6679.541 | 8.8068 | 5.238 |  |  | 167 | 6679.541 | 8.8068 | 5.2407 |  |  | 167 | 6679.541 | 8.8068 | 5.2408
168 | 6719.538 | 8.8128 | 5.2359 |  |  | 168 | 6719.538 | 8.8128 | 5.2385 |  |  | 168 | 6719.538 | 8.8128 | 5.2387
169 | 6759.535 | 8.8187 | 5.2337 |  |  | 169 | 6759.535 | 8.8187 | 5.2363 |  |  | 169 | 6759.535 | 8.8187 | 5.2365
170 | 6799.533 | 8.8246 | 5.2316 |  |  | 170 | 6799.533 | 8.8246 | 5.2342 |  |  | 170 | 6799.533 | 8.8246 | 5.2344
171 | 6839.53 | 8.8305 | 5.2295 |  |  | 171 | 6839.53 | 8.8305 | 5.2321 |  |  | 171 | 6839.53 | 8.8305 | 5.2322
172 | 6879.527 | 8.8363 | 5.2275 |  |  | 172 | 6879.527 | 8.8363 | 5.23 |  |  | 172 | 6879.527 | 8.8363 | 5.2302
173 | 6919.524 | 8.8421 | 5.2254 |  |  | 173 | 6919.524 | 8.8421 | 5.2279 |  |  | 173 | 6919.524 | 8.8421 | 5.2281
174 | 6959.522 | 8.8479 | 5.2234 |  |  | 174 | 6959.522 | 8.8479 | 5.2258 |  |  | 174 | 6959.522 | 8.8479 | 5.226
175 | 6999.519 | 8.8536 | 5.2213 |  |  | 175 | 6999.519 | 8.8536 | 5.2238 |  |  | 175 | 6999.519 | 8.8536 | 5.224
176 | 7039.516 | 8.8593 | 5.2193 |  |  | 176 | 7039.516 | 8.8593 | 5.2217 |  |  | 176 | 7039.516 | 8.8593 | 5.222
177 | 7079.513 | 8.865 | 5.2174 |  |  | 177 | 7079.513 | 8.865 | 5.2197 |  |  | 177 | 7079.513 | 8.865 | 5.22
178 | 7119.511 | 8.8706 | 5.2154 |  |  | 178 | 7119.511 | 8.8706 | 5.2177 |  |  | 178 | 7119.511 | 8.8706 | 5.218
179 | 7159.508 | 8.8762 | 5.2135 |  |  | 179 | 7159.508 | 8.8762 | 5.2158 |  |  | 179 | 7159.508 | 8.8762 | 5.216
180 | 7199.505 | 8.8818 | 5.2115 |  |  | 180 | 7199.505 | 8.8818 | 5.2138 |  |  | 180 | 7199.505 | 8.8818 | 5.2141
181 | 7239.502 | 8.8873 | 5.2096 |  |  | 181 | 7239.502 | 8.8873 | 5.2119 |  |  | 181 | 7239.502 | 8.8873 | 5.2122
182 | 7279.5 | 8.8928 | 5.2078 |  |  | 182 | 7279.5 | 8.8928 | 5.21 |  |  | 182 | 7279.5 | 8.8928 | 5.2103
183 | 7319.497 | 8.8983 | 5.2059 |  |  | 183 | 7319.497 | 8.8983 | 5.2081 |  |  | 183 | 7319.497 | 8.8983 | 5.2084
184 | 7359.494 | 8.9037 | 5.204 |  |  | 184 | 7359.494 | 8.9037 | 5.2062 |  |  | 184 | 7359.494 | 8.9037 | 5.2065
185 | 7399.491 | 8.9092 | 5.2022 |  |  | 185 | 7399.491 | 8.9092 | 5.2043 |  |  | 185 | 7399.491 | 8.9092 | 5.2047
186 | 7439.489 | 8.9146 | 5.2004 |  |  | 186 | 7439.489 | 8.9146 | 5.2025 |  |  | 186 | 7439.489 | 8.9146 | 5.2028
187 | 7479.486 | 8.9199 | 5.1986 |  |  | 187 | 7479.486 | 8.9199 | 5.2007 |  |  | 187 | 7479.486 | 8.9199 | 5.201
188 | 7519.483 | 8.9253 | 5.1968 |  |  | 188 | 7519.483 | 8.9253 | 5.1989 |  |  | 188 | 7519.483 | 8.9253 | 5.1992
189 | 7559.48 | 8.9306 | 5.195 |  |  | 189 | 7559.48 | 8.9306 | 5.1971 |  |  | 189 | 7559.48 | 8.9306 | 5.1974
190 | 7599.478 | 8.9358 | 5.1933 |  |  | 190 | 7599.478 | 8.9358 | 5.1953 |  |  | 190 | 7599.478 | 8.9358 | 5.1957
191 | 7639.475 | 8.9411 | 5.1916 |  |  | 191 | 7639.475 | 8.9411 | 5.1935 |  |  | 191 | 7639.475 | 8.9411 | 5.1939
192 | 7679.472 | 8.9463 | 5.1899 |  |  | 192 | 7679.472 | 8.9463 | 5.1918 |  |  | 192 | 7679.472 | 8.9463 | 5.1922
193 | 7719.469 | 8.9515 | 5.1882 |  |  | 193 | 7719.469 | 8.9515 | 5.19 |  |  | 193 | 7719.469 | 8.9515 | 5.1904
194 | 7759.467 | 8.9567 | 5.1865 |  |  | 194 | 7759.467 | 8.9567 | 5.1883 |  |  | 194 | 7759.467 | 8.9567 | 5.1887
195 | 7799.464 | 8.9618 | 5.1848 |  |  | 195 | 7799.464 | 8.9618 | 5.1866 |  |  | 195 | 7799.464 | 8.9618 | 5.1871
196 | 7839.461 | 8.9669 | 5.1831 |  |  | 196 | 7839.461 | 8.9669 | 5.185 |  |  | 196 | 7839.461 | 8.9669 | 5.1854
197 | 7879.458 | 8.972 | 5.1815 |  |  | 197 | 7879.458 | 8.972 | 5.1833 |  |  | 197 | 7879.458 | 8.972 | 5.1837
198 | 7919.456 | 8.9771 | 5.1799 |  |  | 198 | 7919.456 | 8.9771 | 5.1816 |  |  | 198 | 7919.456 | 8.9771 | 5.1821
199 | 7959.453 | 8.9821 | 5.1783 |  |  | 199 | 7959.453 | 8.9821 | 5.18 |  |  | 199 | 7959.453 | 8.9821 | 5.1805
200 | 7999.45 | 8.9871 | 5.1767 |  |  | 200 | 7999.45 | 8.9871 | 5.1784 |  |  | 200 | 7999.45 | 8.9871 | 5.1789
201 | 8039.447 | 8.9921 | 5.1751 |  |  | 201 | 8039.447 | 8.9921 | 5.1768 |  |  | 201 | 8039.447 | 8.9921 | 5.1773
202 | 8079.445 | 8.9971 | 5.1735 |  |  | 202 | 8079.445 | 8.9971 | 5.1752 |  |  | 202 | 8079.445 | 8.9971 | 5.1757
203 | 8119.442 | 9.002 | 5.172 |  |  | 203 | 8119.442 | 9.002 | 5.1736 |  |  | 203 | 8119.442 | 9.002 | 5.1741
204 | 8159.439 | 9.0069 | 5.1705 |  |  | 204 | 8159.439 | 9.0069 | 5.172 |  |  | 204 | 8159.439 | 9.0069 | 5.1726
205 | 8199.436 | 9.0118 | 5.1689 |  |  | 205 | 8199.436 | 9.0118 | 5.1705 |  |  | 205 | 8199.436 | 9.0118 | 5.171
206 | 8239.434 | 9.0167 | 5.1674 |  |  | 206 | 8239.434 | 9.0167 | 5.169 |  |  | 206 | 8239.434 | 9.0167 | 5.1695
207 | 8279.431 | 9.0215 | 5.1659 |  |  | 207 | 8279.431 | 9.0215 | 5.1674 |  |  | 207 | 8279.431 | 9.0215 | 5.168
208 | 8319.428 | 9.0263 | 5.1645 |  |  | 208 | 8319.428 | 9.0263 | 5.1659 |  |  | 208 | 8319.428 | 9.0263 | 5.1665
209 | 8359.425 | 9.0311 | 5.163 |  |  | 209 | 8359.425 | 9.0311 | 5.1644 |  |  | 209 | 8359.425 | 9.0311 | 5.165
210 | 8399.423 | 9.0359 | 5.1615 |  |  | 210 | 8399.423 | 9.0359 | 5.1629 |  |  | 210 | 8399.423 | 9.0359 | 5.1635
211 | 8439.42 | 9.0407 | 5.1601 |  |  | 211 | 8439.42 | 9.0407 | 5.1615 |  |  | 211 | 8439.42 | 9.0407 | 5.1621
212 | 8479.417 | 9.0454 | 5.1587 |  |  | 212 | 8479.417 | 9.0454 | 5.16 |  |  | 212 | 8479.417 | 9.0454 | 5.1606
213 | 8519.414 | 9.0501 | 5.1572 |  |  | 213 | 8519.414 | 9.0501 | 5.1586 |  |  | 213 | 8519.414 | 9.0501 | 5.1592
214 | 8559.412 | 9.0548 | 5.1558 |  |  | 214 | 8559.412 | 9.0548 | 5.1571 |  |  | 214 | 8559.412 | 9.0548 | 5.1578
215 | 8599.409 | 9.0594 | 5.1545 |  |  | 215 | 8599.409 | 9.0594 | 5.1557 |  |  | 215 | 8599.409 | 9.0594 | 5.1564
216 | 8639.406 | 9.0641 | 5.1531 |  |  | 216 | 8639.406 | 9.0641 | 5.1543 |  |  | 216 | 8639.406 | 9.0641 | 5.155
217 | 8679.403 | 9.0687 | 5.1517 |  |  | 217 | 8679.403 | 9.0687 | 5.1529 |  |  | 217 | 8679.403 | 9.0687 | 5.1536
218 | 8719.401 | 9.0733 | 5.1504 |  |  | 218 | 8719.401 | 9.0733 | 5.1515 |  |  | 218 | 8719.401 | 9.0733 | 5.1522
219 | 8759.398 | 9.0779 | 5.149 |  |  | 219 | 8759.398 | 9.0779 | 5.1502 |  |  | 219 | 8759.398 | 9.0779 | 5.1509
220 | 8799.395 | 9.0824 | 5.1477 |  |  | 220 | 8799.395 | 9.0824 | 5.1488 |  |  | 220 | 8799.395 | 9.0824 | 5.1495
221 | 8839.392 | 9.087 | 5.1464 |  |  | 221 | 8839.392 | 9.087 | 5.1475 |  |  | 221 | 8839.392 | 9.087 | 5.1482
222 | 8879.39 | 9.0915 | 5.1451 |  |  | 222 | 8879.39 | 9.0915 | 5.1461 |  |  | 222 | 8879.39 | 9.0915 | 5.1469
223 | 8919.387 | 9.096 | 5.1438 |  |  | 223 | 8919.387 | 9.096 | 5.1448 |  |  | 223 | 8919.387 | 9.096 | 5.1455
224 | 8959.384 | 9.1005 | 5.1425 |  |  | 224 | 8959.384 | 9.1005 | 5.1435 |  |  | 224 | 8959.384 | 9.1005 | 5.1442
225 | 8999.381 | 9.1049 | 5.1412 |  |  | 225 | 8999.381 | 9.1049 | 5.1422 |  |  | 225 | 8999.381 | 9.1049 | 5.143
226 | 9039.379 | 9.1093 | 5.1399 |  |  | 226 | 9039.379 | 9.1093 | 5.1409 |  |  | 226 | 9039.379 | 9.1093 | 5.1417
227 | 9079.376 | 9.1138 | 5.1387 |  |  | 227 | 9079.376 | 9.1138 | 5.1396 |  |  | 227 | 9079.376 | 9.1138 | 5.1404
228 | 9119.373 | 9.1182 | 5.1375 |  |  | 228 | 9119.373 | 9.1182 | 5.1384 |  |  | 228 | 9119.373 | 9.1182 | 5.1392
229 | 9159.37 | 9.1225 | 5.1362 |  |  | 229 | 9159.37 | 9.1225 | 5.1371 |  |  | 229 | 9159.37 | 9.1225 | 5.1379
230 | 9199.368 | 9.1269 | 5.135 |  |  | 230 | 9199.368 | 9.1269 | 5.1359 |  |  | 230 | 9199.368 | 9.1269 | 5.1367
231 | 9239.365 | 9.1312 | 5.1338 |  |  | 231 | 9239.365 | 9.1312 | 5.1346 |  |  | 231 | 9239.365 | 9.1312 | 5.1355
232 | 9279.362 | 9.1355 | 5.1326 |  |  | 232 | 9279.362 | 9.1355 | 5.1334 |  |  | 232 | 9279.362 | 9.1355 | 5.1342
233 | 9319.359 | 9.1398 | 5.1314 |  |  | 233 | 9319.359 | 9.1398 | 5.1322 |  |  | 233 | 9319.359 | 9.1398 | 5.133
234 | 9359.357 | 9.1441 | 5.1303 |  |  | 234 | 9359.357 | 9.1441 | 5.131 |  |  | 234 | 9359.357 | 9.1441 | 5.1318
235 | 9399.354 | 9.1484 | 5.1291 |  |  | 235 | 9399.354 | 9.1484 | 5.1298 |  |  | 235 | 9399.354 | 9.1484 | 5.1307
236 | 9439.351 | 9.1526 | 5.1279 |  |  | 236 | 9439.351 | 9.1526 | 5.1286 |  |  | 236 | 9439.351 | 9.1526 | 5.1295
237 | 9479.348 | 9.1569 | 5.1268 |  |  | 237 | 9479.348 | 9.1569 | 5.1275 |  |  | 237 | 9479.348 | 9.1569 | 5.1283
238 | 9519.346 | 9.1611 | 5.1256 |  |  | 238 | 9519.346 | 9.1611 | 5.1263 |  |  | 238 | 9519.346 | 9.1611 | 5.1272
239 | 9559.343 | 9.1653 | 5.1245 |  |  | 239 | 9559.343 | 9.1653 | 5.1252 |  |  | 239 | 9559.343 | 9.1653 | 5.126
240 | 9599.34 | 9.1694 | 5.1234 |  |  | 240 | 9599.34 | 9.1694 | 5.124 |  |  | 240 | 9599.34 | 9.1694 | 5.1249
241 | 9639.337 | 9.1736 | 5.1223 |  |  | 241 | 9639.337 | 9.1736 | 5.1229 |  |  | 241 | 9639.337 | 9.1736 | 5.1238
242 | 9679.335 | 9.1777 | 5.1212 |  |  | 242 | 9679.335 | 9.1777 | 5.1218 |  |  | 242 | 9679.335 | 9.1777 | 5.1227
243 | 9719.332 | 9.1819 | 5.1201 |  |  | 243 | 9719.332 | 9.1819 | 5.1206 |  |  | 243 | 9719.332 | 9.1819 | 5.1216
244 | 9759.329 | 9.186 | 5.119 |  |  | 244 | 9759.329 | 9.186 | 5.1195 |  |  | 244 | 9759.329 | 9.186 | 5.1205
245 | 9799.326 | 9.1901 | 5.118 |  |  | 245 | 9799.326 | 9.1901 | 5.1184 |  |  | 245 | 9799.326 | 9.1901 | 5.1194
246 | 9839.324 | 9.1941 | 5.1169 |  |  | 246 | 9839.324 | 9.1941 | 5.1174 |  |  | 246 | 9839.324 | 9.1941 | 5.1183
247 | 9879.321 | 9.1982 | 5.1159 |  |  | 247 | 9879.321 | 9.1982 | 5.1163 |  |  | 247 | 9879.321 | 9.1982 | 5.1173
248 | 9919.318 | 9.2022 | 5.1148 |  |  | 248 | 9919.318 | 9.2022 | 5.1152 |  |  | 248 | 9919.318 | 9.2022 | 5.1162
249 | 9959.315 | 9.2063 | 5.1138 |  |  | 249 | 9959.315 | 9.2063 | 5.1142 |  |  | 249 | 9959.315 | 9.2063 | 5.1151
250 | 9999.313 | 9.2103 | 5.1128 |  |  | 250 | 9999.313 | 9.2103 | 5.1131 |  |  | 250 | 9999.313 | 9.2103 | 5.1141
251 | 10039.31 | 9.2143 | 5.1117 |  |  | 251 | 10039.31 | 9.2143 | 5.1121 |  |  | 251 | 10039.31 | 9.2143 | 5.1131
252 | 10079.307 | 9.2182 | 5.1107 |  |  | 252 | 10079.307 | 9.2182 | 5.111 |  |  | 252 | 10079.307 | 9.2182 | 5.112
253 | 10119.304 | 9.2222 | 5.1097 |  |  | 253 | 10119.304 | 9.2222 | 5.11 |  |  | 253 | 10119.304 | 9.2222 | 5.111
254 | 10159.302 | 9.2261 | 5.1087 |  |  | 254 | 10159.302 | 9.2261 | 5.109 |  |  | 254 | 10159.302 | 9.2261 | 5.11
255 | 10199.299 | 9.2301 | 5.1078 |  |  | 255 | 10199.299 | 9.2301 | 5.108 |  |  | 255 | 10199.299 | 9.2301 | 5.109
256 | 10239.296 | 9.234 | 5.1068 |  |  | 256 | 10239.296 | 9.234 | 5.107 |  |  | 256 | 10239.296 | 9.234 | 5.108
257 | 10279.293 | 9.2379 | 5.1058 |  |  | 257 | 10279.293 | 9.2379 | 5.106 |  |  | 257 | 10279.293 | 9.2379 | 5.1071
258 | 10319.291 | 9.2418 | 5.1049 |  |  | 258 | 10319.291 | 9.2418 | 5.105 |  |  | 258 | 10319.291 | 9.2418 | 5.1061
259 | 10359.288 | 9.2456 | 5.1039 |  |  | 259 | 10359.288 | 9.2456 | 5.104 |  |  | 259 | 10359.288 | 9.2456 | 5.1051
260 | 10399.285 | 9.2495 | 5.103 |  |  | 260 | 10399.285 | 9.2495 | 5.1031 |  |  | 260 | 10399.285 | 9.2495 | 5.1042
261 | 10439.282 | 9.2533 | 5.102 |  |  | 261 | 10439.282 | 9.2533 | 5.1021 |  |  | 261 | 10439.282 | 9.2533 | 5.1032
262 | 10479.28 | 9.2572 | 5.1011 |  |  | 262 | 10479.28 | 9.2572 | 5.1012 |  |  | 262 | 10479.28 | 9.2572 | 5.1023
263 | 10519.277 | 9.261 | 5.1002 |  |  | 263 | 10519.277 | 9.261 | 5.1002 |  |  | 263 | 10519.277 | 9.261 | 5.1013
264 | 10559.274 | 9.2648 | 5.0993 |  |  | 264 | 10559.274 | 9.2648 | 5.0993 |  |  | 264 | 10559.274 | 9.2648 | 5.1004
265 | 10599.271 | 9.2685 | 5.0984 |  |  | 265 | 10599.271 | 9.2685 | 5.0984 |  |  | 265 | 10599.271 | 9.2685 | 5.0995
266 | 10639.269 | 9.2723 | 5.0975 |  |  | 266 | 10639.269 | 9.2723 | 5.0974 |  |  | 266 | 10639.269 | 9.2723 | 5.0986
267 | 10679.266 | 9.2761 | 5.0966 |  |  | 267 | 10679.266 | 9.2761 | 5.0965 |  |  | 267 | 10679.266 | 9.2761 | 5.0977
268 | 10719.263 | 9.2798 | 5.0957 |  |  | 268 | 10719.263 | 9.2798 | 5.0956 |  |  | 268 | 10719.263 | 9.2798 | 5.0968
269 | 10759.26 | 9.2835 | 5.0948 |  |  | 269 | 10759.26 | 9.2835 | 5.0947 |  |  | 269 | 10759.26 | 9.2835 | 5.0959
270 | 10799.258 | 9.2872 | 5.094 |  |  | 270 | 10799.258 | 9.2872 | 5.0938 |  |  | 270 | 10799.258 | 9.2872 | 5.095
271 | 10839.255 | 9.2909 | 5.0931 |  |  | 271 | 10839.255 | 9.2909 | 5.0929 |  |  | 271 | 10839.255 | 9.2909 | 5.0941
272 | 10879.252 | 9.2946 | 5.0923 |  |  | 272 | 10879.252 | 9.2946 | 5.0921 |  |  | 272 | 10879.252 | 9.2946 | 5.0933
273 | 10919.249 | 9.2983 | 5.0914 |  |  | 273 | 10919.249 | 9.2983 | 5.0912 |  |  | 273 | 10919.249 | 9.2983 | 5.0924
274 | 10959.247 | 9.3019 | 5.0906 |  |  | 274 | 10959.247 | 9.3019 | 5.0903 |  |  | 274 | 10959.247 | 9.3019 | 5.0916
275 | 10999.244 | 9.3056 | 5.0898 |  |  | 275 | 10999.244 | 9.3056 | 5.0895 |  |  | 275 | 10999.244 | 9.3056 | 5.0907
276 | 11039.241 | 9.3092 | 5.0889 |  |  | 276 | 11039.241 | 9.3092 | 5.0886 |  |  | 276 | 11039.241 | 9.3092 | 5.0899
277 | 11079.238 | 9.3128 | 5.0881 |  |  | 277 | 11079.238 | 9.3128 | 5.0878 |  |  | 277 | 11079.238 | 9.3128 | 5.089
278 | 11119.236 | 9.3164 | 5.0873 |  |  | 278 | 11119.236 | 9.3164 | 5.087 |  |  | 278 | 11119.236 | 9.3164 | 5.0882
279 | 11159.233 | 9.32 | 5.0865 |  |  | 279 | 11159.233 | 9.32 | 5.0861 |  |  | 279 | 11159.233 | 9.32 | 5.0874
280 | 11199.23 | 9.3236 | 5.0857 |  |  | 280 | 11199.23 | 9.3236 | 5.0853 |  |  | 280 | 11199.23 | 9.3236 | 5.0866
281 | 11239.227 | 9.3272 | 5.0849 |  |  | 281 | 11239.227 | 9.3272 | 5.0845 |  |  | 281 | 11239.227 | 9.3272 | 5.0858
282 | 11279.225 | 9.3307 | 5.0841 |  |  | 282 | 11279.225 | 9.3307 | 5.0837 |  |  | 282 | 11279.225 | 9.3307 | 5.085
283 | 11319.222 | 9.3343 | 5.0833 |  |  | 283 | 11319.222 | 9.3343 | 5.0829 |  |  | 283 | 11319.222 | 9.3343 | 5.0842
284 | 11359.219 | 9.3378 | 5.0826 |  |  | 284 | 11359.219 | 9.3378 | 5.0821 |  |  | 284 | 11359.219 | 9.3378 | 5.0834
285 | 11399.216 | 9.3413 | 5.0818 |  |  | 285 | 11399.216 | 9.3413 | 5.0813 |  |  | 285 | 11399.216 | 9.3413 | 5.0826
286 | 11439.214 | 9.3448 | 5.081 |  |  | 286 | 11439.214 | 9.3448 | 5.0805 |  |  | 286 | 11439.214 | 9.3448 | 5.0818
287 | 11479.211 | 9.3483 | 5.0803 |  |  | 287 | 11479.211 | 9.3483 | 5.0797 |  |  | 287 | 11479.211 | 9.3483 | 5.0811
288 | 11519.208 | 9.3518 | 5.0795 |  |  | 288 | 11519.208 | 9.3518 | 5.079 |  |  | 288 | 11519.208 | 9.3518 | 5.0803
289 | 11559.205 | 9.3552 | 5.0788 |  |  | 289 | 11559.205 | 9.3552 | 5.0782 |  |  | 289 | 11559.205 | 9.3552 | 5.0795
290 | 11599.203 | 9.3587 | 5.0781 |  |  | 290 | 11599.203 | 9.3587 | 5.0774 |  |  | 290 | 11599.203 | 9.3587 | 5.0788
291 | 11639.2 | 9.3621 | 5.0773 |  |  | 291 | 11639.2 | 9.3621 | 5.0767 |  |  | 291 | 11639.2 | 9.3621 | 5.078
292 | 11679.197 | 9.3656 | 5.0766 |  |  | 292 | 11679.197 | 9.3656 | 5.0759 |  |  | 292 | 11679.197 | 9.3656 | 5.0773
293 | 11719.194 | 9.369 | 5.0759 |  |  | 293 | 11719.194 | 9.369 | 5.0752 |  |  | 293 | 11719.194 | 9.369 | 5.0766
294 | 11759.192 | 9.3724 | 5.0752 |  |  | 294 | 11759.192 | 9.3724 | 5.0745 |  |  | 294 | 11759.192 | 9.3724 | 5.0758
295 | 11799.189 | 9.3758 | 5.0745 |  |  | 295 | 11799.189 | 9.3758 | 5.0737 |  |  | 295 | 11799.189 | 9.3758 | 5.0751
296 | 11839.186 | 9.3792 | 5.0738 |  |  | 296 | 11839.186 | 9.3792 | 5.073 |  |  | 296 | 11839.186 | 9.3792 | 5.0744
297 | 11879.183 | 9.3825 | 5.0731 |  |  | 297 | 11879.183 | 9.3825 | 5.0723 |  |  | 297 | 11879.183 | 9.3825 | 5.0737
298 | 11919.181 | 9.3859 | 5.0724 |  |  | 298 | 11919.181 | 9.3859 | 5.0716 |  |  | 298 | 11919.181 | 9.3859 | 5.073
299 | 11959.178 | 9.3893 | 5.0717 |  |  | 299 | 11959.178 | 9.3893 | 5.0709 |  |  | 299 | 11959.178 | 9.3893 | 5.0723
300 | 11999.175 | 9.3926 | 5.071 |  |  | 300 | 11999.175 | 9.3926 | 5.0702 |  |  | 300 | 11999.175 | 9.3926 | 5.0716
301 | 12039.172 | 9.3959 | 5.0703 |  |  | 301 | 12039.172 | 9.3959 | 5.0695 |  |  | 301 | 12039.172 | 9.3959 | 5.0709
302 | 12079.17 | 9.3992 | 5.0697 |  |  | 302 | 12079.17 | 9.3992 | 5.0688 |  |  | 302 | 12079.17 | 9.3992 | 5.0702
303 | 12119.167 | 9.4025 | 5.069 |  |  | 303 | 12119.167 | 9.4025 | 5.0681 |  |  | 303 | 12119.167 | 9.4025 | 5.0695
304 | 12159.164 | 9.4058 | 5.0684 |  |  | 304 | 12159.164 | 9.4058 | 5.0674 |  |  | 304 | 12159.164 | 9.4058 | 5.0689
305 | 12199.161 | 9.4091 | 5.0677 |  |  | 305 | 12199.161 | 9.4091 | 5.0667 |  |  | 305 | 12199.161 | 9.4091 | 5.0682
306 | 12239.159 | 9.4124 | 5.0671 |  |  | 306 | 12239.159 | 9.4124 | 5.0661 |  |  | 306 | 12239.159 | 9.4124 | 5.0675
307 | 12279.156 | 9.4157 | 5.0664 |  |  | 307 | 12279.156 | 9.4157 | 5.0654 |  |  | 307 | 12279.156 | 9.4157 | 5.0669
308 | 12319.153 | 9.4189 | 5.0658 |  |  | 308 | 12319.153 | 9.4189 | 5.0647 |  |  | 308 | 12319.153 | 9.4189 | 5.0662
309 | 12359.15 | 9.4222 | 5.0651 |  |  | 309 | 12359.15 | 9.4222 | 5.0641 |  |  | 309 | 12359.15 | 9.4222 | 5.0656
310 | 12399.148 | 9.4254 | 5.0645 |  |  | 310 | 12399.148 | 9.4254 | 5.0634 |  |  | 310 | 12399.148 | 9.4254 | 5.0649
311 | 12439.145 | 9.4286 | 5.0639 |  |  | 311 | 12439.145 | 9.4286 | 5.0628 |  |  | 311 | 12439.145 | 9.4286 | 5.0643
312 | 12479.142 | 9.4318 | 5.0633 |  |  | 312 | 12479.142 | 9.4318 | 5.0621 |  |  | 312 | 12479.142 | 9.4318 | 5.0637
313 | 12519.139 | 9.435 | 5.0627 |  |  | 313 | 12519.139 | 9.435 | 5.0615 |  |  | 313 | 12519.139 | 9.435 | 5.063
314 | 12559.137 | 9.4382 | 5.062 |  |  | 314 | 12559.137 | 9.4382 | 5.0609 |  |  | 314 | 12559.137 | 9.4382 | 5.0624
315 | 12599.134 | 9.4414 | 5.0614 |  |  | 315 | 12599.134 | 9.4414 | 5.0602 |  |  | 315 | 12599.134 | 9.4414 | 5.0618
316 | 12639.131 | 9.4446 | 5.0608 |  |  | 316 | 12639.131 | 9.4446 | 5.0596 |  |  | 316 | 12639.131 | 9.4446 | 5.0612
317 | 12679.128 | 9.4477 | 5.0603 |  |  | 317 | 12679.128 | 9.4477 | 5.059 |  |  | 317 | 12679.128 | 9.4477 | 5.0606
318 | 12719.126 | 9.4509 | 5.0597 |  |  | 318 | 12719.126 | 9.4509 | 5.0584 |  |  | 318 | 12719.126 | 9.4509 | 5.06
319 | 12759.123 | 9.454 | 5.0591 |  |  | 319 | 12759.123 | 9.454 | 5.0578 |  |  | 319 | 12759.123 | 9.454 | 5.0594
320 | 12799.12 | 9.4571 | 5.0585 |  |  | 320 | 12799.12 | 9.4571 | 5.0572 |  |  | 320 | 12799.12 | 9.4571 | 5.0588
321 | 12839.117 | 9.4603 | 5.0579 |  |  | 321 | 12839.117 | 9.4603 | 5.0566 |  |  | 321 | 12839.117 | 9.4603 | 5.0582
322 | 12879.115 | 9.4634 | 5.0573 |  |  | 322 | 12879.115 | 9.4634 | 5.056 |  |  | 322 | 12879.115 | 9.4634 | 5.0576
323 | 12919.112 | 9.4665 | 5.0568 |  |  | 323 | 12919.112 | 9.4665 | 5.0554 |  |  | 323 | 12919.112 | 9.4665 | 5.057
324 | 12959.109 | 9.4696 | 5.0562 |  |  | 324 | 12959.109 | 9.4696 | 5.0548 |  |  | 324 | 12959.109 | 9.4696 | 5.0564
325 | 12999.106 | 9.4726 | 5.0557 |  |  | 325 | 12999.106 | 9.4726 | 5.0542 |  |  | 325 | 12999.106 | 9.4726 | 5.0558
326 | 13039.104 | 9.4757 | 5.0551 |  |  | 326 | 13039.104 | 9.4757 | 5.0537 |  |  | 326 | 13039.104 | 9.4757 | 5.0553
327 | 13079.101 | 9.4788 | 5.0546 |  |  | 327 | 13079.101 | 9.4788 | 5.0531 |  |  | 327 | 13079.101 | 9.4788 | 5.0547
328 | 13119.098 | 9.4818 | 5.054 |  |  | 328 | 13119.098 | 9.4818 | 5.0525 |  |  | 328 | 13119.098 | 9.4818 | 5.0542
329 | 13159.095 | 9.4849 | 5.0535 |  |  | 329 | 13159.095 | 9.4849 | 5.052 |  |  | 329 | 13159.095 | 9.4849 | 5.0536
330 | 13199.093 | 9.4879 | 5.0529 |  |  | 330 | 13199.093 | 9.4879 | 5.0514 |  |  | 330 | 13199.093 | 9.4879 | 5.053
331 | 13239.09 | 9.4909 | 5.0524 |  |  | 331 | 13239.09 | 9.4909 | 5.0509 |  |  | 331 | 13239.09 | 9.4909 | 5.0525
332 | 13279.087 | 9.4939 | 5.0519 |  |  | 332 | 13279.087 | 9.4939 | 5.0503 |  |  | 332 | 13279.087 | 9.4939 | 5.0519
333 | 13319.084 | 9.497 | 5.0513 |  |  | 333 | 13319.084 | 9.497 | 5.0498 |  |  | 333 | 13319.084 | 9.497 | 5.0514
334 | 13359.082 | 9.5 | 5.0508 |  |  | 334 | 13359.082 | 9.5 | 5.0492 |  |  | 334 | 13359.082 | 9.5 | 5.0509
335 | 13399.079 | 9.5029 | 5.0503 |  |  | 335 | 13399.079 | 9.5029 | 5.0487 |  |  | 335 | 13399.079 | 9.5029 | 5.0503
336 | 13439.076 | 9.5059 | 5.0498 |  |  | 336 | 13439.076 | 9.5059 | 5.0481 |  |  | 336 | 13439.076 | 9.5059 | 5.0498
337 | 13479.073 | 9.5089 | 5.0493 |  |  | 337 | 13479.073 | 9.5089 | 5.0476 |  |  | 337 | 13479.073 | 9.5089 | 5.0493
338 | 13519.071 | 9.5119 | 5.0488 |  |  | 338 | 13519.071 | 9.5119 | 5.0471 |  |  | 338 | 13519.071 | 9.5119 | 5.0488
339 | 13559.068 | 9.5148 | 5.0483 |  |  | 339 | 13559.068 | 9.5148 | 5.0466 |  |  | 339 | 13559.068 | 9.5148 | 5.0482
340 | 13599.065 | 9.5178 | 5.0478 |  |  | 340 | 13599.065 | 9.5178 | 5.046 |  |  | 340 | 13599.065 | 9.5178 | 5.0477
341 | 13639.062 | 9.5207 | 5.0473 |  |  | 341 | 13639.062 | 9.5207 | 5.0455 |  |  | 341 | 13639.062 | 9.5207 | 5.0472
342 | 13679.06 | 9.5236 | 5.0468 |  |  | 342 | 13679.06 | 9.5236 | 5.045 |  |  | 342 | 13679.06 | 9.5236 | 5.0467
343 | 13719.057 | 9.5265 | 5.0463 |  |  | 343 | 13719.057 | 9.5265 | 5.0445 |  |  | 343 | 13719.057 | 9.5265 | 5.0462
344 | 13759.054 | 9.5295 | 5.0458 |  |  | 344 | 13759.054 | 9.5295 | 5.044 |  |  | 344 | 13759.054 | 9.5295 | 5.0457
345 | 13799.051 | 9.5324 | 5.0453 |  |  | 345 | 13799.051 | 9.5324 | 5.0435 |  |  | 345 | 13799.051 | 9.5324 | 5.0452
346 | 13839.049 | 9.5352 | 5.0449 |  |  | 346 | 13839.049 | 9.5352 | 5.043 |  |  | 346 | 13839.049 | 9.5352 | 5.0447
347 | 13879.046 | 9.5381 | 5.0444 |  |  | 347 | 13879.046 | 9.5381 | 5.0425 |  |  | 347 | 13879.046 | 9.5381 | 5.0442
348 | 13919.043 | 9.541 | 5.0439 |  |  | 348 | 13919.043 | 9.541 | 5.042 |  |  | 348 | 13919.043 | 9.541 | 5.0437
349 | 13959.04 | 9.5439 | 5.0434 |  |  | 349 | 13959.04 | 9.5439 | 5.0415 |  |  | 349 | 13959.04 | 9.5439 | 5.0433
350 | 13999.038 | 9.5467 | 5.043 |  |  | 350 | 13999.038 | 9.5467 | 5.041 |  |  | 350 | 13999.038 | 9.5467 | 5.0428
351 | 14039.035 | 9.5496 | 5.0425 |  |  | 351 | 14039.035 | 9.5496 | 5.0406 |  |  | 351 | 14039.035 | 9.5496 | 5.0423
352 | 14079.032 | 9.5524 | 5.0421 |  |  | 352 | 14079.032 | 9.5524 | 5.0401 |  |  | 352 | 14079.032 | 9.5524 | 5.0418
353 | 14119.029 | 9.5553 | 5.0416 |  |  | 353 | 14119.029 | 9.5553 | 5.0396 |  |  | 353 | 14119.029 | 9.5553 | 5.0414
354 | 14159.027 | 9.5581 | 5.0412 |  |  | 354 | 14159.027 | 9.5581 | 5.0392 |  |  | 354 | 14159.027 | 9.5581 | 5.0409
355 | 14199.024 | 9.5609 | 5.0407 |  |  | 355 | 14199.024 | 9.5609 | 5.0387 |  |  | 355 | 14199.024 | 9.5609 | 5.0404
356 | 14239.021 | 9.5637 | 5.0403 |  |  | 356 | 14239.021 | 9.5637 | 5.0382 |  |  | 356 | 14239.021 | 9.5637 | 5.04
357 | 14279.018 | 9.5665 | 5.0398 |  |  | 357 | 14279.018 | 9.5665 | 5.0378 |  |  | 357 | 14279.018 | 9.5665 | 5.0395
358 | 14319.016 | 9.5693 | 5.0394 |  |  | 358 | 14319.016 | 9.5693 | 5.0373 |  |  | 358 | 14319.016 | 9.5693 | 5.0391
359 | 14359.013 | 9.5721 | 5.039 |  |  | 359 | 14359.013 | 9.5721 | 5.0369 |  |  | 359 | 14359.013 | 9.5721 | 5.0386
360 | 14399.01 | 9.5749 | 5.0385 |  |  | 360 | 14399.01 | 9.5749 | 5.0364 |  |  | 360 | 14399.01 | 9.5749 | 5.0382
361 | 14439.007 | 9.5777 | 5.0381 |  |  | 361 | 14439.007 | 9.5777 | 5.036 |  |  | 361 | 14439.007 | 9.5777 | 5.0377
362 | 14479.005 | 9.5805 | 5.0377 |  |  | 362 | 14479.005 | 9.5805 | 5.0355 |  |  | 362 | 14479.005 | 9.5805 | 5.0373
363 | 14519.002 | 9.5832 | 5.0373 |  |  | 363 | 14519.002 | 9.5832 | 5.0351 |  |  | 363 | 14519.002 | 9.5832 | 5.0368
364 | 14558.999 | 9.586 | 5.0369 |  |  | 364 | 14558.999 | 9.586 | 5.0346 |  |  | 364 | 14558.999 | 9.586 | 5.0364
365 | 14598.996 | 9.5887 | 5.0364 |  |  | 365 | 14598.996 | 9.5887 | 5.0342 |  |  | 365 | 14598.996 | 9.5887 | 5.036
366 | 14638.994 | 9.5914 | 5.036 |  |  | 366 | 14638.994 | 9.5914 | 5.0338 |  |  | 366 | 14638.994 | 9.5914 | 5.0356
367 | 14678.991 | 9.5942 | 5.0356 |  |  | 367 | 14678.991 | 9.5942 | 5.0333 |  |  | 367 | 14678.991 | 9.5942 | 5.0351
368 | 14718.988 | 9.5969 | 5.0352 |  |  | 368 | 14718.988 | 9.5969 | 5.0329 |  |  | 368 | 14718.988 | 9.5969 | 5.0347
369 | 14758.985 | 9.5996 | 5.0348 |  |  | 369 | 14758.985 | 9.5996 | 5.0325 |  |  | 369 | 14758.985 | 9.5996 | 5.0343
370 | 14798.983 | 9.6023 | 5.0344 |  |  | 370 | 14798.983 | 9.6023 | 5.0321 |  |  | 370 | 14798.983 | 9.6023 | 5.0339
371 | 14838.98 | 9.605 | 5.034 |  |  | 371 | 14838.98 | 9.605 | 5.0317 |  |  | 371 | 14838.98 | 9.605 | 5.0335
372 | 14878.977 | 9.6077 | 5.0336 |  |  | 372 | 14878.977 | 9.6077 | 5.0312 |  |  | 372 | 14878.977 | 9.6077 | 5.033
373 | 14918.974 | 9.6104 | 5.0332 |  |  | 373 | 14918.974 | 9.6104 | 5.0308 |  |  | 373 | 14918.974 | 9.6104 | 5.0326
374 | 14958.972 | 9.6131 | 5.0328 |  |  | 374 | 14958.972 | 9.6131 | 5.0304 |  |  | 374 | 14958.972 | 9.6131 | 5.0322
375 | 14998.969 | 9.6157 | 5.0325 |  |  | 375 | 14998.969 | 9.6157 | 5.03 |  |  | 375 | 14998.969 | 9.6157 | 5.0318
376 | 15038.966 | 9.6184 | 5.0321 |  |  | 376 | 15038.966 | 9.6184 | 5.0296 |  |  | 376 | 15038.966 | 9.6184 | 5.0314
377 | 15078.963 | 9.6211 | 5.0317 |  |  | 377 | 15078.963 | 9.6211 | 5.0292 |  |  | 377 | 15078.963 | 9.6211 | 5.031
378 | 15118.961 | 9.6237 | 5.0313 |  |  | 378 | 15118.961 | 9.6237 | 5.0288 |  |  | 378 | 15118.961 | 9.6237 | 5.0306
379 | 15158.958 | 9.6263 | 5.0309 |  |  | 379 | 15158.958 | 9.6263 | 5.0284 |  |  | 379 | 15158.958 | 9.6263 | 5.0303
380 | 15198.955 | 9.629 | 5.0306 |  |  | 380 | 15198.955 | 9.629 | 5.028 |  |  | 380 | 15198.955 | 9.629 | 5.0299
381 | 15238.952 | 9.6316 | 5.0302 |  |  | 381 | 15238.952 | 9.6316 | 5.0277 |  |  | 381 | 15238.952 | 9.6316 | 5.0295
382 | 15278.95 | 9.6342 | 5.0298 |  |  | 382 | 15278.95 | 9.6342 | 5.0273 |  |  | 382 | 15278.95 | 9.6342 | 5.0291
383 | 15318.947 | 9.6368 | 5.0295 |  |  | 383 | 15318.947 | 9.6368 | 5.0269 |  |  | 383 | 15318.947 | 9.6368 | 5.0287
384 | 15358.944 | 9.6395 | 5.0291 |  |  | 384 | 15358.944 | 9.6395 | 5.0265 |  |  | 384 | 15358.944 | 9.6395 | 5.0283
385 | 15398.941 | 9.6421 | 5.0287 |  |  | 385 | 15398.941 | 9.6421 | 5.0261 |  |  | 385 | 15398.941 | 9.6421 | 5.028
386 | 15438.939 | 9.6446 | 5.0284 |  |  | 386 | 15438.939 | 9.6446 | 5.0257 |  |  | 386 | 15438.939 | 9.6446 | 5.0276
387 | 15478.936 | 9.6472 | 5.028 |  |  | 387 | 15478.936 | 9.6472 | 5.0254 |  |  | 387 | 15478.936 | 9.6472 | 5.0272
388 | 15518.933 | 9.6498 | 5.0277 |  |  | 388 | 15518.933 | 9.6498 | 5.025 |  |  | 388 | 15518.933 | 9.6498 | 5.0268
389 | 15558.93 | 9.6524 | 5.0273 |  |  | 389 | 15558.93 | 9.6524 | 5.0246 |  |  | 389 | 15558.93 | 9.6524 | 5.0265
390 | 15598.928 | 9.655 | 5.027 |  |  | 390 | 15598.928 | 9.655 | 5.0243 |  |  | 390 | 15598.928 | 9.655 | 5.0261
391 | 15638.925 | 9.6575 | 5.0266 |  |  | 391 | 15638.925 | 9.6575 | 5.0239 |  |  | 391 | 15638.925 | 9.6575 | 5.0257
392 | 15678.922 | 9.6601 | 5.0263 |  |  | 392 | 15678.922 | 9.6601 | 5.0235 |  |  | 392 | 15678.922 | 9.6601 | 5.0254
393 | 15718.919 | 9.6626 | 5.026 |  |  | 393 | 15718.919 | 9.6626 | 5.0232 |  |  | 393 | 15718.919 | 9.6626 | 5.025
394 | 15758.917 | 9.6652 | 5.0256 |  |  | 394 | 15758.917 | 9.6652 | 5.0228 |  |  | 394 | 15758.917 | 9.6652 | 5.0247
395 | 15798.914 | 9.6677 | 5.0253 |  |  | 395 | 15798.914 | 9.6677 | 5.0225 |  |  | 395 | 15798.914 | 9.6677 | 5.0243
396 | 15838.911 | 9.6702 | 5.025 |  |  | 396 | 15838.911 | 9.6702 | 5.0221 |  |  | 396 | 15838.911 | 9.6702 | 5.024
397 | 15878.908 | 9.6727 | 5.0246 |  |  | 397 | 15878.908 | 9.6727 | 5.0218 |  |  | 397 | 15878.908 | 9.6727 | 5.0236
398 | 15918.906 | 9.6753 | 5.0243 |  |  | 398 | 15918.906 | 9.6753 | 5.0214 |  |  | 398 | 15918.906 | 9.6753 | 5.0233
399 | 15958.903 | 9.6778 | 5.024 |  |  | 399 | 15958.903 | 9.6778 | 5.0211 |  |  | 399 | 15958.903 | 9.6778 | 5.0229
400 | 15998.9 | 9.6803 | 5.0236 |  |  | 400 | 15998.9 | 9.6803 | 5.0207 |  |  | 400 | 15998.9 | 9.6803 | 5.0226
401 | 16038.897 | 9.6828 | 5.0233 |  |  | 401 | 16038.897 | 9.6828 | 5.0204 |  |  | 401 | 16038.897 | 9.6828 | 5.0223
402 | 16078.895 | 9.6853 | 5.023 |  |  | 402 | 16078.895 | 9.6853 | 5.0201 |  |  | 402 | 16078.895 | 9.6853 | 5.0219
403 | 16118.892 | 9.6877 | 5.0227 |  |  | 403 | 16118.892 | 9.6877 | 5.0197 |  |  | 403 | 16118.892 | 9.6877 | 5.0216
404 | 16158.889 | 9.6902 | 5.0224 |  |  | 404 | 16158.889 | 9.6902 | 5.0194 |  |  | 404 | 16158.889 | 9.6902 | 5.0213
405 | 16198.886 | 9.6927 | 5.0221 |  |  | 405 | 16198.886 | 9.6927 | 5.0191 |  |  | 405 | 16198.886 | 9.6927 | 5.0209
406 | 16238.884 | 9.6952 | 5.0218 |  |  | 406 | 16238.884 | 9.6952 | 5.0187 |  |  | 406 | 16238.884 | 9.6952 | 5.0206
407 | 16278.881 | 9.6976 | 5.0214 |  |  | 407 | 16278.881 | 9.6976 | 5.0184 |  |  | 407 | 16278.881 | 9.6976 | 5.0203
408 | 16318.878 | 9.7001 | 5.0211 |  |  | 408 | 16318.878 | 9.7001 | 5.0181 |  |  | 408 | 16318.878 | 9.7001 | 5.0199
409 | 16358.875 | 9.7025 | 5.0208 |  |  | 409 | 16358.875 | 9.7025 | 5.0178 |  |  | 409 | 16358.875 | 9.7025 | 5.0196
410 | 16398.873 | 9.705 | 5.0205 |  |  | 410 | 16398.873 | 9.705 | 5.0174 |  |  | 410 | 16398.873 | 9.705 | 5.0193
411 | 16438.87 | 9.7074 | 5.0202 |  |  | 411 | 16438.87 | 9.7074 | 5.0171 |  |  | 411 | 16438.87 | 9.7074 | 5.019
412 | 16478.867 | 9.7098 | 5.0199 |  |  | 412 | 16478.867 | 9.7098 | 5.0168 |  |  | 412 | 16478.867 | 9.7098 | 5.0187
413 | 16518.864 | 9.7123 | 5.0196 |  |  | 413 | 16518.864 | 9.7123 | 5.0165 |  |  | 413 | 16518.864 | 9.7123 | 5.0184
414 | 16558.862 | 9.7147 | 5.0193 |  |  | 414 | 16558.862 | 9.7147 | 5.0162 |  |  | 414 | 16558.862 | 9.7147 | 5.018
415 | 16598.859 | 9.7171 | 5.019 |  |  | 415 | 16598.859 | 9.7171 | 5.0159 |  |  | 415 | 16598.859 | 9.7171 | 5.0177
416 | 16638.856 | 9.7195 | 5.0188 |  |  | 416 | 16638.856 | 9.7195 | 5.0156 |  |  | 416 | 16638.856 | 9.7195 | 5.0174
417 | 16678.853 | 9.7219 | 5.0185 |  |  | 417 | 16678.853 | 9.7219 | 5.0153 |  |  | 417 | 16678.853 | 9.7219 | 5.0171
418 | 16718.851 | 9.7243 | 5.0182 |  |  | 418 | 16718.851 | 9.7243 | 5.0149 |  |  | 418 | 16718.851 | 9.7243 | 5.0168
419 | 16758.848 | 9.7267 | 5.0179 |  |  | 419 | 16758.848 | 9.7267 | 5.0146 |  |  | 419 | 16758.848 | 9.7267 | 5.0165
420 | 16798.845 | 9.7291 | 5.0176 |  |  | 420 | 16798.845 | 9.7291 | 5.0143 |  |  | 420 | 16798.845 | 9.7291 | 5.0162
421 | 16838.842 | 9.7314 | 5.0173 |  |  | 421 | 16838.842 | 9.7314 | 5.014 |  |  | 421 | 16838.842 | 9.7314 | 5.0159
422 | 16878.84 | 9.7338 | 5.0171 |  |  | 422 | 16878.84 | 9.7338 | 5.0137 |  |  | 422 | 16878.84 | 9.7338 | 5.0156
423 | 16918.837 | 9.7362 | 5.0168 |  |  | 423 | 16918.837 | 9.7362 | 5.0135 |  |  | 423 | 16918.837 | 9.7362 | 5.0153
424 | 16958.834 | 9.7385 | 5.0165 |  |  | 424 | 16958.834 | 9.7385 | 5.0132 |  |  | 424 | 16958.834 | 9.7385 | 5.015
425 | 16998.831 | 9.7409 | 5.0162 |  |  | 425 | 16998.831 | 9.7409 | 5.0129 |  |  | 425 | 16998.831 | 9.7409 | 5.0147
426 | 17038.829 | 9.7433 | 5.016 |  |  | 426 | 17038.829 | 9.7433 | 5.0126 |  |  | 426 | 17038.829 | 9.7433 | 5.0144
427 | 17078.826 | 9.7456 | 5.0157 |  |  | 427 | 17078.826 | 9.7456 | 5.0123 |  |  | 427 | 17078.826 | 9.7456 | 5.0141
428 | 17118.823 | 9.7479 | 5.0154 |  |  | 428 | 17118.823 | 9.7479 | 5.012 |  |  | 428 | 17118.823 | 9.7479 | 5.0139
429 | 17158.82 | 9.7503 | 5.0151 |  |  | 429 | 17158.82 | 9.7503 | 5.0117 |  |  | 429 | 17158.82 | 9.7503 | 5.0136
430 | 17198.818 | 9.7526 | 5.0149 |  |  | 430 | 17198.818 | 9.7526 | 5.0114 |  |  | 430 | 17198.818 | 9.7526 | 5.0133
431 | 17238.815 | 9.7549 | 5.0146 |  |  | 431 | 17238.815 | 9.7549 | 5.0112 |  |  | 431 | 17238.815 | 9.7549 | 5.013
432 | 17278.812 | 9.7572 | 5.0144 |  |  | 432 | 17278.812 | 9.7572 | 5.0109 |  |  | 432 | 17278.812 | 9.7572 | 5.0127
433 | 17318.809 | 9.7595 | 5.0141 |  |  | 433 | 17318.809 | 9.7595 | 5.0106 |  |  | 433 | 17318.809 | 9.7595 | 5.0124
434 | 17358.807 | 9.7619 | 5.0138 |  |  | 434 | 17358.807 | 9.7619 | 5.0103 |  |  | 434 | 17358.807 | 9.7619 | 5.0122
435 | 17398.804 | 9.7642 | 5.0136 |  |  | 435 | 17398.804 | 9.7642 | 5.01 |  |  | 435 | 17398.804 | 9.7642 | 5.0119
436 | 17438.801 | 9.7665 | 5.0133 |  |  | 436 | 17438.801 | 9.7665 | 5.0098 |  |  | 436 | 17438.801 | 9.7665 | 5.0116
437 | 17478.798 | 9.7687 | 5.0131 |  |  | 437 | 17478.798 | 9.7687 | 5.0095 |  |  | 437 | 17478.798 | 9.7687 | 5.0113
438 | 17518.796 | 9.771 | 5.0128 |  |  | 438 | 17518.796 | 9.771 | 5.0092 |  |  | 438 | 17518.796 | 9.771 | 5.0111
439 | 17558.793 | 9.7733 | 5.0126 |  |  | 439 | 17558.793 | 9.7733 | 5.009 |  |  | 439 | 17558.793 | 9.7733 | 5.0108
440 | 17598.79 | 9.7756 | 5.0123 |  |  | 440 | 17598.79 | 9.7756 | 5.0087 |  |  | 440 | 17598.79 | 9.7756 | 5.0105
441 | 17638.787 | 9.7779 | 5.0121 |  |  | 441 | 17638.787 | 9.7779 | 5.0084 |  |  | 441 | 17638.787 | 9.7779 | 5.0103
442 | 17678.785 | 9.7801 | 5.0118 |  |  | 442 | 17678.785 | 9.7801 | 5.0082 |  |  | 442 | 17678.785 | 9.7801 | 5.01
443 | 17718.782 | 9.7824 | 5.0116 |  |  | 443 | 17718.782 | 9.7824 | 5.0079 |  |  | 443 | 17718.782 | 9.7824 | 5.0097
444 | 17758.779 | 9.7846 | 5.0113 |  |  | 444 | 17758.779 | 9.7846 | 5.0077 |  |  | 444 | 17758.779 | 9.7846 | 5.0095
445 | 17798.776 | 9.7869 | 5.0111 |  |  | 445 | 17798.776 | 9.7869 | 5.0074 |  |  | 445 | 17798.776 | 9.7869 | 5.0092
446 | 17838.774 | 9.7891 | 5.0109 |  |  | 446 | 17838.774 | 9.7891 | 5.0071 |  |  | 446 | 17838.774 | 9.7891 | 5.009
447 | 17878.771 | 9.7914 | 5.0106 |  |  | 447 | 17878.771 | 9.7914 | 5.0069 |  |  | 447 | 17878.771 | 9.7914 | 5.0087
448 | 17918.768 | 9.7936 | 5.0104 |  |  | 448 | 17918.768 | 9.7936 | 5.0066 |  |  | 448 | 17918.768 | 9.7936 | 5.0085
449 | 17958.765 | 9.7958 | 5.0102 |  |  | 449 | 17958.765 | 9.7958 | 5.0064 |  |  | 449 | 17958.765 | 9.7958 | 5.0082
450 | 17998.763 | 9.7981 | 5.0099 |  |  | 450 | 17998.763 | 9.7981 | 5.0061 |  |  | 450 | 17998.763 | 9.7981 | 5.0079
451 | 18038.76 | 9.8003 | 5.0097 |  |  | 451 | 18038.76 | 9.8003 | 5.0059 |  |  | 451 | 18038.76 | 9.8003 | 5.0077
452 | 18078.757 | 9.8025 | 5.0095 |  |  | 452 | 18078.757 | 9.8025 | 5.0056 |  |  | 452 | 18078.757 | 9.8025 | 5.0074
453 | 18118.755 | 9.8047 | 5.0092 |  |  | 453 | 18118.755 | 9.8047 | 5.0054 |  |  | 453 | 18118.755 | 9.8047 | 5.0072
454 | 18158.752 | 9.8069 | 5.009 |  |  | 454 | 18158.752 | 9.8069 | 5.0051 |  |  | 454 | 18158.752 | 9.8069 | 5.0069
455 | 18198.749 | 9.8091 | 5.0088 |  |  | 455 | 18198.749 | 9.8091 | 5.0049 |  |  | 455 | 18198.749 | 9.8091 | 5.0067
456 | 18238.746 | 9.8113 | 5.0086 |  |  | 456 | 18238.746 | 9.8113 | 5.0047 |  |  | 456 | 18238.746 | 9.8113 | 5.0065
457 | 18278.744 | 9.8135 | 5.0083 |  |  | 457 | 18278.744 | 9.8135 | 5.0044 |  |  | 457 | 18278.744 | 9.8135 | 5.0062
458 | 18318.741 | 9.8157 | 5.0081 |  |  | 458 | 18318.741 | 9.8157 | 5.0042 |  |  | 458 | 18318.741 | 9.8157 | 5.006
459 | 18358.738 | 9.8179 | 5.0079 |  |  | 459 | 18358.738 | 9.8179 | 5.0039 |  |  | 459 | 18358.738 | 9.8179 | 5.0057
460 | 18398.735 | 9.82 | 5.0077 |  |  | 460 | 18398.735 | 9.82 | 5.0037 |  |  | 460 | 18398.735 | 9.82 | 5.0055
461 | 18438.733 | 9.8222 | 5.0074 |  |  | 461 | 18438.733 | 9.8222 | 5.0035 |  |  | 461 | 18438.733 | 9.8222 | 5.0053
462 | 18478.73 | 9.8244 | 5.0072 |  |  | 462 | 18478.73 | 9.8244 | 5.0032 |  |  | 462 | 18478.73 | 9.8244 | 5.005
463 | 18518.727 | 9.8265 | 5.007 |  |  | 463 | 18518.727 | 9.8265 | 5.003 |  |  | 463 | 18518.727 | 9.8265 | 5.0048
464 | 18558.724 | 9.8287 | 5.0068 |  |  | 464 | 18558.724 | 9.8287 | 5.0028 |  |  | 464 | 18558.724 | 9.8287 | 5.0045
465 | 18598.722 | 9.8308 | 5.0066 |  |  | 465 | 18598.722 | 9.8308 | 5.0025 |  |  | 465 | 18598.722 | 9.8308 | 5.0043
466 | 18638.719 | 9.833 | 5.0064 |  |  | 466 | 18638.719 | 9.833 | 5.0023 |  |  | 466 | 18638.719 | 9.833 | 5.0041
467 | 18678.716 | 9.8351 | 5.0062 |  |  | 467 | 18678.716 | 9.8351 | 5.0021 |  |  | 467 | 18678.716 | 9.8351 | 5.0039
468 | 18718.713 | 9.8373 | 5.006 |  |  | 468 | 18718.713 | 9.8373 | 5.0018 |  |  | 468 | 18718.713 | 9.8373 | 5.0036
469 | 18758.711 | 9.8394 | 5.0057 |  |  | 469 | 18758.711 | 9.8394 | 5.0016 |  |  | 469 | 18758.711 | 9.8394 | 5.0034
470 | 18798.708 | 9.8415 | 5.0055 |  |  | 470 | 18798.708 | 9.8415 | 5.0014 |  |  | 470 | 18798.708 | 9.8415 | 5.0032
471 | 18838.705 | 9.8437 | 5.0053 |  |  | 471 | 18838.705 | 9.8437 | 5.0012 |  |  | 471 | 18838.705 | 9.8437 | 5.0029
472 | 18878.702 | 9.8458 | 5.0051 |  |  | 472 | 18878.702 | 9.8458 | 5.001 |  |  | 472 | 18878.702 | 9.8458 | 5.0027
473 | 18918.7 | 9.8479 | 5.0049 |  |  | 473 | 18918.7 | 9.8479 | 5.0007 |  |  | 473 | 18918.7 | 9.8479 | 5.0025
474 | 18958.697 | 9.85 | 5.0047 |  |  | 474 | 18958.697 | 9.85 | 5.0005 |  |  | 474 | 18958.697 | 9.85 | 5.0023
475 | 18998.694 | 9.8521 | 5.0045 |  |  | 475 | 18998.694 | 9.8521 | 5.0003 |  |  | 475 | 18998.694 | 9.8521 | 5.002
476 | 19038.691 | 9.8542 | 5.0043 |  |  | 476 | 19038.691 | 9.8542 | 5.0001 |  |  | 476 | 19038.691 | 9.8542 | 5.0018
477 | 19078.689 | 9.8563 | 5.0041 |  |  | 477 | 19078.689 | 9.8563 | 4.9999 |  |  | 477 | 19078.689 | 9.8563 | 5.0016
478 | 19118.686 | 9.8584 | 5.0039 |  |  | 478 | 19118.686 | 9.8584 | 4.9997 |  |  | 478 | 19118.686 | 9.8584 | 5.0014
479 | 19158.683 | 9.8605 | 5.0037 |  |  | 479 | 19158.683 | 9.8605 | 4.9994 |  |  | 479 | 19158.683 | 9.8605 | 5.0012
480 | 19198.68 | 9.8626 | 5.0035 |  |  | 480 | 19198.68 | 9.8626 | 4.9992 |  |  | 480 | 19198.68 | 9.8626 | 5.001
481 | 19238.678 | 9.8647 | 5.0033 |  |  | 481 | 19238.678 | 9.8647 | 4.999 |  |  | 481 | 19238.678 | 9.8647 | 5.0007
482 | 19278.675 | 9.8668 | 5.0031 |  |  | 482 | 19278.675 | 9.8668 | 4.9988 |  |  | 482 | 19278.675 | 9.8668 | 5.0005
483 | 19318.672 | 9.8688 | 5.0029 |  |  | 483 | 19318.672 | 9.8688 | 4.9986 |  |  | 483 | 19318.672 | 9.8688 | 5.0003
484 | 19358.669 | 9.8709 | 5.0028 |  |  | 484 | 19358.669 | 9.8709 | 4.9984 |  |  | 484 | 19358.669 | 9.8709 | 5.0001
485 | 19398.667 | 9.873 | 5.0026 |  |  | 485 | 19398.667 | 9.873 | 4.9982 |  |  | 485 | 19398.667 | 9.873 | 4.9999
486 | 19438.664 | 9.875 | 5.0024 |  |  | 486 | 19438.664 | 9.875 | 4.998 |  |  | 486 | 19438.664 | 9.875 | 4.9997
487 | 19478.661 | 9.8771 | 5.0022 |  |  | 487 | 19478.661 | 9.8771 | 4.9978 |  |  | 487 | 19478.661 | 9.8771 | 4.9995
488 | 19518.658 | 9.8791 | 5.002 |  |  | 488 | 19518.658 | 9.8791 | 4.9976 |  |  | 488 | 19518.658 | 9.8791 | 4.9993
489 | 19558.656 | 9.8812 | 5.0018 |  |  | 489 | 19558.656 | 9.8812 | 4.9974 |  |  | 489 | 19558.656 | 9.8812 | 4.9991
490 | 19598.653 | 9.8832 | 5.0016 |  |  | 490 | 19598.653 | 9.8832 | 4.9972 |  |  | 490 | 19598.653 | 9.8832 | 4.9989
491 | 19638.65 | 9.8853 | 5.0014 |  |  | 491 | 19638.65 | 9.8853 | 4.997 |  |  | 491 | 19638.65 | 9.8853 | 4.9986
492 | 19678.647 | 9.8873 | 5.0013 |  |  | 492 | 19678.647 | 9.8873 | 4.9968 |  |  | 492 | 19678.647 | 9.8873 | 4.9984
493 | 19718.645 | 9.8893 | 5.0011 |  |  | 493 | 19718.645 | 9.8893 | 4.9966 |  |  | 493 | 19718.645 | 9.8893 | 4.9982
494 | 19758.642 | 9.8913 | 5.0009 |  |  | 494 | 19758.642 | 9.8913 | 4.9964 |  |  | 494 | 19758.642 | 9.8913 | 4.998
495 | 19798.639 | 9.8934 | 5.0007 |  |  | 495 | 19798.639 | 9.8934 | 4.9962 |  |  | 495 | 19798.639 | 9.8934 | 4.9978
496 | 19838.636 | 9.8954 | 5.0005 |  |  | 496 | 19838.636 | 9.8954 | 4.996 |  |  | 496 | 19838.636 | 9.8954 | 4.9976
497 | 19878.634 | 9.8974 | 5.0004 |  |  | 497 | 19878.634 | 9.8974 | 4.9958 |  |  | 497 | 19878.634 | 9.8974 | 4.9974
498 | 19918.631 | 9.8994 | 5.0002 |  |  | 498 | 19918.631 | 9.8994 | 4.9956 |  |  | 498 | 19918.631 | 9.8994 | 4.9972
499 | 19958.628 | 9.9014 | 5 |  |  | 499 | 19958.628 | 9.9014 | 4.9954 |  |  | 499 | 19958.628 | 9.9014 | 4.997
500 | 19998.625 | 9.9034 | 4.9998 |  |  | 500 | 19998.625 | 9.9034 | 4.9952 |  |  | 500 | 19998.625 | 9.9034 | 4.9969
501 | 20038.623 | 9.9054 | 4.9997 |  |  | 501 | 20038.623 | 9.9054 | 4.995 |  |  | 501 | 20038.623 | 9.9054 | 4.9967
502 | 20078.62 | 9.9074 | 4.9995 |  |  | 502 | 20078.62 | 9.9074 | 4.9948 |  |  | 502 | 20078.62 | 9.9074 | 4.9965
503 | 20118.617 | 9.9094 | 4.9993 |  |  | 503 | 20118.617 | 9.9094 | 4.9946 |  |  | 503 | 20118.617 | 9.9094 | 4.9963
504 | 20158.614 | 9.9114 | 4.9991 |  |  | 504 | 20158.614 | 9.9114 | 4.9945 |  |  | 504 | 20158.614 | 9.9114 | 4.9961
505 | 20198.612 | 9.9134 | 4.999 |  |  | 505 | 20198.612 | 9.9134 | 4.9943 |  |  | 505 | 20198.612 | 9.9134 | 4.9959
506 | 20238.609 | 9.9153 | 4.9988 |  |  | 506 | 20238.609 | 9.9153 | 4.9941 |  |  | 506 | 20238.609 | 9.9153 | 4.9957
507 | 20278.606 | 9.9173 | 4.9986 |  |  | 507 | 20278.606 | 9.9173 | 4.9939 |  |  | 507 | 20278.606 | 9.9173 | 4.9955
508 | 20318.603 | 9.9193 | 4.9985 |  |  | 508 | 20318.603 | 9.9193 | 4.9937 |  |  | 508 | 20318.603 | 9.9193 | 4.9953
509 | 20358.601 | 9.9213 | 4.9983 |  |  | 509 | 20358.601 | 9.9213 | 4.9935 |  |  | 509 | 20358.601 | 9.9213 | 4.9951
510 | 20398.598 | 9.9232 | 4.9981 |  |  | 510 | 20398.598 | 9.9232 | 4.9933 |  |  | 510 | 20398.598 | 9.9232 | 4.9949
511 | 20438.595 | 9.9252 | 4.998 |  |  | 511 | 20438.595 | 9.9252 | 4.9932 |  |  | 511 | 20438.595 | 9.9252 | 4.9947
512 | 20478.592 | 9.9271 | 4.9978 |  |  | 512 | 20478.592 | 9.9271 | 4.993 |  |  | 512 | 20478.592 | 9.9271 | 4.9946
513 | 20518.59 | 9.9291 | 4.9976 |  |  | 513 | 20518.59 | 9.9291 | 4.9928 |  |  | 513 | 20518.59 | 9.9291 | 4.9944
514 | 20558.587 | 9.931 | 4.9975 |  |  | 514 | 20558.587 | 9.931 | 4.9926 |  |  | 514 | 20558.587 | 9.931 | 4.9942
515 | 20598.584 | 9.933 | 4.9973 |  |  | 515 | 20598.584 | 9.933 | 4.9924 |  |  | 515 | 20598.584 | 9.933 | 4.994
516 | 20638.581 | 9.9349 | 4.9972 |  |  | 516 | 20638.581 | 9.9349 | 4.9923 |  |  | 516 | 20638.581 | 9.9349 | 4.9938
517 | 20678.579 | 9.9369 | 4.997 |  |  | 517 | 20678.579 | 9.9369 | 4.9921 |  |  | 517 | 20678.579 | 9.9369 | 4.9936
518 | 20718.576 | 9.9388 | 4.9968 |  |  | 518 | 20718.576 | 9.9388 | 4.9919 |  |  | 518 | 20718.576 | 9.9388 | 4.9935
519 | 20758.573 | 9.9407 | 4.9967 |  |  | 519 | 20758.573 | 9.9407 | 4.9917 |  |  | 519 | 20758.573 | 9.9407 | 4.9933
520 | 20798.57 | 9.9426 | 4.9965 |  |  | 520 | 20798.57 | 9.9426 | 4.9916 |  |  | 520 | 20798.57 | 9.9426 | 4.9931
521 | 20838.568 | 9.9446 | 4.9964 |  |  | 521 | 20838.568 | 9.9446 | 4.9914 |  |  | 521 | 20838.568 | 9.9446 | 4.9929
522 | 20878.565 | 9.9465 | 4.9962 |  |  | 522 | 20878.565 | 9.9465 | 4.9912 |  |  | 522 | 20878.565 | 9.9465 | 4.9927
523 | 20918.562 | 9.9484 | 4.9961 |  |  | 523 | 20918.562 | 9.9484 | 4.9911 |  |  | 523 | 20918.562 | 9.9484 | 4.9926
524 | 20958.559 | 9.9503 | 4.9959 |  |  | 524 | 20958.559 | 9.9503 | 4.9909 |  |  | 524 | 20958.559 | 9.9503 | 4.9924
525 | 20998.557 | 9.9522 | 4.9957 |  |  | 525 | 20998.557 | 9.9522 | 4.9907 |  |  | 525 | 20998.557 | 9.9522 | 4.9922
526 | 21038.554 | 9.9541 | 4.9956 |  |  | 526 | 21038.554 | 9.9541 | 4.9905 |  |  | 526 | 21038.554 | 9.9541 | 4.992
527 | 21078.551 | 9.956 | 4.9954 |  |  | 527 | 21078.551 | 9.956 | 4.9904 |  |  | 527 | 21078.551 | 9.956 | 4.9919
528 | 21118.548 | 9.9579 | 4.9953 |  |  | 528 | 21118.548 | 9.9579 | 4.9902 |  |  | 528 | 21118.548 | 9.9579 | 4.9917
529 | 21158.546 | 9.9598 | 4.9951 |  |  | 529 | 21158.546 | 9.9598 | 4.99 |  |  | 529 | 21158.546 | 9.9598 | 4.9915
530 | 21198.543 | 9.9617 | 4.995 |  |  | 530 | 21198.543 | 9.9617 | 4.9899 |  |  | 530 | 21198.543 | 9.9617 | 4.9913
531 | 21238.54 | 9.9636 | 4.9948 |  |  | 531 | 21238.54 | 9.9636 | 4.9897 |  |  | 531 | 21238.54 | 9.9636 | 4.9912
532 | 21278.537 | 9.9655 | 4.9947 |  |  | 532 | 21278.537 | 9.9655 | 4.9895 |  |  | 532 | 21278.537 | 9.9655 | 4.991
533 | 21318.535 | 9.9673 | 4.9945 |  |  | 533 | 21318.535 | 9.9673 | 4.9894 |  |  | 533 | 21318.535 | 9.9673 | 4.9908
534 | 21358.532 | 9.9692 | 4.9944 |  |  | 534 | 21358.532 | 9.9692 | 4.9892 |  |  | 534 | 21358.532 | 9.9692 | 4.9907
535 | 21398.529 | 9.9711 | 4.9942 |  |  | 535 | 21398.529 | 9.9711 | 4.9891 |  |  | 535 | 21398.529 | 9.9711 | 4.9905
536 | 21438.526 | 9.9729 | 4.9941 |  |  | 536 | 21438.526 | 9.9729 | 4.9889 |  |  | 536 | 21438.526 | 9.9729 | 4.9903
537 | 21478.524 | 9.9748 | 4.994 |  |  | 537 | 21478.524 | 9.9748 | 4.9887 |  |  | 537 | 21478.524 | 9.9748 | 4.9902
538 | 21518.521 | 9.9767 | 4.9938 |  |  | 538 | 21518.521 | 9.9767 | 4.9886 |  |  | 538 | 21518.521 | 9.9767 | 4.99
539 | 21558.518 | 9.9785 | 4.9937 |  |  | 539 | 21558.518 | 9.9785 | 4.9884 |  |  | 539 | 21558.518 | 9.9785 | 4.9898
540 | 21598.515 | 9.9804 | 4.9935 |  |  | 540 | 21598.515 | 9.9804 | 4.9883 |  |  | 540 | 21598.515 | 9.9804 | 4.9897
541 | 21638.513 | 9.9822 | 4.9934 |  |  | 541 | 21638.513 | 9.9822 | 4.9881 |  |  | 541 | 21638.513 | 9.9822 | 4.9895
542 | 21678.51 | 9.9841 | 4.9932 |  |  | 542 | 21678.51 | 9.9841 | 4.9879 |  |  | 542 | 21678.51 | 9.9841 | 4.9893
543 | 21718.507 | 9.9859 | 4.9931 |  |  | 543 | 21718.507 | 9.9859 | 4.9878 |  |  | 543 | 21718.507 | 9.9859 | 4.9892
544 | 21758.504 | 9.9878 | 4.993 |  |  | 544 | 21758.504 | 9.9878 | 4.9876 |  |  | 544 | 21758.504 | 9.9878 | 4.989
545 | 21798.502 | 9.9896 | 4.9928 |  |  | 545 | 21798.502 | 9.9896 | 4.9875 |  |  | 545 | 21798.502 | 9.9896 | 4.9888
546 | 21838.499 | 9.9914 | 4.9927 |  |  | 546 | 21838.499 | 9.9914 | 4.9873 |  |  | 546 | 21838.499 | 9.9914 | 4.9887
547 | 21878.496 | 9.9933 | 4.9925 |  |  | 547 | 21878.496 | 9.9933 | 4.9872 |  |  | 547 | 21878.496 | 9.9933 | 4.9885
548 | 21918.493 | 9.9951 | 4.9924 |  |  | 548 | 21918.493 | 9.9951 | 4.987 |  |  | 548 | 21918.493 | 9.9951 | 4.9883
549 | 21958.491 | 9.9969 | 4.9923 |  |  | 549 | 21958.491 | 9.9969 | 4.9868 |  |  | 549 | 21958.491 | 9.9969 | 4.9882
550 | 21998.488 | 9.9987 | 4.9921 |  |  | 550 | 21998.488 | 9.9987 | 4.9867 |  |  | 550 | 21998.488 | 9.9987 | 4.988
551 | 22038.485 | 10.0005 | 4.992 |  |  | 551 | 22038.485 | 10.0005 | 4.9865 |  |  | 551 | 22038.485 | 10.0005 | 4.9879
552 | 22078.482 | 10.0024 | 4.9919 |  |  | 552 | 22078.482 | 10.0024 | 4.9864 |  |  | 552 | 22078.482 | 10.0024 | 4.9877
553 | 22118.48 | 10.0042 | 4.9917 |  |  | 553 | 22118.48 | 10.0042 | 4.9862 |  |  | 553 | 22118.48 | 10.0042 | 4.9876
554 | 22158.477 | 10.006 | 4.9916 |  |  | 554 | 22158.477 | 10.006 | 4.9861 |  |  | 554 | 22158.477 | 10.006 | 4.9874
555 | 22198.474 | 10.0078 | 4.9914 |  |  | 555 | 22198.474 | 10.0078 | 4.9859 |  |  | 555 | 22198.474 | 10.0078 | 4.9872
556 | 22238.471 | 10.0096 | 4.9913 |  |  | 556 | 22238.471 | 10.0096 | 4.9858 |  |  | 556 | 22238.471 | 10.0096 | 4.9871
557 | 22278.469 | 10.0114 | 4.9912 |  |  | 557 | 22278.469 | 10.0114 | 4.9856 |  |  | 557 | 22278.469 | 10.0114 | 4.9869
558 | 22318.466 | 10.0132 | 4.991 |  |  | 558 | 22318.466 | 10.0132 | 4.9855 |  |  | 558 | 22318.466 | 10.0132 | 4.9868
559 | 22358.463 | 10.015 | 4.9909 |  |  | 559 | 22358.463 | 10.015 | 4.9853 |  |  | 559 | 22358.463 | 10.015 | 4.9866
560 | 22398.46 | 10.0167 | 4.9908 |  |  | 560 | 22398.46 | 10.0167 | 4.9852 |  |  | 560 | 22398.46 | 10.0167 | 4.9865
561 | 22438.458 | 10.0185 | 4.9907 |  |  | 561 | 22438.458 | 10.0185 | 4.9851 |  |  | 561 | 22438.458 | 10.0185 | 4.9863
562 | 22478.455 | 10.0203 | 4.9905 |  |  | 562 | 22478.455 | 10.0203 | 4.9849 |  |  | 562 | 22478.455 | 10.0203 | 4.9861
563 | 22518.452 | 10.0221 | 4.9904 |  |  | 563 | 22518.452 | 10.0221 | 4.9848 |  |  | 563 | 22518.452 | 10.0221 | 4.986
564 | 22558.449 | 10.0239 | 4.9903 |  |  | 564 | 22558.449 | 10.0239 | 4.9846 |  |  | 564 | 22558.449 | 10.0239 | 4.9858
565 | 22598.447 | 10.0256 | 4.9901 |  |  | 565 | 22598.447 | 10.0256 | 4.9845 |  |  | 565 | 22598.447 | 10.0256 | 4.9857
566 | 22638.444 | 10.0274 | 4.99 |  |  | 566 | 22638.444 | 10.0274 | 4.9843 |  |  | 566 | 22638.444 | 10.0274 | 4.9855
567 | 22678.441 | 10.0292 | 4.9899 |  |  | 567 | 22678.441 | 10.0292 | 4.9842 |  |  | 567 | 22678.441 | 10.0292 | 4.9854
568 | 22718.438 | 10.0309 | 4.9898 |  |  | 568 | 22718.438 | 10.0309 | 4.984 |  |  | 568 | 22718.438 | 10.0309 | 4.9852
569 | 22758.436 | 10.0327 | 4.9896 |  |  | 569 | 22758.436 | 10.0327 | 4.9839 |  |  | 569 | 22758.436 | 10.0327 | 4.9851
570 | 22798.433 | 10.0344 | 4.9895 |  |  | 570 | 22798.433 | 10.0344 | 4.9838 |  |  | 570 | 22798.433 | 10.0344 | 4.9849
571 | 22838.43 | 10.0362 | 4.9894 |  |  | 571 | 22838.43 | 10.0362 | 4.9836 |  |  | 571 | 22838.43 | 10.0362 | 4.9848
572 | 22878.427 | 10.0379 | 4.9892 |  |  | 572 | 22878.427 | 10.0379 | 4.9835 |  |  | 572 | 22878.427 | 10.0379 | 4.9846
573 | 22918.425 | 10.0397 | 4.9891 |  |  | 573 | 22918.425 | 10.0397 | 4.9833 |  |  | 573 | 22918.425 | 10.0397 | 4.9845
574 | 22958.422 | 10.0414 | 4.989 |  |  | 574 | 22958.422 | 10.0414 | 4.9832 |  |  | 574 | 22958.422 | 10.0414 | 4.9843
575 | 22998.419 | 10.0432 | 4.9889 |  |  | 575 | 22998.419 | 10.0432 | 4.9831 |  |  | 575 | 22998.419 | 10.0432 | 4.9842
576 | 23038.416 | 10.0449 | 4.9888 |  |  | 576 | 23038.416 | 10.0449 | 4.9829 |  |  | 576 | 23038.416 | 10.0449 | 4.984
577 | 23078.414 | 10.0467 | 4.9886 |  |  | 577 | 23078.414 | 10.0467 | 4.9828 |  |  | 577 | 23078.414 | 10.0467 | 4.9839
578 | 23118.411 | 10.0484 | 4.9885 |  |  | 578 | 23118.411 | 10.0484 | 4.9826 |  |  | 578 | 23118.411 | 10.0484 | 4.9838
579 | 23158.408 | 10.0501 | 4.9884 |  |  | 579 | 23158.408 | 10.0501 | 4.9825 |  |  | 579 | 23158.408 | 10.0501 | 4.9836
580 | 23198.405 | 10.0518 | 4.9883 |  |  | 580 | 23198.405 | 10.0518 | 4.9824 |  |  | 580 | 23198.405 | 10.0518 | 4.9835
581 | 23238.403 | 10.0536 | 4.9881 |  |  | 581 | 23238.403 | 10.0536 | 4.9822 |  |  | 581 | 23238.403 | 10.0536 | 4.9833
582 | 23278.4 | 10.0553 | 4.988 |  |  | 582 | 23278.4 | 10.0553 | 4.9821 |  |  | 582 | 23278.4 | 10.0553 | 4.9832
583 | 23318.397 | 10.057 | 4.9879 |  |  | 583 | 23318.397 | 10.057 | 4.982 |  |  | 583 | 23318.397 | 10.057 | 4.983
584 | 23358.394 | 10.0587 | 4.9878 |  |  | 584 | 23358.394 | 10.0587 | 4.9818 |  |  | 584 | 23358.394 | 10.0587 | 4.9829
585 | 23398.392 | 10.0604 | 4.9877 |  |  | 585 | 23398.392 | 10.0604 | 4.9817 |  |  | 585 | 23398.392 | 10.0604 | 4.9827
586 | 23438.389 | 10.0621 | 4.9875 |  |  | 586 | 23438.389 | 10.0621 | 4.9816 |  |  | 586 | 23438.389 | 10.0621 | 4.9826
587 | 23478.386 | 10.0638 | 4.9874 |  |  | 587 | 23478.386 | 10.0638 | 4.9814 |  |  | 587 | 23478.386 | 10.0638 | 4.9825
588 | 23518.383 | 10.0655 | 4.9873 |  |  | 588 | 23518.383 | 10.0655 | 4.9813 |  |  | 588 | 23518.383 | 10.0655 | 4.9823
589 | 23558.381 | 10.0672 | 4.9872 |  |  | 589 | 23558.381 | 10.0672 | 4.9812 |  |  | 589 | 23558.381 | 10.0672 | 4.9822
590 | 23598.378 | 10.0689 | 4.9871 |  |  | 590 | 23598.378 | 10.0689 | 4.981 |  |  | 590 | 23598.378 | 10.0689 | 4.982
591 | 23638.375 | 10.0706 | 4.987 |  |  | 591 | 23638.375 | 10.0706 | 4.9809 |  |  | 591 | 23638.375 | 10.0706 | 4.9819
592 | 23678.372 | 10.0723 | 4.9868 |  |  | 592 | 23678.372 | 10.0723 | 4.9808 |  |  | 592 | 23678.372 | 10.0723 | 4.9817
593 | 23718.37 | 10.074 | 4.9867 |  |  | 593 | 23718.37 | 10.074 | 4.9806 |  |  | 593 | 23718.37 | 10.074 | 4.9816
594 | 23758.367 | 10.0757 | 4.9866 |  |  | 594 | 23758.367 | 10.0757 | 4.9805 |  |  | 594 | 23758.367 | 10.0757 | 4.9815
595 | 23798.364 | 10.0774 | 4.9865 |  |  | 595 | 23798.364 | 10.0774 | 4.9804 |  |  | 595 | 23798.364 | 10.0774 | 4.9813
596 | 23838.361 | 10.0791 | 4.9864 |  |  | 596 | 23838.361 | 10.0791 | 4.9802 |  |  | 596 | 23838.361 | 10.0791 | 4.9812
597 | 23878.359 | 10.0807 | 4.9863 |  |  | 597 | 23878.359 | 10.0807 | 4.9801 |  |  | 597 | 23878.359 | 10.0807 | 4.9811
598 | 23918.356 | 10.0824 | 4.9862 |  |  | 598 | 23918.356 | 10.0824 | 4.98 |  |  | 598 | 23918.356 | 10.0824 | 4.9809
599 | 23958.353 | 10.0841 | 4.986 |  |  | 599 | 23958.353 | 10.0841 | 4.9799 |  |  | 599 | 23958.353 | 10.0841 | 4.9808
600 | 23998.35 | 10.0857 | 4.9859 |  |  | 600 | 23998.35 | 10.0857 | 4.9797 |  |  | 600 | 23998.35 | 10.0857 | 4.9806
601 | 24038.348 | 10.0874 | 4.9858 |  |  | 601 | 24038.348 | 10.0874 | 4.9796 |  |  | 601 | 24038.348 | 10.0874 | 4.9805
602 | 24078.345 | 10.0891 | 4.9857 |  |  | 602 | 24078.345 | 10.0891 | 4.9795 |  |  | 602 | 24078.345 | 10.0891 | 4.9804
603 | 24118.342 | 10.0907 | 4.9856 |  |  | 603 | 24118.342 | 10.0907 | 4.9793 |  |  | 603 | 24118.342 | 10.0907 | 4.9802
604 | 24158.339 | 10.0924 | 4.9855 |  |  | 604 | 24158.339 | 10.0924 | 4.9792 |  |  | 604 | 24158.339 | 10.0924 | 4.9801
605 | 24198.337 | 10.094 | 4.9854 |  |  | 605 | 24198.337 | 10.094 | 4.9791 |  |  | 605 | 24198.337 | 10.094 | 4.98
606 | 24238.334 | 10.0957 | 4.9853 |  |  | 606 | 24238.334 | 10.0957 | 4.979 |  |  | 606 | 24238.334 | 10.0957 | 4.9798
607 | 24278.331 | 10.0973 | 4.9851 |  |  | 607 | 24278.331 | 10.0973 | 4.9788 |  |  | 607 | 24278.331 | 10.0973 | 4.9797
608 | 24318.328 | 10.099 | 4.985 |  |  | 608 | 24318.328 | 10.099 | 4.9787 |  |  | 608 | 24318.328 | 10.099 | 4.9795
609 | 24358.326 | 10.1006 | 4.9849 |  |  | 609 | 24358.326 | 10.1006 | 4.9786 |  |  | 609 | 24358.326 | 10.1006 | 4.9794
610 | 24398.323 | 10.1023 | 4.9848 |  |  | 610 | 24398.323 | 10.1023 | 4.9785 |  |  | 610 | 24398.323 | 10.1023 | 4.9793
611 | 24438.32 | 10.1039 | 4.9847 |  |  | 611 | 24438.32 | 10.1039 | 4.9783 |  |  | 611 | 24438.32 | 10.1039 | 4.9791
612 | 24478.317 | 10.1055 | 4.9846 |  |  | 612 | 24478.317 | 10.1055 | 4.9782 |  |  | 612 | 24478.317 | 10.1055 | 4.979
613 | 24518.315 | 10.1072 | 4.9845 |  |  | 613 | 24518.315 | 10.1072 | 4.9781 |  |  | 613 | 24518.315 | 10.1072 | 4.9789
614 | 24558.312 | 10.1088 | 4.9844 |  |  | 614 | 24558.312 | 10.1088 | 4.978 |  |  | 614 | 24558.312 | 10.1088 | 4.9787
615 | 24598.309 | 10.1104 | 4.9843 |  |  | 615 | 24598.309 | 10.1104 | 4.9779 |  |  | 615 | 24598.309 | 10.1104 | 4.9786
616 | 24638.306 | 10.1121 | 4.9842 |  |  | 616 | 24638.306 | 10.1121 | 4.9777 |  |  | 616 | 24638.306 | 10.1121 | 4.9785
617 | 24678.304 | 10.1137 | 4.9841 |  |  | 617 | 24678.304 | 10.1137 | 4.9776 |  |  | 617 | 24678.304 | 10.1137 | 4.9783
618 | 24718.301 | 10.1153 | 4.984 |  |  | 618 | 24718.301 | 10.1153 | 4.9775 |  |  | 618 | 24718.301 | 10.1153 | 4.9782
619 | 24758.298 | 10.1169 | 4.9838 |  |  | 619 | 24758.298 | 10.1169 | 4.9774 |  |  | 619 | 24758.298 | 10.1169 | 4.9781
620 | 24798.295 | 10.1185 | 4.9837 |  |  | 620 | 24798.295 | 10.1185 | 4.9772 |  |  | 620 | 24798.295 | 10.1185 | 4.978
621 | 24838.293 | 10.1201 | 4.9836 |  |  | 621 | 24838.293 | 10.1201 | 4.9771 |  |  | 621 | 24838.293 | 10.1201 | 4.9778
622 | 24878.29 | 10.1218 | 4.9835 |  |  | 622 | 24878.29 | 10.1218 | 4.977 |  |  | 622 | 24878.29 | 10.1218 | 4.9777
623 | 24918.287 | 10.1234 | 4.9834 |  |  | 623 | 24918.287 | 10.1234 | 4.9769 |  |  | 623 | 24918.287 | 10.1234 | 4.9776
624 | 24958.284 | 10.125 | 4.9833 |  |  | 624 | 24958.284 | 10.125 | 4.9768 |  |  | 624 | 24958.284 | 10.125 | 4.9774
625 | 24998.282 | 10.1266 | 4.9832 |  |  | 625 | 24998.282 | 10.1266 | 4.9766 |  |  | 625 | 24998.282 | 10.1266 | 4.9773
626 | 25038.279 | 10.1282 | 4.9831 |  |  | 626 | 25038.279 | 10.1282 | 4.9765 |  |  | 626 | 25038.279 | 10.1282 | 4.9772
627 | 25078.276 | 10.1298 | 4.983 |  |  | 627 | 25078.276 | 10.1298 | 4.9764 |  |  | 627 | 25078.276 | 10.1298 | 4.977
628 | 25118.273 | 10.1314 | 4.9829 |  |  | 628 | 25118.273 | 10.1314 | 4.9763 |  |  | 628 | 25118.273 | 10.1314 | 4.9769
629 | 25158.271 | 10.1329 | 4.9828 |  |  | 629 | 25158.271 | 10.1329 | 4.9762 |  |  | 629 | 25158.271 | 10.1329 | 4.9768
630 | 25198.268 | 10.1345 | 4.9827 |  |  | 630 | 25198.268 | 10.1345 | 4.976 |  |  | 630 | 25198.268 | 10.1345 | 4.9767
631 | 25238.265 | 10.1361 | 4.9826 |  |  | 631 | 25238.265 | 10.1361 | 4.9759 |  |  | 631 | 25238.265 | 10.1361 | 4.9765
632 | 25278.262 | 10.1377 | 4.9825 |  |  | 632 | 25278.262 | 10.1377 | 4.9758 |  |  | 632 | 25278.262 | 10.1377 | 4.9764
633 | 25318.26 | 10.1393 | 4.9824 |  |  | 633 | 25318.26 | 10.1393 | 4.9757 |  |  | 633 | 25318.26 | 10.1393 | 4.9763
634 | 25358.257 | 10.1409 | 4.9823 |  |  | 634 | 25358.257 | 10.1409 | 4.9756 |  |  | 634 | 25358.257 | 10.1409 | 4.9761
635 | 25398.254 | 10.1424 | 4.9822 |  |  | 635 | 25398.254 | 10.1424 | 4.9755 |  |  | 635 | 25398.254 | 10.1424 | 4.976
636 | 25438.251 | 10.144 | 4.9821 |  |  | 636 | 25438.251 | 10.144 | 4.9753 |  |  | 636 | 25438.251 | 10.144 | 4.9759
637 | 25478.249 | 10.1456 | 4.982 |  |  | 637 | 25478.249 | 10.1456 | 4.9752 |  |  | 637 | 25478.249 | 10.1456 | 4.9758
638 | 25518.246 | 10.1471 | 4.9819 |  |  | 638 | 25518.246 | 10.1471 | 4.9751 |  |  | 638 | 25518.246 | 10.1471 | 4.9756
639 | 25558.243 | 10.1487 | 4.9818 |  |  | 639 | 25558.243 | 10.1487 | 4.975 |  |  | 639 | 25558.243 | 10.1487 | 4.9755
640 | 25598.24 | 10.1503 | 4.9817 |  |  | 640 | 25598.24 | 10.1503 | 4.9749 |  |  | 640 | 25598.24 | 10.1503 | 4.9754
641 | 25638.238 | 10.1518 | 4.9816 |  |  | 641 | 25638.238 | 10.1518 | 4.9748 |  |  | 641 | 25638.238 | 10.1518 | 4.9752
642 | 25678.235 | 10.1534 | 4.9815 |  |  | 642 | 25678.235 | 10.1534 | 4.9746 |  |  | 642 | 25678.235 | 10.1534 | 4.9751
643 | 25718.232 | 10.155 | 4.9814 |  |  | 643 | 25718.232 | 10.155 | 4.9745 |  |  | 643 | 25718.232 | 10.155 | 4.975
644 | 25758.229 | 10.1565 | 4.9813 |  |  | 644 | 25758.229 | 10.1565 | 4.9744 |  |  | 644 | 25758.229 | 10.1565 | 4.9749
645 | 25798.227 | 10.1581 | 4.9812 |  |  | 645 | 25798.227 | 10.1581 | 4.9743 |  |  | 645 | 25798.227 | 10.1581 | 4.9747
646 | 25838.224 | 10.1596 | 4.9811 |  |  | 646 | 25838.224 | 10.1596 | 4.9742 |  |  | 646 | 25838.224 | 10.1596 | 4.9746
647 | 25878.221 | 10.1612 | 4.981 |  |  | 647 | 25878.221 | 10.1612 | 4.9741 |  |  | 647 | 25878.221 | 10.1612 | 4.9745
648 | 25918.218 | 10.1627 | 4.9809 |  |  | 648 | 25918.218 | 10.1627 | 4.974 |  |  | 648 | 25918.218 | 10.1627 | 4.9744
649 | 25958.216 | 10.1642 | 4.9808 |  |  | 649 | 25958.216 | 10.1642 | 4.9739 |  |  | 649 | 25958.216 | 10.1642 | 4.9742
650 | 25998.213 | 10.1658 | 4.9807 |  |  | 650 | 25998.213 | 10.1658 | 4.9737 |  |  | 650 | 25998.213 | 10.1658 | 4.9741
651 | 26038.21 | 10.1673 | 4.9806 |  |  | 651 | 26038.21 | 10.1673 | 4.9736 |  |  | 651 | 26038.21 | 10.1673 | 4.974
652 | 26078.207 | 10.1689 | 4.9805 |  |  | 652 | 26078.207 | 10.1689 | 4.9735 |  |  | 652 | 26078.207 | 10.1689 | 4.9739
653 | 26118.205 | 10.1704 | 4.9804 |  |  | 653 | 26118.205 | 10.1704 | 4.9734 |  |  | 653 | 26118.205 | 10.1704 | 4.9737
654 | 26158.202 | 10.1719 | 4.9803 |  |  | 654 | 26158.202 | 10.1719 | 4.9733 |  |  | 654 | 26158.202 | 10.1719 | 4.9736
655 | 26198.199 | 10.1734 | 4.9802 |  |  | 655 | 26198.199 | 10.1734 | 4.9732 |  |  | 655 | 26198.199 | 10.1734 | 4.9735
656 | 26238.196 | 10.175 | 4.9801 |  |  | 656 | 26238.196 | 10.175 | 4.9731 |  |  | 656 | 26238.196 | 10.175 | 4.9734
657 | 26278.194 | 10.1765 | 4.98 |  |  | 657 | 26278.194 | 10.1765 | 4.973 |  |  | 657 | 26278.194 | 10.1765 | 4.9733
658 | 26318.191 | 10.178 | 4.9799 |  |  | 658 | 26318.191 | 10.178 | 4.9728 |  |  | 658 | 26318.191 | 10.178 | 4.9731
659 | 26358.188 | 10.1795 | 4.9798 |  |  | 659 | 26358.188 | 10.1795 | 4.9727 |  |  | 659 | 26358.188 | 10.1795 | 4.973
660 | 26398.185 | 10.1811 | 4.9797 |  |  | 660 | 26398.185 | 10.1811 | 4.9726 |  |  | 660 | 26398.185 | 10.1811 | 4.9729
661 | 26438.183 | 10.1826 | 4.9796 |  |  | 661 | 26438.183 | 10.1826 | 4.9725 |  |  | 661 | 26438.183 | 10.1826 | 4.9728
662 | 26478.18 | 10.1841 | 4.9795 |  |  | 662 | 26478.18 | 10.1841 | 4.9724 |  |  | 662 | 26478.18 | 10.1841 | 4.9726
663 | 26518.177 | 10.1856 | 4.9794 |  |  | 663 | 26518.177 | 10.1856 | 4.9723 |  |  | 663 | 26518.177 | 10.1856 | 4.9725
664 | 26558.174 | 10.1871 | 4.9793 |  |  | 664 | 26558.174 | 10.1871 | 4.9722 |  |  | 664 | 26558.174 | 10.1871 | 4.9724
665 | 26598.172 | 10.1886 | 4.9792 |  |  | 665 | 26598.172 | 10.1886 | 4.9721 |  |  | 665 | 26598.172 | 10.1886 | 4.9723
666 | 26638.169 | 10.1901 | 4.9791 |  |  | 666 | 26638.169 | 10.1901 | 4.972 |  |  | 666 | 26638.169 | 10.1901 | 4.9722
667 | 26678.166 | 10.1916 | 4.9791 |  |  | 667 | 26678.166 | 10.1916 | 4.9718 |  |  | 667 | 26678.166 | 10.1916 | 4.972
668 | 26718.163 | 10.1931 | 4.979 |  |  | 668 | 26718.163 | 10.1931 | 4.9717 |  |  | 668 | 26718.163 | 10.1931 | 4.9719
669 | 26758.161 | 10.1946 | 4.9789 |  |  | 669 | 26758.161 | 10.1946 | 4.9716 |  |  | 669 | 26758.161 | 10.1946 | 4.9718
670 | 26798.158 | 10.1961 | 4.9788 |  |  | 670 | 26798.158 | 10.1961 | 4.9715 |  |  | 670 | 26798.158 | 10.1961 | 4.9717
671 | 26838.155 | 10.1976 | 4.9787 |  |  | 671 | 26838.155 | 10.1976 | 4.9714 |  |  | 671 | 26838.155 | 10.1976 | 4.9716
672 | 26878.152 | 10.1991 | 4.9786 |  |  | 672 | 26878.152 | 10.1991 | 4.9713 |  |  | 672 | 26878.152 | 10.1991 | 4.9714
673 | 26918.15 | 10.2006 | 4.9785 |  |  | 673 | 26918.15 | 10.2006 | 4.9712 |  |  | 673 | 26918.15 | 10.2006 | 4.9713
674 | 26958.147 | 10.202 | 4.9784 |  |  | 674 | 26958.147 | 10.202 | 4.9711 |  |  | 674 | 26958.147 | 10.202 | 4.9712
675 | 26998.144 | 10.2035 | 4.9783 |  |  | 675 | 26998.144 | 10.2035 | 4.971 |  |  | 675 | 26998.144 | 10.2035 | 4.9711
676 | 27038.141 | 10.205 | 4.9782 |  |  | 676 | 27038.141 | 10.205 | 4.9709 |  |  | 676 | 27038.141 | 10.205 | 4.9709
677 | 27078.139 | 10.2065 | 4.9781 |  |  | 677 | 27078.139 | 10.2065 | 4.9708 |  |  | 677 | 27078.139 | 10.2065 | 4.9708
678 | 27118.136 | 10.208 | 4.978 |  |  | 678 | 27118.136 | 10.208 | 4.9707 |  |  | 678 | 27118.136 | 10.208 | 4.9707
679 | 27158.133 | 10.2094 | 4.9779 |  |  | 679 | 27158.133 | 10.2094 | 4.9705 |  |  | 679 | 27158.133 | 10.2094 | 4.9706
680 | 27198.13 | 10.2109 | 4.9778 |  |  | 680 | 27198.13 | 10.2109 | 4.9704 |  |  | 680 | 27198.13 | 10.2109 | 4.9705
681 | 27238.128 | 10.2124 | 4.9777 |  |  | 681 | 27238.128 | 10.2124 | 4.9703 |  |  | 681 | 27238.128 | 10.2124 | 4.9704
682 | 27278.125 | 10.2138 | 4.9777 |  |  | 682 | 27278.125 | 10.2138 | 4.9702 |  |  | 682 | 27278.125 | 10.2138 | 4.9702
683 | 27318.122 | 10.2153 | 4.9776 |  |  | 683 | 27318.122 | 10.2153 | 4.9701 |  |  | 683 | 27318.122 | 10.2153 | 4.9701
684 | 27358.119 | 10.2168 | 4.9775 |  |  | 684 | 27358.119 | 10.2168 | 4.97 |  |  | 684 | 27358.119 | 10.2168 | 4.97
685 | 27398.117 | 10.2182 | 4.9774 |  |  | 685 | 27398.117 | 10.2182 | 4.9699 |  |  | 685 | 27398.117 | 10.2182 | 4.9699
686 | 27438.114 | 10.2197 | 4.9773 |  |  | 686 | 27438.114 | 10.2197 | 4.9698 |  |  | 686 | 27438.114 | 10.2197 | 4.9698
687 | 27478.111 | 10.2211 | 4.9772 |  |  | 687 | 27478.111 | 10.2211 | 4.9697 |  |  | 687 | 27478.111 | 10.2211 | 4.9696
688 | 27518.108 | 10.2226 | 4.9771 |  |  | 688 | 27518.108 | 10.2226 | 4.9696 |  |  | 688 | 27518.108 | 10.2226 | 4.9695
689 | 27558.106 | 10.2241 | 4.977 |  |  | 689 | 27558.106 | 10.2241 | 4.9695 |  |  | 689 | 27558.106 | 10.2241 | 4.9694
690 | 27598.103 | 10.2255 | 4.9769 |  |  | 690 | 27598.103 | 10.2255 | 4.9694 |  |  | 690 | 27598.103 | 10.2255 | 4.9693
691 | 27638.1 | 10.227 | 4.9768 |  |  | 691 | 27638.1 | 10.227 | 4.9693 |  |  | 691 | 27638.1 | 10.227 | 4.9692
692 | 27678.097 | 10.2284 | 4.9767 |  |  | 692 | 27678.097 | 10.2284 | 4.9692 |  |  | 692 | 27678.097 | 10.2284 | 4.969
693 | 27718.095 | 10.2298 | 4.9767 |  |  | 693 | 27718.095 | 10.2298 | 4.9691 |  |  | 693 | 27718.095 | 10.2298 | 4.9689
694 | 27758.092 | 10.2313 | 4.9766 |  |  | 694 | 27758.092 | 10.2313 | 4.969 |  |  | 694 | 27758.092 | 10.2313 | 4.9688
695 | 27798.089 | 10.2327 | 4.9765 |  |  | 695 | 27798.089 | 10.2327 | 4.9689 |  |  | 695 | 27798.089 | 10.2327 | 4.9687
696 | 27838.086 | 10.2342 | 4.9764 |  |  | 696 | 27838.086 | 10.2342 | 4.9687 |  |  | 696 | 27838.086 | 10.2342 | 4.9686
697 | 27878.084 | 10.2356 | 4.9763 |  |  | 697 | 27878.084 | 10.2356 | 4.9686 |  |  | 697 | 27878.084 | 10.2356 | 4.9685
698 | 27918.081 | 10.237 | 4.9762 |  |  | 698 | 27918.081 | 10.237 | 4.9685 |  |  | 698 | 27918.081 | 10.237 | 4.9683
699 | 27958.078 | 10.2385 | 4.9761 |  |  | 699 | 27958.078 | 10.2385 | 4.9684 |  |  | 699 | 27958.078 | 10.2385 | 4.9682
700 | 27998.075 | 10.2399 | 4.976 |  |  | 700 | 27998.075 | 10.2399 | 4.9683 |  |  | 700 | 27998.075 | 10.2399 | 4.9681
701 | 28038.073 | 10.2413 | 4.9759 |  |  | 701 | 28038.073 | 10.2413 | 4.9682 |  |  | 701 | 28038.073 | 10.2413 | 4.968
702 | 28078.07 | 10.2427 | 4.9758 |  |  | 702 | 28078.07 | 10.2427 | 4.9681 |  |  | 702 | 28078.07 | 10.2427 | 4.9679
703 | 28118.067 | 10.2442 | 4.9758 |  |  | 703 | 28118.067 | 10.2442 | 4.968 |  |  | 703 | 28118.067 | 10.2442 | 4.9678
704 | 28158.064 | 10.2456 | 4.9757 |  |  | 704 | 28158.064 | 10.2456 | 4.9679 |  |  | 704 | 28158.064 | 10.2456 | 4.9676
705 | 28198.062 | 10.247 | 4.9756 |  |  | 705 | 28198.062 | 10.247 | 4.9678 |  |  | 705 | 28198.062 | 10.247 | 4.9675
706 | 28238.059 | 10.2484 | 4.9755 |  |  | 706 | 28238.059 | 10.2484 | 4.9677 |  |  | 706 | 28238.059 | 10.2484 | 4.9674
707 | 28278.056 | 10.2498 | 4.9754 |  |  | 707 | 28278.056 | 10.2498 | 4.9676 |  |  | 707 | 28278.056 | 10.2498 | 4.9673
708 | 28318.053 | 10.2513 | 4.9753 |  |  | 708 | 28318.053 | 10.2513 | 4.9675 |  |  | 708 | 28318.053 | 10.2513 | 4.9672
709 | 28358.051 | 10.2527 | 4.9752 |  |  | 709 | 28358.051 | 10.2527 | 4.9674 |  |  | 709 | 28358.051 | 10.2527 | 4.9671
710 | 28398.048 | 10.2541 | 4.9751 |  |  | 710 | 28398.048 | 10.2541 | 4.9673 |  |  | 710 | 28398.048 | 10.2541 | 4.9669
711 | 28438.045 | 10.2555 | 4.9751 |  |  | 711 | 28438.045 | 10.2555 | 4.9672 |  |  | 711 | 28438.045 | 10.2555 | 4.9668
712 | 28478.042 | 10.2569 | 4.975 |  |  | 712 | 28478.042 | 10.2569 | 4.9671 |  |  | 712 | 28478.042 | 10.2569 | 4.9667
713 | 28518.04 | 10.2583 | 4.9749 |  |  | 713 | 28518.04 | 10.2583 | 4.967 |  |  | 713 | 28518.04 | 10.2583 | 4.9666
714 | 28558.037 | 10.2597 | 4.9748 |  |  | 714 | 28558.037 | 10.2597 | 4.9669 |  |  | 714 | 28558.037 | 10.2597 | 4.9665
715 | 28598.034 | 10.2611 | 4.9747 |  |  | 715 | 28598.034 | 10.2611 | 4.9668 |  |  | 715 | 28598.034 | 10.2611 | 4.9664
716 | 28638.031 | 10.2625 | 4.9746 |  |  | 716 | 28638.031 | 10.2625 | 4.9667 |  |  | 716 | 28638.031 | 10.2625 | 4.9662
717 | 28678.029 | 10.2639 | 4.9745 |  |  | 717 | 28678.029 | 10.2639 | 4.9666 |  |  | 717 | 28678.029 | 10.2639 | 4.9661
718 | 28718.026 | 10.2653 | 4.9744 |  |  | 718 | 28718.026 | 10.2653 | 4.9665 |  |  | 718 | 28718.026 | 10.2653 | 4.966
719 | 28758.023 | 10.2667 | 4.9744 |  |  | 719 | 28758.023 | 10.2667 | 4.9664 |  |  | 719 | 28758.023 | 10.2667 | 4.9659
720 | 28798.02 | 10.2681 | 4.9743 |  |  | 720 | 28798.02 | 10.2681 | 4.9663 |  |  | 720 | 28798.02 | 10.2681 | 4.9658
721 | 28838.018 | 10.2694 | 4.9742 |  |  | 721 | 28838.018 | 10.2694 | 4.9662 |  |  | 721 | 28838.018 | 10.2694 | 4.9657
722 | 28878.015 | 10.2708 | 4.9741 |  |  | 722 | 28878.015 | 10.2708 | 4.9661 |  |  | 722 | 28878.015 | 10.2708 | 4.9656
723 | 28918.012 | 10.2722 | 4.974 |  |  | 723 | 28918.012 | 10.2722 | 4.966 |  |  | 723 | 28918.012 | 10.2722 | 4.9654
724 | 28958.009 | 10.2736 | 4.9739 |  |  | 724 | 28958.009 | 10.2736 | 4.9659 |  |  | 724 | 28958.009 | 10.2736 | 4.9653
725 | 28998.007 | 10.275 | 4.9738 |  |  | 725 | 28998.007 | 10.275 | 4.9658 |  |  | 725 | 28998.007 | 10.275 | 4.9652
726 | 29038.004 | 10.2764 | 4.9737 |  |  | 726 | 29038.004 | 10.2764 | 4.9657 |  |  | 726 | 29038.004 | 10.2764 | 4.9651
727 | 29078.001 | 10.2777 | 4.9737 |  |  | 727 | 29078.001 | 10.2777 | 4.9656 |  |  | 727 | 29078.001 | 10.2777 | 4.965
728 | 29117.998 | 10.2791 | 4.9736 |  |  | 728 | 29117.998 | 10.2791 | 4.9655 |  |  | 728 | 29117.998 | 10.2791 | 4.9649
729 | 29157.996 | 10.2805 | 4.9735 |  |  | 729 | 29157.996 | 10.2805 | 4.9654 |  |  | 729 | 29157.996 | 10.2805 | 4.9647
730 | 29197.993 | 10.2819 | 4.9734 |  |  | 730 | 29197.993 | 10.2819 | 4.9653 |  |  | 730 | 29197.993 | 10.2819 | 4.9646
731 | 29237.99 | 10.2832 | 4.9733 |  |  | 731 | 29237.99 | 10.2832 | 4.9652 |  |  | 731 | 29237.99 | 10.2832 | 4.9645
732 | 29277.987 | 10.2846 | 4.9732 |  |  | 732 | 29277.987 | 10.2846 | 4.9651 |  |  | 732 | 29277.987 | 10.2846 | 4.9644
733 | 29317.985 | 10.286 | 4.9731 |  |  | 733 | 29317.985 | 10.286 | 4.965 |  |  | 733 | 29317.985 | 10.286 | 4.9643
734 | 29357.982 | 10.2873 | 4.9731 |  |  | 734 | 29357.982 | 10.2873 | 4.9649 |  |  | 734 | 29357.982 | 10.2873 | 4.9642
735 | 29397.979 | 10.2887 | 4.973 |  |  | 735 | 29397.979 | 10.2887 | 4.9648 |  |  | 735 | 29397.979 | 10.2887 | 4.9641
736 | 29437.976 | 10.29 | 4.9729 |  |  | 736 | 29437.976 | 10.29 | 4.9647 |  |  | 736 | 29437.976 | 10.29 | 4.9639
737 | 29477.974 | 10.2914 | 4.9728 |  |  | 737 | 29477.974 | 10.2914 | 4.9646 |  |  | 737 | 29477.974 | 10.2914 | 4.9638
738 | 29517.971 | 10.2928 | 4.9727 |  |  | 738 | 29517.971 | 10.2928 | 4.9645 |  |  | 738 | 29517.971 | 10.2928 | 4.9637
739 | 29557.968 | 10.2941 | 4.9726 |  |  | 739 | 29557.968 | 10.2941 | 4.9644 |  |  | 739 | 29557.968 | 10.2941 | 4.9636
740 | 29597.965 | 10.2955 | 4.9725 |  |  | 740 | 29597.965 | 10.2955 | 4.9643 |  |  | 740 | 29597.965 | 10.2955 | 4.9635
741 | 29637.963 | 10.2968 | 4.9725 |  |  | 741 | 29637.963 | 10.2968 | 4.9642 |  |  | 741 | 29637.963 | 10.2968 | 4.9634
742 | 29677.96 | 10.2982 | 4.9724 |  |  | 742 | 29677.96 | 10.2982 | 4.9641 |  |  | 742 | 29677.96 | 10.2982 | 4.9633
743 | 29717.957 | 10.2995 | 4.9723 |  |  | 743 | 29717.957 | 10.2995 | 4.964 |  |  | 743 | 29717.957 | 10.2995 | 4.9631
744 | 29757.954 | 10.3009 | 4.9722 |  |  | 744 | 29757.954 | 10.3009 | 4.9639 |  |  | 744 | 29757.954 | 10.3009 | 4.963
745 | 29797.952 | 10.3022 | 4.9721 |  |  | 745 | 29797.952 | 10.3022 | 4.9638 |  |  | 745 | 29797.952 | 10.3022 | 4.9629
746 | 29837.949 | 10.3035 | 4.972 |  |  | 746 | 29837.949 | 10.3035 | 4.9637 |  |  | 746 | 29837.949 | 10.3035 | 4.9628
747 | 29877.946 | 10.3049 | 4.9719 |  |  | 747 | 29877.946 | 10.3049 | 4.9636 |  |  | 747 | 29877.946 | 10.3049 | 4.9627
748 | 29917.943 | 10.3062 | 4.9719 |  |  | 748 | 29917.943 | 10.3062 | 4.9635 |  |  | 748 | 29917.943 | 10.3062 | 4.9626
749 | 29957.941 | 10.3075 | 4.9718 |  |  | 749 | 29957.941 | 10.3075 | 4.9634 |  |  | 749 | 29957.941 | 10.3075 | 4.9625
750 | 29997.938 | 10.3089 | 4.9717 |  |  | 750 | 29997.938 | 10.3089 | 4.9633 |  |  | 750 | 29997.938 | 10.3089 | 4.9623
751 | 30037.935 | 10.3102 | 4.9716 |  |  | 751 | 30037.935 | 10.3102 | 4.9632 |  |  | 751 | 30037.935 | 10.3102 | 4.9622
752 | 30077.932 | 10.3115 | 4.9715 |  |  | 752 | 30077.932 | 10.3115 | 4.9631 |  |  | 752 | 30077.932 | 10.3115 | 4.9621
753 | 30117.93 | 10.3129 | 4.9714 |  |  | 753 | 30117.93 | 10.3129 | 4.963 |  |  | 753 | 30117.93 | 10.3129 | 4.962
754 | 30157.927 | 10.3142 | 4.9714 |  |  | 754 | 30157.927 | 10.3142 | 4.9629 |  |  | 754 | 30157.927 | 10.3142 | 4.9619
755 | 30197.924 | 10.3155 | 4.9713 |  |  | 755 | 30197.924 | 10.3155 | 4.9628 |  |  | 755 | 30197.924 | 10.3155 | 4.9618
756 | 30237.921 | 10.3169 | 4.9712 |  |  | 756 | 30237.921 | 10.3169 | 4.9627 |  |  | 756 | 30237.921 | 10.3169 | 4.9617
757 | 30277.919 | 10.3182 | 4.9711 |  |  | 757 | 30277.919 | 10.3182 | 4.9626 |  |  | 757 | 30277.919 | 10.3182 | 4.9616
758 | 30317.916 | 10.3195 | 4.971 |  |  | 758 | 30317.916 | 10.3195 | 4.9625 |  |  | 758 | 30317.916 | 10.3195 | 4.9614
759 | 30357.913 | 10.3208 | 4.9709 |  |  | 759 | 30357.913 | 10.3208 | 4.9624 |  |  | 759 | 30357.913 | 10.3208 | 4.9613
760 | 30397.91 | 10.3221 | 4.9708 |  |  | 760 | 30397.91 | 10.3221 | 4.9623 |  |  | 760 | 30397.91 | 10.3221 | 4.9612
761 | 30437.908 | 10.3234 | 4.9708 |  |  | 761 | 30437.908 | 10.3234 | 4.9622 |  |  | 761 | 30437.908 | 10.3234 | 4.9611
762 | 30477.905 | 10.3248 | 4.9707 |  |  | 762 | 30477.905 | 10.3248 | 4.9621 |  |  | 762 | 30477.905 | 10.3248 | 4.961
763 | 30517.902 | 10.3261 | 4.9706 |  |  | 763 | 30517.902 | 10.3261 | 4.962 |  |  | 763 | 30517.902 | 10.3261 | 4.9609
764 | 30557.899 | 10.3274 | 4.9705 |  |  | 764 | 30557.899 | 10.3274 | 4.9619 |  |  | 764 | 30557.899 | 10.3274 | 4.9608
765 | 30597.897 | 10.3287 | 4.9704 |  |  | 765 | 30597.897 | 10.3287 | 4.9618 |  |  | 765 | 30597.897 | 10.3287 | 4.9606
766 | 30637.894 | 10.33 | 4.9703 |  |  | 766 | 30637.894 | 10.33 | 4.9617 |  |  | 766 | 30637.894 | 10.33 | 4.9605
767 | 30677.891 | 10.3313 | 4.9703 |  |  | 767 | 30677.891 | 10.3313 | 4.9616 |  |  | 767 | 30677.891 | 10.3313 | 4.9604
768 | 30717.888 | 10.3326 | 4.9702 |  |  | 768 | 30717.888 | 10.3326 | 4.9615 |  |  | 768 | 30717.888 | 10.3326 | 4.9603
769 | 30757.886 | 10.3339 | 4.9701 |  |  | 769 | 30757.886 | 10.3339 | 4.9614 |  |  | 769 | 30757.886 | 10.3339 | 4.9602
770 | 30797.883 | 10.3352 | 4.97 |  |  | 770 | 30797.883 | 10.3352 | 4.9613 |  |  | 770 | 30797.883 | 10.3352 | 4.9601
771 | 30837.88 | 10.3365 | 4.9699 |  |  | 771 | 30837.88 | 10.3365 | 4.9612 |  |  | 771 | 30837.88 | 10.3365 | 4.96
772 | 30877.877 | 10.3378 | 4.9698 |  |  | 772 | 30877.877 | 10.3378 | 4.9611 |  |  | 772 | 30877.877 | 10.3378 | 4.9599
773 | 30917.875 | 10.3391 | 4.9698 |  |  | 773 | 30917.875 | 10.3391 | 4.961 |  |  | 773 | 30917.875 | 10.3391 | 4.9597
774 | 30957.872 | 10.3404 | 4.9697 |  |  | 774 | 30957.872 | 10.3404 | 4.9609 |  |  | 774 | 30957.872 | 10.3404 | 4.9596
775 | 30997.869 | 10.3417 | 4.9696 |  |  | 775 | 30997.869 | 10.3417 | 4.9608 |  |  | 775 | 30997.869 | 10.3417 | 4.9595
776 | 31037.866 | 10.343 | 4.9695 |  |  | 776 | 31037.866 | 10.343 | 4.9607 |  |  | 776 | 31037.866 | 10.343 | 4.9594
777 | 31077.864 | 10.3443 | 4.9694 |  |  | 777 | 31077.864 | 10.3443 | 4.9606 |  |  | 777 | 31077.864 | 10.3443 | 4.9593
778 | 31117.861 | 10.3455 | 4.9693 |  |  | 778 | 31117.861 | 10.3455 | 4.9605 |  |  | 778 | 31117.861 | 10.3455 | 4.9592
779 | 31157.858 | 10.3468 | 4.9693 |  |  | 779 | 31157.858 | 10.3468 | 4.9604 |  |  | 779 | 31157.858 | 10.3468 | 4.9591
780 | 31197.855 | 10.3481 | 4.9692 |  |  | 780 | 31197.855 | 10.3481 | 4.9603 |  |  | 780 | 31197.855 | 10.3481 | 4.9589
781 | 31237.853 | 10.3494 | 4.9691 |  |  | 781 | 31237.853 | 10.3494 | 4.9602 |  |  | 781 | 31237.853 | 10.3494 | 4.9588
782 | 31277.85 | 10.3507 | 4.969 |  |  | 782 | 31277.85 | 10.3507 | 4.9601 |  |  | 782 | 31277.85 | 10.3507 | 4.9587
783 | 31317.847 | 10.3519 | 4.9689 |  |  | 783 | 31317.847 | 10.3519 | 4.96 |  |  | 783 | 31317.847 | 10.3519 | 4.9586
784 | 31357.844 | 10.3532 | 4.9688 |  |  | 784 | 31357.844 | 10.3532 | 4.9599 |  |  | 784 | 31357.844 | 10.3532 | 4.9585
785 | 31397.842 | 10.3545 | 4.9688 |  |  | 785 | 31397.842 | 10.3545 | 4.9598 |  |  | 785 | 31397.842 | 10.3545 | 4.9584
786 | 31437.839 | 10.3558 | 4.9687 |  |  | 786 | 31437.839 | 10.3558 | 4.9597 |  |  | 786 | 31437.839 | 10.3558 | 4.9583
787 | 31477.836 | 10.357 | 4.9686 |  |  | 787 | 31477.836 | 10.357 | 4.9596 |  |  | 787 | 31477.836 | 10.357 | 4.9582
788 | 31517.833 | 10.3583 | 4.9685 |  |  | 788 | 31517.833 | 10.3583 | 4.9595 |  |  | 788 | 31517.833 | 10.3583 | 4.958
789 | 31557.831 | 10.3596 | 4.9684 |  |  | 789 | 31557.831 | 10.3596 | 4.9594 |  |  | 789 | 31557.831 | 10.3596 | 4.9579
790 | 31597.828 | 10.3608 | 4.9683 |  |  | 790 | 31597.828 | 10.3608 | 4.9593 |  |  | 790 | 31597.828 | 10.3608 | 4.9578
791 | 31637.825 | 10.3621 | 4.9683 |  |  | 791 | 31637.825 | 10.3621 | 4.9592 |  |  | 791 | 31637.825 | 10.3621 | 4.9577
792 | 31677.822 | 10.3634 | 4.9682 |  |  | 792 | 31677.822 | 10.3634 | 4.9591 |  |  | 792 | 31677.822 | 10.3634 | 4.9576
793 | 31717.82 | 10.3646 | 4.9681 |  |  | 793 | 31717.82 | 10.3646 | 4.959 |  |  | 793 | 31717.82 | 10.3646 | 4.9575
794 | 31757.817 | 10.3659 | 4.968 |  |  | 794 | 31757.817 | 10.3659 | 4.9589 |  |  | 794 | 31757.817 | 10.3659 | 4.9574
795 | 31797.814 | 10.3672 | 4.9679 |  |  | 795 | 31797.814 | 10.3672 | 4.9588 |  |  | 795 | 31797.814 | 10.3672 | 4.9572
796 | 31837.811 | 10.3684 | 4.9678 |  |  | 796 | 31837.811 | 10.3684 | 4.9587 |  |  | 796 | 31837.811 | 10.3684 | 4.9571
797 | 31877.809 | 10.3697 | 4.9678 |  |  | 797 | 31877.809 | 10.3697 | 4.9586 |  |  | 797 | 31877.809 | 10.3697 | 4.957
798 | 31917.806 | 10.3709 | 4.9677 |  |  | 798 | 31917.806 | 10.3709 | 4.9585 |  |  | 798 | 31917.806 | 10.3709 | 4.9569
799 | 31957.803 | 10.3722 | 4.9676 |  |  | 799 | 31957.803 | 10.3722 | 4.9584 |  |  | 799 | 31957.803 | 10.3722 | 4.9568
800 | 31997.8 | 10.3734 | 4.9675 |  |  | 800 | 31997.8 | 10.3734 | 4.9583 |  |  | 800 | 31997.8 | 10.3734 | 4.9567
801 | 32037.798 | 10.3747 | 4.9674 |  |  | 801 | 32037.798 | 10.3747 | 4.9582 |  |  | 801 | 32037.798 | 10.3747 | 4.9566
802 | 32077.795 | 10.3759 | 4.9674 |  |  | 802 | 32077.795 | 10.3759 | 4.9581 |  |  | 802 | 32077.795 | 10.3759 | 4.9565
803 | 32117.792 | 10.3772 | 4.9673 |  |  | 803 | 32117.792 | 10.3772 | 4.958 |  |  | 803 | 32117.792 | 10.3772 | 4.9563
804 | 32157.789 | 10.3784 | 4.9672 |  |  | 804 | 32157.789 | 10.3784 | 4.9579 |  |  | 804 | 32157.789 | 10.3784 | 4.9562
805 | 32197.787 | 10.3797 | 4.9671 |  |  | 805 | 32197.787 | 10.3797 | 4.9578 |  |  | 805 | 32197.787 | 10.3797 | 4.9561
806 | 32237.784 | 10.3809 | 4.967 |  |  | 806 | 32237.784 | 10.3809 | 4.9577 |  |  | 806 | 32237.784 | 10.3809 | 4.956
807 | 32277.781 | 10.3821 | 4.9669 |  |  | 807 | 32277.781 | 10.3821 | 4.9576 |  |  | 807 | 32277.781 | 10.3821 | 4.9559
808 | 32317.778 | 10.3834 | 4.9669 |  |  | 808 | 32317.778 | 10.3834 | 4.9575 |  |  | 808 | 32317.778 | 10.3834 | 4.9558
809 | 32357.776 | 10.3846 | 4.9668 |  |  | 809 | 32357.776 | 10.3846 | 4.9575 |  |  | 809 | 32357.776 | 10.3846 | 4.9557
810 | 32397.773 | 10.3858 | 4.9667 |  |  | 810 | 32397.773 | 10.3858 | 4.9574 |  |  | 810 | 32397.773 | 10.3858 | 4.9556
811 | 32437.77 | 10.3871 | 4.9666 |  |  | 811 | 32437.77 | 10.3871 | 4.9573 |  |  | 811 | 32437.77 | 10.3871 | 4.9554
812 | 32477.767 | 10.3883 | 4.9665 |  |  | 812 | 32477.767 | 10.3883 | 4.9572 |  |  | 812 | 32477.767 | 10.3883 | 4.9553
813 | 32517.765 | 10.3895 | 4.9664 |  |  | 813 | 32517.765 | 10.3895 | 4.9571 |  |  | 813 | 32517.765 | 10.3895 | 4.9552
814 | 32557.762 | 10.3908 | 4.9664 |  |  | 814 | 32557.762 | 10.3908 | 4.957 |  |  | 814 | 32557.762 | 10.3908 | 4.9551
815 | 32597.759 | 10.392 | 4.9663 |  |  | 815 | 32597.759 | 10.392 | 4.9569 |  |  | 815 | 32597.759 | 10.392 | 4.955
816 | 32637.756 | 10.3932 | 4.9662 |  |  | 816 | 32637.756 | 10.3932 | 4.9568 |  |  | 816 | 32637.756 | 10.3932 | 4.9549
817 | 32677.754 | 10.3944 | 4.9661 |  |  | 817 | 32677.754 | 10.3944 | 4.9567 |  |  | 817 | 32677.754 | 10.3944 | 4.9548
818 | 32717.751 | 10.3957 | 4.966 |  |  | 818 | 32717.751 | 10.3957 | 4.9566 |  |  | 818 | 32717.751 | 10.3957 | 4.9546
819 | 32757.748 | 10.3969 | 4.9659 |  |  | 819 | 32757.748 | 10.3969 | 4.9565 |  |  | 819 | 32757.748 | 10.3969 | 4.9545
820 | 32797.745 | 10.3981 | 4.9659 |  |  | 820 | 32797.745 | 10.3981 | 4.9564 |  |  | 820 | 32797.745 | 10.3981 | 4.9544
821 | 32837.743 | 10.3993 | 4.9658 |  |  | 821 | 32837.743 | 10.3993 | 4.9563 |  |  | 821 | 32837.743 | 10.3993 | 4.9543
822 | 32877.74 | 10.4006 | 4.9657 |  |  | 822 | 32877.74 | 10.4006 | 4.9562 |  |  | 822 | 32877.74 | 10.4006 | 4.9542
823 | 32917.737 | 10.4018 | 4.9656 |  |  | 823 | 32917.737 | 10.4018 | 4.9561 |  |  | 823 | 32917.737 | 10.4018 | 4.9541
824 | 32957.734 | 10.403 | 4.9655 |  |  | 824 | 32957.734 | 10.403 | 4.956 |  |  | 824 | 32957.734 | 10.403 | 4.954
825 | 32997.732 | 10.4042 | 4.9655 |  |  | 825 | 32997.732 | 10.4042 | 4.9559 |  |  | 825 | 32997.732 | 10.4042 | 4.9539
826 | 33037.729 | 10.4054 | 4.9654 |  |  | 826 | 33037.729 | 10.4054 | 4.9558 |  |  | 826 | 33037.729 | 10.4054 | 4.9537
827 | 33077.726 | 10.4066 | 4.9653 |  |  | 827 | 33077.726 | 10.4066 | 4.9557 |  |  | 827 | 33077.726 | 10.4066 | 4.9536
828 | 33117.723 | 10.4078 | 4.9652 |  |  | 828 | 33117.723 | 10.4078 | 4.9556 |  |  | 828 | 33117.723 | 10.4078 | 4.9535
829 | 33157.721 | 10.409 | 4.9651 |  |  | 829 | 33157.721 | 10.409 | 4.9555 |  |  | 829 | 33157.721 | 10.409 | 4.9534
830 | 33197.718 | 10.4102 | 4.965 |  |  | 830 | 33197.718 | 10.4102 | 4.9554 |  |  | 830 | 33197.718 | 10.4102 | 4.9533
831 | 33237.715 | 10.4114 | 4.965 |  |  | 831 | 33237.715 | 10.4114 | 4.9553 |  |  | 831 | 33237.715 | 10.4114 | 4.9532
832 | 33277.712 | 10.4126 | 4.9649 |  |  | 832 | 33277.712 | 10.4126 | 4.9552 |  |  | 832 | 33277.712 | 10.4126 | 4.9531
833 | 33317.71 | 10.4138 | 4.9648 |  |  | 833 | 33317.71 | 10.4138 | 4.9551 |  |  | 833 | 33317.71 | 10.4138 | 4.9529
834 | 33357.707 | 10.415 | 4.9647 |  |  | 834 | 33357.707 | 10.415 | 4.955 |  |  | 834 | 33357.707 | 10.415 | 4.9528
835 | 33397.704 | 10.4162 | 4.9646 |  |  | 835 | 33397.704 | 10.4162 | 4.9549 |  |  | 835 | 33397.704 | 10.4162 | 4.9527
836 | 33437.701 | 10.4174 | 4.9645 |  |  | 836 | 33437.701 | 10.4174 | 4.9548 |  |  | 836 | 33437.701 | 10.4174 | 4.9526
837 | 33477.699 | 10.4186 | 4.9645 |  |  | 837 | 33477.699 | 10.4186 | 4.9547 |  |  | 837 | 33477.699 | 10.4186 | 4.9525
838 | 33517.696 | 10.4198 | 4.9644 |  |  | 838 | 33517.696 | 10.4198 | 4.9546 |  |  | 838 | 33517.696 | 10.4198 | 4.9524
839 | 33557.693 | 10.421 | 4.9643 |  |  | 839 | 33557.693 | 10.421 | 4.9545 |  |  | 839 | 33557.693 | 10.421 | 4.9523
840 | 33597.69 | 10.4222 | 4.9642 |  |  | 840 | 33597.69 | 10.4222 | 4.9544 |  |  | 840 | 33597.69 | 10.4222 | 4.9522
841 | 33637.688 | 10.4234 | 4.9641 |  |  | 841 | 33637.688 | 10.4234 | 4.9543 |  |  | 841 | 33637.688 | 10.4234 | 4.952
842 | 33677.685 | 10.4246 | 4.964 |  |  | 842 | 33677.685 | 10.4246 | 4.9542 |  |  | 842 | 33677.685 | 10.4246 | 4.9519
843 | 33717.682 | 10.4258 | 4.964 |  |  | 843 | 33717.682 | 10.4258 | 4.9541 |  |  | 843 | 33717.682 | 10.4258 | 4.9518
844 | 33757.679 | 10.427 | 4.9639 |  |  | 844 | 33757.679 | 10.427 | 4.954 |  |  | 844 | 33757.679 | 10.427 | 4.9517
845 | 33797.677 | 10.4281 | 4.9638 |  |  | 845 | 33797.677 | 10.4281 | 4.9539 |  |  | 845 | 33797.677 | 10.4281 | 4.9516
846 | 33837.674 | 10.4293 | 4.9637 |  |  | 846 | 33837.674 | 10.4293 | 4.9538 |  |  | 846 | 33837.674 | 10.4293 | 4.9515
847 | 33877.671 | 10.4305 | 4.9636 |  |  | 847 | 33877.671 | 10.4305 | 4.9537 |  |  | 847 | 33877.671 | 10.4305 | 4.9514
848 | 33917.668 | 10.4317 | 4.9636 |  |  | 848 | 33917.668 | 10.4317 | 4.9536 |  |  | 848 | 33917.668 | 10.4317 | 4.9512
849 | 33957.666 | 10.4329 | 4.9635 |  |  | 849 | 33957.666 | 10.4329 | 4.9535 |  |  | 849 | 33957.666 | 10.4329 | 4.9511
850 | 33997.663 | 10.434 | 4.9634 |  |  | 850 | 33997.663 | 10.434 | 4.9535 |  |  | 850 | 33997.663 | 10.434 | 4.951
851 | 34037.66 | 10.4352 | 4.9633 |  |  | 851 | 34037.66 | 10.4352 | 4.9534 |  |  | 851 | 34037.66 | 10.4352 | 4.9509
852 | 34077.657 | 10.4364 | 4.9632 |  |  | 852 | 34077.657 | 10.4364 | 4.9533 |  |  | 852 | 34077.657 | 10.4364 | 4.9508
853 | 34117.655 | 10.4376 | 4.9631 |  |  | 853 | 34117.655 | 10.4376 | 4.9532 |  |  | 853 | 34117.655 | 10.4376 | 4.9507
854 | 34157.652 | 10.4387 | 4.9631 |  |  | 854 | 34157.652 | 10.4387 | 4.9531 |  |  | 854 | 34157.652 | 10.4387 | 4.9506
855 | 34197.649 | 10.4399 | 4.963 |  |  | 855 | 34197.649 | 10.4399 | 4.953 |  |  | 855 | 34197.649 | 10.4399 | 4.9504
856 | 34237.646 | 10.4411 | 4.9629 |  |  | 856 | 34237.646 | 10.4411 | 4.9529 |  |  | 856 | 34237.646 | 10.4411 | 4.9503
857 | 34277.644 | 10.4422 | 4.9628 |  |  | 857 | 34277.644 | 10.4422 | 4.9528 |  |  | 857 | 34277.644 | 10.4422 | 4.9502
858 | 34317.641 | 10.4434 | 4.9627 |  |  | 858 | 34317.641 | 10.4434 | 4.9527 |  |  | 858 | 34317.641 | 10.4434 | 4.9501
859 | 34357.638 | 10.4446 | 4.9626 |  |  | 859 | 34357.638 | 10.4446 | 4.9526 |  |  | 859 | 34357.638 | 10.4446 | 4.95
860 | 34397.635 | 10.4457 | 4.9626 |  |  | 860 | 34397.635 | 10.4457 | 4.9525 |  |  | 860 | 34397.635 | 10.4457 | 4.9499
861 | 34437.633 | 10.4469 | 4.9625 |  |  | 861 | 34437.633 | 10.4469 | 4.9524 |  |  | 861 | 34437.633 | 10.4469 | 4.9498
862 | 34477.63 | 10.4481 | 4.9624 |  |  | 862 | 34477.63 | 10.4481 | 4.9523 |  |  | 862 | 34477.63 | 10.4481 | 4.9496
863 | 34517.627 | 10.4492 | 4.9623 |  |  | 863 | 34517.627 | 10.4492 | 4.9522 |  |  | 863 | 34517.627 | 10.4492 | 4.9495
864 | 34557.624 | 10.4504 | 4.9622 |  |  | 864 | 34557.624 | 10.4504 | 4.9521 |  |  | 864 | 34557.624 | 10.4504 | 4.9494
865 | 34597.622 | 10.4515 | 4.9621 |  |  | 865 | 34597.622 | 10.4515 | 4.952 |  |  | 865 | 34597.622 | 10.4515 | 4.9493
866 | 34637.619 | 10.4527 | 4.9621 |  |  | 866 | 34637.619 | 10.4527 | 4.9519 |  |  | 866 | 34637.619 | 10.4527 | 4.9492
867 | 34677.616 | 10.4538 | 4.962 |  |  | 867 | 34677.616 | 10.4538 | 4.9518 |  |  | 867 | 34677.616 | 10.4538 | 4.9491
868 | 34717.613 | 10.455 | 4.9619 |  |  | 868 | 34717.613 | 10.455 | 4.9517 |  |  | 868 | 34717.613 | 10.455 | 4.949
869 | 34757.611 | 10.4562 | 4.9618 |  |  | 869 | 34757.611 | 10.4562 | 4.9516 |  |  | 869 | 34757.611 | 10.4562 | 4.9488
870 | 34797.608 | 10.4573 | 4.9617 |  |  | 870 | 34797.608 | 10.4573 | 4.9515 |  |  | 870 | 34797.608 | 10.4573 | 4.9487
871 | 34837.605 | 10.4585 | 4.9617 |  |  | 871 | 34837.605 | 10.4585 | 4.9514 |  |  | 871 | 34837.605 | 10.4585 | 4.9486
872 | 34877.602 | 10.4596 | 4.9616 |  |  | 872 | 34877.602 | 10.4596 | 4.9513 |  |  | 872 | 34877.602 | 10.4596 | 4.9485
873 | 34917.6 | 10.4607 | 4.9615 |  |  | 873 | 34917.6 | 10.4607 | 4.9512 |  |  | 873 | 34917.6 | 10.4607 | 4.9484
874 | 34957.597 | 10.4619 | 4.9614 |  |  | 874 | 34957.597 | 10.4619 | 4.9511 |  |  | 874 | 34957.597 | 10.4619 | 4.9483
875 | 34997.594 | 10.463 | 4.9613 |  |  | 875 | 34997.594 | 10.463 | 4.951 |  |  | 875 | 34997.594 | 10.463 | 4.9482
876 | 35037.591 | 10.4642 | 4.9612 |  |  | 876 | 35037.591 | 10.4642 | 4.9509 |  |  | 876 | 35037.591 | 10.4642 | 4.948
877 | 35077.589 | 10.4653 | 4.9612 |  |  | 877 | 35077.589 | 10.4653 | 4.9508 |  |  | 877 | 35077.589 | 10.4653 | 4.9479
878 | 35117.586 | 10.4665 | 4.9611 |  |  | 878 | 35117.586 | 10.4665 | 4.9507 |  |  | 878 | 35117.586 | 10.4665 | 4.9478
879 | 35157.583 | 10.4676 | 4.961 |  |  | 879 | 35157.583 | 10.4676 | 4.9506 |  |  | 879 | 35157.583 | 10.4676 | 4.9477
880 | 35197.58 | 10.4687 | 4.9609 |  |  | 880 | 35197.58 | 10.4687 | 4.9505 |  |  | 880 | 35197.58 | 10.4687 | 4.9476
881 | 35237.578 | 10.4699 | 4.9608 |  |  | 881 | 35237.578 | 10.4699 | 4.9504 |  |  | 881 | 35237.578 | 10.4699 | 4.9475
882 | 35277.575 | 10.471 | 4.9607 |  |  | 882 | 35277.575 | 10.471 | 4.9503 |  |  | 882 | 35277.575 | 10.471 | 4.9473
883 | 35317.572 | 10.4721 | 4.9607 |  |  | 883 | 35317.572 | 10.4721 | 4.9502 |  |  | 883 | 35317.572 | 10.4721 | 4.9472
884 | 35357.569 | 10.4733 | 4.9606 |  |  | 884 | 35357.569 | 10.4733 | 4.9501 |  |  | 884 | 35357.569 | 10.4733 | 4.9471
885 | 35397.567 | 10.4744 | 4.9605 |  |  | 885 | 35397.567 | 10.4744 | 4.95 |  |  | 885 | 35397.567 | 10.4744 | 4.947
886 | 35437.564 | 10.4755 | 4.9604 |  |  | 886 | 35437.564 | 10.4755 | 4.9499 |  |  | 886 | 35437.564 | 10.4755 | 4.9469
887 | 35477.561 | 10.4767 | 4.9603 |  |  | 887 | 35477.561 | 10.4767 | 4.9498 |  |  | 887 | 35477.561 | 10.4767 | 4.9468
888 | 35517.558 | 10.4778 | 4.9602 |  |  | 888 | 35517.558 | 10.4778 | 4.9497 |  |  | 888 | 35517.558 | 10.4778 | 4.9467
889 | 35557.556 | 10.4789 | 4.9602 |  |  | 889 | 35557.556 | 10.4789 | 4.9496 |  |  | 889 | 35557.556 | 10.4789 | 4.9465
890 | 35597.553 | 10.48 | 4.9601 |  |  | 890 | 35597.553 | 10.48 | 4.9495 |  |  | 890 | 35597.553 | 10.48 | 4.9464
891 | 35637.55 | 10.4812 | 4.96 |  |  | 891 | 35637.55 | 10.4812 | 4.9494 |  |  | 891 | 35637.55 | 10.4812 | 4.9463
892 | 35677.547 | 10.4823 | 4.9599 |  |  | 892 | 35677.547 | 10.4823 | 4.9493 |  |  | 892 | 35677.547 | 10.4823 | 4.9462
893 | 35717.545 | 10.4834 | 4.9598 |  |  | 893 | 35717.545 | 10.4834 | 4.9492 |  |  | 893 | 35717.545 | 10.4834 | 4.9461
894 | 35757.542 | 10.4845 | 4.9597 |  |  | 894 | 35757.542 | 10.4845 | 4.9491 |  |  | 894 | 35757.542 | 10.4845 | 4.946
895 | 35797.539 | 10.4856 | 4.9597 |  |  | 895 | 35797.539 | 10.4856 | 4.949 |  |  | 895 | 35797.539 | 10.4856 | 4.9458
896 | 35837.536 | 10.4868 | 4.9596 |  |  | 896 | 35837.536 | 10.4868 | 4.9489 |  |  | 896 | 35837.536 | 10.4868 | 4.9457
897 | 35877.534 | 10.4879 | 4.9595 |  |  | 897 | 35877.534 | 10.4879 | 4.9488 |  |  | 897 | 35877.534 | 10.4879 | 4.9456
898 | 35917.531 | 10.489 | 4.9594 |  |  | 898 | 35917.531 | 10.489 | 4.9487 |  |  | 898 | 35917.531 | 10.489 | 4.9455
899 | 35957.528 | 10.4901 | 4.9593 |  |  | 899 | 35957.528 | 10.4901 | 4.9486 |  |  | 899 | 35957.528 | 10.4901 | 4.9454
900 | 35997.525 | 10.4912 | 4.9592 |  |  | 900 | 35997.525 | 10.4912 | 4.9485 |  |  | 900 | 35997.525 | 10.4912 | 4.9453
901 | 36037.523 | 10.4923 | 4.9592 |  |  | 901 | 36037.523 | 10.4923 | 4.9484 |  |  | 901 | 36037.523 | 10.4923 | 4.9452
902 | 36077.52 | 10.4934 | 4.9591 |  |  | 902 | 36077.52 | 10.4934 | 4.9483 |  |  | 902 | 36077.52 | 10.4934 | 4.945
903 | 36117.517 | 10.4945 | 4.959 |  |  | 903 | 36117.517 | 10.4945 | 4.9482 |  |  | 903 | 36117.517 | 10.4945 | 4.9449
904 | 36157.515 | 10.4956 | 4.9589 |  |  | 904 | 36157.515 | 10.4956 | 4.9481 |  |  | 904 | 36157.515 | 10.4956 | 4.9448
905 | 36197.512 | 10.4967 | 4.9588 |  |  | 905 | 36197.512 | 10.4967 | 4.948 |  |  | 905 | 36197.512 | 10.4967 | 4.9447
906 | 36237.509 | 10.4978 | 4.9587 |  |  | 906 | 36237.509 | 10.4978 | 4.948 |  |  | 906 | 36237.509 | 10.4978 | 4.9446
907 | 36277.506 | 10.499 | 4.9587 |  |  | 907 | 36277.506 | 10.499 | 4.9479 |  |  | 907 | 36277.506 | 10.499 | 4.9445
908 | 36317.504 | 10.5001 | 4.9586 |  |  | 908 | 36317.504 | 10.5001 | 4.9478 |  |  | 908 | 36317.504 | 10.5001 | 4.9443
909 | 36357.501 | 10.5012 | 4.9585 |  |  | 909 | 36357.501 | 10.5012 | 4.9477 |  |  | 909 | 36357.501 | 10.5012 | 4.9442
910 | 36397.498 | 10.5023 | 4.9584 |  |  | 910 | 36397.498 | 10.5023 | 4.9476 |  |  | 910 | 36397.498 | 10.5023 | 4.9441
911 | 36437.495 | 10.5034 | 4.9583 |  |  | 911 | 36437.495 | 10.5034 | 4.9475 |  |  | 911 | 36437.495 | 10.5034 | 4.944
912 | 36477.493 | 10.5045 | 4.9582 |  |  | 912 | 36477.493 | 10.5045 | 4.9474 |  |  | 912 | 36477.493 | 10.5045 | 4.9439
913 | 36517.49 | 10.5055 | 4.9581 |  |  | 913 | 36517.49 | 10.5055 | 4.9473 |  |  | 913 | 36517.49 | 10.5055 | 4.9438
914 | 36557.487 | 10.5066 | 4.9581 |  |  | 914 | 36557.487 | 10.5066 | 4.9472 |  |  | 914 | 36557.487 | 10.5066 | 4.9436
915 | 36597.484 | 10.5077 | 4.958 |  |  | 915 | 36597.484 | 10.5077 | 4.9471 |  |  | 915 | 36597.484 | 10.5077 | 4.9435
916 | 36637.482 | 10.5088 | 4.9579 |  |  | 916 | 36637.482 | 10.5088 | 4.947 |  |  | 916 | 36637.482 | 10.5088 | 4.9434
917 | 36677.479 | 10.5099 | 4.9578 |  |  | 917 | 36677.479 | 10.5099 | 4.9469 |  |  | 917 | 36677.479 | 10.5099 | 4.9433
918 | 36717.476 | 10.511 | 4.9577 |  |  | 918 | 36717.476 | 10.511 | 4.9468 |  |  | 918 | 36717.476 | 10.511 | 4.9432
919 | 36757.473 | 10.5121 | 4.9576 |  |  | 919 | 36757.473 | 10.5121 | 4.9467 |  |  | 919 | 36757.473 | 10.5121 | 4.9431
920 | 36797.471 | 10.5132 | 4.9576 |  |  | 920 | 36797.471 | 10.5132 | 4.9466 |  |  | 920 | 36797.471 | 10.5132 | 4.9429
921 | 36837.468 | 10.5143 | 4.9575 |  |  | 921 | 36837.468 | 10.5143 | 4.9465 |  |  | 921 | 36837.468 | 10.5143 | 4.9428
922 | 36877.465 | 10.5154 | 4.9574 |  |  | 922 | 36877.465 | 10.5154 | 4.9464 |  |  | 922 | 36877.465 | 10.5154 | 4.9427
923 | 36917.462 | 10.5164 | 4.9573 |  |  | 923 | 36917.462 | 10.5164 | 4.9463 |  |  | 923 | 36917.462 | 10.5164 | 4.9426
924 | 36957.46 | 10.5175 | 4.9572 |  |  | 924 | 36957.46 | 10.5175 | 4.9462 |  |  | 924 | 36957.46 | 10.5175 | 4.9425
925 | 36997.457 | 10.5186 | 4.9571 |  |  | 925 | 36997.457 | 10.5186 | 4.9461 |  |  | 925 | 36997.457 | 10.5186 | 4.9424
926 | 37037.454 | 10.5197 | 4.9571 |  |  | 926 | 37037.454 | 10.5197 | 4.946 |  |  | 926 | 37037.454 | 10.5197 | 4.9422
927 | 37077.451 | 10.5208 | 4.957 |  |  | 927 | 37077.451 | 10.5208 | 4.9459 |  |  | 927 | 37077.451 | 10.5208 | 4.9421
928 | 37117.449 | 10.5218 | 4.9569 |  |  | 928 | 37117.449 | 10.5218 | 4.9458 |  |  | 928 | 37117.449 | 10.5218 | 4.942
929 | 37157.446 | 10.5229 | 4.9568 |  |  | 929 | 37157.446 | 10.5229 | 4.9457 |  |  | 929 | 37157.446 | 10.5229 | 4.9419
930 | 37197.443 | 10.524 | 4.9567 |  |  | 930 | 37197.443 | 10.524 | 4.9456 |  |  | 930 | 37197.443 | 10.524 | 4.9418
931 | 37237.44 | 10.5251 | 4.9566 |  |  | 931 | 37237.44 | 10.5251 | 4.9455 |  |  | 931 | 37237.44 | 10.5251 | 4.9417
932 | 37277.438 | 10.5261 | 4.9566 |  |  | 932 | 37277.438 | 10.5261 | 4.9454 |  |  | 932 | 37277.438 | 10.5261 | 4.9415
933 | 37317.435 | 10.5272 | 4.9565 |  |  | 933 | 37317.435 | 10.5272 | 4.9453 |  |  | 933 | 37317.435 | 10.5272 | 4.9414
934 | 37357.432 | 10.5283 | 4.9564 |  |  | 934 | 37357.432 | 10.5283 | 4.9452 |  |  | 934 | 37357.432 | 10.5283 | 4.9413
935 | 37397.429 | 10.5294 | 4.9563 |  |  | 935 | 37397.429 | 10.5294 | 4.9451 |  |  | 935 | 37397.429 | 10.5294 | 4.9412
936 | 37437.427 | 10.5304 | 4.9562 |  |  | 936 | 37437.427 | 10.5304 | 4.945 |  |  | 936 | 37437.427 | 10.5304 | 4.9411
937 | 37477.424 | 10.5315 | 4.9561 |  |  | 937 | 37477.424 | 10.5315 | 4.9449 |  |  | 937 | 37477.424 | 10.5315 | 4.941
938 | 37517.421 | 10.5326 | 4.956 |  |  | 938 | 37517.421 | 10.5326 | 4.9448 |  |  | 938 | 37517.421 | 10.5326 | 4.9408
939 | 37557.418 | 10.5336 | 4.956 |  |  | 939 | 37557.418 | 10.5336 | 4.9447 |  |  | 939 | 37557.418 | 10.5336 | 4.9407
940 | 37597.416 | 10.5347 | 4.9559 |  |  | 940 | 37597.416 | 10.5347 | 4.9446 |  |  | 940 | 37597.416 | 10.5347 | 4.9406
941 | 37637.413 | 10.5358 | 4.9558 |  |  | 941 | 37637.413 | 10.5358 | 4.9445 |  |  | 941 | 37637.413 | 10.5358 | 4.9405
942 | 37677.41 | 10.5368 | 4.9557 |  |  | 942 | 37677.41 | 10.5368 | 4.9444 |  |  | 942 | 37677.41 | 10.5368 | 4.9404
943 | 37717.407 | 10.5379 | 4.9556 |  |  | 943 | 37717.407 | 10.5379 | 4.9443 |  |  | 943 | 37717.407 | 10.5379 | 4.9402
944 | 37757.405 | 10.5389 | 4.9555 |  |  | 944 | 37757.405 | 10.5389 | 4.9442 |  |  | 944 | 37757.405 | 10.5389 | 4.9401
945 | 37797.402 | 10.54 | 4.9554 |  |  | 945 | 37797.402 | 10.54 | 4.9441 |  |  | 945 | 37797.402 | 10.54 | 4.94
946 | 37837.399 | 10.5411 | 4.9554 |  |  | 946 | 37837.399 | 10.5411 | 4.944 |  |  | 946 | 37837.399 | 10.5411 | 4.9399
947 | 37877.396 | 10.5421 | 4.9553 |  |  | 947 | 37877.396 | 10.5421 | 4.9439 |  |  | 947 | 37877.396 | 10.5421 | 4.9398
948 | 37917.394 | 10.5432 | 4.9552 |  |  | 948 | 37917.394 | 10.5432 | 4.9438 |  |  | 948 | 37917.394 | 10.5432 | 4.9397
949 | 37957.391 | 10.5442 | 4.9551 |  |  | 949 | 37957.391 | 10.5442 | 4.9437 |  |  | 949 | 37957.391 | 10.5442 | 4.9395
950 | 37997.388 | 10.5453 | 4.955 |  |  | 950 | 37997.388 | 10.5453 | 4.9436 |  |  | 950 | 37997.388 | 10.5453 | 4.9394
951 | 38037.385 | 10.5463 | 4.9549 |  |  | 951 | 38037.385 | 10.5463 | 4.9435 |  |  | 951 | 38037.385 | 10.5463 | 4.9393
952 | 38077.383 | 10.5474 | 4.9549 |  |  | 952 | 38077.383 | 10.5474 | 4.9434 |  |  | 952 | 38077.383 | 10.5474 | 4.9392
953 | 38117.38 | 10.5484 | 4.9548 |  |  | 953 | 38117.38 | 10.5484 | 4.9433 |  |  | 953 | 38117.38 | 10.5484 | 4.9391
954 | 38157.377 | 10.5495 | 4.9547 |  |  | 954 | 38157.377 | 10.5495 | 4.9432 |  |  | 954 | 38157.377 | 10.5495 | 4.9389
955 | 38197.374 | 10.5505 | 4.9546 |  |  | 955 | 38197.374 | 10.5505 | 4.9431 |  |  | 955 | 38197.374 | 10.5505 | 4.9388
956 | 38237.372 | 10.5516 | 4.9545 |  |  | 956 | 38237.372 | 10.5516 | 4.943 |  |  | 956 | 38237.372 | 10.5516 | 4.9387
957 | 38277.369 | 10.5526 | 4.9544 |  |  | 957 | 38277.369 | 10.5526 | 4.9429 |  |  | 957 | 38277.369 | 10.5526 | 4.9386
958 | 38317.366 | 10.5537 | 4.9543 |  |  | 958 | 38317.366 | 10.5537 | 4.9428 |  |  | 958 | 38317.366 | 10.5537 | 4.9385
959 | 38357.363 | 10.5547 | 4.9543 |  |  | 959 | 38357.363 | 10.5547 | 4.9427 |  |  | 959 | 38357.363 | 10.5547 | 4.9384
960 | 38397.361 | 10.5557 | 4.9542 |  |  | 960 | 38397.361 | 10.5557 | 4.9426 |  |  | 960 | 38397.361 | 10.5557 | 4.9382
961 | 38437.358 | 10.5568 | 4.9541 |  |  | 961 | 38437.358 | 10.5568 | 4.9425 |  |  | 961 | 38437.358 | 10.5568 | 4.9381
962 | 38477.355 | 10.5578 | 4.954 |  |  | 962 | 38477.355 | 10.5578 | 4.9424 |  |  | 962 | 38477.355 | 10.5578 | 4.938
963 | 38517.352 | 10.5589 | 4.9539 |  |  | 963 | 38517.352 | 10.5589 | 4.9423 |  |  | 963 | 38517.352 | 10.5589 | 4.9379
964 | 38557.35 | 10.5599 | 4.9538 |  |  | 964 | 38557.35 | 10.5599 | 4.9422 |  |  | 964 | 38557.35 | 10.5599 | 4.9378
965 | 38597.347 | 10.5609 | 4.9537 |  |  | 965 | 38597.347 | 10.5609 | 4.9421 |  |  | 965 | 38597.347 | 10.5609 | 4.9376
966 | 38637.344 | 10.562 | 4.9537 |  |  | 966 | 38637.344 | 10.562 | 4.942 |  |  | 966 | 38637.344 | 10.562 | 4.9375
967 | 38677.341 | 10.563 | 4.9536 |  |  | 967 | 38677.341 | 10.563 | 4.9419 |  |  | 967 | 38677.341 | 10.563 | 4.9374
968 | 38717.339 | 10.564 | 4.9535 |  |  | 968 | 38717.339 | 10.564 | 4.9418 |  |  | 968 | 38717.339 | 10.564 | 4.9373
969 | 38757.336 | 10.5651 | 4.9534 |  |  | 969 | 38757.336 | 10.5651 | 4.9416 |  |  | 969 | 38757.336 | 10.5651 | 4.9372
970 | 38797.333 | 10.5661 | 4.9533 |  |  | 970 | 38797.333 | 10.5661 | 4.9415 |  |  | 970 | 38797.333 | 10.5661 | 4.937
971 | 38837.33 | 10.5671 | 4.9532 |  |  | 971 | 38837.33 | 10.5671 | 4.9414 |  |  | 971 | 38837.33 | 10.5671 | 4.9369
972 | 38877.328 | 10.5682 | 4.9531 |  |  | 972 | 38877.328 | 10.5682 | 4.9413 |  |  | 972 | 38877.328 | 10.5682 | 4.9368
973 | 38917.325 | 10.5692 | 4.9531 |  |  | 973 | 38917.325 | 10.5692 | 4.9412 |  |  | 973 | 38917.325 | 10.5692 | 4.9367
974 | 38957.322 | 10.5702 | 4.953 |  |  | 974 | 38957.322 | 10.5702 | 4.9411 |  |  | 974 | 38957.322 | 10.5702 | 4.9366
975 | 38997.319 | 10.5712 | 4.9529 |  |  | 975 | 38997.319 | 10.5712 | 4.941 |  |  | 975 | 38997.319 | 10.5712 | 4.9365
976 | 39037.317 | 10.5723 | 4.9528 |  |  | 976 | 39037.317 | 10.5723 | 4.9409 |  |  | 976 | 39037.317 | 10.5723 | 4.9363
977 | 39077.314 | 10.5733 | 4.9527 |  |  | 977 | 39077.314 | 10.5733 | 4.9408 |  |  | 977 | 39077.314 | 10.5733 | 4.9362
978 | 39117.311 | 10.5743 | 4.9526 |  |  | 978 | 39117.311 | 10.5743 | 4.9407 |  |  | 978 | 39117.311 | 10.5743 | 4.9361
979 | 39157.308 | 10.5753 | 4.9525 |  |  | 979 | 39157.308 | 10.5753 | 4.9406 |  |  | 979 | 39157.308 | 10.5753 | 4.936
980 | 39197.306 | 10.5764 | 4.9525 |  |  | 980 | 39197.306 | 10.5764 | 4.9405 |  |  | 980 | 39197.306 | 10.5764 | 4.9359
981 | 39237.303 | 10.5774 | 4.9524 |  |  | 981 | 39237.303 | 10.5774 | 4.9404 |  |  | 981 | 39237.303 | 10.5774 | 4.9357
982 | 39277.3 | 10.5784 | 4.9523 |  |  | 982 | 39277.3 | 10.5784 | 4.9403 |  |  | 982 | 39277.3 | 10.5784 | 4.9356
983 | 39317.297 | 10.5794 | 4.9522 |  |  | 983 | 39317.297 | 10.5794 | 4.9402 |  |  | 983 | 39317.297 | 10.5794 | 4.9355
984 | 39357.295 | 10.5804 | 4.9521 |  |  | 984 | 39357.295 | 10.5804 | 4.9401 |  |  | 984 | 39357.295 | 10.5804 | 4.9354
985 | 39397.292 | 10.5815 | 4.952 |  |  | 985 | 39397.292 | 10.5815 | 4.94 |  |  | 985 | 39397.292 | 10.5815 | 4.9353
986 | 39437.289 | 10.5825 | 4.9519 |  |  | 986 | 39437.289 | 10.5825 | 4.9399 |  |  | 986 | 39437.289 | 10.5825 | 4.9351
987 | 39477.286 | 10.5835 | 4.9519 |  |  | 987 | 39477.286 | 10.5835 | 4.9398 |  |  | 987 | 39477.286 | 10.5835 | 4.935
988 | 39517.284 | 10.5845 | 4.9518 |  |  | 988 | 39517.284 | 10.5845 | 4.9397 |  |  | 988 | 39517.284 | 10.5845 | 4.9349
989 | 39557.281 | 10.5855 | 4.9517 |  |  | 989 | 39557.281 | 10.5855 | 4.9396 |  |  | 989 | 39557.281 | 10.5855 | 4.9348
990 | 39597.278 | 10.5865 | 4.9516 |  |  | 990 | 39597.278 | 10.5865 | 4.9395 |  |  | 990 | 39597.278 | 10.5865 | 4.9347

## Run8
MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR
File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run8\1 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run8\2 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run8\3 01 001.RAW
Cursor Location: |  |  | (179, 193) |  |  | Cursor Location: |  |  | (179, 192) |  |  | Cursor Location: |  |  | (178, 193)
Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15)
Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev
1 | 39.997 | 3.6888 | 7.7415 | 5.1407 | 0.32414 | 1 | 39.997 | 3.6888 | 7.7445 | 5.1396 | 0.32355 | 1 | 39.997 | 3.6888 | 7.7393 | 5.1433 | 0.32206
2 | 79.995 | 4.382 | 7.4124 |  |  | 2 | 79.995 | 4.382 | 7.4113 |  |  | 2 | 79.995 | 4.382 | 7.4101
3 | 119.992 | 4.7874 | 7.2205 |  |  | 3 | 119.992 | 4.7874 | 7.2183 |  |  | 3 | 119.992 | 4.7874 | 7.2169
4 | 159.989 | 5.0751 | 7.0818 |  |  | 4 | 159.989 | 5.0751 | 7.0792 |  |  | 4 | 159.989 | 5.0751 | 7.0775
5 | 199.986 | 5.2982 | 6.9724 |  |  | 5 | 199.986 | 5.2982 | 6.9697 |  |  | 5 | 199.986 | 5.2982 | 6.9678
6 | 239.984 | 5.4806 | 6.8819 |  |  | 6 | 239.984 | 5.4806 | 6.8793 |  |  | 6 | 239.984 | 5.4806 | 6.8772
7 | 279.981 | 5.6347 | 6.8047 |  |  | 7 | 279.981 | 5.6347 | 6.8021 |  |  | 7 | 279.981 | 5.6347 | 6.8
8 | 319.978 | 5.7683 | 6.7373 |  |  | 8 | 319.978 | 5.7683 | 6.7348 |  |  | 8 | 319.978 | 5.7683 | 6.7328
9 | 359.975 | 5.886 | 6.6775 |  |  | 9 | 359.975 | 5.886 | 6.6752 |  |  | 9 | 359.975 | 5.886 | 6.6732
10 | 399.973 | 5.9914 | 6.6238 |  |  | 10 | 399.973 | 5.9914 | 6.6218 |  |  | 10 | 399.973 | 5.9914 | 6.6198
11 | 439.97 | 6.0867 | 6.5751 |  |  | 11 | 439.97 | 6.0867 | 6.5732 |  |  | 11 | 439.97 | 6.0867 | 6.5714
12 | 479.967 | 6.1737 | 6.5306 |  |  | 12 | 479.967 | 6.1737 | 6.5289 |  |  | 12 | 479.967 | 6.1737 | 6.5271
13 | 519.964 | 6.2538 | 6.4895 |  |  | 13 | 519.964 | 6.2538 | 6.488 |  |  | 13 | 519.964 | 6.2538 | 6.4863
14 | 559.962 | 6.3279 | 6.4515 |  |  | 14 | 559.962 | 6.3279 | 6.4501 |  |  | 14 | 559.962 | 6.3279 | 6.4485
15 | 599.959 | 6.3969 | 6.416 |  |  | 15 | 599.959 | 6.3969 | 6.4148 |  |  | 15 | 599.959 | 6.3969 | 6.4133
16 | 639.956 | 6.4614 | 6.3829 |  |  | 16 | 639.956 | 6.4614 | 6.3818 |  |  | 16 | 639.956 | 6.4614 | 6.3804
17 | 679.953 | 6.522 | 6.3517 |  |  | 17 | 679.953 | 6.522 | 6.3508 |  |  | 17 | 679.953 | 6.522 | 6.3494
18 | 719.951 | 6.5792 | 6.3223 |  |  | 18 | 719.951 | 6.5792 | 6.3215 |  |  | 18 | 719.951 | 6.5792 | 6.3202
19 | 759.948 | 6.6332 | 6.2945 |  |  | 19 | 759.948 | 6.6332 | 6.2939 |  |  | 19 | 759.948 | 6.6332 | 6.2926
20 | 799.945 | 6.6845 | 6.2681 |  |  | 20 | 799.945 | 6.6845 | 6.2676 |  |  | 20 | 799.945 | 6.6845 | 6.2664
21 | 839.942 | 6.7333 | 6.2431 |  |  | 21 | 839.942 | 6.7333 | 6.2427 |  |  | 21 | 839.942 | 6.7333 | 6.2415
22 | 879.94 | 6.7799 | 6.2192 |  |  | 22 | 879.94 | 6.7799 | 6.2189 |  |  | 22 | 879.94 | 6.7799 | 6.2178
23 | 919.937 | 6.8243 | 6.1963 |  |  | 23 | 919.937 | 6.8243 | 6.1962 |  |  | 23 | 919.937 | 6.8243 | 6.1951
24 | 959.934 | 6.8669 | 6.1745 |  |  | 24 | 959.934 | 6.8669 | 6.1744 |  |  | 24 | 959.934 | 6.8669 | 6.1734
25 | 999.931 | 6.9077 | 6.1536 |  |  | 25 | 999.931 | 6.9077 | 6.1536 |  |  | 25 | 999.931 | 6.9077 | 6.1526
26 | 1039.929 | 6.9469 | 6.1335 |  |  | 26 | 1039.929 | 6.9469 | 6.1335 |  |  | 26 | 1039.929 | 6.9469 | 6.1326
27 | 1079.926 | 6.9846 | 6.1141 |  |  | 27 | 1079.926 | 6.9846 | 6.1143 |  |  | 27 | 1079.926 | 6.9846 | 6.1134
28 | 1119.923 | 7.021 | 6.0955 |  |  | 28 | 1119.923 | 7.021 | 6.0957 |  |  | 28 | 1119.923 | 7.021 | 6.0949
29 | 1159.92 | 7.0561 | 6.0776 |  |  | 29 | 1159.92 | 7.0561 | 6.0778 |  |  | 29 | 1159.92 | 7.0561 | 6.077
30 | 1199.918 | 7.09 | 6.0602 |  |  | 30 | 1199.918 | 7.09 | 6.0605 |  |  | 30 | 1199.918 | 7.09 | 6.0598
31 | 1239.915 | 7.1228 | 6.0435 |  |  | 31 | 1239.915 | 7.1228 | 6.0438 |  |  | 31 | 1239.915 | 7.1228 | 6.0431
32 | 1279.912 | 7.1545 | 6.0273 |  |  | 32 | 1279.912 | 7.1545 | 6.0277 |  |  | 32 | 1279.912 | 7.1545 | 6.0269
33 | 1319.909 | 7.1853 | 6.0116 |  |  | 33 | 1319.909 | 7.1853 | 6.012 |  |  | 33 | 1319.909 | 7.1853 | 6.0113
34 | 1359.907 | 7.2152 | 5.9964 |  |  | 34 | 1359.907 | 7.2152 | 5.9969 |  |  | 34 | 1359.907 | 7.2152 | 5.9962
35 | 1399.904 | 7.2442 | 5.9817 |  |  | 35 | 1399.904 | 7.2442 | 5.9821 |  |  | 35 | 1399.904 | 7.2442 | 5.9815
36 | 1439.901 | 7.2723 | 5.9674 |  |  | 36 | 1439.901 | 7.2723 | 5.9678 |  |  | 36 | 1439.901 | 7.2723 | 5.9672
37 | 1479.898 | 7.2997 | 5.9535 |  |  | 37 | 1479.898 | 7.2997 | 5.9539 |  |  | 37 | 1479.898 | 7.2997 | 5.9533
38 | 1519.896 | 7.3264 | 5.9399 |  |  | 38 | 1519.896 | 7.3264 | 5.9404 |  |  | 38 | 1519.896 | 7.3264 | 5.9398
39 | 1559.893 | 7.3524 | 5.9268 |  |  | 39 | 1559.893 | 7.3524 | 5.9273 |  |  | 39 | 1559.893 | 7.3524 | 5.9267
40 | 1599.89 | 7.3777 | 5.914 |  |  | 40 | 1599.89 | 7.3777 | 5.9144 |  |  | 40 | 1599.89 | 7.3777 | 5.9139
41 | 1639.887 | 7.4024 | 5.9015 |  |  | 41 | 1639.887 | 7.4024 | 5.902 |  |  | 41 | 1639.887 | 7.4024 | 5.9014
42 | 1679.885 | 7.4265 | 5.8893 |  |  | 42 | 1679.885 | 7.4265 | 5.8898 |  |  | 42 | 1679.885 | 7.4265 | 5.8893
43 | 1719.882 | 7.45 | 5.8774 |  |  | 43 | 1719.882 | 7.45 | 5.8779 |  |  | 43 | 1719.882 | 7.45 | 5.8774
44 | 1759.879 | 7.473 | 5.8659 |  |  | 44 | 1759.879 | 7.473 | 5.8663 |  |  | 44 | 1759.879 | 7.473 | 5.8659
45 | 1799.876 | 7.4955 | 5.8546 |  |  | 45 | 1799.876 | 7.4955 | 5.855 |  |  | 45 | 1799.876 | 7.4955 | 5.8546
46 | 1839.874 | 7.5175 | 5.8435 |  |  | 46 | 1839.874 | 7.5175 | 5.8439 |  |  | 46 | 1839.874 | 7.5175 | 5.8435
47 | 1879.871 | 7.539 | 5.8327 |  |  | 47 | 1879.871 | 7.539 | 5.8331 |  |  | 47 | 1879.871 | 7.539 | 5.8327
48 | 1919.868 | 7.56 | 5.8222 |  |  | 48 | 1919.868 | 7.56 | 5.8225 |  |  | 48 | 1919.868 | 7.56 | 5.8222
49 | 1959.865 | 7.5806 | 5.8118 |  |  | 49 | 1959.865 | 7.5806 | 5.8122 |  |  | 49 | 1959.865 | 7.5806 | 5.8119
50 | 1999.863 | 7.6008 | 5.8018 |  |  | 50 | 1999.863 | 7.6008 | 5.802 |  |  | 50 | 1999.863 | 7.6008 | 5.8018
51 | 2039.86 | 7.6206 | 5.7919 |  |  | 51 | 2039.86 | 7.6206 | 5.7921 |  |  | 51 | 2039.86 | 7.6206 | 5.7919
52 | 2079.857 | 7.6401 | 5.7822 |  |  | 52 | 2079.857 | 7.6401 | 5.7824 |  |  | 52 | 2079.857 | 7.6401 | 5.7822
53 | 2119.854 | 7.6591 | 5.7727 |  |  | 53 | 2119.854 | 7.6591 | 5.7729 |  |  | 53 | 2119.854 | 7.6591 | 5.7727
54 | 2159.852 | 7.6778 | 5.7634 |  |  | 54 | 2159.852 | 7.6778 | 5.7636 |  |  | 54 | 2159.852 | 7.6778 | 5.7634
55 | 2199.849 | 7.6961 | 5.7543 |  |  | 55 | 2199.849 | 7.6961 | 5.7544 |  |  | 55 | 2199.849 | 7.6961 | 5.7543
56 | 2239.846 | 7.7142 | 5.7454 |  |  | 56 | 2239.846 | 7.7142 | 5.7455 |  |  | 56 | 2239.846 | 7.7142 | 5.7453
57 | 2279.843 | 7.7319 | 5.7366 |  |  | 57 | 2279.843 | 7.7319 | 5.7367 |  |  | 57 | 2279.843 | 7.7319 | 5.7366
58 | 2319.841 | 7.7493 | 5.7281 |  |  | 58 | 2319.841 | 7.7493 | 5.728 |  |  | 58 | 2319.841 | 7.7493 | 5.728
59 | 2359.838 | 7.7663 | 5.7196 |  |  | 59 | 2359.838 | 7.7663 | 5.7196 |  |  | 59 | 2359.838 | 7.7663 | 5.7195
60 | 2399.835 | 7.7832 | 5.7114 |  |  | 60 | 2399.835 | 7.7832 | 5.7113 |  |  | 60 | 2399.835 | 7.7832 | 5.7112
61 | 2439.832 | 7.7997 | 5.7032 |  |  | 61 | 2439.832 | 7.7997 | 5.7031 |  |  | 61 | 2439.832 | 7.7997 | 5.7031
62 | 2479.83 | 7.8159 | 5.6953 |  |  | 62 | 2479.83 | 7.8159 | 5.6951 |  |  | 62 | 2479.83 | 7.8159 | 5.6951
63 | 2519.827 | 7.8319 | 5.6874 |  |  | 63 | 2519.827 | 7.8319 | 5.6872 |  |  | 63 | 2519.827 | 7.8319 | 5.6872
64 | 2559.824 | 7.8477 | 5.6797 |  |  | 64 | 2559.824 | 7.8477 | 5.6794 |  |  | 64 | 2559.824 | 7.8477 | 5.6795
65 | 2599.821 | 7.8632 | 5.6722 |  |  | 65 | 2599.821 | 7.8632 | 5.6718 |  |  | 65 | 2599.821 | 7.8632 | 5.6719
66 | 2639.819 | 7.8785 | 5.6647 |  |  | 66 | 2639.819 | 7.8785 | 5.6643 |  |  | 66 | 2639.819 | 7.8785 | 5.6645
67 | 2679.816 | 7.8935 | 5.6574 |  |  | 67 | 2679.816 | 7.8935 | 5.657 |  |  | 67 | 2679.816 | 7.8935 | 5.6571
68 | 2719.813 | 7.9083 | 5.6503 |  |  | 68 | 2719.813 | 7.9083 | 5.6497 |  |  | 68 | 2719.813 | 7.9083 | 5.6499
69 | 2759.81 | 7.9229 | 5.6432 |  |  | 69 | 2759.81 | 7.9229 | 5.6426 |  |  | 69 | 2759.81 | 7.9229 | 5.6428
70 | 2799.808 | 7.9373 | 5.6362 |  |  | 70 | 2799.808 | 7.9373 | 5.6356 |  |  | 70 | 2799.808 | 7.9373 | 5.6359
71 | 2839.805 | 7.9515 | 5.6294 |  |  | 71 | 2839.805 | 7.9515 | 5.6287 |  |  | 71 | 2839.805 | 7.9515 | 5.629
72 | 2879.802 | 7.9655 | 5.6227 |  |  | 72 | 2879.802 | 7.9655 | 5.6219 |  |  | 72 | 2879.802 | 7.9655 | 5.6222
73 | 2919.799 | 7.9793 | 5.616 |  |  | 73 | 2919.799 | 7.9793 | 5.6152 |  |  | 73 | 2919.799 | 7.9793 | 5.6156
74 | 2959.797 | 7.9929 | 5.6095 |  |  | 74 | 2959.797 | 7.9929 | 5.6086 |  |  | 74 | 2959.797 | 7.9929 | 5.609
75 | 2999.794 | 8.0063 | 5.6031 |  |  | 75 | 2999.794 | 8.0063 | 5.6022 |  |  | 75 | 2999.794 | 8.0063 | 5.6026
76 | 3039.791 | 8.0195 | 5.5968 |  |  | 76 | 3039.791 | 8.0195 | 5.5958 |  |  | 76 | 3039.791 | 8.0195 | 5.5962
77 | 3079.788 | 8.0326 | 5.5906 |  |  | 77 | 3079.788 | 8.0326 | 5.5895 |  |  | 77 | 3079.788 | 8.0326 | 5.59
78 | 3119.786 | 8.0455 | 5.5844 |  |  | 78 | 3119.786 | 8.0455 | 5.5833 |  |  | 78 | 3119.786 | 8.0455 | 5.5838
79 | 3159.783 | 8.0583 | 5.5784 |  |  | 79 | 3159.783 | 8.0583 | 5.5772 |  |  | 79 | 3159.783 | 8.0583 | 5.5777
80 | 3199.78 | 8.0708 | 5.5724 |  |  | 80 | 3199.78 | 8.0708 | 5.5712 |  |  | 80 | 3199.78 | 8.0708 | 5.5718
81 | 3239.777 | 8.0833 | 5.5666 |  |  | 81 | 3239.777 | 8.0833 | 5.5653 |  |  | 81 | 3239.777 | 8.0833 | 5.5659
82 | 3279.775 | 8.0955 | 5.5608 |  |  | 82 | 3279.775 | 8.0955 | 5.5594 |  |  | 82 | 3279.775 | 8.0955 | 5.5601
83 | 3319.772 | 8.1077 | 5.5551 |  |  | 83 | 3319.772 | 8.1077 | 5.5537 |  |  | 83 | 3319.772 | 8.1077 | 5.5543
84 | 3359.769 | 8.1196 | 5.5495 |  |  | 84 | 3359.769 | 8.1196 | 5.548 |  |  | 84 | 3359.769 | 8.1196 | 5.5487
85 | 3399.766 | 8.1315 | 5.5439 |  |  | 85 | 3399.766 | 8.1315 | 5.5424 |  |  | 85 | 3399.766 | 8.1315 | 5.5431
86 | 3439.764 | 8.1432 | 5.5385 |  |  | 86 | 3439.764 | 8.1432 | 5.5369 |  |  | 86 | 3439.764 | 8.1432 | 5.5376
87 | 3479.761 | 8.1547 | 5.5331 |  |  | 87 | 3479.761 | 8.1547 | 5.5314 |  |  | 87 | 3479.761 | 8.1547 | 5.5322
88 | 3519.758 | 8.1661 | 5.5278 |  |  | 88 | 3519.758 | 8.1661 | 5.5261 |  |  | 88 | 3519.758 | 8.1661 | 5.5269
89 | 3559.755 | 8.1774 | 5.5226 |  |  | 89 | 3559.755 | 8.1774 | 5.5208 |  |  | 89 | 3559.755 | 8.1774 | 5.5216
90 | 3599.753 | 8.1886 | 5.5174 |  |  | 90 | 3599.753 | 8.1886 | 5.5156 |  |  | 90 | 3599.753 | 8.1886 | 5.5164
91 | 3639.75 | 8.1997 | 5.5123 |  |  | 91 | 3639.75 | 8.1997 | 5.5104 |  |  | 91 | 3639.75 | 8.1997 | 5.5113
92 | 3679.747 | 8.2106 | 5.5073 |  |  | 92 | 3679.747 | 8.2106 | 5.5053 |  |  | 92 | 3679.747 | 8.2106 | 5.5063
93 | 3719.744 | 8.2214 | 5.5023 |  |  | 93 | 3719.744 | 8.2214 | 5.5003 |  |  | 93 | 3719.744 | 8.2214 | 5.5013
94 | 3759.742 | 8.2321 | 5.4974 |  |  | 94 | 3759.742 | 8.2321 | 5.4954 |  |  | 94 | 3759.742 | 8.2321 | 5.4964
95 | 3799.739 | 8.2427 | 5.4926 |  |  | 95 | 3799.739 | 8.2427 | 5.4905 |  |  | 95 | 3799.739 | 8.2427 | 5.4915
96 | 3839.736 | 8.2532 | 5.4878 |  |  | 96 | 3839.736 | 8.2532 | 5.4856 |  |  | 96 | 3839.736 | 8.2532 | 5.4867
97 | 3879.733 | 8.2635 | 5.4831 |  |  | 97 | 3879.733 | 8.2635 | 5.4809 |  |  | 97 | 3879.733 | 8.2635 | 5.482
98 | 3919.731 | 8.2738 | 5.4785 |  |  | 98 | 3919.731 | 8.2738 | 5.4762 |  |  | 98 | 3919.731 | 8.2738 | 5.4773
99 | 3959.728 | 8.2839 | 5.4739 |  |  | 99 | 3959.728 | 8.2839 | 5.4715 |  |  | 99 | 3959.728 | 8.2839 | 5.4727
100 | 3999.725 | 8.294 | 5.4694 |  |  | 100 | 3999.725 | 8.294 | 5.467 |  |  | 100 | 3999.725 | 8.294 | 5.4681
101 | 4039.722 | 8.3039 | 5.4649 |  |  | 101 | 4039.722 | 8.3039 | 5.4624 |  |  | 101 | 4039.722 | 8.3039 | 5.4636
102 | 4079.72 | 8.3138 | 5.4605 |  |  | 102 | 4079.72 | 8.3138 | 5.458 |  |  | 102 | 4079.72 | 8.3138 | 5.4592
103 | 4119.717 | 8.3235 | 5.4561 |  |  | 103 | 4119.717 | 8.3235 | 5.4536 |  |  | 103 | 4119.717 | 8.3235 | 5.4548
104 | 4159.714 | 8.3332 | 5.4518 |  |  | 104 | 4159.714 | 8.3332 | 5.4492 |  |  | 104 | 4159.714 | 8.3332 | 5.4505
105 | 4199.711 | 8.3428 | 5.4476 |  |  | 105 | 4199.711 | 8.3428 | 5.4449 |  |  | 105 | 4199.711 | 8.3428 | 5.4462
106 | 4239.709 | 8.3522 | 5.4434 |  |  | 106 | 4239.709 | 8.3522 | 5.4406 |  |  | 106 | 4239.709 | 8.3522 | 5.442
107 | 4279.706 | 8.3616 | 5.4392 |  |  | 107 | 4279.706 | 8.3616 | 5.4364 |  |  | 107 | 4279.706 | 8.3616 | 5.4378
108 | 4319.703 | 8.3709 | 5.4351 |  |  | 108 | 4319.703 | 8.3709 | 5.4323 |  |  | 108 | 4319.703 | 8.3709 | 5.4337
109 | 4359.7 | 8.3802 | 5.4311 |  |  | 109 | 4359.7 | 8.3802 | 5.4282 |  |  | 109 | 4359.7 | 8.3802 | 5.4296
110 | 4399.698 | 8.3893 | 5.4271 |  |  | 110 | 4399.698 | 8.3893 | 5.4241 |  |  | 110 | 4399.698 | 8.3893 | 5.4256
111 | 4439.695 | 8.3983 | 5.4231 |  |  | 111 | 4439.695 | 8.3983 | 5.4201 |  |  | 111 | 4439.695 | 8.3983 | 5.4216
112 | 4479.692 | 8.4073 | 5.4192 |  |  | 112 | 4479.692 | 8.4073 | 5.4162 |  |  | 112 | 4479.692 | 8.4073 | 5.4177
113 | 4519.689 | 8.4162 | 5.4154 |  |  | 113 | 4519.689 | 8.4162 | 5.4123 |  |  | 113 | 4519.689 | 8.4162 | 5.4138
114 | 4559.687 | 8.425 | 5.4116 |  |  | 114 | 4559.687 | 8.425 | 5.4084 |  |  | 114 | 4559.687 | 8.425 | 5.41
115 | 4599.684 | 8.4337 | 5.4078 |  |  | 115 | 4599.684 | 8.4337 | 5.4046 |  |  | 115 | 4599.684 | 8.4337 | 5.4062
116 | 4639.681 | 8.4424 | 5.4041 |  |  | 116 | 4639.681 | 8.4424 | 5.4008 |  |  | 116 | 4639.681 | 8.4424 | 5.4025
117 | 4679.678 | 8.451 | 5.4004 |  |  | 117 | 4679.678 | 8.451 | 5.3971 |  |  | 117 | 4679.678 | 8.451 | 5.3988
118 | 4719.676 | 8.4595 | 5.3968 |  |  | 118 | 4719.676 | 8.4595 | 5.3934 |  |  | 118 | 4719.676 | 8.4595 | 5.3951
119 | 4759.673 | 8.4679 | 5.3932 |  |  | 119 | 4759.673 | 8.4679 | 5.3897 |  |  | 119 | 4759.673 | 8.4679 | 5.3915
120 | 4799.67 | 8.4763 | 5.3896 |  |  | 120 | 4799.67 | 8.4763 | 5.3861 |  |  | 120 | 4799.67 | 8.4763 | 5.3879
121 | 4839.667 | 8.4846 | 5.3861 |  |  | 121 | 4839.667 | 8.4846 | 5.3826 |  |  | 121 | 4839.667 | 8.4846 | 5.3844
122 | 4879.665 | 8.4928 | 5.3826 |  |  | 122 | 4879.665 | 8.4928 | 5.379 |  |  | 122 | 4879.665 | 8.4928 | 5.3809
123 | 4919.662 | 8.501 | 5.3792 |  |  | 123 | 4919.662 | 8.501 | 5.3756 |  |  | 123 | 4919.662 | 8.501 | 5.3774
124 | 4959.659 | 8.5091 | 5.3758 |  |  | 124 | 4959.659 | 8.5091 | 5.3721 |  |  | 124 | 4959.659 | 8.5091 | 5.374
125 | 4999.656 | 8.5171 | 5.3724 |  |  | 125 | 4999.656 | 8.5171 | 5.3687 |  |  | 125 | 4999.656 | 8.5171 | 5.3706
126 | 5039.654 | 8.5251 | 5.3691 |  |  | 126 | 5039.654 | 8.5251 | 5.3653 |  |  | 126 | 5039.654 | 8.5251 | 5.3673
127 | 5079.651 | 8.533 | 5.3658 |  |  | 127 | 5079.651 | 8.533 | 5.362 |  |  | 127 | 5079.651 | 8.533 | 5.364
128 | 5119.648 | 8.5408 | 5.3625 |  |  | 128 | 5119.648 | 8.5408 | 5.3587 |  |  | 128 | 5119.648 | 8.5408 | 5.3607
129 | 5159.645 | 8.5486 | 5.3593 |  |  | 129 | 5159.645 | 8.5486 | 5.3554 |  |  | 129 | 5159.645 | 8.5486 | 5.3575
130 | 5199.643 | 8.5563 | 5.3562 |  |  | 130 | 5199.643 | 8.5563 | 5.3522 |  |  | 130 | 5199.643 | 8.5563 | 5.3543
131 | 5239.64 | 8.564 | 5.353 |  |  | 131 | 5239.64 | 8.564 | 5.349 |  |  | 131 | 5239.64 | 8.564 | 5.3511
132 | 5279.637 | 8.5716 | 5.3499 |  |  | 132 | 5279.637 | 8.5716 | 5.3459 |  |  | 132 | 5279.637 | 8.5716 | 5.348
133 | 5319.634 | 8.5792 | 5.3468 |  |  | 133 | 5319.634 | 8.5792 | 5.3428 |  |  | 133 | 5319.634 | 8.5792 | 5.3449
134 | 5359.632 | 8.5867 | 5.3438 |  |  | 134 | 5359.632 | 8.5867 | 5.3397 |  |  | 134 | 5359.632 | 8.5867 | 5.3418
135 | 5399.629 | 8.5941 | 5.3408 |  |  | 135 | 5399.629 | 8.5941 | 5.3366 |  |  | 135 | 5399.629 | 8.5941 | 5.3388
136 | 5439.626 | 8.6015 | 5.3378 |  |  | 136 | 5439.626 | 8.6015 | 5.3336 |  |  | 136 | 5439.626 | 8.6015 | 5.3358
137 | 5479.623 | 8.6088 | 5.3348 |  |  | 137 | 5479.623 | 8.6088 | 5.3306 |  |  | 137 | 5479.623 | 8.6088 | 5.3329
138 | 5519.621 | 8.6161 | 5.3319 |  |  | 138 | 5519.621 | 8.6161 | 5.3277 |  |  | 138 | 5519.621 | 8.6161 | 5.3299
139 | 5559.618 | 8.6233 | 5.329 |  |  | 139 | 5559.618 | 8.6233 | 5.3247 |  |  | 139 | 5559.618 | 8.6233 | 5.327
140 | 5599.615 | 8.6305 | 5.3262 |  |  | 140 | 5599.615 | 8.6305 | 5.3218 |  |  | 140 | 5599.615 | 8.6305 | 5.3242
141 | 5639.612 | 8.6376 | 5.3234 |  |  | 141 | 5639.612 | 8.6376 | 5.319 |  |  | 141 | 5639.612 | 8.6376 | 5.3213
142 | 5679.61 | 8.6446 | 5.3206 |  |  | 142 | 5679.61 | 8.6446 | 5.3161 |  |  | 142 | 5679.61 | 8.6446 | 5.3185
143 | 5719.607 | 8.6517 | 5.3178 |  |  | 143 | 5719.607 | 8.6517 | 5.3133 |  |  | 143 | 5719.607 | 8.6517 | 5.3157
144 | 5759.604 | 8.6586 | 5.3151 |  |  | 144 | 5759.604 | 8.6586 | 5.3106 |  |  | 144 | 5759.604 | 8.6586 | 5.313
145 | 5799.601 | 8.6655 | 5.3124 |  |  | 145 | 5799.601 | 8.6655 | 5.3078 |  |  | 145 | 5799.601 | 8.6655 | 5.3103
146 | 5839.599 | 8.6724 | 5.3097 |  |  | 146 | 5839.599 | 8.6724 | 5.3051 |  |  | 146 | 5839.599 | 8.6724 | 5.3076
147 | 5879.596 | 8.6792 | 5.307 |  |  | 147 | 5879.596 | 8.6792 | 5.3024 |  |  | 147 | 5879.596 | 8.6792 | 5.3049
148 | 5919.593 | 8.686 | 5.3044 |  |  | 148 | 5919.593 | 8.686 | 5.2998 |  |  | 148 | 5919.593 | 8.686 | 5.3023
149 | 5959.59 | 8.6928 | 5.3018 |  |  | 149 | 5959.59 | 8.6928 | 5.2971 |  |  | 149 | 5959.59 | 8.6928 | 5.2997
150 | 5999.588 | 8.6994 | 5.2992 |  |  | 150 | 5999.588 | 8.6994 | 5.2945 |  |  | 150 | 5999.588 | 8.6994 | 5.2971
151 | 6039.585 | 8.7061 | 5.2967 |  |  | 151 | 6039.585 | 8.7061 | 5.2919 |  |  | 151 | 6039.585 | 8.7061 | 5.2945
152 | 6079.582 | 8.7127 | 5.2941 |  |  | 152 | 6079.582 | 8.7127 | 5.2894 |  |  | 152 | 6079.582 | 8.7127 | 5.292
153 | 6119.579 | 8.7192 | 5.2917 |  |  | 153 | 6119.579 | 8.7192 | 5.2869 |  |  | 153 | 6119.579 | 8.7192 | 5.2895
154 | 6159.577 | 8.7258 | 5.2892 |  |  | 154 | 6159.577 | 8.7258 | 5.2844 |  |  | 154 | 6159.577 | 8.7258 | 5.287
155 | 6199.574 | 8.7322 | 5.2867 |  |  | 155 | 6199.574 | 8.7322 | 5.2819 |  |  | 155 | 6199.574 | 8.7322 | 5.2846
156 | 6239.571 | 8.7387 | 5.2843 |  |  | 156 | 6239.571 | 8.7387 | 5.2794 |  |  | 156 | 6239.571 | 8.7387 | 5.2821
157 | 6279.568 | 8.7451 | 5.2819 |  |  | 157 | 6279.568 | 8.7451 | 5.277 |  |  | 157 | 6279.568 | 8.7451 | 5.2797
158 | 6319.566 | 8.7514 | 5.2795 |  |  | 158 | 6319.566 | 8.7514 | 5.2746 |  |  | 158 | 6319.566 | 8.7514 | 5.2773
159 | 6359.563 | 8.7577 | 5.2772 |  |  | 159 | 6359.563 | 8.7577 | 5.2722 |  |  | 159 | 6359.563 | 8.7577 | 5.275
160 | 6399.56 | 8.764 | 5.2749 |  |  | 160 | 6399.56 | 8.764 | 5.2699 |  |  | 160 | 6399.56 | 8.764 | 5.2726
161 | 6439.557 | 8.7702 | 5.2726 |  |  | 161 | 6439.557 | 8.7702 | 5.2675 |  |  | 161 | 6439.557 | 8.7702 | 5.2703
162 | 6479.555 | 8.7764 | 5.2703 |  |  | 162 | 6479.555 | 8.7764 | 5.2652 |  |  | 162 | 6479.555 | 8.7764 | 5.268
163 | 6519.552 | 8.7826 | 5.268 |  |  | 163 | 6519.552 | 8.7826 | 5.2629 |  |  | 163 | 6519.552 | 8.7826 | 5.2658
164 | 6559.549 | 8.7887 | 5.2658 |  |  | 164 | 6559.549 | 8.7887 | 5.2607 |  |  | 164 | 6559.549 | 8.7887 | 5.2635
165 | 6599.546 | 8.7948 | 5.2636 |  |  | 165 | 6599.546 | 8.7948 | 5.2584 |  |  | 165 | 6599.546 | 8.7948 | 5.2613
166 | 6639.544 | 8.8008 | 5.2614 |  |  | 166 | 6639.544 | 8.8008 | 5.2562 |  |  | 166 | 6639.544 | 8.8008 | 5.2591
167 | 6679.541 | 8.8068 | 5.2592 |  |  | 167 | 6679.541 | 8.8068 | 5.254 |  |  | 167 | 6679.541 | 8.8068 | 5.2569
168 | 6719.538 | 8.8128 | 5.257 |  |  | 168 | 6719.538 | 8.8128 | 5.2518 |  |  | 168 | 6719.538 | 8.8128 | 5.2548
169 | 6759.535 | 8.8187 | 5.2549 |  |  | 169 | 6759.535 | 8.8187 | 5.2497 |  |  | 169 | 6759.535 | 8.8187 | 5.2527
170 | 6799.533 | 8.8246 | 5.2528 |  |  | 170 | 6799.533 | 8.8246 | 5.2475 |  |  | 170 | 6799.533 | 8.8246 | 5.2505
171 | 6839.53 | 8.8305 | 5.2507 |  |  | 171 | 6839.53 | 8.8305 | 5.2454 |  |  | 171 | 6839.53 | 8.8305 | 5.2484
172 | 6879.527 | 8.8363 | 5.2486 |  |  | 172 | 6879.527 | 8.8363 | 5.2433 |  |  | 172 | 6879.527 | 8.8363 | 5.2464
173 | 6919.524 | 8.8421 | 5.2466 |  |  | 173 | 6919.524 | 8.8421 | 5.2413 |  |  | 173 | 6919.524 | 8.8421 | 5.2443
174 | 6959.522 | 8.8479 | 5.2446 |  |  | 174 | 6959.522 | 8.8479 | 5.2392 |  |  | 174 | 6959.522 | 8.8479 | 5.2423
175 | 6999.519 | 8.8536 | 5.2426 |  |  | 175 | 6999.519 | 8.8536 | 5.2372 |  |  | 175 | 6999.519 | 8.8536 | 5.2403
176 | 7039.516 | 8.8593 | 5.2406 |  |  | 176 | 7039.516 | 8.8593 | 5.2352 |  |  | 176 | 7039.516 | 8.8593 | 5.2383
177 | 7079.513 | 8.865 | 5.2386 |  |  | 177 | 7079.513 | 8.865 | 5.2332 |  |  | 177 | 7079.513 | 8.865 | 5.2363
178 | 7119.511 | 8.8706 | 5.2366 |  |  | 178 | 7119.511 | 8.8706 | 5.2312 |  |  | 178 | 7119.511 | 8.8706 | 5.2343
179 | 7159.508 | 8.8762 | 5.2347 |  |  | 179 | 7159.508 | 8.8762 | 5.2292 |  |  | 179 | 7159.508 | 8.8762 | 5.2324
180 | 7199.505 | 8.8818 | 5.2328 |  |  | 180 | 7199.505 | 8.8818 | 5.2273 |  |  | 180 | 7199.505 | 8.8818 | 5.2305
181 | 7239.502 | 8.8873 | 5.2309 |  |  | 181 | 7239.502 | 8.8873 | 5.2254 |  |  | 181 | 7239.502 | 8.8873 | 5.2286
182 | 7279.5 | 8.8928 | 5.229 |  |  | 182 | 7279.5 | 8.8928 | 5.2235 |  |  | 182 | 7279.5 | 8.8928 | 5.2267
183 | 7319.497 | 8.8983 | 5.2271 |  |  | 183 | 7319.497 | 8.8983 | 5.2216 |  |  | 183 | 7319.497 | 8.8983 | 5.2248
184 | 7359.494 | 8.9037 | 5.2253 |  |  | 184 | 7359.494 | 8.9037 | 5.2197 |  |  | 184 | 7359.494 | 8.9037 | 5.223
185 | 7399.491 | 8.9092 | 5.2235 |  |  | 185 | 7399.491 | 8.9092 | 5.2179 |  |  | 185 | 7399.491 | 8.9092 | 5.2212
186 | 7439.489 | 8.9146 | 5.2216 |  |  | 186 | 7439.489 | 8.9146 | 5.2161 |  |  | 186 | 7439.489 | 8.9146 | 5.2193
187 | 7479.486 | 8.9199 | 5.2198 |  |  | 187 | 7479.486 | 8.9199 | 5.2143 |  |  | 187 | 7479.486 | 8.9199 | 5.2175
188 | 7519.483 | 8.9253 | 5.2181 |  |  | 188 | 7519.483 | 8.9253 | 5.2125 |  |  | 188 | 7519.483 | 8.9253 | 5.2158
189 | 7559.48 | 8.9306 | 5.2163 |  |  | 189 | 7559.48 | 8.9306 | 5.2107 |  |  | 189 | 7559.48 | 8.9306 | 5.214
190 | 7599.478 | 8.9358 | 5.2145 |  |  | 190 | 7599.478 | 8.9358 | 5.2089 |  |  | 190 | 7599.478 | 8.9358 | 5.2123
191 | 7639.475 | 8.9411 | 5.2128 |  |  | 191 | 7639.475 | 8.9411 | 5.2072 |  |  | 191 | 7639.475 | 8.9411 | 5.2105
192 | 7679.472 | 8.9463 | 5.2111 |  |  | 192 | 7679.472 | 8.9463 | 5.2054 |  |  | 192 | 7679.472 | 8.9463 | 5.2088
193 | 7719.469 | 8.9515 | 5.2094 |  |  | 193 | 7719.469 | 8.9515 | 5.2037 |  |  | 193 | 7719.469 | 8.9515 | 5.2071
194 | 7759.467 | 8.9567 | 5.2077 |  |  | 194 | 7759.467 | 8.9567 | 5.202 |  |  | 194 | 7759.467 | 8.9567 | 5.2054
195 | 7799.464 | 8.9618 | 5.206 |  |  | 195 | 7799.464 | 8.9618 | 5.2003 |  |  | 195 | 7799.464 | 8.9618 | 5.2038
196 | 7839.461 | 8.9669 | 5.2044 |  |  | 196 | 7839.461 | 8.9669 | 5.1987 |  |  | 196 | 7839.461 | 8.9669 | 5.2021
197 | 7879.458 | 8.972 | 5.2027 |  |  | 197 | 7879.458 | 8.972 | 5.197 |  |  | 197 | 7879.458 | 8.972 | 5.2005
198 | 7919.456 | 8.9771 | 5.2011 |  |  | 198 | 7919.456 | 8.9771 | 5.1954 |  |  | 198 | 7919.456 | 8.9771 | 5.1988
199 | 7959.453 | 8.9821 | 5.1995 |  |  | 199 | 7959.453 | 8.9821 | 5.1938 |  |  | 199 | 7959.453 | 8.9821 | 5.1972
200 | 7999.45 | 8.9871 | 5.1979 |  |  | 200 | 7999.45 | 8.9871 | 5.1922 |  |  | 200 | 7999.45 | 8.9871 | 5.1956
201 | 8039.447 | 8.9921 | 5.1963 |  |  | 201 | 8039.447 | 8.9921 | 5.1906 |  |  | 201 | 8039.447 | 8.9921 | 5.1941
202 | 8079.445 | 8.9971 | 5.1948 |  |  | 202 | 8079.445 | 8.9971 | 5.189 |  |  | 202 | 8079.445 | 8.9971 | 5.1925
203 | 8119.442 | 9.002 | 5.1932 |  |  | 203 | 8119.442 | 9.002 | 5.1874 |  |  | 203 | 8119.442 | 9.002 | 5.191
204 | 8159.439 | 9.0069 | 5.1917 |  |  | 204 | 8159.439 | 9.0069 | 5.1859 |  |  | 204 | 8159.439 | 9.0069 | 5.1894
205 | 8199.436 | 9.0118 | 5.1901 |  |  | 205 | 8199.436 | 9.0118 | 5.1843 |  |  | 205 | 8199.436 | 9.0118 | 5.1879
206 | 8239.434 | 9.0167 | 5.1886 |  |  | 206 | 8239.434 | 9.0167 | 5.1828 |  |  | 206 | 8239.434 | 9.0167 | 5.1864
207 | 8279.431 | 9.0215 | 5.1871 |  |  | 207 | 8279.431 | 9.0215 | 5.1813 |  |  | 207 | 8279.431 | 9.0215 | 5.1849
208 | 8319.428 | 9.0263 | 5.1856 |  |  | 208 | 8319.428 | 9.0263 | 5.1798 |  |  | 208 | 8319.428 | 9.0263 | 5.1834
209 | 8359.425 | 9.0311 | 5.1842 |  |  | 209 | 8359.425 | 9.0311 | 5.1783 |  |  | 209 | 8359.425 | 9.0311 | 5.1819
210 | 8399.423 | 9.0359 | 5.1827 |  |  | 210 | 8399.423 | 9.0359 | 5.1769 |  |  | 210 | 8399.423 | 9.0359 | 5.1805
211 | 8439.42 | 9.0407 | 5.1813 |  |  | 211 | 8439.42 | 9.0407 | 5.1754 |  |  | 211 | 8439.42 | 9.0407 | 5.1791
212 | 8479.417 | 9.0454 | 5.1798 |  |  | 212 | 8479.417 | 9.0454 | 5.174 |  |  | 212 | 8479.417 | 9.0454 | 5.1776
213 | 8519.414 | 9.0501 | 5.1784 |  |  | 213 | 8519.414 | 9.0501 | 5.1726 |  |  | 213 | 8519.414 | 9.0501 | 5.1762
214 | 8559.412 | 9.0548 | 5.177 |  |  | 214 | 8559.412 | 9.0548 | 5.1711 |  |  | 214 | 8559.412 | 9.0548 | 5.1748
215 | 8599.409 | 9.0594 | 5.1756 |  |  | 215 | 8599.409 | 9.0594 | 5.1697 |  |  | 215 | 8599.409 | 9.0594 | 5.1734
216 | 8639.406 | 9.0641 | 5.1742 |  |  | 216 | 8639.406 | 9.0641 | 5.1684 |  |  | 216 | 8639.406 | 9.0641 | 5.172
217 | 8679.403 | 9.0687 | 5.1728 |  |  | 217 | 8679.403 | 9.0687 | 5.167 |  |  | 217 | 8679.403 | 9.0687 | 5.1707
218 | 8719.401 | 9.0733 | 5.1715 |  |  | 218 | 8719.401 | 9.0733 | 5.1656 |  |  | 218 | 8719.401 | 9.0733 | 5.1693
219 | 8759.398 | 9.0779 | 5.1701 |  |  | 219 | 8759.398 | 9.0779 | 5.1643 |  |  | 219 | 8759.398 | 9.0779 | 5.168
220 | 8799.395 | 9.0824 | 5.1688 |  |  | 220 | 8799.395 | 9.0824 | 5.1629 |  |  | 220 | 8799.395 | 9.0824 | 5.1666
221 | 8839.392 | 9.087 | 5.1675 |  |  | 221 | 8839.392 | 9.087 | 5.1616 |  |  | 221 | 8839.392 | 9.087 | 5.1653
222 | 8879.39 | 9.0915 | 5.1662 |  |  | 222 | 8879.39 | 9.0915 | 5.1603 |  |  | 222 | 8879.39 | 9.0915 | 5.164
223 | 8919.387 | 9.096 | 5.1649 |  |  | 223 | 8919.387 | 9.096 | 5.159 |  |  | 223 | 8919.387 | 9.096 | 5.1627
224 | 8959.384 | 9.1005 | 5.1636 |  |  | 224 | 8959.384 | 9.1005 | 5.1577 |  |  | 224 | 8959.384 | 9.1005 | 5.1614
225 | 8999.381 | 9.1049 | 5.1623 |  |  | 225 | 8999.381 | 9.1049 | 5.1564 |  |  | 225 | 8999.381 | 9.1049 | 5.1602
226 | 9039.379 | 9.1093 | 5.161 |  |  | 226 | 9039.379 | 9.1093 | 5.1551 |  |  | 226 | 9039.379 | 9.1093 | 5.1589
227 | 9079.376 | 9.1138 | 5.1597 |  |  | 227 | 9079.376 | 9.1138 | 5.1539 |  |  | 227 | 9079.376 | 9.1138 | 5.1576
228 | 9119.373 | 9.1182 | 5.1585 |  |  | 228 | 9119.373 | 9.1182 | 5.1526 |  |  | 228 | 9119.373 | 9.1182 | 5.1564
229 | 9159.37 | 9.1225 | 5.1573 |  |  | 229 | 9159.37 | 9.1225 | 5.1514 |  |  | 229 | 9159.37 | 9.1225 | 5.1552
230 | 9199.368 | 9.1269 | 5.156 |  |  | 230 | 9199.368 | 9.1269 | 5.1501 |  |  | 230 | 9199.368 | 9.1269 | 5.154
231 | 9239.365 | 9.1312 | 5.1548 |  |  | 231 | 9239.365 | 9.1312 | 5.1489 |  |  | 231 | 9239.365 | 9.1312 | 5.1527
232 | 9279.362 | 9.1355 | 5.1536 |  |  | 232 | 9279.362 | 9.1355 | 5.1477 |  |  | 232 | 9279.362 | 9.1355 | 5.1515
233 | 9319.359 | 9.1398 | 5.1524 |  |  | 233 | 9319.359 | 9.1398 | 5.1465 |  |  | 233 | 9319.359 | 9.1398 | 5.1504
234 | 9359.357 | 9.1441 | 5.1512 |  |  | 234 | 9359.357 | 9.1441 | 5.1453 |  |  | 234 | 9359.357 | 9.1441 | 5.1492
235 | 9399.354 | 9.1484 | 5.15 |  |  | 235 | 9399.354 | 9.1484 | 5.1442 |  |  | 235 | 9399.354 | 9.1484 | 5.148
236 | 9439.351 | 9.1526 | 5.1489 |  |  | 236 | 9439.351 | 9.1526 | 5.143 |  |  | 236 | 9439.351 | 9.1526 | 5.1469
237 | 9479.348 | 9.1569 | 5.1477 |  |  | 237 | 9479.348 | 9.1569 | 5.1418 |  |  | 237 | 9479.348 | 9.1569 | 5.1457
238 | 9519.346 | 9.1611 | 5.1466 |  |  | 238 | 9519.346 | 9.1611 | 5.1407 |  |  | 238 | 9519.346 | 9.1611 | 5.1446
239 | 9559.343 | 9.1653 | 5.1454 |  |  | 239 | 9559.343 | 9.1653 | 5.1396 |  |  | 239 | 9559.343 | 9.1653 | 5.1434
240 | 9599.34 | 9.1694 | 5.1443 |  |  | 240 | 9599.34 | 9.1694 | 5.1384 |  |  | 240 | 9599.34 | 9.1694 | 5.1423
241 | 9639.337 | 9.1736 | 5.1432 |  |  | 241 | 9639.337 | 9.1736 | 5.1373 |  |  | 241 | 9639.337 | 9.1736 | 5.1412
242 | 9679.335 | 9.1777 | 5.1421 |  |  | 242 | 9679.335 | 9.1777 | 5.1362 |  |  | 242 | 9679.335 | 9.1777 | 5.1401
243 | 9719.332 | 9.1819 | 5.141 |  |  | 243 | 9719.332 | 9.1819 | 5.1351 |  |  | 243 | 9719.332 | 9.1819 | 5.139
244 | 9759.329 | 9.186 | 5.1399 |  |  | 244 | 9759.329 | 9.186 | 5.134 |  |  | 244 | 9759.329 | 9.186 | 5.1379
245 | 9799.326 | 9.1901 | 5.1388 |  |  | 245 | 9799.326 | 9.1901 | 5.1329 |  |  | 245 | 9799.326 | 9.1901 | 5.1369
246 | 9839.324 | 9.1941 | 5.1377 |  |  | 246 | 9839.324 | 9.1941 | 5.1319 |  |  | 246 | 9839.324 | 9.1941 | 5.1358
247 | 9879.321 | 9.1982 | 5.1367 |  |  | 247 | 9879.321 | 9.1982 | 5.1308 |  |  | 247 | 9879.321 | 9.1982 | 5.1348
248 | 9919.318 | 9.2022 | 5.1356 |  |  | 248 | 9919.318 | 9.2022 | 5.1298 |  |  | 248 | 9919.318 | 9.2022 | 5.1337
249 | 9959.315 | 9.2063 | 5.1346 |  |  | 249 | 9959.315 | 9.2063 | 5.1287 |  |  | 249 | 9959.315 | 9.2063 | 5.1327
250 | 9999.313 | 9.2103 | 5.1335 |  |  | 250 | 9999.313 | 9.2103 | 5.1277 |  |  | 250 | 9999.313 | 9.2103 | 5.1316
251 | 10039.31 | 9.2143 | 5.1325 |  |  | 251 | 10039.31 | 9.2143 | 5.1267 |  |  | 251 | 10039.31 | 9.2143 | 5.1306
252 | 10079.307 | 9.2182 | 5.1315 |  |  | 252 | 10079.307 | 9.2182 | 5.1256 |  |  | 252 | 10079.307 | 9.2182 | 5.1296
253 | 10119.304 | 9.2222 | 5.1305 |  |  | 253 | 10119.304 | 9.2222 | 5.1246 |  |  | 253 | 10119.304 | 9.2222 | 5.1286
254 | 10159.302 | 9.2261 | 5.1295 |  |  | 254 | 10159.302 | 9.2261 | 5.1236 |  |  | 254 | 10159.302 | 9.2261 | 5.1276
255 | 10199.299 | 9.2301 | 5.1285 |  |  | 255 | 10199.299 | 9.2301 | 5.1226 |  |  | 255 | 10199.299 | 9.2301 | 5.1266
256 | 10239.296 | 9.234 | 5.1275 |  |  | 256 | 10239.296 | 9.234 | 5.1217 |  |  | 256 | 10239.296 | 9.234 | 5.1257
257 | 10279.293 | 9.2379 | 5.1265 |  |  | 257 | 10279.293 | 9.2379 | 5.1207 |  |  | 257 | 10279.293 | 9.2379 | 5.1247
258 | 10319.291 | 9.2418 | 5.1255 |  |  | 258 | 10319.291 | 9.2418 | 5.1197 |  |  | 258 | 10319.291 | 9.2418 | 5.1237
259 | 10359.288 | 9.2456 | 5.1245 |  |  | 259 | 10359.288 | 9.2456 | 5.1188 |  |  | 259 | 10359.288 | 9.2456 | 5.1228
260 | 10399.285 | 9.2495 | 5.1236 |  |  | 260 | 10399.285 | 9.2495 | 5.1178 |  |  | 260 | 10399.285 | 9.2495 | 5.1218
261 | 10439.282 | 9.2533 | 5.1226 |  |  | 261 | 10439.282 | 9.2533 | 5.1169 |  |  | 261 | 10439.282 | 9.2533 | 5.1209
262 | 10479.28 | 9.2572 | 5.1217 |  |  | 262 | 10479.28 | 9.2572 | 5.1159 |  |  | 262 | 10479.28 | 9.2572 | 5.12
263 | 10519.277 | 9.261 | 5.1208 |  |  | 263 | 10519.277 | 9.261 | 5.115 |  |  | 263 | 10519.277 | 9.261 | 5.119
264 | 10559.274 | 9.2648 | 5.1198 |  |  | 264 | 10559.274 | 9.2648 | 5.1141 |  |  | 264 | 10559.274 | 9.2648 | 5.1181
265 | 10599.271 | 9.2685 | 5.1189 |  |  | 265 | 10599.271 | 9.2685 | 5.1132 |  |  | 265 | 10599.271 | 9.2685 | 5.1172
266 | 10639.269 | 9.2723 | 5.118 |  |  | 266 | 10639.269 | 9.2723 | 5.1123 |  |  | 266 | 10639.269 | 9.2723 | 5.1163
267 | 10679.266 | 9.2761 | 5.1171 |  |  | 267 | 10679.266 | 9.2761 | 5.1114 |  |  | 267 | 10679.266 | 9.2761 | 5.1154
268 | 10719.263 | 9.2798 | 5.1162 |  |  | 268 | 10719.263 | 9.2798 | 5.1105 |  |  | 268 | 10719.263 | 9.2798 | 5.1145
269 | 10759.26 | 9.2835 | 5.1153 |  |  | 269 | 10759.26 | 9.2835 | 5.1096 |  |  | 269 | 10759.26 | 9.2835 | 5.1137
270 | 10799.258 | 9.2872 | 5.1144 |  |  | 270 | 10799.258 | 9.2872 | 5.1087 |  |  | 270 | 10799.258 | 9.2872 | 5.1128
271 | 10839.255 | 9.2909 | 5.1136 |  |  | 271 | 10839.255 | 9.2909 | 5.1079 |  |  | 271 | 10839.255 | 9.2909 | 5.1119
272 | 10879.252 | 9.2946 | 5.1127 |  |  | 272 | 10879.252 | 9.2946 | 5.107 |  |  | 272 | 10879.252 | 9.2946 | 5.1111
273 | 10919.249 | 9.2983 | 5.1118 |  |  | 273 | 10919.249 | 9.2983 | 5.1062 |  |  | 273 | 10919.249 | 9.2983 | 5.1102
274 | 10959.247 | 9.3019 | 5.111 |  |  | 274 | 10959.247 | 9.3019 | 5.1053 |  |  | 274 | 10959.247 | 9.3019 | 5.1094
275 | 10999.244 | 9.3056 | 5.1101 |  |  | 275 | 10999.244 | 9.3056 | 5.1045 |  |  | 275 | 10999.244 | 9.3056 | 5.1086
276 | 11039.241 | 9.3092 | 5.1093 |  |  | 276 | 11039.241 | 9.3092 | 5.1036 |  |  | 276 | 11039.241 | 9.3092 | 5.1077
277 | 11079.238 | 9.3128 | 5.1084 |  |  | 277 | 11079.238 | 9.3128 | 5.1028 |  |  | 277 | 11079.238 | 9.3128 | 5.1069
278 | 11119.236 | 9.3164 | 5.1076 |  |  | 278 | 11119.236 | 9.3164 | 5.102 |  |  | 278 | 11119.236 | 9.3164 | 5.1061
279 | 11159.233 | 9.32 | 5.1068 |  |  | 279 | 11159.233 | 9.32 | 5.1012 |  |  | 279 | 11159.233 | 9.32 | 5.1053
280 | 11199.23 | 9.3236 | 5.106 |  |  | 280 | 11199.23 | 9.3236 | 5.1004 |  |  | 280 | 11199.23 | 9.3236 | 5.1045
281 | 11239.227 | 9.3272 | 5.1052 |  |  | 281 | 11239.227 | 9.3272 | 5.0996 |  |  | 281 | 11239.227 | 9.3272 | 5.1037
282 | 11279.225 | 9.3307 | 5.1044 |  |  | 282 | 11279.225 | 9.3307 | 5.0988 |  |  | 282 | 11279.225 | 9.3307 | 5.1029
283 | 11319.222 | 9.3343 | 5.1036 |  |  | 283 | 11319.222 | 9.3343 | 5.098 |  |  | 283 | 11319.222 | 9.3343 | 5.1021
284 | 11359.219 | 9.3378 | 5.1028 |  |  | 284 | 11359.219 | 9.3378 | 5.0972 |  |  | 284 | 11359.219 | 9.3378 | 5.1013
285 | 11399.216 | 9.3413 | 5.102 |  |  | 285 | 11399.216 | 9.3413 | 5.0965 |  |  | 285 | 11399.216 | 9.3413 | 5.1006
286 | 11439.214 | 9.3448 | 5.1012 |  |  | 286 | 11439.214 | 9.3448 | 5.0957 |  |  | 286 | 11439.214 | 9.3448 | 5.0998
287 | 11479.211 | 9.3483 | 5.1004 |  |  | 287 | 11479.211 | 9.3483 | 5.0949 |  |  | 287 | 11479.211 | 9.3483 | 5.099
288 | 11519.208 | 9.3518 | 5.0997 |  |  | 288 | 11519.208 | 9.3518 | 5.0942 |  |  | 288 | 11519.208 | 9.3518 | 5.0983
289 | 11559.205 | 9.3552 | 5.0989 |  |  | 289 | 11559.205 | 9.3552 | 5.0934 |  |  | 289 | 11559.205 | 9.3552 | 5.0975
290 | 11599.203 | 9.3587 | 5.0981 |  |  | 290 | 11599.203 | 9.3587 | 5.0927 |  |  | 290 | 11599.203 | 9.3587 | 5.0968
291 | 11639.2 | 9.3621 | 5.0974 |  |  | 291 | 11639.2 | 9.3621 | 5.0919 |  |  | 291 | 11639.2 | 9.3621 | 5.0961
292 | 11679.197 | 9.3656 | 5.0966 |  |  | 292 | 11679.197 | 9.3656 | 5.0912 |  |  | 292 | 11679.197 | 9.3656 | 5.0953
293 | 11719.194 | 9.369 | 5.0959 |  |  | 293 | 11719.194 | 9.369 | 5.0905 |  |  | 293 | 11719.194 | 9.369 | 5.0946
294 | 11759.192 | 9.3724 | 5.0952 |  |  | 294 | 11759.192 | 9.3724 | 5.0898 |  |  | 294 | 11759.192 | 9.3724 | 5.0939
295 | 11799.189 | 9.3758 | 5.0944 |  |  | 295 | 11799.189 | 9.3758 | 5.0891 |  |  | 295 | 11799.189 | 9.3758 | 5.0932
296 | 11839.186 | 9.3792 | 5.0937 |  |  | 296 | 11839.186 | 9.3792 | 5.0883 |  |  | 296 | 11839.186 | 9.3792 | 5.0925
297 | 11879.183 | 9.3825 | 5.093 |  |  | 297 | 11879.183 | 9.3825 | 5.0876 |  |  | 297 | 11879.183 | 9.3825 | 5.0918
298 | 11919.181 | 9.3859 | 5.0923 |  |  | 298 | 11919.181 | 9.3859 | 5.0869 |  |  | 298 | 11919.181 | 9.3859 | 5.0911
299 | 11959.178 | 9.3893 | 5.0916 |  |  | 299 | 11959.178 | 9.3893 | 5.0863 |  |  | 299 | 11959.178 | 9.3893 | 5.0904
300 | 11999.175 | 9.3926 | 5.0909 |  |  | 300 | 11999.175 | 9.3926 | 5.0856 |  |  | 300 | 11999.175 | 9.3926 | 5.0897
301 | 12039.172 | 9.3959 | 5.0902 |  |  | 301 | 12039.172 | 9.3959 | 5.0849 |  |  | 301 | 12039.172 | 9.3959 | 5.089
302 | 12079.17 | 9.3992 | 5.0895 |  |  | 302 | 12079.17 | 9.3992 | 5.0842 |  |  | 302 | 12079.17 | 9.3992 | 5.0884
303 | 12119.167 | 9.4025 | 5.0888 |  |  | 303 | 12119.167 | 9.4025 | 5.0835 |  |  | 303 | 12119.167 | 9.4025 | 5.0877
304 | 12159.164 | 9.4058 | 5.0881 |  |  | 304 | 12159.164 | 9.4058 | 5.0829 |  |  | 304 | 12159.164 | 9.4058 | 5.087
305 | 12199.161 | 9.4091 | 5.0875 |  |  | 305 | 12199.161 | 9.4091 | 5.0822 |  |  | 305 | 12199.161 | 9.4091 | 5.0864
306 | 12239.159 | 9.4124 | 5.0868 |  |  | 306 | 12239.159 | 9.4124 | 5.0816 |  |  | 306 | 12239.159 | 9.4124 | 5.0857
307 | 12279.156 | 9.4157 | 5.0861 |  |  | 307 | 12279.156 | 9.4157 | 5.0809 |  |  | 307 | 12279.156 | 9.4157 | 5.0851
308 | 12319.153 | 9.4189 | 5.0855 |  |  | 308 | 12319.153 | 9.4189 | 5.0803 |  |  | 308 | 12319.153 | 9.4189 | 5.0844
309 | 12359.15 | 9.4222 | 5.0848 |  |  | 309 | 12359.15 | 9.4222 | 5.0796 |  |  | 309 | 12359.15 | 9.4222 | 5.0838
310 | 12399.148 | 9.4254 | 5.0842 |  |  | 310 | 12399.148 | 9.4254 | 5.079 |  |  | 310 | 12399.148 | 9.4254 | 5.0831
311 | 12439.145 | 9.4286 | 5.0835 |  |  | 311 | 12439.145 | 9.4286 | 5.0784 |  |  | 311 | 12439.145 | 9.4286 | 5.0825
312 | 12479.142 | 9.4318 | 5.0829 |  |  | 312 | 12479.142 | 9.4318 | 5.0778 |  |  | 312 | 12479.142 | 9.4318 | 5.0819
313 | 12519.139 | 9.435 | 5.0823 |  |  | 313 | 12519.139 | 9.435 | 5.0771 |  |  | 313 | 12519.139 | 9.435 | 5.0813
314 | 12559.137 | 9.4382 | 5.0816 |  |  | 314 | 12559.137 | 9.4382 | 5.0765 |  |  | 314 | 12559.137 | 9.4382 | 5.0807
315 | 12599.134 | 9.4414 | 5.081 |  |  | 315 | 12599.134 | 9.4414 | 5.0759 |  |  | 315 | 12599.134 | 9.4414 | 5.0801
316 | 12639.131 | 9.4446 | 5.0804 |  |  | 316 | 12639.131 | 9.4446 | 5.0753 |  |  | 316 | 12639.131 | 9.4446 | 5.0794
317 | 12679.128 | 9.4477 | 5.0798 |  |  | 317 | 12679.128 | 9.4477 | 5.0747 |  |  | 317 | 12679.128 | 9.4477 | 5.0788
318 | 12719.126 | 9.4509 | 5.0792 |  |  | 318 | 12719.126 | 9.4509 | 5.0741 |  |  | 318 | 12719.126 | 9.4509 | 5.0783
319 | 12759.123 | 9.454 | 5.0785 |  |  | 319 | 12759.123 | 9.454 | 5.0735 |  |  | 319 | 12759.123 | 9.454 | 5.0777
320 | 12799.12 | 9.4571 | 5.0779 |  |  | 320 | 12799.12 | 9.4571 | 5.0729 |  |  | 320 | 12799.12 | 9.4571 | 5.0771
321 | 12839.117 | 9.4603 | 5.0773 |  |  | 321 | 12839.117 | 9.4603 | 5.0724 |  |  | 321 | 12839.117 | 9.4603 | 5.0765
322 | 12879.115 | 9.4634 | 5.0768 |  |  | 322 | 12879.115 | 9.4634 | 5.0718 |  |  | 322 | 12879.115 | 9.4634 | 5.0759
323 | 12919.112 | 9.4665 | 5.0762 |  |  | 323 | 12919.112 | 9.4665 | 5.0712 |  |  | 323 | 12919.112 | 9.4665 | 5.0753
324 | 12959.109 | 9.4696 | 5.0756 |  |  | 324 | 12959.109 | 9.4696 | 5.0706 |  |  | 324 | 12959.109 | 9.4696 | 5.0748
325 | 12999.106 | 9.4726 | 5.075 |  |  | 325 | 12999.106 | 9.4726 | 5.0701 |  |  | 325 | 12999.106 | 9.4726 | 5.0742
326 | 13039.104 | 9.4757 | 5.0744 |  |  | 326 | 13039.104 | 9.4757 | 5.0695 |  |  | 326 | 13039.104 | 9.4757 | 5.0737
327 | 13079.101 | 9.4788 | 5.0738 |  |  | 327 | 13079.101 | 9.4788 | 5.069 |  |  | 327 | 13079.101 | 9.4788 | 5.0731
328 | 13119.098 | 9.4818 | 5.0733 |  |  | 328 | 13119.098 | 9.4818 | 5.0684 |  |  | 328 | 13119.098 | 9.4818 | 5.0725
329 | 13159.095 | 9.4849 | 5.0727 |  |  | 329 | 13159.095 | 9.4849 | 5.0679 |  |  | 329 | 13159.095 | 9.4849 | 5.072
330 | 13199.093 | 9.4879 | 5.0722 |  |  | 330 | 13199.093 | 9.4879 | 5.0673 |  |  | 330 | 13199.093 | 9.4879 | 5.0715
331 | 13239.09 | 9.4909 | 5.0716 |  |  | 331 | 13239.09 | 9.4909 | 5.0668 |  |  | 331 | 13239.09 | 9.4909 | 5.0709
332 | 13279.087 | 9.4939 | 5.071 |  |  | 332 | 13279.087 | 9.4939 | 5.0662 |  |  | 332 | 13279.087 | 9.4939 | 5.0704
333 | 13319.084 | 9.497 | 5.0705 |  |  | 333 | 13319.084 | 9.497 | 5.0657 |  |  | 333 | 13319.084 | 9.497 | 5.0698
334 | 13359.082 | 9.5 | 5.0699 |  |  | 334 | 13359.082 | 9.5 | 5.0652 |  |  | 334 | 13359.082 | 9.5 | 5.0693
335 | 13399.079 | 9.5029 | 5.0694 |  |  | 335 | 13399.079 | 9.5029 | 5.0647 |  |  | 335 | 13399.079 | 9.5029 | 5.0688
336 | 13439.076 | 9.5059 | 5.0689 |  |  | 336 | 13439.076 | 9.5059 | 5.0641 |  |  | 336 | 13439.076 | 9.5059 | 5.0683
337 | 13479.073 | 9.5089 | 5.0683 |  |  | 337 | 13479.073 | 9.5089 | 5.0636 |  |  | 337 | 13479.073 | 9.5089 | 5.0678
338 | 13519.071 | 9.5119 | 5.0678 |  |  | 338 | 13519.071 | 9.5119 | 5.0631 |  |  | 338 | 13519.071 | 9.5119 | 5.0672
339 | 13559.068 | 9.5148 | 5.0673 |  |  | 339 | 13559.068 | 9.5148 | 5.0626 |  |  | 339 | 13559.068 | 9.5148 | 5.0667
340 | 13599.065 | 9.5178 | 5.0668 |  |  | 340 | 13599.065 | 9.5178 | 5.0621 |  |  | 340 | 13599.065 | 9.5178 | 5.0662
341 | 13639.062 | 9.5207 | 5.0662 |  |  | 341 | 13639.062 | 9.5207 | 5.0616 |  |  | 341 | 13639.062 | 9.5207 | 5.0657
342 | 13679.06 | 9.5236 | 5.0657 |  |  | 342 | 13679.06 | 9.5236 | 5.0611 |  |  | 342 | 13679.06 | 9.5236 | 5.0652
343 | 13719.057 | 9.5265 | 5.0652 |  |  | 343 | 13719.057 | 9.5265 | 5.0606 |  |  | 343 | 13719.057 | 9.5265 | 5.0647
344 | 13759.054 | 9.5295 | 5.0647 |  |  | 344 | 13759.054 | 9.5295 | 5.0601 |  |  | 344 | 13759.054 | 9.5295 | 5.0643
345 | 13799.051 | 9.5324 | 5.0642 |  |  | 345 | 13799.051 | 9.5324 | 5.0596 |  |  | 345 | 13799.051 | 9.5324 | 5.0638
346 | 13839.049 | 9.5352 | 5.0637 |  |  | 346 | 13839.049 | 9.5352 | 5.0592 |  |  | 346 | 13839.049 | 9.5352 | 5.0633
347 | 13879.046 | 9.5381 | 5.0632 |  |  | 347 | 13879.046 | 9.5381 | 5.0587 |  |  | 347 | 13879.046 | 9.5381 | 5.0628
348 | 13919.043 | 9.541 | 5.0627 |  |  | 348 | 13919.043 | 9.541 | 5.0582 |  |  | 348 | 13919.043 | 9.541 | 5.0623
349 | 13959.04 | 9.5439 | 5.0622 |  |  | 349 | 13959.04 | 9.5439 | 5.0577 |  |  | 349 | 13959.04 | 9.5439 | 5.0619
350 | 13999.038 | 9.5467 | 5.0617 |  |  | 350 | 13999.038 | 9.5467 | 5.0573 |  |  | 350 | 13999.038 | 9.5467 | 5.0614
351 | 14039.035 | 9.5496 | 5.0613 |  |  | 351 | 14039.035 | 9.5496 | 5.0568 |  |  | 351 | 14039.035 | 9.5496 | 5.0609
352 | 14079.032 | 9.5524 | 5.0608 |  |  | 352 | 14079.032 | 9.5524 | 5.0563 |  |  | 352 | 14079.032 | 9.5524 | 5.0605
353 | 14119.029 | 9.5553 | 5.0603 |  |  | 353 | 14119.029 | 9.5553 | 5.0559 |  |  | 353 | 14119.029 | 9.5553 | 5.06
354 | 14159.027 | 9.5581 | 5.0598 |  |  | 354 | 14159.027 | 9.5581 | 5.0554 |  |  | 354 | 14159.027 | 9.5581 | 5.0595
355 | 14199.024 | 9.5609 | 5.0594 |  |  | 355 | 14199.024 | 9.5609 | 5.055 |  |  | 355 | 14199.024 | 9.5609 | 5.0591
356 | 14239.021 | 9.5637 | 5.0589 |  |  | 356 | 14239.021 | 9.5637 | 5.0545 |  |  | 356 | 14239.021 | 9.5637 | 5.0586
357 | 14279.018 | 9.5665 | 5.0584 |  |  | 357 | 14279.018 | 9.5665 | 5.0541 |  |  | 357 | 14279.018 | 9.5665 | 5.0582
358 | 14319.016 | 9.5693 | 5.058 |  |  | 358 | 14319.016 | 9.5693 | 5.0536 |  |  | 358 | 14319.016 | 9.5693 | 5.0578
359 | 14359.013 | 9.5721 | 5.0575 |  |  | 359 | 14359.013 | 9.5721 | 5.0532 |  |  | 359 | 14359.013 | 9.5721 | 5.0573
360 | 14399.01 | 9.5749 | 5.0571 |  |  | 360 | 14399.01 | 9.5749 | 5.0528 |  |  | 360 | 14399.01 | 9.5749 | 5.0569
361 | 14439.007 | 9.5777 | 5.0566 |  |  | 361 | 14439.007 | 9.5777 | 5.0523 |  |  | 361 | 14439.007 | 9.5777 | 5.0564
362 | 14479.005 | 9.5805 | 5.0562 |  |  | 362 | 14479.005 | 9.5805 | 5.0519 |  |  | 362 | 14479.005 | 9.5805 | 5.056
363 | 14519.002 | 9.5832 | 5.0557 |  |  | 363 | 14519.002 | 9.5832 | 5.0515 |  |  | 363 | 14519.002 | 9.5832 | 5.0556
364 | 14558.999 | 9.586 | 5.0553 |  |  | 364 | 14558.999 | 9.586 | 5.0511 |  |  | 364 | 14558.999 | 9.586 | 5.0552
365 | 14598.996 | 9.5887 | 5.0548 |  |  | 365 | 14598.996 | 9.5887 | 5.0507 |  |  | 365 | 14598.996 | 9.5887 | 5.0547
366 | 14638.994 | 9.5914 | 5.0544 |  |  | 366 | 14638.994 | 9.5914 | 5.0502 |  |  | 366 | 14638.994 | 9.5914 | 5.0543
367 | 14678.991 | 9.5942 | 5.054 |  |  | 367 | 14678.991 | 9.5942 | 5.0498 |  |  | 367 | 14678.991 | 9.5942 | 5.0539
368 | 14718.988 | 9.5969 | 5.0535 |  |  | 368 | 14718.988 | 9.5969 | 5.0494 |  |  | 368 | 14718.988 | 9.5969 | 5.0535
369 | 14758.985 | 9.5996 | 5.0531 |  |  | 369 | 14758.985 | 9.5996 | 5.049 |  |  | 369 | 14758.985 | 9.5996 | 5.0531
370 | 14798.983 | 9.6023 | 5.0527 |  |  | 370 | 14798.983 | 9.6023 | 5.0486 |  |  | 370 | 14798.983 | 9.6023 | 5.0527
371 | 14838.98 | 9.605 | 5.0523 |  |  | 371 | 14838.98 | 9.605 | 5.0482 |  |  | 371 | 14838.98 | 9.605 | 5.0523
372 | 14878.977 | 9.6077 | 5.0518 |  |  | 372 | 14878.977 | 9.6077 | 5.0478 |  |  | 372 | 14878.977 | 9.6077 | 5.0519
373 | 14918.974 | 9.6104 | 5.0514 |  |  | 373 | 14918.974 | 9.6104 | 5.0474 |  |  | 373 | 14918.974 | 9.6104 | 5.0515
374 | 14958.972 | 9.6131 | 5.051 |  |  | 374 | 14958.972 | 9.6131 | 5.047 |  |  | 374 | 14958.972 | 9.6131 | 5.0511
375 | 14998.969 | 9.6157 | 5.0506 |  |  | 375 | 14998.969 | 9.6157 | 5.0466 |  |  | 375 | 14998.969 | 9.6157 | 5.0507
376 | 15038.966 | 9.6184 | 5.0502 |  |  | 376 | 15038.966 | 9.6184 | 5.0462 |  |  | 376 | 15038.966 | 9.6184 | 5.0503
377 | 15078.963 | 9.6211 | 5.0498 |  |  | 377 | 15078.963 | 9.6211 | 5.0458 |  |  | 377 | 15078.963 | 9.6211 | 5.0499
378 | 15118.961 | 9.6237 | 5.0494 |  |  | 378 | 15118.961 | 9.6237 | 5.0455 |  |  | 378 | 15118.961 | 9.6237 | 5.0495
379 | 15158.958 | 9.6263 | 5.049 |  |  | 379 | 15158.958 | 9.6263 | 5.0451 |  |  | 379 | 15158.958 | 9.6263 | 5.0492
380 | 15198.955 | 9.629 | 5.0486 |  |  | 380 | 15198.955 | 9.629 | 5.0447 |  |  | 380 | 15198.955 | 9.629 | 5.0488
381 | 15238.952 | 9.6316 | 5.0482 |  |  | 381 | 15238.952 | 9.6316 | 5.0443 |  |  | 381 | 15238.952 | 9.6316 | 5.0484
382 | 15278.95 | 9.6342 | 5.0478 |  |  | 382 | 15278.95 | 9.6342 | 5.044 |  |  | 382 | 15278.95 | 9.6342 | 5.048
383 | 15318.947 | 9.6368 | 5.0474 |  |  | 383 | 15318.947 | 9.6368 | 5.0436 |  |  | 383 | 15318.947 | 9.6368 | 5.0477
384 | 15358.944 | 9.6395 | 5.047 |  |  | 384 | 15358.944 | 9.6395 | 5.0432 |  |  | 384 | 15358.944 | 9.6395 | 5.0473
385 | 15398.941 | 9.6421 | 5.0466 |  |  | 385 | 15398.941 | 9.6421 | 5.0429 |  |  | 385 | 15398.941 | 9.6421 | 5.0469
386 | 15438.939 | 9.6446 | 5.0463 |  |  | 386 | 15438.939 | 9.6446 | 5.0425 |  |  | 386 | 15438.939 | 9.6446 | 5.0466
387 | 15478.936 | 9.6472 | 5.0459 |  |  | 387 | 15478.936 | 9.6472 | 5.0421 |  |  | 387 | 15478.936 | 9.6472 | 5.0462
388 | 15518.933 | 9.6498 | 5.0455 |  |  | 388 | 15518.933 | 9.6498 | 5.0418 |  |  | 388 | 15518.933 | 9.6498 | 5.0458
389 | 15558.93 | 9.6524 | 5.0451 |  |  | 389 | 15558.93 | 9.6524 | 5.0414 |  |  | 389 | 15558.93 | 9.6524 | 5.0455
390 | 15598.928 | 9.655 | 5.0448 |  |  | 390 | 15598.928 | 9.655 | 5.0411 |  |  | 390 | 15598.928 | 9.655 | 5.0451
391 | 15638.925 | 9.6575 | 5.0444 |  |  | 391 | 15638.925 | 9.6575 | 5.0407 |  |  | 391 | 15638.925 | 9.6575 | 5.0448
392 | 15678.922 | 9.6601 | 5.044 |  |  | 392 | 15678.922 | 9.6601 | 5.0404 |  |  | 392 | 15678.922 | 9.6601 | 5.0444
393 | 15718.919 | 9.6626 | 5.0437 |  |  | 393 | 15718.919 | 9.6626 | 5.04 |  |  | 393 | 15718.919 | 9.6626 | 5.0441
394 | 15758.917 | 9.6652 | 5.0433 |  |  | 394 | 15758.917 | 9.6652 | 5.0397 |  |  | 394 | 15758.917 | 9.6652 | 5.0437
395 | 15798.914 | 9.6677 | 5.0429 |  |  | 395 | 15798.914 | 9.6677 | 5.0394 |  |  | 395 | 15798.914 | 9.6677 | 5.0434
396 | 15838.911 | 9.6702 | 5.0426 |  |  | 396 | 15838.911 | 9.6702 | 5.039 |  |  | 396 | 15838.911 | 9.6702 | 5.043
397 | 15878.908 | 9.6727 | 5.0422 |  |  | 397 | 15878.908 | 9.6727 | 5.0387 |  |  | 397 | 15878.908 | 9.6727 | 5.0427
398 | 15918.906 | 9.6753 | 5.0419 |  |  | 398 | 15918.906 | 9.6753 | 5.0384 |  |  | 398 | 15918.906 | 9.6753 | 5.0424
399 | 15958.903 | 9.6778 | 5.0415 |  |  | 399 | 15958.903 | 9.6778 | 5.038 |  |  | 399 | 15958.903 | 9.6778 | 5.042
400 | 15998.9 | 9.6803 | 5.0412 |  |  | 400 | 15998.9 | 9.6803 | 5.0377 |  |  | 400 | 15998.9 | 9.6803 | 5.0417
401 | 16038.897 | 9.6828 | 5.0408 |  |  | 401 | 16038.897 | 9.6828 | 5.0374 |  |  | 401 | 16038.897 | 9.6828 | 5.0414
402 | 16078.895 | 9.6853 | 5.0405 |  |  | 402 | 16078.895 | 9.6853 | 5.037 |  |  | 402 | 16078.895 | 9.6853 | 5.0411
403 | 16118.892 | 9.6877 | 5.0401 |  |  | 403 | 16118.892 | 9.6877 | 5.0367 |  |  | 403 | 16118.892 | 9.6877 | 5.0407
404 | 16158.889 | 9.6902 | 5.0398 |  |  | 404 | 16158.889 | 9.6902 | 5.0364 |  |  | 404 | 16158.889 | 9.6902 | 5.0404
405 | 16198.886 | 9.6927 | 5.0395 |  |  | 405 | 16198.886 | 9.6927 | 5.0361 |  |  | 405 | 16198.886 | 9.6927 | 5.0401
406 | 16238.884 | 9.6952 | 5.0391 |  |  | 406 | 16238.884 | 9.6952 | 5.0358 |  |  | 406 | 16238.884 | 9.6952 | 5.0398
407 | 16278.881 | 9.6976 | 5.0388 |  |  | 407 | 16278.881 | 9.6976 | 5.0355 |  |  | 407 | 16278.881 | 9.6976 | 5.0395
408 | 16318.878 | 9.7001 | 5.0385 |  |  | 408 | 16318.878 | 9.7001 | 5.0351 |  |  | 408 | 16318.878 | 9.7001 | 5.0391
409 | 16358.875 | 9.7025 | 5.0381 |  |  | 409 | 16358.875 | 9.7025 | 5.0348 |  |  | 409 | 16358.875 | 9.7025 | 5.0388
410 | 16398.873 | 9.705 | 5.0378 |  |  | 410 | 16398.873 | 9.705 | 5.0345 |  |  | 410 | 16398.873 | 9.705 | 5.0385
411 | 16438.87 | 9.7074 | 5.0375 |  |  | 411 | 16438.87 | 9.7074 | 5.0342 |  |  | 411 | 16438.87 | 9.7074 | 5.0382
412 | 16478.867 | 9.7098 | 5.0371 |  |  | 412 | 16478.867 | 9.7098 | 5.0339 |  |  | 412 | 16478.867 | 9.7098 | 5.0379
413 | 16518.864 | 9.7123 | 5.0368 |  |  | 413 | 16518.864 | 9.7123 | 5.0336 |  |  | 413 | 16518.864 | 9.7123 | 5.0376
414 | 16558.862 | 9.7147 | 5.0365 |  |  | 414 | 16558.862 | 9.7147 | 5.0333 |  |  | 414 | 16558.862 | 9.7147 | 5.0373
415 | 16598.859 | 9.7171 | 5.0362 |  |  | 415 | 16598.859 | 9.7171 | 5.033 |  |  | 415 | 16598.859 | 9.7171 | 5.037
416 | 16638.856 | 9.7195 | 5.0359 |  |  | 416 | 16638.856 | 9.7195 | 5.0327 |  |  | 416 | 16638.856 | 9.7195 | 5.0367
417 | 16678.853 | 9.7219 | 5.0356 |  |  | 417 | 16678.853 | 9.7219 | 5.0324 |  |  | 417 | 16678.853 | 9.7219 | 5.0364
418 | 16718.851 | 9.7243 | 5.0352 |  |  | 418 | 16718.851 | 9.7243 | 5.0321 |  |  | 418 | 16718.851 | 9.7243 | 5.0361
419 | 16758.848 | 9.7267 | 5.0349 |  |  | 419 | 16758.848 | 9.7267 | 5.0319 |  |  | 419 | 16758.848 | 9.7267 | 5.0358
420 | 16798.845 | 9.7291 | 5.0346 |  |  | 420 | 16798.845 | 9.7291 | 5.0316 |  |  | 420 | 16798.845 | 9.7291 | 5.0355
421 | 16838.842 | 9.7314 | 5.0343 |  |  | 421 | 16838.842 | 9.7314 | 5.0313 |  |  | 421 | 16838.842 | 9.7314 | 5.0352
422 | 16878.84 | 9.7338 | 5.034 |  |  | 422 | 16878.84 | 9.7338 | 5.031 |  |  | 422 | 16878.84 | 9.7338 | 5.035
423 | 16918.837 | 9.7362 | 5.0337 |  |  | 423 | 16918.837 | 9.7362 | 5.0307 |  |  | 423 | 16918.837 | 9.7362 | 5.0347
424 | 16958.834 | 9.7385 | 5.0334 |  |  | 424 | 16958.834 | 9.7385 | 5.0304 |  |  | 424 | 16958.834 | 9.7385 | 5.0344
425 | 16998.831 | 9.7409 | 5.0331 |  |  | 425 | 16998.831 | 9.7409 | 5.0302 |  |  | 425 | 16998.831 | 9.7409 | 5.0341
426 | 17038.829 | 9.7433 | 5.0328 |  |  | 426 | 17038.829 | 9.7433 | 5.0299 |  |  | 426 | 17038.829 | 9.7433 | 5.0338
427 | 17078.826 | 9.7456 | 5.0325 |  |  | 427 | 17078.826 | 9.7456 | 5.0296 |  |  | 427 | 17078.826 | 9.7456 | 5.0336
428 | 17118.823 | 9.7479 | 5.0322 |  |  | 428 | 17118.823 | 9.7479 | 5.0293 |  |  | 428 | 17118.823 | 9.7479 | 5.0333
429 | 17158.82 | 9.7503 | 5.0319 |  |  | 429 | 17158.82 | 9.7503 | 5.0291 |  |  | 429 | 17158.82 | 9.7503 | 5.033
430 | 17198.818 | 9.7526 | 5.0316 |  |  | 430 | 17198.818 | 9.7526 | 5.0288 |  |  | 430 | 17198.818 | 9.7526 | 5.0327
431 | 17238.815 | 9.7549 | 5.0314 |  |  | 431 | 17238.815 | 9.7549 | 5.0285 |  |  | 431 | 17238.815 | 9.7549 | 5.0325
432 | 17278.812 | 9.7572 | 5.0311 |  |  | 432 | 17278.812 | 9.7572 | 5.0283 |  |  | 432 | 17278.812 | 9.7572 | 5.0322
433 | 17318.809 | 9.7595 | 5.0308 |  |  | 433 | 17318.809 | 9.7595 | 5.028 |  |  | 433 | 17318.809 | 9.7595 | 5.0319
434 | 17358.807 | 9.7619 | 5.0305 |  |  | 434 | 17358.807 | 9.7619 | 5.0277 |  |  | 434 | 17358.807 | 9.7619 | 5.0317
435 | 17398.804 | 9.7642 | 5.0302 |  |  | 435 | 17398.804 | 9.7642 | 5.0275 |  |  | 435 | 17398.804 | 9.7642 | 5.0314
436 | 17438.801 | 9.7665 | 5.0299 |  |  | 436 | 17438.801 | 9.7665 | 5.0272 |  |  | 436 | 17438.801 | 9.7665 | 5.0311
437 | 17478.798 | 9.7687 | 5.0297 |  |  | 437 | 17478.798 | 9.7687 | 5.0269 |  |  | 437 | 17478.798 | 9.7687 | 5.0309
438 | 17518.796 | 9.771 | 5.0294 |  |  | 438 | 17518.796 | 9.771 | 5.0267 |  |  | 438 | 17518.796 | 9.771 | 5.0306
439 | 17558.793 | 9.7733 | 5.0291 |  |  | 439 | 17558.793 | 9.7733 | 5.0264 |  |  | 439 | 17558.793 | 9.7733 | 5.0304
440 | 17598.79 | 9.7756 | 5.0288 |  |  | 440 | 17598.79 | 9.7756 | 5.0262 |  |  | 440 | 17598.79 | 9.7756 | 5.0301
441 | 17638.787 | 9.7779 | 5.0286 |  |  | 441 | 17638.787 | 9.7779 | 5.0259 |  |  | 441 | 17638.787 | 9.7779 | 5.0298
442 | 17678.785 | 9.7801 | 5.0283 |  |  | 442 | 17678.785 | 9.7801 | 5.0257 |  |  | 442 | 17678.785 | 9.7801 | 5.0296
443 | 17718.782 | 9.7824 | 5.028 |  |  | 443 | 17718.782 | 9.7824 | 5.0254 |  |  | 443 | 17718.782 | 9.7824 | 5.0293
444 | 17758.779 | 9.7846 | 5.0277 |  |  | 444 | 17758.779 | 9.7846 | 5.0252 |  |  | 444 | 17758.779 | 9.7846 | 5.0291
445 | 17798.776 | 9.7869 | 5.0275 |  |  | 445 | 17798.776 | 9.7869 | 5.0249 |  |  | 445 | 17798.776 | 9.7869 | 5.0288
446 | 17838.774 | 9.7891 | 5.0272 |  |  | 446 | 17838.774 | 9.7891 | 5.0247 |  |  | 446 | 17838.774 | 9.7891 | 5.0286
447 | 17878.771 | 9.7914 | 5.0269 |  |  | 447 | 17878.771 | 9.7914 | 5.0244 |  |  | 447 | 17878.771 | 9.7914 | 5.0283
448 | 17918.768 | 9.7936 | 5.0267 |  |  | 448 | 17918.768 | 9.7936 | 5.0242 |  |  | 448 | 17918.768 | 9.7936 | 5.0281
449 | 17958.765 | 9.7958 | 5.0264 |  |  | 449 | 17958.765 | 9.7958 | 5.024 |  |  | 449 | 17958.765 | 9.7958 | 5.0279
450 | 17998.763 | 9.7981 | 5.0262 |  |  | 450 | 17998.763 | 9.7981 | 5.0237 |  |  | 450 | 17998.763 | 9.7981 | 5.0276
451 | 18038.76 | 9.8003 | 5.0259 |  |  | 451 | 18038.76 | 9.8003 | 5.0235 |  |  | 451 | 18038.76 | 9.8003 | 5.0274
452 | 18078.757 | 9.8025 | 5.0257 |  |  | 452 | 18078.757 | 9.8025 | 5.0232 |  |  | 452 | 18078.757 | 9.8025 | 5.0271
453 | 18118.755 | 9.8047 | 5.0254 |  |  | 453 | 18118.755 | 9.8047 | 5.023 |  |  | 453 | 18118.755 | 9.8047 | 5.0269
454 | 18158.752 | 9.8069 | 5.0251 |  |  | 454 | 18158.752 | 9.8069 | 5.0228 |  |  | 454 | 18158.752 | 9.8069 | 5.0267
455 | 18198.749 | 9.8091 | 5.0249 |  |  | 455 | 18198.749 | 9.8091 | 5.0225 |  |  | 455 | 18198.749 | 9.8091 | 5.0264
456 | 18238.746 | 9.8113 | 5.0246 |  |  | 456 | 18238.746 | 9.8113 | 5.0223 |  |  | 456 | 18238.746 | 9.8113 | 5.0262
457 | 18278.744 | 9.8135 | 5.0244 |  |  | 457 | 18278.744 | 9.8135 | 5.0221 |  |  | 457 | 18278.744 | 9.8135 | 5.026
458 | 18318.741 | 9.8157 | 5.0241 |  |  | 458 | 18318.741 | 9.8157 | 5.0219 |  |  | 458 | 18318.741 | 9.8157 | 5.0257
459 | 18358.738 | 9.8179 | 5.0239 |  |  | 459 | 18358.738 | 9.8179 | 5.0216 |  |  | 459 | 18358.738 | 9.8179 | 5.0255
460 | 18398.735 | 9.82 | 5.0237 |  |  | 460 | 18398.735 | 9.82 | 5.0214 |  |  | 460 | 18398.735 | 9.82 | 5.0253
461 | 18438.733 | 9.8222 | 5.0234 |  |  | 461 | 18438.733 | 9.8222 | 5.0212 |  |  | 461 | 18438.733 | 9.8222 | 5.0251
462 | 18478.73 | 9.8244 | 5.0232 |  |  | 462 | 18478.73 | 9.8244 | 5.021 |  |  | 462 | 18478.73 | 9.8244 | 5.0248
463 | 18518.727 | 9.8265 | 5.0229 |  |  | 463 | 18518.727 | 9.8265 | 5.0207 |  |  | 463 | 18518.727 | 9.8265 | 5.0246
464 | 18558.724 | 9.8287 | 5.0227 |  |  | 464 | 18558.724 | 9.8287 | 5.0205 |  |  | 464 | 18558.724 | 9.8287 | 5.0244
465 | 18598.722 | 9.8308 | 5.0224 |  |  | 465 | 18598.722 | 9.8308 | 5.0203 |  |  | 465 | 18598.722 | 9.8308 | 5.0242
466 | 18638.719 | 9.833 | 5.0222 |  |  | 466 | 18638.719 | 9.833 | 5.0201 |  |  | 466 | 18638.719 | 9.833 | 5.0239
467 | 18678.716 | 9.8351 | 5.022 |  |  | 467 | 18678.716 | 9.8351 | 5.0199 |  |  | 467 | 18678.716 | 9.8351 | 5.0237
468 | 18718.713 | 9.8373 | 5.0217 |  |  | 468 | 18718.713 | 9.8373 | 5.0196 |  |  | 468 | 18718.713 | 9.8373 | 5.0235
469 | 18758.711 | 9.8394 | 5.0215 |  |  | 469 | 18758.711 | 9.8394 | 5.0194 |  |  | 469 | 18758.711 | 9.8394 | 5.0233
470 | 18798.708 | 9.8415 | 5.0213 |  |  | 470 | 18798.708 | 9.8415 | 5.0192 |  |  | 470 | 18798.708 | 9.8415 | 5.0231
471 | 18838.705 | 9.8437 | 5.021 |  |  | 471 | 18838.705 | 9.8437 | 5.019 |  |  | 471 | 18838.705 | 9.8437 | 5.0229
472 | 18878.702 | 9.8458 | 5.0208 |  |  | 472 | 18878.702 | 9.8458 | 5.0188 |  |  | 472 | 18878.702 | 9.8458 | 5.0227
473 | 18918.7 | 9.8479 | 5.0206 |  |  | 473 | 18918.7 | 9.8479 | 5.0186 |  |  | 473 | 18918.7 | 9.8479 | 5.0224
474 | 18958.697 | 9.85 | 5.0203 |  |  | 474 | 18958.697 | 9.85 | 5.0184 |  |  | 474 | 18958.697 | 9.85 | 5.0222
475 | 18998.694 | 9.8521 | 5.0201 |  |  | 475 | 18998.694 | 9.8521 | 5.0182 |  |  | 475 | 18998.694 | 9.8521 | 5.022
476 | 19038.691 | 9.8542 | 5.0199 |  |  | 476 | 19038.691 | 9.8542 | 5.018 |  |  | 476 | 19038.691 | 9.8542 | 5.0218
477 | 19078.689 | 9.8563 | 5.0197 |  |  | 477 | 19078.689 | 9.8563 | 5.0178 |  |  | 477 | 19078.689 | 9.8563 | 5.0216
478 | 19118.686 | 9.8584 | 5.0194 |  |  | 478 | 19118.686 | 9.8584 | 5.0176 |  |  | 478 | 19118.686 | 9.8584 | 5.0214
479 | 19158.683 | 9.8605 | 5.0192 |  |  | 479 | 19158.683 | 9.8605 | 5.0174 |  |  | 479 | 19158.683 | 9.8605 | 5.0212
480 | 19198.68 | 9.8626 | 5.019 |  |  | 480 | 19198.68 | 9.8626 | 5.0172 |  |  | 480 | 19198.68 | 9.8626 | 5.021
481 | 19238.678 | 9.8647 | 5.0188 |  |  | 481 | 19238.678 | 9.8647 | 5.017 |  |  | 481 | 19238.678 | 9.8647 | 5.0208
482 | 19278.675 | 9.8668 | 5.0186 |  |  | 482 | 19278.675 | 9.8668 | 5.0168 |  |  | 482 | 19278.675 | 9.8668 | 5.0206
483 | 19318.672 | 9.8688 | 5.0183 |  |  | 483 | 19318.672 | 9.8688 | 5.0166 |  |  | 483 | 19318.672 | 9.8688 | 5.0204
484 | 19358.669 | 9.8709 | 5.0181 |  |  | 484 | 19358.669 | 9.8709 | 5.0164 |  |  | 484 | 19358.669 | 9.8709 | 5.0202
485 | 19398.667 | 9.873 | 5.0179 |  |  | 485 | 19398.667 | 9.873 | 5.0162 |  |  | 485 | 19398.667 | 9.873 | 5.02
486 | 19438.664 | 9.875 | 5.0177 |  |  | 486 | 19438.664 | 9.875 | 5.016 |  |  | 486 | 19438.664 | 9.875 | 5.0198
487 | 19478.661 | 9.8771 | 5.0175 |  |  | 487 | 19478.661 | 9.8771 | 5.0158 |  |  | 487 | 19478.661 | 9.8771 | 5.0196
488 | 19518.658 | 9.8791 | 5.0173 |  |  | 488 | 19518.658 | 9.8791 | 5.0156 |  |  | 488 | 19518.658 | 9.8791 | 5.0194
489 | 19558.656 | 9.8812 | 5.0171 |  |  | 489 | 19558.656 | 9.8812 | 5.0154 |  |  | 489 | 19558.656 | 9.8812 | 5.0192
490 | 19598.653 | 9.8832 | 5.0169 |  |  | 490 | 19598.653 | 9.8832 | 5.0152 |  |  | 490 | 19598.653 | 9.8832 | 5.019
491 | 19638.65 | 9.8853 | 5.0166 |  |  | 491 | 19638.65 | 9.8853 | 5.015 |  |  | 491 | 19638.65 | 9.8853 | 5.0188
492 | 19678.647 | 9.8873 | 5.0164 |  |  | 492 | 19678.647 | 9.8873 | 5.0148 |  |  | 492 | 19678.647 | 9.8873 | 5.0186
493 | 19718.645 | 9.8893 | 5.0162 |  |  | 493 | 19718.645 | 9.8893 | 5.0146 |  |  | 493 | 19718.645 | 9.8893 | 5.0184
494 | 19758.642 | 9.8913 | 5.016 |  |  | 494 | 19758.642 | 9.8913 | 5.0144 |  |  | 494 | 19758.642 | 9.8913 | 5.0183
495 | 19798.639 | 9.8934 | 5.0158 |  |  | 495 | 19798.639 | 9.8934 | 5.0143 |  |  | 495 | 19798.639 | 9.8934 | 5.0181
496 | 19838.636 | 9.8954 | 5.0156 |  |  | 496 | 19838.636 | 9.8954 | 5.0141 |  |  | 496 | 19838.636 | 9.8954 | 5.0179
497 | 19878.634 | 9.8974 | 5.0154 |  |  | 497 | 19878.634 | 9.8974 | 5.0139 |  |  | 497 | 19878.634 | 9.8974 | 5.0177
498 | 19918.631 | 9.8994 | 5.0152 |  |  | 498 | 19918.631 | 9.8994 | 5.0137 |  |  | 498 | 19918.631 | 9.8994 | 5.0175
499 | 19958.628 | 9.9014 | 5.015 |  |  | 499 | 19958.628 | 9.9014 | 5.0135 |  |  | 499 | 19958.628 | 9.9014 | 5.0173
500 | 19998.625 | 9.9034 | 5.0148 |  |  | 500 | 19998.625 | 9.9034 | 5.0133 |  |  | 500 | 19998.625 | 9.9034 | 5.0171
501 | 20038.623 | 9.9054 | 5.0146 |  |  | 501 | 20038.623 | 9.9054 | 5.0132 |  |  | 501 | 20038.623 | 9.9054 | 5.017
502 | 20078.62 | 9.9074 | 5.0144 |  |  | 502 | 20078.62 | 9.9074 | 5.013 |  |  | 502 | 20078.62 | 9.9074 | 5.0168
503 | 20118.617 | 9.9094 | 5.0142 |  |  | 503 | 20118.617 | 9.9094 | 5.0128 |  |  | 503 | 20118.617 | 9.9094 | 5.0166
504 | 20158.614 | 9.9114 | 5.014 |  |  | 504 | 20158.614 | 9.9114 | 5.0126 |  |  | 504 | 20158.614 | 9.9114 | 5.0164
505 | 20198.612 | 9.9134 | 5.0138 |  |  | 505 | 20198.612 | 9.9134 | 5.0124 |  |  | 505 | 20198.612 | 9.9134 | 5.0162
506 | 20238.609 | 9.9153 | 5.0136 |  |  | 506 | 20238.609 | 9.9153 | 5.0123 |  |  | 506 | 20238.609 | 9.9153 | 5.0161
507 | 20278.606 | 9.9173 | 5.0134 |  |  | 507 | 20278.606 | 9.9173 | 5.0121 |  |  | 507 | 20278.606 | 9.9173 | 5.0159
508 | 20318.603 | 9.9193 | 5.0132 |  |  | 508 | 20318.603 | 9.9193 | 5.0119 |  |  | 508 | 20318.603 | 9.9193 | 5.0157
509 | 20358.601 | 9.9213 | 5.013 |  |  | 509 | 20358.601 | 9.9213 | 5.0117 |  |  | 509 | 20358.601 | 9.9213 | 5.0155
510 | 20398.598 | 9.9232 | 5.0128 |  |  | 510 | 20398.598 | 9.9232 | 5.0116 |  |  | 510 | 20398.598 | 9.9232 | 5.0154
511 | 20438.595 | 9.9252 | 5.0127 |  |  | 511 | 20438.595 | 9.9252 | 5.0114 |  |  | 511 | 20438.595 | 9.9252 | 5.0152
512 | 20478.592 | 9.9271 | 5.0125 |  |  | 512 | 20478.592 | 9.9271 | 5.0112 |  |  | 512 | 20478.592 | 9.9271 | 5.015
513 | 20518.59 | 9.9291 | 5.0123 |  |  | 513 | 20518.59 | 9.9291 | 5.0111 |  |  | 513 | 20518.59 | 9.9291 | 5.0149
514 | 20558.587 | 9.931 | 5.0121 |  |  | 514 | 20558.587 | 9.931 | 5.0109 |  |  | 514 | 20558.587 | 9.931 | 5.0147
515 | 20598.584 | 9.933 | 5.0119 |  |  | 515 | 20598.584 | 9.933 | 5.0107 |  |  | 515 | 20598.584 | 9.933 | 5.0145
516 | 20638.581 | 9.9349 | 5.0117 |  |  | 516 | 20638.581 | 9.9349 | 5.0106 |  |  | 516 | 20638.581 | 9.9349 | 5.0143
517 | 20678.579 | 9.9369 | 5.0115 |  |  | 517 | 20678.579 | 9.9369 | 5.0104 |  |  | 517 | 20678.579 | 9.9369 | 5.0142
518 | 20718.576 | 9.9388 | 5.0113 |  |  | 518 | 20718.576 | 9.9388 | 5.0102 |  |  | 518 | 20718.576 | 9.9388 | 5.014
519 | 20758.573 | 9.9407 | 5.0112 |  |  | 519 | 20758.573 | 9.9407 | 5.0101 |  |  | 519 | 20758.573 | 9.9407 | 5.0138
520 | 20798.57 | 9.9426 | 5.011 |  |  | 520 | 20798.57 | 9.9426 | 5.0099 |  |  | 520 | 20798.57 | 9.9426 | 5.0137
521 | 20838.568 | 9.9446 | 5.0108 |  |  | 521 | 20838.568 | 9.9446 | 5.0097 |  |  | 521 | 20838.568 | 9.9446 | 5.0135
522 | 20878.565 | 9.9465 | 5.0106 |  |  | 522 | 20878.565 | 9.9465 | 5.0096 |  |  | 522 | 20878.565 | 9.9465 | 5.0133
523 | 20918.562 | 9.9484 | 5.0104 |  |  | 523 | 20918.562 | 9.9484 | 5.0094 |  |  | 523 | 20918.562 | 9.9484 | 5.0132
524 | 20958.559 | 9.9503 | 5.0103 |  |  | 524 | 20958.559 | 9.9503 | 5.0092 |  |  | 524 | 20958.559 | 9.9503 | 5.013
525 | 20998.557 | 9.9522 | 5.0101 |  |  | 525 | 20998.557 | 9.9522 | 5.0091 |  |  | 525 | 20998.557 | 9.9522 | 5.0129
526 | 21038.554 | 9.9541 | 5.0099 |  |  | 526 | 21038.554 | 9.9541 | 5.0089 |  |  | 526 | 21038.554 | 9.9541 | 5.0127
527 | 21078.551 | 9.956 | 5.0097 |  |  | 527 | 21078.551 | 9.956 | 5.0088 |  |  | 527 | 21078.551 | 9.956 | 5.0125
528 | 21118.548 | 9.9579 | 5.0095 |  |  | 528 | 21118.548 | 9.9579 | 5.0086 |  |  | 528 | 21118.548 | 9.9579 | 5.0124
529 | 21158.546 | 9.9598 | 5.0094 |  |  | 529 | 21158.546 | 9.9598 | 5.0084 |  |  | 529 | 21158.546 | 9.9598 | 5.0122
530 | 21198.543 | 9.9617 | 5.0092 |  |  | 530 | 21198.543 | 9.9617 | 5.0083 |  |  | 530 | 21198.543 | 9.9617 | 5.0121
531 | 21238.54 | 9.9636 | 5.009 |  |  | 531 | 21238.54 | 9.9636 | 5.0081 |  |  | 531 | 21238.54 | 9.9636 | 5.0119
532 | 21278.537 | 9.9655 | 5.0088 |  |  | 532 | 21278.537 | 9.9655 | 5.008 |  |  | 532 | 21278.537 | 9.9655 | 5.0117
533 | 21318.535 | 9.9673 | 5.0087 |  |  | 533 | 21318.535 | 9.9673 | 5.0078 |  |  | 533 | 21318.535 | 9.9673 | 5.0116
534 | 21358.532 | 9.9692 | 5.0085 |  |  | 534 | 21358.532 | 9.9692 | 5.0077 |  |  | 534 | 21358.532 | 9.9692 | 5.0114
535 | 21398.529 | 9.9711 | 5.0083 |  |  | 535 | 21398.529 | 9.9711 | 5.0075 |  |  | 535 | 21398.529 | 9.9711 | 5.0113
536 | 21438.526 | 9.9729 | 5.0081 |  |  | 536 | 21438.526 | 9.9729 | 5.0073 |  |  | 536 | 21438.526 | 9.9729 | 5.0111
537 | 21478.524 | 9.9748 | 5.008 |  |  | 537 | 21478.524 | 9.9748 | 5.0072 |  |  | 537 | 21478.524 | 9.9748 | 5.011
538 | 21518.521 | 9.9767 | 5.0078 |  |  | 538 | 21518.521 | 9.9767 | 5.007 |  |  | 538 | 21518.521 | 9.9767 | 5.0108
539 | 21558.518 | 9.9785 | 5.0076 |  |  | 539 | 21558.518 | 9.9785 | 5.0069 |  |  | 539 | 21558.518 | 9.9785 | 5.0107
540 | 21598.515 | 9.9804 | 5.0075 |  |  | 540 | 21598.515 | 9.9804 | 5.0067 |  |  | 540 | 21598.515 | 9.9804 | 5.0105
541 | 21638.513 | 9.9822 | 5.0073 |  |  | 541 | 21638.513 | 9.9822 | 5.0066 |  |  | 541 | 21638.513 | 9.9822 | 5.0104
542 | 21678.51 | 9.9841 | 5.0071 |  |  | 542 | 21678.51 | 9.9841 | 5.0064 |  |  | 542 | 21678.51 | 9.9841 | 5.0102
543 | 21718.507 | 9.9859 | 5.007 |  |  | 543 | 21718.507 | 9.9859 | 5.0063 |  |  | 543 | 21718.507 | 9.9859 | 5.0101
544 | 21758.504 | 9.9878 | 5.0068 |  |  | 544 | 21758.504 | 9.9878 | 5.0061 |  |  | 544 | 21758.504 | 9.9878 | 5.0099
545 | 21798.502 | 9.9896 | 5.0066 |  |  | 545 | 21798.502 | 9.9896 | 5.006 |  |  | 545 | 21798.502 | 9.9896 | 5.0098
546 | 21838.499 | 9.9914 | 5.0065 |  |  | 546 | 21838.499 | 9.9914 | 5.0058 |  |  | 546 | 21838.499 | 9.9914 | 5.0096
547 | 21878.496 | 9.9933 | 5.0063 |  |  | 547 | 21878.496 | 9.9933 | 5.0057 |  |  | 547 | 21878.496 | 9.9933 | 5.0095
548 | 21918.493 | 9.9951 | 5.0061 |  |  | 548 | 21918.493 | 9.9951 | 5.0055 |  |  | 548 | 21918.493 | 9.9951 | 5.0093
549 | 21958.491 | 9.9969 | 5.006 |  |  | 549 | 21958.491 | 9.9969 | 5.0054 |  |  | 549 | 21958.491 | 9.9969 | 5.0092
550 | 21998.488 | 9.9987 | 5.0058 |  |  | 550 | 21998.488 | 9.9987 | 5.0053 |  |  | 550 | 21998.488 | 9.9987 | 5.009
551 | 22038.485 | 10.0005 | 5.0057 |  |  | 551 | 22038.485 | 10.0005 | 5.0051 |  |  | 551 | 22038.485 | 10.0005 | 5.0089
552 | 22078.482 | 10.0024 | 5.0055 |  |  | 552 | 22078.482 | 10.0024 | 5.005 |  |  | 552 | 22078.482 | 10.0024 | 5.0087
553 | 22118.48 | 10.0042 | 5.0053 |  |  | 553 | 22118.48 | 10.0042 | 5.0048 |  |  | 553 | 22118.48 | 10.0042 | 5.0086
554 | 22158.477 | 10.006 | 5.0052 |  |  | 554 | 22158.477 | 10.006 | 5.0047 |  |  | 554 | 22158.477 | 10.006 | 5.0085
555 | 22198.474 | 10.0078 | 5.005 |  |  | 555 | 22198.474 | 10.0078 | 5.0045 |  |  | 555 | 22198.474 | 10.0078 | 5.0083
556 | 22238.471 | 10.0096 | 5.0049 |  |  | 556 | 22238.471 | 10.0096 | 5.0044 |  |  | 556 | 22238.471 | 10.0096 | 5.0082
557 | 22278.469 | 10.0114 | 5.0047 |  |  | 557 | 22278.469 | 10.0114 | 5.0043 |  |  | 557 | 22278.469 | 10.0114 | 5.008
558 | 22318.466 | 10.0132 | 5.0045 |  |  | 558 | 22318.466 | 10.0132 | 5.0041 |  |  | 558 | 22318.466 | 10.0132 | 5.0079
559 | 22358.463 | 10.015 | 5.0044 |  |  | 559 | 22358.463 | 10.015 | 5.004 |  |  | 559 | 22358.463 | 10.015 | 5.0077
560 | 22398.46 | 10.0167 | 5.0042 |  |  | 560 | 22398.46 | 10.0167 | 5.0038 |  |  | 560 | 22398.46 | 10.0167 | 5.0076
561 | 22438.458 | 10.0185 | 5.0041 |  |  | 561 | 22438.458 | 10.0185 | 5.0037 |  |  | 561 | 22438.458 | 10.0185 | 5.0075
562 | 22478.455 | 10.0203 | 5.0039 |  |  | 562 | 22478.455 | 10.0203 | 5.0036 |  |  | 562 | 22478.455 | 10.0203 | 5.0073
563 | 22518.452 | 10.0221 | 5.0038 |  |  | 563 | 22518.452 | 10.0221 | 5.0034 |  |  | 563 | 22518.452 | 10.0221 | 5.0072
564 | 22558.449 | 10.0239 | 5.0036 |  |  | 564 | 22558.449 | 10.0239 | 5.0033 |  |  | 564 | 22558.449 | 10.0239 | 5.0071
565 | 22598.447 | 10.0256 | 5.0035 |  |  | 565 | 22598.447 | 10.0256 | 5.0031 |  |  | 565 | 22598.447 | 10.0256 | 5.0069
566 | 22638.444 | 10.0274 | 5.0033 |  |  | 566 | 22638.444 | 10.0274 | 5.003 |  |  | 566 | 22638.444 | 10.0274 | 5.0068
567 | 22678.441 | 10.0292 | 5.0031 |  |  | 567 | 22678.441 | 10.0292 | 5.0029 |  |  | 567 | 22678.441 | 10.0292 | 5.0066
568 | 22718.438 | 10.0309 | 5.003 |  |  | 568 | 22718.438 | 10.0309 | 5.0027 |  |  | 568 | 22718.438 | 10.0309 | 5.0065
569 | 22758.436 | 10.0327 | 5.0028 |  |  | 569 | 22758.436 | 10.0327 | 5.0026 |  |  | 569 | 22758.436 | 10.0327 | 5.0064
570 | 22798.433 | 10.0344 | 5.0027 |  |  | 570 | 22798.433 | 10.0344 | 5.0025 |  |  | 570 | 22798.433 | 10.0344 | 5.0062
571 | 22838.43 | 10.0362 | 5.0025 |  |  | 571 | 22838.43 | 10.0362 | 5.0023 |  |  | 571 | 22838.43 | 10.0362 | 5.0061
572 | 22878.427 | 10.0379 | 5.0024 |  |  | 572 | 22878.427 | 10.0379 | 5.0022 |  |  | 572 | 22878.427 | 10.0379 | 5.006
573 | 22918.425 | 10.0397 | 5.0022 |  |  | 573 | 22918.425 | 10.0397 | 5.0021 |  |  | 573 | 22918.425 | 10.0397 | 5.0058
574 | 22958.422 | 10.0414 | 5.0021 |  |  | 574 | 22958.422 | 10.0414 | 5.0019 |  |  | 574 | 22958.422 | 10.0414 | 5.0057
575 | 22998.419 | 10.0432 | 5.0019 |  |  | 575 | 22998.419 | 10.0432 | 5.0018 |  |  | 575 | 22998.419 | 10.0432 | 5.0056
576 | 23038.416 | 10.0449 | 5.0018 |  |  | 576 | 23038.416 | 10.0449 | 5.0017 |  |  | 576 | 23038.416 | 10.0449 | 5.0054
577 | 23078.414 | 10.0467 | 5.0017 |  |  | 577 | 23078.414 | 10.0467 | 5.0015 |  |  | 577 | 23078.414 | 10.0467 | 5.0053
578 | 23118.411 | 10.0484 | 5.0015 |  |  | 578 | 23118.411 | 10.0484 | 5.0014 |  |  | 578 | 23118.411 | 10.0484 | 5.0052
579 | 23158.408 | 10.0501 | 5.0014 |  |  | 579 | 23158.408 | 10.0501 | 5.0013 |  |  | 579 | 23158.408 | 10.0501 | 5.0051
580 | 23198.405 | 10.0518 | 5.0012 |  |  | 580 | 23198.405 | 10.0518 | 5.0011 |  |  | 580 | 23198.405 | 10.0518 | 5.0049
581 | 23238.403 | 10.0536 | 5.0011 |  |  | 581 | 23238.403 | 10.0536 | 5.001 |  |  | 581 | 23238.403 | 10.0536 | 5.0048
582 | 23278.4 | 10.0553 | 5.0009 |  |  | 582 | 23278.4 | 10.0553 | 5.0009 |  |  | 582 | 23278.4 | 10.0553 | 5.0047
583 | 23318.397 | 10.057 | 5.0008 |  |  | 583 | 23318.397 | 10.057 | 5.0008 |  |  | 583 | 23318.397 | 10.057 | 5.0045
584 | 23358.394 | 10.0587 | 5.0006 |  |  | 584 | 23358.394 | 10.0587 | 5.0006 |  |  | 584 | 23358.394 | 10.0587 | 5.0044
585 | 23398.392 | 10.0604 | 5.0005 |  |  | 585 | 23398.392 | 10.0604 | 5.0005 |  |  | 585 | 23398.392 | 10.0604 | 5.0043
586 | 23438.389 | 10.0621 | 5.0003 |  |  | 586 | 23438.389 | 10.0621 | 5.0004 |  |  | 586 | 23438.389 | 10.0621 | 5.0042
587 | 23478.386 | 10.0638 | 5.0002 |  |  | 587 | 23478.386 | 10.0638 | 5.0002 |  |  | 587 | 23478.386 | 10.0638 | 5.004
588 | 23518.383 | 10.0655 | 5.0001 |  |  | 588 | 23518.383 | 10.0655 | 5.0001 |  |  | 588 | 23518.383 | 10.0655 | 5.0039
589 | 23558.381 | 10.0672 | 4.9999 |  |  | 589 | 23558.381 | 10.0672 | 5 |  |  | 589 | 23558.381 | 10.0672 | 5.0038
590 | 23598.378 | 10.0689 | 4.9998 |  |  | 590 | 23598.378 | 10.0689 | 4.9999 |  |  | 590 | 23598.378 | 10.0689 | 5.0037
591 | 23638.375 | 10.0706 | 4.9996 |  |  | 591 | 23638.375 | 10.0706 | 4.9997 |  |  | 591 | 23638.375 | 10.0706 | 5.0035
592 | 23678.372 | 10.0723 | 4.9995 |  |  | 592 | 23678.372 | 10.0723 | 4.9996 |  |  | 592 | 23678.372 | 10.0723 | 5.0034
593 | 23718.37 | 10.074 | 4.9994 |  |  | 593 | 23718.37 | 10.074 | 4.9995 |  |  | 593 | 23718.37 | 10.074 | 5.0033
594 | 23758.367 | 10.0757 | 4.9992 |  |  | 594 | 23758.367 | 10.0757 | 4.9994 |  |  | 594 | 23758.367 | 10.0757 | 5.0032
595 | 23798.364 | 10.0774 | 4.9991 |  |  | 595 | 23798.364 | 10.0774 | 4.9992 |  |  | 595 | 23798.364 | 10.0774 | 5.003
596 | 23838.361 | 10.0791 | 4.9989 |  |  | 596 | 23838.361 | 10.0791 | 4.9991 |  |  | 596 | 23838.361 | 10.0791 | 5.0029
597 | 23878.359 | 10.0807 | 4.9988 |  |  | 597 | 23878.359 | 10.0807 | 4.999 |  |  | 597 | 23878.359 | 10.0807 | 5.0028
598 | 23918.356 | 10.0824 | 4.9987 |  |  | 598 | 23918.356 | 10.0824 | 4.9989 |  |  | 598 | 23918.356 | 10.0824 | 5.0027
599 | 23958.353 | 10.0841 | 4.9985 |  |  | 599 | 23958.353 | 10.0841 | 4.9987 |  |  | 599 | 23958.353 | 10.0841 | 5.0025
600 | 23998.35 | 10.0857 | 4.9984 |  |  | 600 | 23998.35 | 10.0857 | 4.9986 |  |  | 600 | 23998.35 | 10.0857 | 5.0024
601 | 24038.348 | 10.0874 | 4.9982 |  |  | 601 | 24038.348 | 10.0874 | 4.9985 |  |  | 601 | 24038.348 | 10.0874 | 5.0023
602 | 24078.345 | 10.0891 | 4.9981 |  |  | 602 | 24078.345 | 10.0891 | 4.9984 |  |  | 602 | 24078.345 | 10.0891 | 5.0022
603 | 24118.342 | 10.0907 | 4.998 |  |  | 603 | 24118.342 | 10.0907 | 4.9983 |  |  | 603 | 24118.342 | 10.0907 | 5.0021
604 | 24158.339 | 10.0924 | 4.9978 |  |  | 604 | 24158.339 | 10.0924 | 4.9981 |  |  | 604 | 24158.339 | 10.0924 | 5.0019
605 | 24198.337 | 10.094 | 4.9977 |  |  | 605 | 24198.337 | 10.094 | 4.998 |  |  | 605 | 24198.337 | 10.094 | 5.0018
606 | 24238.334 | 10.0957 | 4.9976 |  |  | 606 | 24238.334 | 10.0957 | 4.9979 |  |  | 606 | 24238.334 | 10.0957 | 5.0017
607 | 24278.331 | 10.0973 | 4.9974 |  |  | 607 | 24278.331 | 10.0973 | 4.9978 |  |  | 607 | 24278.331 | 10.0973 | 5.0016
608 | 24318.328 | 10.099 | 4.9973 |  |  | 608 | 24318.328 | 10.099 | 4.9977 |  |  | 608 | 24318.328 | 10.099 | 5.0015
609 | 24358.326 | 10.1006 | 4.9972 |  |  | 609 | 24358.326 | 10.1006 | 4.9975 |  |  | 609 | 24358.326 | 10.1006 | 5.0013
610 | 24398.323 | 10.1023 | 4.997 |  |  | 610 | 24398.323 | 10.1023 | 4.9974 |  |  | 610 | 24398.323 | 10.1023 | 5.0012
611 | 24438.32 | 10.1039 | 4.9969 |  |  | 611 | 24438.32 | 10.1039 | 4.9973 |  |  | 611 | 24438.32 | 10.1039 | 5.0011
612 | 24478.317 | 10.1055 | 4.9968 |  |  | 612 | 24478.317 | 10.1055 | 4.9972 |  |  | 612 | 24478.317 | 10.1055 | 5.001
613 | 24518.315 | 10.1072 | 4.9966 |  |  | 613 | 24518.315 | 10.1072 | 4.9971 |  |  | 613 | 24518.315 | 10.1072 | 5.0009
614 | 24558.312 | 10.1088 | 4.9965 |  |  | 614 | 24558.312 | 10.1088 | 4.9969 |  |  | 614 | 24558.312 | 10.1088 | 5.0008
615 | 24598.309 | 10.1104 | 4.9964 |  |  | 615 | 24598.309 | 10.1104 | 4.9968 |  |  | 615 | 24598.309 | 10.1104 | 5.0006
616 | 24638.306 | 10.1121 | 4.9962 |  |  | 616 | 24638.306 | 10.1121 | 4.9967 |  |  | 616 | 24638.306 | 10.1121 | 5.0005
617 | 24678.304 | 10.1137 | 4.9961 |  |  | 617 | 24678.304 | 10.1137 | 4.9966 |  |  | 617 | 24678.304 | 10.1137 | 5.0004
618 | 24718.301 | 10.1153 | 4.996 |  |  | 618 | 24718.301 | 10.1153 | 4.9965 |  |  | 618 | 24718.301 | 10.1153 | 5.0003
619 | 24758.298 | 10.1169 | 4.9958 |  |  | 619 | 24758.298 | 10.1169 | 4.9964 |  |  | 619 | 24758.298 | 10.1169 | 5.0002
620 | 24798.295 | 10.1185 | 4.9957 |  |  | 620 | 24798.295 | 10.1185 | 4.9962 |  |  | 620 | 24798.295 | 10.1185 | 5.0001
621 | 24838.293 | 10.1201 | 4.9956 |  |  | 621 | 24838.293 | 10.1201 | 4.9961 |  |  | 621 | 24838.293 | 10.1201 | 5
622 | 24878.29 | 10.1218 | 4.9954 |  |  | 622 | 24878.29 | 10.1218 | 4.996 |  |  | 622 | 24878.29 | 10.1218 | 4.9998
623 | 24918.287 | 10.1234 | 4.9953 |  |  | 623 | 24918.287 | 10.1234 | 4.9959 |  |  | 623 | 24918.287 | 10.1234 | 4.9997
624 | 24958.284 | 10.125 | 4.9952 |  |  | 624 | 24958.284 | 10.125 | 4.9958 |  |  | 624 | 24958.284 | 10.125 | 4.9996
625 | 24998.282 | 10.1266 | 4.9951 |  |  | 625 | 24998.282 | 10.1266 | 4.9957 |  |  | 625 | 24998.282 | 10.1266 | 4.9995
626 | 25038.279 | 10.1282 | 4.9949 |  |  | 626 | 25038.279 | 10.1282 | 4.9955 |  |  | 626 | 25038.279 | 10.1282 | 4.9994
627 | 25078.276 | 10.1298 | 4.9948 |  |  | 627 | 25078.276 | 10.1298 | 4.9954 |  |  | 627 | 25078.276 | 10.1298 | 4.9993
628 | 25118.273 | 10.1314 | 4.9947 |  |  | 628 | 25118.273 | 10.1314 | 4.9953 |  |  | 628 | 25118.273 | 10.1314 | 4.9992
629 | 25158.271 | 10.1329 | 4.9945 |  |  | 629 | 25158.271 | 10.1329 | 4.9952 |  |  | 629 | 25158.271 | 10.1329 | 4.9991
630 | 25198.268 | 10.1345 | 4.9944 |  |  | 630 | 25198.268 | 10.1345 | 4.9951 |  |  | 630 | 25198.268 | 10.1345 | 4.9989
631 | 25238.265 | 10.1361 | 4.9943 |  |  | 631 | 25238.265 | 10.1361 | 4.995 |  |  | 631 | 25238.265 | 10.1361 | 4.9988
632 | 25278.262 | 10.1377 | 4.9942 |  |  | 632 | 25278.262 | 10.1377 | 4.9949 |  |  | 632 | 25278.262 | 10.1377 | 4.9987
633 | 25318.26 | 10.1393 | 4.994 |  |  | 633 | 25318.26 | 10.1393 | 4.9947 |  |  | 633 | 25318.26 | 10.1393 | 4.9986
634 | 25358.257 | 10.1409 | 4.9939 |  |  | 634 | 25358.257 | 10.1409 | 4.9946 |  |  | 634 | 25358.257 | 10.1409 | 4.9985
635 | 25398.254 | 10.1424 | 4.9938 |  |  | 635 | 25398.254 | 10.1424 | 4.9945 |  |  | 635 | 25398.254 | 10.1424 | 4.9984
636 | 25438.251 | 10.144 | 4.9937 |  |  | 636 | 25438.251 | 10.144 | 4.9944 |  |  | 636 | 25438.251 | 10.144 | 4.9983
637 | 25478.249 | 10.1456 | 4.9935 |  |  | 637 | 25478.249 | 10.1456 | 4.9943 |  |  | 637 | 25478.249 | 10.1456 | 4.9982
638 | 25518.246 | 10.1471 | 4.9934 |  |  | 638 | 25518.246 | 10.1471 | 4.9942 |  |  | 638 | 25518.246 | 10.1471 | 4.9981
639 | 25558.243 | 10.1487 | 4.9933 |  |  | 639 | 25558.243 | 10.1487 | 4.9941 |  |  | 639 | 25558.243 | 10.1487 | 4.998
640 | 25598.24 | 10.1503 | 4.9932 |  |  | 640 | 25598.24 | 10.1503 | 4.994 |  |  | 640 | 25598.24 | 10.1503 | 4.9978
641 | 25638.238 | 10.1518 | 4.993 |  |  | 641 | 25638.238 | 10.1518 | 4.9938 |  |  | 641 | 25638.238 | 10.1518 | 4.9977
642 | 25678.235 | 10.1534 | 4.9929 |  |  | 642 | 25678.235 | 10.1534 | 4.9937 |  |  | 642 | 25678.235 | 10.1534 | 4.9976
643 | 25718.232 | 10.155 | 4.9928 |  |  | 643 | 25718.232 | 10.155 | 4.9936 |  |  | 643 | 25718.232 | 10.155 | 4.9975
644 | 25758.229 | 10.1565 | 4.9927 |  |  | 644 | 25758.229 | 10.1565 | 4.9935 |  |  | 644 | 25758.229 | 10.1565 | 4.9974
645 | 25798.227 | 10.1581 | 4.9925 |  |  | 645 | 25798.227 | 10.1581 | 4.9934 |  |  | 645 | 25798.227 | 10.1581 | 4.9973
646 | 25838.224 | 10.1596 | 4.9924 |  |  | 646 | 25838.224 | 10.1596 | 4.9933 |  |  | 646 | 25838.224 | 10.1596 | 4.9972
647 | 25878.221 | 10.1612 | 4.9923 |  |  | 647 | 25878.221 | 10.1612 | 4.9932 |  |  | 647 | 25878.221 | 10.1612 | 4.9971
648 | 25918.218 | 10.1627 | 4.9922 |  |  | 648 | 25918.218 | 10.1627 | 4.9931 |  |  | 648 | 25918.218 | 10.1627 | 4.997
649 | 25958.216 | 10.1642 | 4.992 |  |  | 649 | 25958.216 | 10.1642 | 4.993 |  |  | 649 | 25958.216 | 10.1642 | 4.9969
650 | 25998.213 | 10.1658 | 4.9919 |  |  | 650 | 25998.213 | 10.1658 | 4.9929 |  |  | 650 | 25998.213 | 10.1658 | 4.9968
651 | 26038.21 | 10.1673 | 4.9918 |  |  | 651 | 26038.21 | 10.1673 | 4.9927 |  |  | 651 | 26038.21 | 10.1673 | 4.9967
652 | 26078.207 | 10.1689 | 4.9917 |  |  | 652 | 26078.207 | 10.1689 | 4.9926 |  |  | 652 | 26078.207 | 10.1689 | 4.9966
653 | 26118.205 | 10.1704 | 4.9916 |  |  | 653 | 26118.205 | 10.1704 | 4.9925 |  |  | 653 | 26118.205 | 10.1704 | 4.9965
654 | 26158.202 | 10.1719 | 4.9914 |  |  | 654 | 26158.202 | 10.1719 | 4.9924 |  |  | 654 | 26158.202 | 10.1719 | 4.9963
655 | 26198.199 | 10.1734 | 4.9913 |  |  | 655 | 26198.199 | 10.1734 | 4.9923 |  |  | 655 | 26198.199 | 10.1734 | 4.9962
656 | 26238.196 | 10.175 | 4.9912 |  |  | 656 | 26238.196 | 10.175 | 4.9922 |  |  | 656 | 26238.196 | 10.175 | 4.9961
657 | 26278.194 | 10.1765 | 4.9911 |  |  | 657 | 26278.194 | 10.1765 | 4.9921 |  |  | 657 | 26278.194 | 10.1765 | 4.996
658 | 26318.191 | 10.178 | 4.991 |  |  | 658 | 26318.191 | 10.178 | 4.992 |  |  | 658 | 26318.191 | 10.178 | 4.9959
659 | 26358.188 | 10.1795 | 4.9908 |  |  | 659 | 26358.188 | 10.1795 | 4.9919 |  |  | 659 | 26358.188 | 10.1795 | 4.9958
660 | 26398.185 | 10.1811 | 4.9907 |  |  | 660 | 26398.185 | 10.1811 | 4.9918 |  |  | 660 | 26398.185 | 10.1811 | 4.9957
661 | 26438.183 | 10.1826 | 4.9906 |  |  | 661 | 26438.183 | 10.1826 | 4.9917 |  |  | 661 | 26438.183 | 10.1826 | 4.9956
662 | 26478.18 | 10.1841 | 4.9905 |  |  | 662 | 26478.18 | 10.1841 | 4.9916 |  |  | 662 | 26478.18 | 10.1841 | 4.9955
663 | 26518.177 | 10.1856 | 4.9904 |  |  | 663 | 26518.177 | 10.1856 | 4.9915 |  |  | 663 | 26518.177 | 10.1856 | 4.9954
664 | 26558.174 | 10.1871 | 4.9902 |  |  | 664 | 26558.174 | 10.1871 | 4.9913 |  |  | 664 | 26558.174 | 10.1871 | 4.9953
665 | 26598.172 | 10.1886 | 4.9901 |  |  | 665 | 26598.172 | 10.1886 | 4.9912 |  |  | 665 | 26598.172 | 10.1886 | 4.9952
666 | 26638.169 | 10.1901 | 4.99 |  |  | 666 | 26638.169 | 10.1901 | 4.9911 |  |  | 666 | 26638.169 | 10.1901 | 4.9951
667 | 26678.166 | 10.1916 | 4.9899 |  |  | 667 | 26678.166 | 10.1916 | 4.991 |  |  | 667 | 26678.166 | 10.1916 | 4.995
668 | 26718.163 | 10.1931 | 4.9898 |  |  | 668 | 26718.163 | 10.1931 | 4.9909 |  |  | 668 | 26718.163 | 10.1931 | 4.9949
669 | 26758.161 | 10.1946 | 4.9896 |  |  | 669 | 26758.161 | 10.1946 | 4.9908 |  |  | 669 | 26758.161 | 10.1946 | 4.9948
670 | 26798.158 | 10.1961 | 4.9895 |  |  | 670 | 26798.158 | 10.1961 | 4.9907 |  |  | 670 | 26798.158 | 10.1961 | 4.9947
671 | 26838.155 | 10.1976 | 4.9894 |  |  | 671 | 26838.155 | 10.1976 | 4.9906 |  |  | 671 | 26838.155 | 10.1976 | 4.9946
672 | 26878.152 | 10.1991 | 4.9893 |  |  | 672 | 26878.152 | 10.1991 | 4.9905 |  |  | 672 | 26878.152 | 10.1991 | 4.9945
673 | 26918.15 | 10.2006 | 4.9892 |  |  | 673 | 26918.15 | 10.2006 | 4.9904 |  |  | 673 | 26918.15 | 10.2006 | 4.9944
674 | 26958.147 | 10.202 | 4.9891 |  |  | 674 | 26958.147 | 10.202 | 4.9903 |  |  | 674 | 26958.147 | 10.202 | 4.9943
675 | 26998.144 | 10.2035 | 4.9889 |  |  | 675 | 26998.144 | 10.2035 | 4.9902 |  |  | 675 | 26998.144 | 10.2035 | 4.9942
676 | 27038.141 | 10.205 | 4.9888 |  |  | 676 | 27038.141 | 10.205 | 4.9901 |  |  | 676 | 27038.141 | 10.205 | 4.9941
677 | 27078.139 | 10.2065 | 4.9887 |  |  | 677 | 27078.139 | 10.2065 | 4.99 |  |  | 677 | 27078.139 | 10.2065 | 4.994
678 | 27118.136 | 10.208 | 4.9886 |  |  | 678 | 27118.136 | 10.208 | 4.9899 |  |  | 678 | 27118.136 | 10.208 | 4.9939
679 | 27158.133 | 10.2094 | 4.9885 |  |  | 679 | 27158.133 | 10.2094 | 4.9898 |  |  | 679 | 27158.133 | 10.2094 | 4.9938
680 | 27198.13 | 10.2109 | 4.9884 |  |  | 680 | 27198.13 | 10.2109 | 4.9897 |  |  | 680 | 27198.13 | 10.2109 | 4.9937
681 | 27238.128 | 10.2124 | 4.9882 |  |  | 681 | 27238.128 | 10.2124 | 4.9896 |  |  | 681 | 27238.128 | 10.2124 | 4.9936
682 | 27278.125 | 10.2138 | 4.9881 |  |  | 682 | 27278.125 | 10.2138 | 4.9894 |  |  | 682 | 27278.125 | 10.2138 | 4.9935
683 | 27318.122 | 10.2153 | 4.988 |  |  | 683 | 27318.122 | 10.2153 | 4.9893 |  |  | 683 | 27318.122 | 10.2153 | 4.9934
684 | 27358.119 | 10.2168 | 4.9879 |  |  | 684 | 27358.119 | 10.2168 | 4.9892 |  |  | 684 | 27358.119 | 10.2168 | 4.9933
685 | 27398.117 | 10.2182 | 4.9878 |  |  | 685 | 27398.117 | 10.2182 | 4.9891 |  |  | 685 | 27398.117 | 10.2182 | 4.9932
686 | 27438.114 | 10.2197 | 4.9877 |  |  | 686 | 27438.114 | 10.2197 | 4.989 |  |  | 686 | 27438.114 | 10.2197 | 4.9931
687 | 27478.111 | 10.2211 | 4.9876 |  |  | 687 | 27478.111 | 10.2211 | 4.9889 |  |  | 687 | 27478.111 | 10.2211 | 4.993
688 | 27518.108 | 10.2226 | 4.9874 |  |  | 688 | 27518.108 | 10.2226 | 4.9888 |  |  | 688 | 27518.108 | 10.2226 | 4.9929
689 | 27558.106 | 10.2241 | 4.9873 |  |  | 689 | 27558.106 | 10.2241 | 4.9887 |  |  | 689 | 27558.106 | 10.2241 | 4.9928
690 | 27598.103 | 10.2255 | 4.9872 |  |  | 690 | 27598.103 | 10.2255 | 4.9886 |  |  | 690 | 27598.103 | 10.2255 | 4.9927
691 | 27638.1 | 10.227 | 4.9871 |  |  | 691 | 27638.1 | 10.227 | 4.9885 |  |  | 691 | 27638.1 | 10.227 | 4.9926
692 | 27678.097 | 10.2284 | 4.987 |  |  | 692 | 27678.097 | 10.2284 | 4.9884 |  |  | 692 | 27678.097 | 10.2284 | 4.9925
693 | 27718.095 | 10.2298 | 4.9869 |  |  | 693 | 27718.095 | 10.2298 | 4.9883 |  |  | 693 | 27718.095 | 10.2298 | 4.9924
694 | 27758.092 | 10.2313 | 4.9868 |  |  | 694 | 27758.092 | 10.2313 | 4.9882 |  |  | 694 | 27758.092 | 10.2313 | 4.9923
695 | 27798.089 | 10.2327 | 4.9866 |  |  | 695 | 27798.089 | 10.2327 | 4.9881 |  |  | 695 | 27798.089 | 10.2327 | 4.9922
696 | 27838.086 | 10.2342 | 4.9865 |  |  | 696 | 27838.086 | 10.2342 | 4.988 |  |  | 696 | 27838.086 | 10.2342 | 4.9921
697 | 27878.084 | 10.2356 | 4.9864 |  |  | 697 | 27878.084 | 10.2356 | 4.9879 |  |  | 697 | 27878.084 | 10.2356 | 4.992
698 | 27918.081 | 10.237 | 4.9863 |  |  | 698 | 27918.081 | 10.237 | 4.9878 |  |  | 698 | 27918.081 | 10.237 | 4.9919
699 | 27958.078 | 10.2385 | 4.9862 |  |  | 699 | 27958.078 | 10.2385 | 4.9877 |  |  | 699 | 27958.078 | 10.2385 | 4.9918
700 | 27998.075 | 10.2399 | 4.9861 |  |  | 700 | 27998.075 | 10.2399 | 4.9876 |  |  | 700 | 27998.075 | 10.2399 | 4.9917
701 | 28038.073 | 10.2413 | 4.986 |  |  | 701 | 28038.073 | 10.2413 | 4.9875 |  |  | 701 | 28038.073 | 10.2413 | 4.9916
702 | 28078.07 | 10.2427 | 4.9859 |  |  | 702 | 28078.07 | 10.2427 | 4.9874 |  |  | 702 | 28078.07 | 10.2427 | 4.9915
703 | 28118.067 | 10.2442 | 4.9857 |  |  | 703 | 28118.067 | 10.2442 | 4.9873 |  |  | 703 | 28118.067 | 10.2442 | 4.9914
704 | 28158.064 | 10.2456 | 4.9856 |  |  | 704 | 28158.064 | 10.2456 | 4.9872 |  |  | 704 | 28158.064 | 10.2456 | 4.9913
705 | 28198.062 | 10.247 | 4.9855 |  |  | 705 | 28198.062 | 10.247 | 4.9871 |  |  | 705 | 28198.062 | 10.247 | 4.9912
706 | 28238.059 | 10.2484 | 4.9854 |  |  | 706 | 28238.059 | 10.2484 | 4.987 |  |  | 706 | 28238.059 | 10.2484 | 4.9911
707 | 28278.056 | 10.2498 | 4.9853 |  |  | 707 | 28278.056 | 10.2498 | 4.9869 |  |  | 707 | 28278.056 | 10.2498 | 4.991
708 | 28318.053 | 10.2513 | 4.9852 |  |  | 708 | 28318.053 | 10.2513 | 4.9868 |  |  | 708 | 28318.053 | 10.2513 | 4.9909
709 | 28358.051 | 10.2527 | 4.9851 |  |  | 709 | 28358.051 | 10.2527 | 4.9867 |  |  | 709 | 28358.051 | 10.2527 | 4.9908
710 | 28398.048 | 10.2541 | 4.985 |  |  | 710 | 28398.048 | 10.2541 | 4.9866 |  |  | 710 | 28398.048 | 10.2541 | 4.9907
711 | 28438.045 | 10.2555 | 4.9848 |  |  | 711 | 28438.045 | 10.2555 | 4.9865 |  |  | 711 | 28438.045 | 10.2555 | 4.9906
712 | 28478.042 | 10.2569 | 4.9847 |  |  | 712 | 28478.042 | 10.2569 | 4.9864 |  |  | 712 | 28478.042 | 10.2569 | 4.9905
713 | 28518.04 | 10.2583 | 4.9846 |  |  | 713 | 28518.04 | 10.2583 | 4.9863 |  |  | 713 | 28518.04 | 10.2583 | 4.9904
714 | 28558.037 | 10.2597 | 4.9845 |  |  | 714 | 28558.037 | 10.2597 | 4.9862 |  |  | 714 | 28558.037 | 10.2597 | 4.9903
715 | 28598.034 | 10.2611 | 4.9844 |  |  | 715 | 28598.034 | 10.2611 | 4.9861 |  |  | 715 | 28598.034 | 10.2611 | 4.9902
716 | 28638.031 | 10.2625 | 4.9843 |  |  | 716 | 28638.031 | 10.2625 | 4.986 |  |  | 716 | 28638.031 | 10.2625 | 4.9901
717 | 28678.029 | 10.2639 | 4.9842 |  |  | 717 | 28678.029 | 10.2639 | 4.9859 |  |  | 717 | 28678.029 | 10.2639 | 4.9901
718 | 28718.026 | 10.2653 | 4.9841 |  |  | 718 | 28718.026 | 10.2653 | 4.9858 |  |  | 718 | 28718.026 | 10.2653 | 4.99
719 | 28758.023 | 10.2667 | 4.984 |  |  | 719 | 28758.023 | 10.2667 | 4.9857 |  |  | 719 | 28758.023 | 10.2667 | 4.9899
720 | 28798.02 | 10.2681 | 4.9839 |  |  | 720 | 28798.02 | 10.2681 | 4.9856 |  |  | 720 | 28798.02 | 10.2681 | 4.9898
721 | 28838.018 | 10.2694 | 4.9837 |  |  | 721 | 28838.018 | 10.2694 | 4.9855 |  |  | 721 | 28838.018 | 10.2694 | 4.9897
722 | 28878.015 | 10.2708 | 4.9836 |  |  | 722 | 28878.015 | 10.2708 | 4.9854 |  |  | 722 | 28878.015 | 10.2708 | 4.9896
723 | 28918.012 | 10.2722 | 4.9835 |  |  | 723 | 28918.012 | 10.2722 | 4.9853 |  |  | 723 | 28918.012 | 10.2722 | 4.9895
724 | 28958.009 | 10.2736 | 4.9834 |  |  | 724 | 28958.009 | 10.2736 | 4.9852 |  |  | 724 | 28958.009 | 10.2736 | 4.9894
725 | 28998.007 | 10.275 | 4.9833 |  |  | 725 | 28998.007 | 10.275 | 4.9851 |  |  | 725 | 28998.007 | 10.275 | 4.9893
726 | 29038.004 | 10.2764 | 4.9832 |  |  | 726 | 29038.004 | 10.2764 | 4.985 |  |  | 726 | 29038.004 | 10.2764 | 4.9892
727 | 29078.001 | 10.2777 | 4.9831 |  |  | 727 | 29078.001 | 10.2777 | 4.9849 |  |  | 727 | 29078.001 | 10.2777 | 4.9891
728 | 29117.998 | 10.2791 | 4.983 |  |  | 728 | 29117.998 | 10.2791 | 4.9848 |  |  | 728 | 29117.998 | 10.2791 | 4.989
729 | 29157.996 | 10.2805 | 4.9829 |  |  | 729 | 29157.996 | 10.2805 | 4.9847 |  |  | 729 | 29157.996 | 10.2805 | 4.9889
730 | 29197.993 | 10.2819 | 4.9828 |  |  | 730 | 29197.993 | 10.2819 | 4.9846 |  |  | 730 | 29197.993 | 10.2819 | 4.9888
731 | 29237.99 | 10.2832 | 4.9826 |  |  | 731 | 29237.99 | 10.2832 | 4.9845 |  |  | 731 | 29237.99 | 10.2832 | 4.9887
732 | 29277.987 | 10.2846 | 4.9825 |  |  | 732 | 29277.987 | 10.2846 | 4.9844 |  |  | 732 | 29277.987 | 10.2846 | 4.9886
733 | 29317.985 | 10.286 | 4.9824 |  |  | 733 | 29317.985 | 10.286 | 4.9843 |  |  | 733 | 29317.985 | 10.286 | 4.9885
734 | 29357.982 | 10.2873 | 4.9823 |  |  | 734 | 29357.982 | 10.2873 | 4.9842 |  |  | 734 | 29357.982 | 10.2873 | 4.9884
735 | 29397.979 | 10.2887 | 4.9822 |  |  | 735 | 29397.979 | 10.2887 | 4.9841 |  |  | 735 | 29397.979 | 10.2887 | 4.9883
736 | 29437.976 | 10.29 | 4.9821 |  |  | 736 | 29437.976 | 10.29 | 4.984 |  |  | 736 | 29437.976 | 10.29 | 4.9883
737 | 29477.974 | 10.2914 | 4.982 |  |  | 737 | 29477.974 | 10.2914 | 4.9839 |  |  | 737 | 29477.974 | 10.2914 | 4.9882
738 | 29517.971 | 10.2928 | 4.9819 |  |  | 738 | 29517.971 | 10.2928 | 4.9838 |  |  | 738 | 29517.971 | 10.2928 | 4.9881
739 | 29557.968 | 10.2941 | 4.9818 |  |  | 739 | 29557.968 | 10.2941 | 4.9837 |  |  | 739 | 29557.968 | 10.2941 | 4.988
740 | 29597.965 | 10.2955 | 4.9817 |  |  | 740 | 29597.965 | 10.2955 | 4.9836 |  |  | 740 | 29597.965 | 10.2955 | 4.9879
741 | 29637.963 | 10.2968 | 4.9816 |  |  | 741 | 29637.963 | 10.2968 | 4.9835 |  |  | 741 | 29637.963 | 10.2968 | 4.9878
742 | 29677.96 | 10.2982 | 4.9814 |  |  | 742 | 29677.96 | 10.2982 | 4.9834 |  |  | 742 | 29677.96 | 10.2982 | 4.9877
743 | 29717.957 | 10.2995 | 4.9813 |  |  | 743 | 29717.957 | 10.2995 | 4.9833 |  |  | 743 | 29717.957 | 10.2995 | 4.9876
744 | 29757.954 | 10.3009 | 4.9812 |  |  | 744 | 29757.954 | 10.3009 | 4.9832 |  |  | 744 | 29757.954 | 10.3009 | 4.9875
745 | 29797.952 | 10.3022 | 4.9811 |  |  | 745 | 29797.952 | 10.3022 | 4.9831 |  |  | 745 | 29797.952 | 10.3022 | 4.9874
746 | 29837.949 | 10.3035 | 4.981 |  |  | 746 | 29837.949 | 10.3035 | 4.983 |  |  | 746 | 29837.949 | 10.3035 | 4.9873
747 | 29877.946 | 10.3049 | 4.9809 |  |  | 747 | 29877.946 | 10.3049 | 4.9829 |  |  | 747 | 29877.946 | 10.3049 | 4.9872
748 | 29917.943 | 10.3062 | 4.9808 |  |  | 748 | 29917.943 | 10.3062 | 4.9828 |  |  | 748 | 29917.943 | 10.3062 | 4.9871
749 | 29957.941 | 10.3075 | 4.9807 |  |  | 749 | 29957.941 | 10.3075 | 4.9827 |  |  | 749 | 29957.941 | 10.3075 | 4.987
750 | 29997.938 | 10.3089 | 4.9806 |  |  | 750 | 29997.938 | 10.3089 | 4.9826 |  |  | 750 | 29997.938 | 10.3089 | 4.9869
751 | 30037.935 | 10.3102 | 4.9805 |  |  | 751 | 30037.935 | 10.3102 | 4.9825 |  |  | 751 | 30037.935 | 10.3102 | 4.9868
752 | 30077.932 | 10.3115 | 4.9804 |  |  | 752 | 30077.932 | 10.3115 | 4.9824 |  |  | 752 | 30077.932 | 10.3115 | 4.9868
753 | 30117.93 | 10.3129 | 4.9803 |  |  | 753 | 30117.93 | 10.3129 | 4.9823 |  |  | 753 | 30117.93 | 10.3129 | 4.9867
754 | 30157.927 | 10.3142 | 4.9802 |  |  | 754 | 30157.927 | 10.3142 | 4.9822 |  |  | 754 | 30157.927 | 10.3142 | 4.9866
755 | 30197.924 | 10.3155 | 4.98 |  |  | 755 | 30197.924 | 10.3155 | 4.9821 |  |  | 755 | 30197.924 | 10.3155 | 4.9865
756 | 30237.921 | 10.3169 | 4.9799 |  |  | 756 | 30237.921 | 10.3169 | 4.982 |  |  | 756 | 30237.921 | 10.3169 | 4.9864
757 | 30277.919 | 10.3182 | 4.9798 |  |  | 757 | 30277.919 | 10.3182 | 4.9819 |  |  | 757 | 30277.919 | 10.3182 | 4.9863
758 | 30317.916 | 10.3195 | 4.9797 |  |  | 758 | 30317.916 | 10.3195 | 4.9818 |  |  | 758 | 30317.916 | 10.3195 | 4.9862
759 | 30357.913 | 10.3208 | 4.9796 |  |  | 759 | 30357.913 | 10.3208 | 4.9817 |  |  | 759 | 30357.913 | 10.3208 | 4.9861
760 | 30397.91 | 10.3221 | 4.9795 |  |  | 760 | 30397.91 | 10.3221 | 4.9816 |  |  | 760 | 30397.91 | 10.3221 | 4.986
761 | 30437.908 | 10.3234 | 4.9794 |  |  | 761 | 30437.908 | 10.3234 | 4.9815 |  |  | 761 | 30437.908 | 10.3234 | 4.9859
762 | 30477.905 | 10.3248 | 4.9793 |  |  | 762 | 30477.905 | 10.3248 | 4.9814 |  |  | 762 | 30477.905 | 10.3248 | 4.9858
763 | 30517.902 | 10.3261 | 4.9792 |  |  | 763 | 30517.902 | 10.3261 | 4.9813 |  |  | 763 | 30517.902 | 10.3261 | 4.9857
764 | 30557.899 | 10.3274 | 4.9791 |  |  | 764 | 30557.899 | 10.3274 | 4.9812 |  |  | 764 | 30557.899 | 10.3274 | 4.9856
765 | 30597.897 | 10.3287 | 4.979 |  |  | 765 | 30597.897 | 10.3287 | 4.9811 |  |  | 765 | 30597.897 | 10.3287 | 4.9856
766 | 30637.894 | 10.33 | 4.9789 |  |  | 766 | 30637.894 | 10.33 | 4.981 |  |  | 766 | 30637.894 | 10.33 | 4.9855
767 | 30677.891 | 10.3313 | 4.9788 |  |  | 767 | 30677.891 | 10.3313 | 4.9809 |  |  | 767 | 30677.891 | 10.3313 | 4.9854
768 | 30717.888 | 10.3326 | 4.9787 |  |  | 768 | 30717.888 | 10.3326 | 4.9808 |  |  | 768 | 30717.888 | 10.3326 | 4.9853
769 | 30757.886 | 10.3339 | 4.9786 |  |  | 769 | 30757.886 | 10.3339 | 4.9807 |  |  | 769 | 30757.886 | 10.3339 | 4.9852
770 | 30797.883 | 10.3352 | 4.9784 |  |  | 770 | 30797.883 | 10.3352 | 4.9806 |  |  | 770 | 30797.883 | 10.3352 | 4.9851
771 | 30837.88 | 10.3365 | 4.9783 |  |  | 771 | 30837.88 | 10.3365 | 4.9805 |  |  | 771 | 30837.88 | 10.3365 | 4.985
772 | 30877.877 | 10.3378 | 4.9782 |  |  | 772 | 30877.877 | 10.3378 | 4.9804 |  |  | 772 | 30877.877 | 10.3378 | 4.9849
773 | 30917.875 | 10.3391 | 4.9781 |  |  | 773 | 30917.875 | 10.3391 | 4.9803 |  |  | 773 | 30917.875 | 10.3391 | 4.9848
774 | 30957.872 | 10.3404 | 4.978 |  |  | 774 | 30957.872 | 10.3404 | 4.9802 |  |  | 774 | 30957.872 | 10.3404 | 4.9847
775 | 30997.869 | 10.3417 | 4.9779 |  |  | 775 | 30997.869 | 10.3417 | 4.9801 |  |  | 775 | 30997.869 | 10.3417 | 4.9846
776 | 31037.866 | 10.343 | 4.9778 |  |  | 776 | 31037.866 | 10.343 | 4.98 |  |  | 776 | 31037.866 | 10.343 | 4.9845
777 | 31077.864 | 10.3443 | 4.9777 |  |  | 777 | 31077.864 | 10.3443 | 4.9799 |  |  | 777 | 31077.864 | 10.3443 | 4.9844
778 | 31117.861 | 10.3455 | 4.9776 |  |  | 778 | 31117.861 | 10.3455 | 4.9798 |  |  | 778 | 31117.861 | 10.3455 | 4.9844
779 | 31157.858 | 10.3468 | 4.9775 |  |  | 779 | 31157.858 | 10.3468 | 4.9797 |  |  | 779 | 31157.858 | 10.3468 | 4.9843
780 | 31197.855 | 10.3481 | 4.9774 |  |  | 780 | 31197.855 | 10.3481 | 4.9796 |  |  | 780 | 31197.855 | 10.3481 | 4.9842
781 | 31237.853 | 10.3494 | 4.9773 |  |  | 781 | 31237.853 | 10.3494 | 4.9795 |  |  | 781 | 31237.853 | 10.3494 | 4.9841
782 | 31277.85 | 10.3507 | 4.9772 |  |  | 782 | 31277.85 | 10.3507 | 4.9794 |  |  | 782 | 31277.85 | 10.3507 | 4.984
783 | 31317.847 | 10.3519 | 4.9771 |  |  | 783 | 31317.847 | 10.3519 | 4.9793 |  |  | 783 | 31317.847 | 10.3519 | 4.9839
784 | 31357.844 | 10.3532 | 4.977 |  |  | 784 | 31357.844 | 10.3532 | 4.9792 |  |  | 784 | 31357.844 | 10.3532 | 4.9838
785 | 31397.842 | 10.3545 | 4.9769 |  |  | 785 | 31397.842 | 10.3545 | 4.9791 |  |  | 785 | 31397.842 | 10.3545 | 4.9837
786 | 31437.839 | 10.3558 | 4.9768 |  |  | 786 | 31437.839 | 10.3558 | 4.979 |  |  | 786 | 31437.839 | 10.3558 | 4.9836
787 | 31477.836 | 10.357 | 4.9766 |  |  | 787 | 31477.836 | 10.357 | 4.9789 |  |  | 787 | 31477.836 | 10.357 | 4.9835
788 | 31517.833 | 10.3583 | 4.9765 |  |  | 788 | 31517.833 | 10.3583 | 4.9788 |  |  | 788 | 31517.833 | 10.3583 | 4.9834
789 | 31557.831 | 10.3596 | 4.9764 |  |  | 789 | 31557.831 | 10.3596 | 4.9787 |  |  | 789 | 31557.831 | 10.3596 | 4.9833
790 | 31597.828 | 10.3608 | 4.9763 |  |  | 790 | 31597.828 | 10.3608 | 4.9786 |  |  | 790 | 31597.828 | 10.3608 | 4.9833
791 | 31637.825 | 10.3621 | 4.9762 |  |  | 791 | 31637.825 | 10.3621 | 4.9785 |  |  | 791 | 31637.825 | 10.3621 | 4.9832
792 | 31677.822 | 10.3634 | 4.9761 |  |  | 792 | 31677.822 | 10.3634 | 4.9784 |  |  | 792 | 31677.822 | 10.3634 | 4.9831
793 | 31717.82 | 10.3646 | 4.976 |  |  | 793 | 31717.82 | 10.3646 | 4.9783 |  |  | 793 | 31717.82 | 10.3646 | 4.983
794 | 31757.817 | 10.3659 | 4.9759 |  |  | 794 | 31757.817 | 10.3659 | 4.9782 |  |  | 794 | 31757.817 | 10.3659 | 4.9829
795 | 31797.814 | 10.3672 | 4.9758 |  |  | 795 | 31797.814 | 10.3672 | 4.9781 |  |  | 795 | 31797.814 | 10.3672 | 4.9828
796 | 31837.811 | 10.3684 | 4.9757 |  |  | 796 | 31837.811 | 10.3684 | 4.978 |  |  | 796 | 31837.811 | 10.3684 | 4.9827
797 | 31877.809 | 10.3697 | 4.9756 |  |  | 797 | 31877.809 | 10.3697 | 4.9779 |  |  | 797 | 31877.809 | 10.3697 | 4.9826
798 | 31917.806 | 10.3709 | 4.9755 |  |  | 798 | 31917.806 | 10.3709 | 4.9778 |  |  | 798 | 31917.806 | 10.3709 | 4.9825
799 | 31957.803 | 10.3722 | 4.9754 |  |  | 799 | 31957.803 | 10.3722 | 4.9777 |  |  | 799 | 31957.803 | 10.3722 | 4.9824
800 | 31997.8 | 10.3734 | 4.9753 |  |  | 800 | 31997.8 | 10.3734 | 4.9776 |  |  | 800 | 31997.8 | 10.3734 | 4.9823
801 | 32037.798 | 10.3747 | 4.9752 |  |  | 801 | 32037.798 | 10.3747 | 4.9775 |  |  | 801 | 32037.798 | 10.3747 | 4.9823
802 | 32077.795 | 10.3759 | 4.9751 |  |  | 802 | 32077.795 | 10.3759 | 4.9774 |  |  | 802 | 32077.795 | 10.3759 | 4.9822
803 | 32117.792 | 10.3772 | 4.975 |  |  | 803 | 32117.792 | 10.3772 | 4.9773 |  |  | 803 | 32117.792 | 10.3772 | 4.9821
804 | 32157.789 | 10.3784 | 4.9749 |  |  | 804 | 32157.789 | 10.3784 | 4.9772 |  |  | 804 | 32157.789 | 10.3784 | 4.982
805 | 32197.787 | 10.3797 | 4.9747 |  |  | 805 | 32197.787 | 10.3797 | 4.9771 |  |  | 805 | 32197.787 | 10.3797 | 4.9819
806 | 32237.784 | 10.3809 | 4.9746 |  |  | 806 | 32237.784 | 10.3809 | 4.977 |  |  | 806 | 32237.784 | 10.3809 | 4.9818
807 | 32277.781 | 10.3821 | 4.9745 |  |  | 807 | 32277.781 | 10.3821 | 4.9769 |  |  | 807 | 32277.781 | 10.3821 | 4.9817
808 | 32317.778 | 10.3834 | 4.9744 |  |  | 808 | 32317.778 | 10.3834 | 4.9768 |  |  | 808 | 32317.778 | 10.3834 | 4.9816
809 | 32357.776 | 10.3846 | 4.9743 |  |  | 809 | 32357.776 | 10.3846 | 4.9767 |  |  | 809 | 32357.776 | 10.3846 | 4.9815
810 | 32397.773 | 10.3858 | 4.9742 |  |  | 810 | 32397.773 | 10.3858 | 4.9766 |  |  | 810 | 32397.773 | 10.3858 | 4.9814
811 | 32437.77 | 10.3871 | 4.9741 |  |  | 811 | 32437.77 | 10.3871 | 4.9765 |  |  | 811 | 32437.77 | 10.3871 | 4.9813
812 | 32477.767 | 10.3883 | 4.974 |  |  | 812 | 32477.767 | 10.3883 | 4.9764 |  |  | 812 | 32477.767 | 10.3883 | 4.9813
813 | 32517.765 | 10.3895 | 4.9739 |  |  | 813 | 32517.765 | 10.3895 | 4.9763 |  |  | 813 | 32517.765 | 10.3895 | 4.9812
814 | 32557.762 | 10.3908 | 4.9738 |  |  | 814 | 32557.762 | 10.3908 | 4.9762 |  |  | 814 | 32557.762 | 10.3908 | 4.9811
815 | 32597.759 | 10.392 | 4.9737 |  |  | 815 | 32597.759 | 10.392 | 4.9762 |  |  | 815 | 32597.759 | 10.392 | 4.981
816 | 32637.756 | 10.3932 | 4.9736 |  |  | 816 | 32637.756 | 10.3932 | 4.9761 |  |  | 816 | 32637.756 | 10.3932 | 4.9809
817 | 32677.754 | 10.3944 | 4.9735 |  |  | 817 | 32677.754 | 10.3944 | 4.976 |  |  | 817 | 32677.754 | 10.3944 | 4.9808
818 | 32717.751 | 10.3957 | 4.9734 |  |  | 818 | 32717.751 | 10.3957 | 4.9759 |  |  | 818 | 32717.751 | 10.3957 | 4.9807
819 | 32757.748 | 10.3969 | 4.9733 |  |  | 819 | 32757.748 | 10.3969 | 4.9758 |  |  | 819 | 32757.748 | 10.3969 | 4.9806
820 | 32797.745 | 10.3981 | 4.9732 |  |  | 820 | 32797.745 | 10.3981 | 4.9757 |  |  | 820 | 32797.745 | 10.3981 | 4.9805
821 | 32837.743 | 10.3993 | 4.9731 |  |  | 821 | 32837.743 | 10.3993 | 4.9756 |  |  | 821 | 32837.743 | 10.3993 | 4.9804
822 | 32877.74 | 10.4006 | 4.973 |  |  | 822 | 32877.74 | 10.4006 | 4.9755 |  |  | 822 | 32877.74 | 10.4006 | 4.9803
823 | 32917.737 | 10.4018 | 4.9729 |  |  | 823 | 32917.737 | 10.4018 | 4.9754 |  |  | 823 | 32917.737 | 10.4018 | 4.9803
824 | 32957.734 | 10.403 | 4.9728 |  |  | 824 | 32957.734 | 10.403 | 4.9753 |  |  | 824 | 32957.734 | 10.403 | 4.9802
825 | 32997.732 | 10.4042 | 4.9727 |  |  | 825 | 32997.732 | 10.4042 | 4.9752 |  |  | 825 | 32997.732 | 10.4042 | 4.9801
826 | 33037.729 | 10.4054 | 4.9725 |  |  | 826 | 33037.729 | 10.4054 | 4.9751 |  |  | 826 | 33037.729 | 10.4054 | 4.98
827 | 33077.726 | 10.4066 | 4.9724 |  |  | 827 | 33077.726 | 10.4066 | 4.975 |  |  | 827 | 33077.726 | 10.4066 | 4.9799
828 | 33117.723 | 10.4078 | 4.9723 |  |  | 828 | 33117.723 | 10.4078 | 4.9749 |  |  | 828 | 33117.723 | 10.4078 | 4.9798
829 | 33157.721 | 10.409 | 4.9722 |  |  | 829 | 33157.721 | 10.409 | 4.9748 |  |  | 829 | 33157.721 | 10.409 | 4.9797
830 | 33197.718 | 10.4102 | 4.9721 |  |  | 830 | 33197.718 | 10.4102 | 4.9747 |  |  | 830 | 33197.718 | 10.4102 | 4.9796
831 | 33237.715 | 10.4114 | 4.972 |  |  | 831 | 33237.715 | 10.4114 | 4.9746 |  |  | 831 | 33237.715 | 10.4114 | 4.9795
832 | 33277.712 | 10.4126 | 4.9719 |  |  | 832 | 33277.712 | 10.4126 | 4.9745 |  |  | 832 | 33277.712 | 10.4126 | 4.9794
833 | 33317.71 | 10.4138 | 4.9718 |  |  | 833 | 33317.71 | 10.4138 | 4.9744 |  |  | 833 | 33317.71 | 10.4138 | 4.9793
834 | 33357.707 | 10.415 | 4.9717 |  |  | 834 | 33357.707 | 10.415 | 4.9743 |  |  | 834 | 33357.707 | 10.415 | 4.9793
835 | 33397.704 | 10.4162 | 4.9716 |  |  | 835 | 33397.704 | 10.4162 | 4.9742 |  |  | 835 | 33397.704 | 10.4162 | 4.9792
836 | 33437.701 | 10.4174 | 4.9715 |  |  | 836 | 33437.701 | 10.4174 | 4.9741 |  |  | 836 | 33437.701 | 10.4174 | 4.9791
837 | 33477.699 | 10.4186 | 4.9714 |  |  | 837 | 33477.699 | 10.4186 | 4.974 |  |  | 837 | 33477.699 | 10.4186 | 4.979
838 | 33517.696 | 10.4198 | 4.9713 |  |  | 838 | 33517.696 | 10.4198 | 4.9739 |  |  | 838 | 33517.696 | 10.4198 | 4.9789
839 | 33557.693 | 10.421 | 4.9712 |  |  | 839 | 33557.693 | 10.421 | 4.9738 |  |  | 839 | 33557.693 | 10.421 | 4.9788
840 | 33597.69 | 10.4222 | 4.9711 |  |  | 840 | 33597.69 | 10.4222 | 4.9737 |  |  | 840 | 33597.69 | 10.4222 | 4.9787
841 | 33637.688 | 10.4234 | 4.971 |  |  | 841 | 33637.688 | 10.4234 | 4.9736 |  |  | 841 | 33637.688 | 10.4234 | 4.9786
842 | 33677.685 | 10.4246 | 4.9709 |  |  | 842 | 33677.685 | 10.4246 | 4.9735 |  |  | 842 | 33677.685 | 10.4246 | 4.9785
843 | 33717.682 | 10.4258 | 4.9708 |  |  | 843 | 33717.682 | 10.4258 | 4.9734 |  |  | 843 | 33717.682 | 10.4258 | 4.9784
844 | 33757.679 | 10.427 | 4.9707 |  |  | 844 | 33757.679 | 10.427 | 4.9733 |  |  | 844 | 33757.679 | 10.427 | 4.9783
845 | 33797.677 | 10.4281 | 4.9706 |  |  | 845 | 33797.677 | 10.4281 | 4.9732 |  |  | 845 | 33797.677 | 10.4281 | 4.9783
846 | 33837.674 | 10.4293 | 4.9705 |  |  | 846 | 33837.674 | 10.4293 | 4.9731 |  |  | 846 | 33837.674 | 10.4293 | 4.9782
847 | 33877.671 | 10.4305 | 4.9703 |  |  | 847 | 33877.671 | 10.4305 | 4.973 |  |  | 847 | 33877.671 | 10.4305 | 4.9781
848 | 33917.668 | 10.4317 | 4.9702 |  |  | 848 | 33917.668 | 10.4317 | 4.9729 |  |  | 848 | 33917.668 | 10.4317 | 4.978
849 | 33957.666 | 10.4329 | 4.9701 |  |  | 849 | 33957.666 | 10.4329 | 4.9728 |  |  | 849 | 33957.666 | 10.4329 | 4.9779
850 | 33997.663 | 10.434 | 4.97 |  |  | 850 | 33997.663 | 10.434 | 4.9727 |  |  | 850 | 33997.663 | 10.434 | 4.9778
851 | 34037.66 | 10.4352 | 4.9699 |  |  | 851 | 34037.66 | 10.4352 | 4.9726 |  |  | 851 | 34037.66 | 10.4352 | 4.9777
852 | 34077.657 | 10.4364 | 4.9698 |  |  | 852 | 34077.657 | 10.4364 | 4.9725 |  |  | 852 | 34077.657 | 10.4364 | 4.9776
853 | 34117.655 | 10.4376 | 4.9697 |  |  | 853 | 34117.655 | 10.4376 | 4.9724 |  |  | 853 | 34117.655 | 10.4376 | 4.9775
854 | 34157.652 | 10.4387 | 4.9696 |  |  | 854 | 34157.652 | 10.4387 | 4.9723 |  |  | 854 | 34157.652 | 10.4387 | 4.9774
855 | 34197.649 | 10.4399 | 4.9695 |  |  | 855 | 34197.649 | 10.4399 | 4.9722 |  |  | 855 | 34197.649 | 10.4399 | 4.9773
856 | 34237.646 | 10.4411 | 4.9694 |  |  | 856 | 34237.646 | 10.4411 | 4.9721 |  |  | 856 | 34237.646 | 10.4411 | 4.9773
857 | 34277.644 | 10.4422 | 4.9693 |  |  | 857 | 34277.644 | 10.4422 | 4.972 |  |  | 857 | 34277.644 | 10.4422 | 4.9772
858 | 34317.641 | 10.4434 | 4.9692 |  |  | 858 | 34317.641 | 10.4434 | 4.9719 |  |  | 858 | 34317.641 | 10.4434 | 4.9771
859 | 34357.638 | 10.4446 | 4.9691 |  |  | 859 | 34357.638 | 10.4446 | 4.9718 |  |  | 859 | 34357.638 | 10.4446 | 4.977
860 | 34397.635 | 10.4457 | 4.969 |  |  | 860 | 34397.635 | 10.4457 | 4.9717 |  |  | 860 | 34397.635 | 10.4457 | 4.9769
861 | 34437.633 | 10.4469 | 4.9689 |  |  | 861 | 34437.633 | 10.4469 | 4.9716 |  |  | 861 | 34437.633 | 10.4469 | 4.9768
862 | 34477.63 | 10.4481 | 4.9688 |  |  | 862 | 34477.63 | 10.4481 | 4.9715 |  |  | 862 | 34477.63 | 10.4481 | 4.9767
863 | 34517.627 | 10.4492 | 4.9687 |  |  | 863 | 34517.627 | 10.4492 | 4.9714 |  |  | 863 | 34517.627 | 10.4492 | 4.9766
864 | 34557.624 | 10.4504 | 4.9686 |  |  | 864 | 34557.624 | 10.4504 | 4.9713 |  |  | 864 | 34557.624 | 10.4504 | 4.9765
865 | 34597.622 | 10.4515 | 4.9685 |  |  | 865 | 34597.622 | 10.4515 | 4.9712 |  |  | 865 | 34597.622 | 10.4515 | 4.9764
866 | 34637.619 | 10.4527 | 4.9684 |  |  | 866 | 34637.619 | 10.4527 | 4.9711 |  |  | 866 | 34637.619 | 10.4527 | 4.9763
867 | 34677.616 | 10.4538 | 4.9683 |  |  | 867 | 34677.616 | 10.4538 | 4.971 |  |  | 867 | 34677.616 | 10.4538 | 4.9762
868 | 34717.613 | 10.455 | 4.9681 |  |  | 868 | 34717.613 | 10.455 | 4.9709 |  |  | 868 | 34717.613 | 10.455 | 4.9762
869 | 34757.611 | 10.4562 | 4.968 |  |  | 869 | 34757.611 | 10.4562 | 4.9708 |  |  | 869 | 34757.611 | 10.4562 | 4.9761
870 | 34797.608 | 10.4573 | 4.9679 |  |  | 870 | 34797.608 | 10.4573 | 4.9707 |  |  | 870 | 34797.608 | 10.4573 | 4.976
871 | 34837.605 | 10.4585 | 4.9678 |  |  | 871 | 34837.605 | 10.4585 | 4.9706 |  |  | 871 | 34837.605 | 10.4585 | 4.9759
872 | 34877.602 | 10.4596 | 4.9677 |  |  | 872 | 34877.602 | 10.4596 | 4.9705 |  |  | 872 | 34877.602 | 10.4596 | 4.9758
873 | 34917.6 | 10.4607 | 4.9676 |  |  | 873 | 34917.6 | 10.4607 | 4.9704 |  |  | 873 | 34917.6 | 10.4607 | 4.9757
874 | 34957.597 | 10.4619 | 4.9675 |  |  | 874 | 34957.597 | 10.4619 | 4.9703 |  |  | 874 | 34957.597 | 10.4619 | 4.9756
875 | 34997.594 | 10.463 | 4.9674 |  |  | 875 | 34997.594 | 10.463 | 4.9702 |  |  | 875 | 34997.594 | 10.463 | 4.9755
876 | 35037.591 | 10.4642 | 4.9673 |  |  | 876 | 35037.591 | 10.4642 | 4.9701 |  |  | 876 | 35037.591 | 10.4642 | 4.9754
877 | 35077.589 | 10.4653 | 4.9672 |  |  | 877 | 35077.589 | 10.4653 | 4.97 |  |  | 877 | 35077.589 | 10.4653 | 4.9753
878 | 35117.586 | 10.4665 | 4.9671 |  |  | 878 | 35117.586 | 10.4665 | 4.9699 |  |  | 878 | 35117.586 | 10.4665 | 4.9752
879 | 35157.583 | 10.4676 | 4.967 |  |  | 879 | 35157.583 | 10.4676 | 4.9698 |  |  | 879 | 35157.583 | 10.4676 | 4.9752
880 | 35197.58 | 10.4687 | 4.9669 |  |  | 880 | 35197.58 | 10.4687 | 4.9697 |  |  | 880 | 35197.58 | 10.4687 | 4.9751
881 | 35237.578 | 10.4699 | 4.9668 |  |  | 881 | 35237.578 | 10.4699 | 4.9696 |  |  | 881 | 35237.578 | 10.4699 | 4.975
882 | 35277.575 | 10.471 | 4.9667 |  |  | 882 | 35277.575 | 10.471 | 4.9694 |  |  | 882 | 35277.575 | 10.471 | 4.9749
883 | 35317.572 | 10.4721 | 4.9666 |  |  | 883 | 35317.572 | 10.4721 | 4.9693 |  |  | 883 | 35317.572 | 10.4721 | 4.9748
884 | 35357.569 | 10.4733 | 4.9665 |  |  | 884 | 35357.569 | 10.4733 | 4.9692 |  |  | 884 | 35357.569 | 10.4733 | 4.9747
885 | 35397.567 | 10.4744 | 4.9664 |  |  | 885 | 35397.567 | 10.4744 | 4.9691 |  |  | 885 | 35397.567 | 10.4744 | 4.9746
886 | 35437.564 | 10.4755 | 4.9663 |  |  | 886 | 35437.564 | 10.4755 | 4.969 |  |  | 886 | 35437.564 | 10.4755 | 4.9745
887 | 35477.561 | 10.4767 | 4.9662 |  |  | 887 | 35477.561 | 10.4767 | 4.9689 |  |  | 887 | 35477.561 | 10.4767 | 4.9744
888 | 35517.558 | 10.4778 | 4.9661 |  |  | 888 | 35517.558 | 10.4778 | 4.9688 |  |  | 888 | 35517.558 | 10.4778 | 4.9743
889 | 35557.556 | 10.4789 | 4.9659 |  |  | 889 | 35557.556 | 10.4789 | 4.9687 |  |  | 889 | 35557.556 | 10.4789 | 4.9742
890 | 35597.553 | 10.48 | 4.9658 |  |  | 890 | 35597.553 | 10.48 | 4.9686 |  |  | 890 | 35597.553 | 10.48 | 4.9741
891 | 35637.55 | 10.4812 | 4.9657 |  |  | 891 | 35637.55 | 10.4812 | 4.9685 |  |  | 891 | 35637.55 | 10.4812 | 4.9741
892 | 35677.547 | 10.4823 | 4.9656 |  |  | 892 | 35677.547 | 10.4823 | 4.9684 |  |  | 892 | 35677.547 | 10.4823 | 4.974
893 | 35717.545 | 10.4834 | 4.9655 |  |  | 893 | 35717.545 | 10.4834 | 4.9683 |  |  | 893 | 35717.545 | 10.4834 | 4.9739
894 | 35757.542 | 10.4845 | 4.9654 |  |  | 894 | 35757.542 | 10.4845 | 4.9682 |  |  | 894 | 35757.542 | 10.4845 | 4.9738
895 | 35797.539 | 10.4856 | 4.9653 |  |  | 895 | 35797.539 | 10.4856 | 4.9681 |  |  | 895 | 35797.539 | 10.4856 | 4.9737
896 | 35837.536 | 10.4868 | 4.9652 |  |  | 896 | 35837.536 | 10.4868 | 4.968 |  |  | 896 | 35837.536 | 10.4868 | 4.9736
897 | 35877.534 | 10.4879 | 4.9651 |  |  | 897 | 35877.534 | 10.4879 | 4.9679 |  |  | 897 | 35877.534 | 10.4879 | 4.9735
898 | 35917.531 | 10.489 | 4.965 |  |  | 898 | 35917.531 | 10.489 | 4.9678 |  |  | 898 | 35917.531 | 10.489 | 4.9734
899 | 35957.528 | 10.4901 | 4.9649 |  |  | 899 | 35957.528 | 10.4901 | 4.9677 |  |  | 899 | 35957.528 | 10.4901 | 4.9733
900 | 35997.525 | 10.4912 | 4.9648 |  |  | 900 | 35997.525 | 10.4912 | 4.9676 |  |  | 900 | 35997.525 | 10.4912 | 4.9732
901 | 36037.523 | 10.4923 | 4.9647 |  |  | 901 | 36037.523 | 10.4923 | 4.9675 |  |  | 901 | 36037.523 | 10.4923 | 4.9731
902 | 36077.52 | 10.4934 | 4.9646 |  |  | 902 | 36077.52 | 10.4934 | 4.9674 |  |  | 902 | 36077.52 | 10.4934 | 4.973
903 | 36117.517 | 10.4945 | 4.9645 |  |  | 903 | 36117.517 | 10.4945 | 4.9673 |  |  | 903 | 36117.517 | 10.4945 | 4.973
904 | 36157.515 | 10.4956 | 4.9644 |  |  | 904 | 36157.515 | 10.4956 | 4.9672 |  |  | 904 | 36157.515 | 10.4956 | 4.9729
905 | 36197.512 | 10.4967 | 4.9643 |  |  | 905 | 36197.512 | 10.4967 | 4.9671 |  |  | 905 | 36197.512 | 10.4967 | 4.9728
906 | 36237.509 | 10.4978 | 4.9642 |  |  | 906 | 36237.509 | 10.4978 | 4.967 |  |  | 906 | 36237.509 | 10.4978 | 4.9727
907 | 36277.506 | 10.499 | 4.9641 |  |  | 907 | 36277.506 | 10.499 | 4.9669 |  |  | 907 | 36277.506 | 10.499 | 4.9726
908 | 36317.504 | 10.5001 | 4.9639 |  |  | 908 | 36317.504 | 10.5001 | 4.9668 |  |  | 908 | 36317.504 | 10.5001 | 4.9725
909 | 36357.501 | 10.5012 | 4.9638 |  |  | 909 | 36357.501 | 10.5012 | 4.9667 |  |  | 909 | 36357.501 | 10.5012 | 4.9724
910 | 36397.498 | 10.5023 | 4.9637 |  |  | 910 | 36397.498 | 10.5023 | 4.9666 |  |  | 910 | 36397.498 | 10.5023 | 4.9723
911 | 36437.495 | 10.5034 | 4.9636 |  |  | 911 | 36437.495 | 10.5034 | 4.9665 |  |  | 911 | 36437.495 | 10.5034 | 4.9722
912 | 36477.493 | 10.5045 | 4.9635 |  |  | 912 | 36477.493 | 10.5045 | 4.9664 |  |  | 912 | 36477.493 | 10.5045 | 4.9721
913 | 36517.49 | 10.5055 | 4.9634 |  |  | 913 | 36517.49 | 10.5055 | 4.9663 |  |  | 913 | 36517.49 | 10.5055 | 4.972
914 | 36557.487 | 10.5066 | 4.9633 |  |  | 914 | 36557.487 | 10.5066 | 4.9662 |  |  | 914 | 36557.487 | 10.5066 | 4.9719
915 | 36597.484 | 10.5077 | 4.9632 |  |  | 915 | 36597.484 | 10.5077 | 4.9661 |  |  | 915 | 36597.484 | 10.5077 | 4.9718
916 | 36637.482 | 10.5088 | 4.9631 |  |  | 916 | 36637.482 | 10.5088 | 4.966 |  |  | 916 | 36637.482 | 10.5088 | 4.9718
917 | 36677.479 | 10.5099 | 4.963 |  |  | 917 | 36677.479 | 10.5099 | 4.9659 |  |  | 917 | 36677.479 | 10.5099 | 4.9717
918 | 36717.476 | 10.511 | 4.9629 |  |  | 918 | 36717.476 | 10.511 | 4.9658 |  |  | 918 | 36717.476 | 10.511 | 4.9716
919 | 36757.473 | 10.5121 | 4.9628 |  |  | 919 | 36757.473 | 10.5121 | 4.9657 |  |  | 919 | 36757.473 | 10.5121 | 4.9715
920 | 36797.471 | 10.5132 | 4.9627 |  |  | 920 | 36797.471 | 10.5132 | 4.9656 |  |  | 920 | 36797.471 | 10.5132 | 4.9714
921 | 36837.468 | 10.5143 | 4.9626 |  |  | 921 | 36837.468 | 10.5143 | 4.9655 |  |  | 921 | 36837.468 | 10.5143 | 4.9713
922 | 36877.465 | 10.5154 | 4.9625 |  |  | 922 | 36877.465 | 10.5154 | 4.9654 |  |  | 922 | 36877.465 | 10.5154 | 4.9712
923 | 36917.462 | 10.5164 | 4.9624 |  |  | 923 | 36917.462 | 10.5164 | 4.9653 |  |  | 923 | 36917.462 | 10.5164 | 4.9711
924 | 36957.46 | 10.5175 | 4.9623 |  |  | 924 | 36957.46 | 10.5175 | 4.9652 |  |  | 924 | 36957.46 | 10.5175 | 4.971
925 | 36997.457 | 10.5186 | 4.9622 |  |  | 925 | 36997.457 | 10.5186 | 4.965 |  |  | 925 | 36997.457 | 10.5186 | 4.9709
926 | 37037.454 | 10.5197 | 4.962 |  |  | 926 | 37037.454 | 10.5197 | 4.9649 |  |  | 926 | 37037.454 | 10.5197 | 4.9708
927 | 37077.451 | 10.5208 | 4.9619 |  |  | 927 | 37077.451 | 10.5208 | 4.9648 |  |  | 927 | 37077.451 | 10.5208 | 4.9707
928 | 37117.449 | 10.5218 | 4.9618 |  |  | 928 | 37117.449 | 10.5218 | 4.9647 |  |  | 928 | 37117.449 | 10.5218 | 4.9706
929 | 37157.446 | 10.5229 | 4.9617 |  |  | 929 | 37157.446 | 10.5229 | 4.9646 |  |  | 929 | 37157.446 | 10.5229 | 4.9705
930 | 37197.443 | 10.524 | 4.9616 |  |  | 930 | 37197.443 | 10.524 | 4.9645 |  |  | 930 | 37197.443 | 10.524 | 4.9705
931 | 37237.44 | 10.5251 | 4.9615 |  |  | 931 | 37237.44 | 10.5251 | 4.9644 |  |  | 931 | 37237.44 | 10.5251 | 4.9704
932 | 37277.438 | 10.5261 | 4.9614 |  |  | 932 | 37277.438 | 10.5261 | 4.9643 |  |  | 932 | 37277.438 | 10.5261 | 4.9703
933 | 37317.435 | 10.5272 | 4.9613 |  |  | 933 | 37317.435 | 10.5272 | 4.9642 |  |  | 933 | 37317.435 | 10.5272 | 4.9702
934 | 37357.432 | 10.5283 | 4.9612 |  |  | 934 | 37357.432 | 10.5283 | 4.9641 |  |  | 934 | 37357.432 | 10.5283 | 4.9701
935 | 37397.429 | 10.5294 | 4.9611 |  |  | 935 | 37397.429 | 10.5294 | 4.964 |  |  | 935 | 37397.429 | 10.5294 | 4.97
936 | 37437.427 | 10.5304 | 4.961 |  |  | 936 | 37437.427 | 10.5304 | 4.9639 |  |  | 936 | 37437.427 | 10.5304 | 4.9699
937 | 37477.424 | 10.5315 | 4.9609 |  |  | 937 | 37477.424 | 10.5315 | 4.9638 |  |  | 937 | 37477.424 | 10.5315 | 4.9698
938 | 37517.421 | 10.5326 | 4.9608 |  |  | 938 | 37517.421 | 10.5326 | 4.9637 |  |  | 938 | 37517.421 | 10.5326 | 4.9697
939 | 37557.418 | 10.5336 | 4.9607 |  |  | 939 | 37557.418 | 10.5336 | 4.9636 |  |  | 939 | 37557.418 | 10.5336 | 4.9696
940 | 37597.416 | 10.5347 | 4.9606 |  |  | 940 | 37597.416 | 10.5347 | 4.9635 |  |  | 940 | 37597.416 | 10.5347 | 4.9695
941 | 37637.413 | 10.5358 | 4.9605 |  |  | 941 | 37637.413 | 10.5358 | 4.9634 |  |  | 941 | 37637.413 | 10.5358 | 4.9694
942 | 37677.41 | 10.5368 | 4.9604 |  |  | 942 | 37677.41 | 10.5368 | 4.9633 |  |  | 942 | 37677.41 | 10.5368 | 4.9693
943 | 37717.407 | 10.5379 | 4.9602 |  |  | 943 | 37717.407 | 10.5379 | 4.9632 |  |  | 943 | 37717.407 | 10.5379 | 4.9692
944 | 37757.405 | 10.5389 | 4.9601 |  |  | 944 | 37757.405 | 10.5389 | 4.9631 |  |  | 944 | 37757.405 | 10.5389 | 4.9692
945 | 37797.402 | 10.54 | 4.96 |  |  | 945 | 37797.402 | 10.54 | 4.963 |  |  | 945 | 37797.402 | 10.54 | 4.9691
946 | 37837.399 | 10.5411 | 4.9599 |  |  | 946 | 37837.399 | 10.5411 | 4.9629 |  |  | 946 | 37837.399 | 10.5411 | 4.969
947 | 37877.396 | 10.5421 | 4.9598 |  |  | 947 | 37877.396 | 10.5421 | 4.9628 |  |  | 947 | 37877.396 | 10.5421 | 4.9689
948 | 37917.394 | 10.5432 | 4.9597 |  |  | 948 | 37917.394 | 10.5432 | 4.9627 |  |  | 948 | 37917.394 | 10.5432 | 4.9688
949 | 37957.391 | 10.5442 | 4.9596 |  |  | 949 | 37957.391 | 10.5442 | 4.9625 |  |  | 949 | 37957.391 | 10.5442 | 4.9687
950 | 37997.388 | 10.5453 | 4.9595 |  |  | 950 | 37997.388 | 10.5453 | 4.9624 |  |  | 950 | 37997.388 | 10.5453 | 4.9686
951 | 38037.385 | 10.5463 | 4.9594 |  |  | 951 | 38037.385 | 10.5463 | 4.9623 |  |  | 951 | 38037.385 | 10.5463 | 4.9685
952 | 38077.383 | 10.5474 | 4.9593 |  |  | 952 | 38077.383 | 10.5474 | 4.9622 |  |  | 952 | 38077.383 | 10.5474 | 4.9684
953 | 38117.38 | 10.5484 | 4.9592 |  |  | 953 | 38117.38 | 10.5484 | 4.9621 |  |  | 953 | 38117.38 | 10.5484 | 4.9683
954 | 38157.377 | 10.5495 | 4.9591 |  |  | 954 | 38157.377 | 10.5495 | 4.962 |  |  | 954 | 38157.377 | 10.5495 | 4.9682
955 | 38197.374 | 10.5505 | 4.959 |  |  | 955 | 38197.374 | 10.5505 | 4.9619 |  |  | 955 | 38197.374 | 10.5505 | 4.9681
956 | 38237.372 | 10.5516 | 4.9589 |  |  | 956 | 38237.372 | 10.5516 | 4.9618 |  |  | 956 | 38237.372 | 10.5516 | 4.968
957 | 38277.369 | 10.5526 | 4.9588 |  |  | 957 | 38277.369 | 10.5526 | 4.9617 |  |  | 957 | 38277.369 | 10.5526 | 4.9679
958 | 38317.366 | 10.5537 | 4.9587 |  |  | 958 | 38317.366 | 10.5537 | 4.9616 |  |  | 958 | 38317.366 | 10.5537 | 4.9678
959 | 38357.363 | 10.5547 | 4.9585 |  |  | 959 | 38357.363 | 10.5547 | 4.9615 |  |  | 959 | 38357.363 | 10.5547 | 4.9678
960 | 38397.361 | 10.5557 | 4.9584 |  |  | 960 | 38397.361 | 10.5557 | 4.9614 |  |  | 960 | 38397.361 | 10.5557 | 4.9677
961 | 38437.358 | 10.5568 | 4.9583 |  |  | 961 | 38437.358 | 10.5568 | 4.9613 |  |  | 961 | 38437.358 | 10.5568 | 4.9676
962 | 38477.355 | 10.5578 | 4.9582 |  |  | 962 | 38477.355 | 10.5578 | 4.9612 |  |  | 962 | 38477.355 | 10.5578 | 4.9675
963 | 38517.352 | 10.5589 | 4.9581 |  |  | 963 | 38517.352 | 10.5589 | 4.9611 |  |  | 963 | 38517.352 | 10.5589 | 4.9674
964 | 38557.35 | 10.5599 | 4.958 |  |  | 964 | 38557.35 | 10.5599 | 4.961 |  |  | 964 | 38557.35 | 10.5599 | 4.9673
965 | 38597.347 | 10.5609 | 4.9579 |  |  | 965 | 38597.347 | 10.5609 | 4.9609 |  |  | 965 | 38597.347 | 10.5609 | 4.9672
966 | 38637.344 | 10.562 | 4.9578 |  |  | 966 | 38637.344 | 10.562 | 4.9608 |  |  | 966 | 38637.344 | 10.562 | 4.9671
967 | 38677.341 | 10.563 | 4.9577 |  |  | 967 | 38677.341 | 10.563 | 4.9607 |  |  | 967 | 38677.341 | 10.563 | 4.967
968 | 38717.339 | 10.564 | 4.9576 |  |  | 968 | 38717.339 | 10.564 | 4.9605 |  |  | 968 | 38717.339 | 10.564 | 4.9669
969 | 38757.336 | 10.5651 | 4.9575 |  |  | 969 | 38757.336 | 10.5651 | 4.9604 |  |  | 969 | 38757.336 | 10.5651 | 4.9668
970 | 38797.333 | 10.5661 | 4.9574 |  |  | 970 | 38797.333 | 10.5661 | 4.9603 |  |  | 970 | 38797.333 | 10.5661 | 4.9667
971 | 38837.33 | 10.5671 | 4.9573 |  |  | 971 | 38837.33 | 10.5671 | 4.9602 |  |  | 971 | 38837.33 | 10.5671 | 4.9666
972 | 38877.328 | 10.5682 | 4.9572 |  |  | 972 | 38877.328 | 10.5682 | 4.9601 |  |  | 972 | 38877.328 | 10.5682 | 4.9665
973 | 38917.325 | 10.5692 | 4.957 |  |  | 973 | 38917.325 | 10.5692 | 4.96 |  |  | 973 | 38917.325 | 10.5692 | 4.9664
974 | 38957.322 | 10.5702 | 4.9569 |  |  | 974 | 38957.322 | 10.5702 | 4.9599 |  |  | 974 | 38957.322 | 10.5702 | 4.9663
975 | 38997.319 | 10.5712 | 4.9568 |  |  | 975 | 38997.319 | 10.5712 | 4.9598 |  |  | 975 | 38997.319 | 10.5712 | 4.9662
976 | 39037.317 | 10.5723 | 4.9567 |  |  | 976 | 39037.317 | 10.5723 | 4.9597 |  |  | 976 | 39037.317 | 10.5723 | 4.9662
977 | 39077.314 | 10.5733 | 4.9566 |  |  | 977 | 39077.314 | 10.5733 | 4.9596 |  |  | 977 | 39077.314 | 10.5733 | 4.9661
978 | 39117.311 | 10.5743 | 4.9565 |  |  | 978 | 39117.311 | 10.5743 | 4.9595 |  |  | 978 | 39117.311 | 10.5743 | 4.966
979 | 39157.308 | 10.5753 | 4.9564 |  |  | 979 | 39157.308 | 10.5753 | 4.9594 |  |  | 979 | 39157.308 | 10.5753 | 4.9659
980 | 39197.306 | 10.5764 | 4.9563 |  |  | 980 | 39197.306 | 10.5764 | 4.9593 |  |  | 980 | 39197.306 | 10.5764 | 4.9658
981 | 39237.303 | 10.5774 | 4.9562 |  |  | 981 | 39237.303 | 10.5774 | 4.9592 |  |  | 981 | 39237.303 | 10.5774 | 4.9657
982 | 39277.3 | 10.5784 | 4.9561 |  |  | 982 | 39277.3 | 10.5784 | 4.9591 |  |  | 982 | 39277.3 | 10.5784 | 4.9656
983 | 39317.297 | 10.5794 | 4.956 |  |  | 983 | 39317.297 | 10.5794 | 4.959 |  |  | 983 | 39317.297 | 10.5794 | 4.9655
984 | 39357.295 | 10.5804 | 4.9559 |  |  | 984 | 39357.295 | 10.5804 | 4.9588 |  |  | 984 | 39357.295 | 10.5804 | 4.9654
985 | 39397.292 | 10.5815 | 4.9558 |  |  | 985 | 39397.292 | 10.5815 | 4.9587 |  |  | 985 | 39397.292 | 10.5815 | 4.9653
986 | 39437.289 | 10.5825 | 4.9557 |  |  | 986 | 39437.289 | 10.5825 | 4.9586 |  |  | 986 | 39437.289 | 10.5825 | 4.9652
987 | 39477.286 | 10.5835 | 4.9555 |  |  | 987 | 39477.286 | 10.5835 | 4.9585 |  |  | 987 | 39477.286 | 10.5835 | 4.9651
988 | 39517.284 | 10.5845 | 4.9554 |  |  | 988 | 39517.284 | 10.5845 | 4.9584 |  |  | 988 | 39517.284 | 10.5845 | 4.965
989 | 39557.281 | 10.5855 | 4.9553 |  |  | 989 | 39557.281 | 10.5855 | 4.9583 |  |  | 989 | 39557.281 | 10.5855 | 4.9649
990 | 39597.278 | 10.5865 | 4.9552 |  |  | 990 | 39597.278 | 10.5865 | 4.9582 |  |  | 990 | 39597.278 | 10.5865 | 4.9648

## Run9
MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR
File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run9\1 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run9\2 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run9\3 01 001.RAW
Cursor Location: |  |  | (178, 193) |  |  | Cursor Location: |  |  | (179, 193) |  |  | Cursor Location: |  |  | (179, 195)
Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15)
Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev
1 | 39.997 | 3.6888 | 7.7384 | 5.1311 | 0.32376 | 1 | 39.997 | 3.6888 | 7.743 | 5.1358 | 0.32514 | 1 | 39.997 | 3.6888 | 7.7513 | 5.1356 | 0.32573
2 | 79.995 | 4.382 | 7.4065 |  |  | 2 | 79.995 | 4.382 | 7.4117 |  |  | 2 | 79.995 | 4.382 | 7.4152
3 | 119.992 | 4.7874 | 7.2121 |  |  | 3 | 119.992 | 4.7874 | 7.2178 |  |  | 3 | 119.992 | 4.7874 | 7.22
4 | 159.989 | 5.0751 | 7.0719 |  |  | 4 | 159.989 | 5.0751 | 7.0781 |  |  | 4 | 159.989 | 5.0751 | 7.0797
5 | 199.986 | 5.2982 | 6.9618 |  |  | 5 | 199.986 | 5.2982 | 6.9684 |  |  | 5 | 199.986 | 5.2982 | 6.9696
6 | 239.984 | 5.4806 | 6.871 |  |  | 6 | 239.984 | 5.4806 | 6.8779 |  |  | 6 | 239.984 | 5.4806 | 6.8788
7 | 279.981 | 5.6347 | 6.7936 |  |  | 7 | 279.981 | 5.6347 | 6.8008 |  |  | 7 | 279.981 | 5.6347 | 6.8015
8 | 319.978 | 5.7683 | 6.7263 |  |  | 8 | 319.978 | 5.7683 | 6.7337 |  |  | 8 | 319.978 | 5.7683 | 6.7342
9 | 359.975 | 5.886 | 6.6667 |  |  | 9 | 359.975 | 5.886 | 6.6742 |  |  | 9 | 359.975 | 5.886 | 6.6746
10 | 399.973 | 5.9914 | 6.6132 |  |  | 10 | 399.973 | 5.9914 | 6.6209 |  |  | 10 | 399.973 | 5.9914 | 6.6211
11 | 439.97 | 6.0867 | 6.5647 |  |  | 11 | 439.97 | 6.0867 | 6.5725 |  |  | 11 | 439.97 | 6.0867 | 6.5727
12 | 479.967 | 6.1737 | 6.5203 |  |  | 12 | 479.967 | 6.1737 | 6.5283 |  |  | 12 | 479.967 | 6.1737 | 6.5284
13 | 519.964 | 6.2538 | 6.4795 |  |  | 13 | 519.964 | 6.2538 | 6.4876 |  |  | 13 | 519.964 | 6.2538 | 6.4876
14 | 559.962 | 6.3279 | 6.4417 |  |  | 14 | 559.962 | 6.3279 | 6.4499 |  |  | 14 | 559.962 | 6.3279 | 6.4498
15 | 599.959 | 6.3969 | 6.4064 |  |  | 15 | 599.959 | 6.3969 | 6.4147 |  |  | 15 | 599.959 | 6.3969 | 6.4146
16 | 639.956 | 6.4614 | 6.3734 |  |  | 16 | 639.956 | 6.4614 | 6.3818 |  |  | 16 | 639.956 | 6.4614 | 6.3817
17 | 679.953 | 6.522 | 6.3424 |  |  | 17 | 679.953 | 6.522 | 6.3509 |  |  | 17 | 679.953 | 6.522 | 6.3508
18 | 719.951 | 6.5792 | 6.3132 |  |  | 18 | 719.951 | 6.5792 | 6.3217 |  |  | 18 | 719.951 | 6.5792 | 6.3216
19 | 759.948 | 6.6332 | 6.2856 |  |  | 19 | 759.948 | 6.6332 | 6.2941 |  |  | 19 | 759.948 | 6.6332 | 6.294
20 | 799.945 | 6.6845 | 6.2593 |  |  | 20 | 799.945 | 6.6845 | 6.268 |  |  | 20 | 799.945 | 6.6845 | 6.2678
21 | 839.942 | 6.7333 | 6.2344 |  |  | 21 | 839.942 | 6.7333 | 6.2431 |  |  | 21 | 839.942 | 6.7333 | 6.2429
22 | 879.94 | 6.7799 | 6.2106 |  |  | 22 | 879.94 | 6.7799 | 6.2193 |  |  | 22 | 879.94 | 6.7799 | 6.2192
23 | 919.937 | 6.8243 | 6.1879 |  |  | 23 | 919.937 | 6.8243 | 6.1966 |  |  | 23 | 919.937 | 6.8243 | 6.1965
24 | 959.934 | 6.8669 | 6.1661 |  |  | 24 | 959.934 | 6.8669 | 6.1749 |  |  | 24 | 959.934 | 6.8669 | 6.1748
25 | 999.931 | 6.9077 | 6.1453 |  |  | 25 | 999.931 | 6.9077 | 6.1541 |  |  | 25 | 999.931 | 6.9077 | 6.154
26 | 1039.929 | 6.9469 | 6.1253 |  |  | 26 | 1039.929 | 6.9469 | 6.1341 |  |  | 26 | 1039.929 | 6.9469 | 6.134
27 | 1079.926 | 6.9846 | 6.106 |  |  | 27 | 1079.926 | 6.9846 | 6.1149 |  |  | 27 | 1079.926 | 6.9846 | 6.1148
28 | 1119.923 | 7.021 | 6.0874 |  |  | 28 | 1119.923 | 7.021 | 6.0963 |  |  | 28 | 1119.923 | 7.021 | 6.0963
29 | 1159.92 | 7.0561 | 6.0695 |  |  | 29 | 1159.92 | 7.0561 | 6.0785 |  |  | 29 | 1159.92 | 7.0561 | 6.0784
30 | 1199.918 | 7.09 | 6.0523 |  |  | 30 | 1199.918 | 7.09 | 6.0612 |  |  | 30 | 1199.918 | 7.09 | 6.0612
31 | 1239.915 | 7.1228 | 6.0355 |  |  | 31 | 1239.915 | 7.1228 | 6.0445 |  |  | 31 | 1239.915 | 7.1228 | 6.0445
32 | 1279.912 | 7.1545 | 6.0194 |  |  | 32 | 1279.912 | 7.1545 | 6.0284 |  |  | 32 | 1279.912 | 7.1545 | 6.0283
33 | 1319.909 | 7.1853 | 6.0037 |  |  | 33 | 1319.909 | 7.1853 | 6.0127 |  |  | 33 | 1319.909 | 7.1853 | 6.0127
34 | 1359.907 | 7.2152 | 5.9885 |  |  | 34 | 1359.907 | 7.2152 | 5.9975 |  |  | 34 | 1359.907 | 7.2152 | 5.9976
35 | 1399.904 | 7.2442 | 5.9738 |  |  | 35 | 1399.904 | 7.2442 | 5.9828 |  |  | 35 | 1399.904 | 7.2442 | 5.9828
36 | 1439.901 | 7.2723 | 5.9595 |  |  | 36 | 1439.901 | 7.2723 | 5.9685 |  |  | 36 | 1439.901 | 7.2723 | 5.9686
37 | 1479.898 | 7.2997 | 5.9455 |  |  | 37 | 1479.898 | 7.2997 | 5.9546 |  |  | 37 | 1479.898 | 7.2997 | 5.9547
38 | 1519.896 | 7.3264 | 5.932 |  |  | 38 | 1519.896 | 7.3264 | 5.9411 |  |  | 38 | 1519.896 | 7.3264 | 5.9412
39 | 1559.893 | 7.3524 | 5.9188 |  |  | 39 | 1559.893 | 7.3524 | 5.9279 |  |  | 39 | 1559.893 | 7.3524 | 5.928
40 | 1599.89 | 7.3777 | 5.906 |  |  | 40 | 1599.89 | 7.3777 | 5.9151 |  |  | 40 | 1599.89 | 7.3777 | 5.9152
41 | 1639.887 | 7.4024 | 5.8935 |  |  | 41 | 1639.887 | 7.4024 | 5.9026 |  |  | 41 | 1639.887 | 7.4024 | 5.9027
42 | 1679.885 | 7.4265 | 5.8813 |  |  | 42 | 1679.885 | 7.4265 | 5.8904 |  |  | 42 | 1679.885 | 7.4265 | 5.8906
43 | 1719.882 | 7.45 | 5.8694 |  |  | 43 | 1719.882 | 7.45 | 5.8785 |  |  | 43 | 1719.882 | 7.45 | 5.8787
44 | 1759.879 | 7.473 | 5.8578 |  |  | 44 | 1759.879 | 7.473 | 5.8669 |  |  | 44 | 1759.879 | 7.473 | 5.8671
45 | 1799.876 | 7.4955 | 5.8465 |  |  | 45 | 1799.876 | 7.4955 | 5.8556 |  |  | 45 | 1799.876 | 7.4955 | 5.8558
46 | 1839.874 | 7.5175 | 5.8354 |  |  | 46 | 1839.874 | 7.5175 | 5.8445 |  |  | 46 | 1839.874 | 7.5175 | 5.8447
47 | 1879.871 | 7.539 | 5.8246 |  |  | 47 | 1879.871 | 7.539 | 5.8336 |  |  | 47 | 1879.871 | 7.539 | 5.8339
48 | 1919.868 | 7.56 | 5.814 |  |  | 48 | 1919.868 | 7.56 | 5.8231 |  |  | 48 | 1919.868 | 7.56 | 5.8233
49 | 1959.865 | 7.5806 | 5.8036 |  |  | 49 | 1959.865 | 7.5806 | 5.8127 |  |  | 49 | 1959.865 | 7.5806 | 5.813
50 | 1999.863 | 7.6008 | 5.7935 |  |  | 50 | 1999.863 | 7.6008 | 5.8026 |  |  | 50 | 1999.863 | 7.6008 | 5.8028
51 | 2039.86 | 7.6206 | 5.7836 |  |  | 51 | 2039.86 | 7.6206 | 5.7926 |  |  | 51 | 2039.86 | 7.6206 | 5.7929
52 | 2079.857 | 7.6401 | 5.7739 |  |  | 52 | 2079.857 | 7.6401 | 5.7829 |  |  | 52 | 2079.857 | 7.6401 | 5.7832
53 | 2119.854 | 7.6591 | 5.7643 |  |  | 53 | 2119.854 | 7.6591 | 5.7734 |  |  | 53 | 2119.854 | 7.6591 | 5.7737
54 | 2159.852 | 7.6778 | 5.755 |  |  | 54 | 2159.852 | 7.6778 | 5.764 |  |  | 54 | 2159.852 | 7.6778 | 5.7644
55 | 2199.849 | 7.6961 | 5.7459 |  |  | 55 | 2199.849 | 7.6961 | 5.7549 |  |  | 55 | 2199.849 | 7.6961 | 5.7552
56 | 2239.846 | 7.7142 | 5.7369 |  |  | 56 | 2239.846 | 7.7142 | 5.7459 |  |  | 56 | 2239.846 | 7.7142 | 5.7463
57 | 2279.843 | 7.7319 | 5.7281 |  |  | 57 | 2279.843 | 7.7319 | 5.7371 |  |  | 57 | 2279.843 | 7.7319 | 5.7375
58 | 2319.841 | 7.7493 | 5.7194 |  |  | 58 | 2319.841 | 7.7493 | 5.7284 |  |  | 58 | 2319.841 | 7.7493 | 5.7288
59 | 2359.838 | 7.7663 | 5.711 |  |  | 59 | 2359.838 | 7.7663 | 5.7199 |  |  | 59 | 2359.838 | 7.7663 | 5.7203
60 | 2399.835 | 7.7832 | 5.7026 |  |  | 60 | 2399.835 | 7.7832 | 5.7116 |  |  | 60 | 2399.835 | 7.7832 | 5.712
61 | 2439.832 | 7.7997 | 5.6945 |  |  | 61 | 2439.832 | 7.7997 | 5.7034 |  |  | 61 | 2439.832 | 7.7997 | 5.7038
62 | 2479.83 | 7.8159 | 5.6864 |  |  | 62 | 2479.83 | 7.8159 | 5.6954 |  |  | 62 | 2479.83 | 7.8159 | 5.6958
63 | 2519.827 | 7.8319 | 5.6786 |  |  | 63 | 2519.827 | 7.8319 | 5.6875 |  |  | 63 | 2519.827 | 7.8319 | 5.6879
64 | 2559.824 | 7.8477 | 5.6708 |  |  | 64 | 2559.824 | 7.8477 | 5.6797 |  |  | 64 | 2559.824 | 7.8477 | 5.6802
65 | 2599.821 | 7.8632 | 5.6632 |  |  | 65 | 2599.821 | 7.8632 | 5.6721 |  |  | 65 | 2599.821 | 7.8632 | 5.6726
66 | 2639.819 | 7.8785 | 5.6557 |  |  | 66 | 2639.819 | 7.8785 | 5.6646 |  |  | 66 | 2639.819 | 7.8785 | 5.6651
67 | 2679.816 | 7.8935 | 5.6483 |  |  | 67 | 2679.816 | 7.8935 | 5.6572 |  |  | 67 | 2679.816 | 7.8935 | 5.6577
68 | 2719.813 | 7.9083 | 5.6411 |  |  | 68 | 2719.813 | 7.9083 | 5.6499 |  |  | 68 | 2719.813 | 7.9083 | 5.6505
69 | 2759.81 | 7.9229 | 5.634 |  |  | 69 | 2759.81 | 7.9229 | 5.6428 |  |  | 69 | 2759.81 | 7.9229 | 5.6433
70 | 2799.808 | 7.9373 | 5.627 |  |  | 70 | 2799.808 | 7.9373 | 5.6358 |  |  | 70 | 2799.808 | 7.9373 | 5.6363
71 | 2839.805 | 7.9515 | 5.6201 |  |  | 71 | 2839.805 | 7.9515 | 5.6289 |  |  | 71 | 2839.805 | 7.9515 | 5.6294
72 | 2879.802 | 7.9655 | 5.6133 |  |  | 72 | 2879.802 | 7.9655 | 5.6221 |  |  | 72 | 2879.802 | 7.9655 | 5.6226
73 | 2919.799 | 7.9793 | 5.6066 |  |  | 73 | 2919.799 | 7.9793 | 5.6154 |  |  | 73 | 2919.799 | 7.9793 | 5.6159
74 | 2959.797 | 7.9929 | 5.6 |  |  | 74 | 2959.797 | 7.9929 | 5.6088 |  |  | 74 | 2959.797 | 7.9929 | 5.6094
75 | 2999.794 | 8.0063 | 5.5936 |  |  | 75 | 2999.794 | 8.0063 | 5.6023 |  |  | 75 | 2999.794 | 8.0063 | 5.6029
76 | 3039.791 | 8.0195 | 5.5872 |  |  | 76 | 3039.791 | 8.0195 | 5.5959 |  |  | 76 | 3039.791 | 8.0195 | 5.5965
77 | 3079.788 | 8.0326 | 5.5809 |  |  | 77 | 3079.788 | 8.0326 | 5.5896 |  |  | 77 | 3079.788 | 8.0326 | 5.5902
78 | 3119.786 | 8.0455 | 5.5747 |  |  | 78 | 3119.786 | 8.0455 | 5.5834 |  |  | 78 | 3119.786 | 8.0455 | 5.584
79 | 3159.783 | 8.0583 | 5.5686 |  |  | 79 | 3159.783 | 8.0583 | 5.5773 |  |  | 79 | 3159.783 | 8.0583 | 5.5779
80 | 3199.78 | 8.0708 | 5.5626 |  |  | 80 | 3199.78 | 8.0708 | 5.5713 |  |  | 80 | 3199.78 | 8.0708 | 5.5719
81 | 3239.777 | 8.0833 | 5.5567 |  |  | 81 | 3239.777 | 8.0833 | 5.5653 |  |  | 81 | 3239.777 | 8.0833 | 5.566
82 | 3279.775 | 8.0955 | 5.5509 |  |  | 82 | 3279.775 | 8.0955 | 5.5595 |  |  | 82 | 3279.775 | 8.0955 | 5.5601
83 | 3319.772 | 8.1077 | 5.5451 |  |  | 83 | 3319.772 | 8.1077 | 5.5537 |  |  | 83 | 3319.772 | 8.1077 | 5.5544
84 | 3359.769 | 8.1196 | 5.5394 |  |  | 84 | 3359.769 | 8.1196 | 5.548 |  |  | 84 | 3359.769 | 8.1196 | 5.5487
85 | 3399.766 | 8.1315 | 5.5339 |  |  | 85 | 3399.766 | 8.1315 | 5.5424 |  |  | 85 | 3399.766 | 8.1315 | 5.5431
86 | 3439.764 | 8.1432 | 5.5283 |  |  | 86 | 3439.764 | 8.1432 | 5.5369 |  |  | 86 | 3439.764 | 8.1432 | 5.5376
87 | 3479.761 | 8.1547 | 5.5229 |  |  | 87 | 3479.761 | 8.1547 | 5.5314 |  |  | 87 | 3479.761 | 8.1547 | 5.5321
88 | 3519.758 | 8.1661 | 5.5175 |  |  | 88 | 3519.758 | 8.1661 | 5.5261 |  |  | 88 | 3519.758 | 8.1661 | 5.5267
89 | 3559.755 | 8.1774 | 5.5123 |  |  | 89 | 3559.755 | 8.1774 | 5.5208 |  |  | 89 | 3559.755 | 8.1774 | 5.5214
90 | 3599.753 | 8.1886 | 5.507 |  |  | 90 | 3599.753 | 8.1886 | 5.5155 |  |  | 90 | 3599.753 | 8.1886 | 5.5162
91 | 3639.75 | 8.1997 | 5.5019 |  |  | 91 | 3639.75 | 8.1997 | 5.5104 |  |  | 91 | 3639.75 | 8.1997 | 5.5111
92 | 3679.747 | 8.2106 | 5.4968 |  |  | 92 | 3679.747 | 8.2106 | 5.5053 |  |  | 92 | 3679.747 | 8.2106 | 5.506
93 | 3719.744 | 8.2214 | 5.4918 |  |  | 93 | 3719.744 | 8.2214 | 5.5002 |  |  | 93 | 3719.744 | 8.2214 | 5.501
94 | 3759.742 | 8.2321 | 5.4869 |  |  | 94 | 3759.742 | 8.2321 | 5.4953 |  |  | 94 | 3759.742 | 8.2321 | 5.496
95 | 3799.739 | 8.2427 | 5.482 |  |  | 95 | 3799.739 | 8.2427 | 5.4904 |  |  | 95 | 3799.739 | 8.2427 | 5.4911
96 | 3839.736 | 8.2532 | 5.4772 |  |  | 96 | 3839.736 | 8.2532 | 5.4855 |  |  | 96 | 3839.736 | 8.2532 | 5.4863
97 | 3879.733 | 8.2635 | 5.4724 |  |  | 97 | 3879.733 | 8.2635 | 5.4808 |  |  | 97 | 3879.733 | 8.2635 | 5.4815
98 | 3919.731 | 8.2738 | 5.4677 |  |  | 98 | 3919.731 | 8.2738 | 5.4761 |  |  | 98 | 3919.731 | 8.2738 | 5.4768
99 | 3959.728 | 8.2839 | 5.4631 |  |  | 99 | 3959.728 | 8.2839 | 5.4714 |  |  | 99 | 3959.728 | 8.2839 | 5.4722
100 | 3999.725 | 8.294 | 5.4585 |  |  | 100 | 3999.725 | 8.294 | 5.4668 |  |  | 100 | 3999.725 | 8.294 | 5.4676
101 | 4039.722 | 8.3039 | 5.454 |  |  | 101 | 4039.722 | 8.3039 | 5.4623 |  |  | 101 | 4039.722 | 8.3039 | 5.4631
102 | 4079.72 | 8.3138 | 5.4496 |  |  | 102 | 4079.72 | 8.3138 | 5.4578 |  |  | 102 | 4079.72 | 8.3138 | 5.4586
103 | 4119.717 | 8.3235 | 5.4452 |  |  | 103 | 4119.717 | 8.3235 | 5.4534 |  |  | 103 | 4119.717 | 8.3235 | 5.4542
104 | 4159.714 | 8.3332 | 5.4408 |  |  | 104 | 4159.714 | 8.3332 | 5.449 |  |  | 104 | 4159.714 | 8.3332 | 5.4498
105 | 4199.711 | 8.3428 | 5.4365 |  |  | 105 | 4199.711 | 8.3428 | 5.4447 |  |  | 105 | 4199.711 | 8.3428 | 5.4455
106 | 4239.709 | 8.3522 | 5.4323 |  |  | 106 | 4239.709 | 8.3522 | 5.4404 |  |  | 106 | 4239.709 | 8.3522 | 5.4412
107 | 4279.706 | 8.3616 | 5.4281 |  |  | 107 | 4279.706 | 8.3616 | 5.4362 |  |  | 107 | 4279.706 | 8.3616 | 5.437
108 | 4319.703 | 8.3709 | 5.424 |  |  | 108 | 4319.703 | 8.3709 | 5.4321 |  |  | 108 | 4319.703 | 8.3709 | 5.4329
109 | 4359.7 | 8.3802 | 5.4199 |  |  | 109 | 4359.7 | 8.3802 | 5.428 |  |  | 109 | 4359.7 | 8.3802 | 5.4288
110 | 4399.698 | 8.3893 | 5.4158 |  |  | 110 | 4399.698 | 8.3893 | 5.4239 |  |  | 110 | 4399.698 | 8.3893 | 5.4247
111 | 4439.695 | 8.3983 | 5.4118 |  |  | 111 | 4439.695 | 8.3983 | 5.4199 |  |  | 111 | 4439.695 | 8.3983 | 5.4207
112 | 4479.692 | 8.4073 | 5.4079 |  |  | 112 | 4479.692 | 8.4073 | 5.4159 |  |  | 112 | 4479.692 | 8.4073 | 5.4168
113 | 4519.689 | 8.4162 | 5.404 |  |  | 113 | 4519.689 | 8.4162 | 5.412 |  |  | 113 | 4519.689 | 8.4162 | 5.4128
114 | 4559.687 | 8.425 | 5.4001 |  |  | 114 | 4559.687 | 8.425 | 5.4081 |  |  | 114 | 4559.687 | 8.425 | 5.409
115 | 4599.684 | 8.4337 | 5.3963 |  |  | 115 | 4599.684 | 8.4337 | 5.4043 |  |  | 115 | 4599.684 | 8.4337 | 5.4051
116 | 4639.681 | 8.4424 | 5.3926 |  |  | 116 | 4639.681 | 8.4424 | 5.4005 |  |  | 116 | 4639.681 | 8.4424 | 5.4014
117 | 4679.678 | 8.451 | 5.3889 |  |  | 117 | 4679.678 | 8.451 | 5.3968 |  |  | 117 | 4679.678 | 8.451 | 5.3976
118 | 4719.676 | 8.4595 | 5.3852 |  |  | 118 | 4719.676 | 8.4595 | 5.3931 |  |  | 118 | 4719.676 | 8.4595 | 5.3939
119 | 4759.673 | 8.4679 | 5.3815 |  |  | 119 | 4759.673 | 8.4679 | 5.3894 |  |  | 119 | 4759.673 | 8.4679 | 5.3903
120 | 4799.67 | 8.4763 | 5.378 |  |  | 120 | 4799.67 | 8.4763 | 5.3858 |  |  | 120 | 4799.67 | 8.4763 | 5.3867
121 | 4839.667 | 8.4846 | 5.3744 |  |  | 121 | 4839.667 | 8.4846 | 5.3822 |  |  | 121 | 4839.667 | 8.4846 | 5.3831
122 | 4879.665 | 8.4928 | 5.3709 |  |  | 122 | 4879.665 | 8.4928 | 5.3787 |  |  | 122 | 4879.665 | 8.4928 | 5.3796
123 | 4919.662 | 8.501 | 5.3674 |  |  | 123 | 4919.662 | 8.501 | 5.3752 |  |  | 123 | 4919.662 | 8.501 | 5.3761
124 | 4959.659 | 8.5091 | 5.364 |  |  | 124 | 4959.659 | 8.5091 | 5.3718 |  |  | 124 | 4959.659 | 8.5091 | 5.3727
125 | 4999.656 | 8.5171 | 5.3606 |  |  | 125 | 4999.656 | 8.5171 | 5.3683 |  |  | 125 | 4999.656 | 8.5171 | 5.3692
126 | 5039.654 | 8.5251 | 5.3572 |  |  | 126 | 5039.654 | 8.5251 | 5.365 |  |  | 126 | 5039.654 | 8.5251 | 5.3659
127 | 5079.651 | 8.533 | 5.3539 |  |  | 127 | 5079.651 | 8.533 | 5.3616 |  |  | 127 | 5079.651 | 8.533 | 5.3625
128 | 5119.648 | 8.5408 | 5.3506 |  |  | 128 | 5119.648 | 8.5408 | 5.3583 |  |  | 128 | 5119.648 | 8.5408 | 5.3592
129 | 5159.645 | 8.5486 | 5.3474 |  |  | 129 | 5159.645 | 8.5486 | 5.355 |  |  | 129 | 5159.645 | 8.5486 | 5.356
130 | 5199.643 | 8.5563 | 5.3442 |  |  | 130 | 5199.643 | 8.5563 | 5.3518 |  |  | 130 | 5199.643 | 8.5563 | 5.3528
131 | 5239.64 | 8.564 | 5.341 |  |  | 131 | 5239.64 | 8.564 | 5.3486 |  |  | 131 | 5239.64 | 8.564 | 5.3496
132 | 5279.637 | 8.5716 | 5.3379 |  |  | 132 | 5279.637 | 8.5716 | 5.3455 |  |  | 132 | 5279.637 | 8.5716 | 5.3464
133 | 5319.634 | 8.5792 | 5.3348 |  |  | 133 | 5319.634 | 8.5792 | 5.3423 |  |  | 133 | 5319.634 | 8.5792 | 5.3433
134 | 5359.632 | 8.5867 | 5.3317 |  |  | 134 | 5359.632 | 8.5867 | 5.3392 |  |  | 134 | 5359.632 | 8.5867 | 5.3402
135 | 5399.629 | 8.5941 | 5.3286 |  |  | 135 | 5399.629 | 8.5941 | 5.3362 |  |  | 135 | 5399.629 | 8.5941 | 5.3371
136 | 5439.626 | 8.6015 | 5.3256 |  |  | 136 | 5439.626 | 8.6015 | 5.3331 |  |  | 136 | 5439.626 | 8.6015 | 5.3341
137 | 5479.623 | 8.6088 | 5.3227 |  |  | 137 | 5479.623 | 8.6088 | 5.3302 |  |  | 137 | 5479.623 | 8.6088 | 5.3311
138 | 5519.621 | 8.6161 | 5.3197 |  |  | 138 | 5519.621 | 8.6161 | 5.3272 |  |  | 138 | 5519.621 | 8.6161 | 5.3282
139 | 5559.618 | 8.6233 | 5.3168 |  |  | 139 | 5559.618 | 8.6233 | 5.3243 |  |  | 139 | 5559.618 | 8.6233 | 5.3252
140 | 5599.615 | 8.6305 | 5.3139 |  |  | 140 | 5599.615 | 8.6305 | 5.3214 |  |  | 140 | 5599.615 | 8.6305 | 5.3223
141 | 5639.612 | 8.6376 | 5.3111 |  |  | 141 | 5639.612 | 8.6376 | 5.3185 |  |  | 141 | 5639.612 | 8.6376 | 5.3195
142 | 5679.61 | 8.6446 | 5.3082 |  |  | 142 | 5679.61 | 8.6446 | 5.3156 |  |  | 142 | 5679.61 | 8.6446 | 5.3166
143 | 5719.607 | 8.6517 | 5.3055 |  |  | 143 | 5719.607 | 8.6517 | 5.3128 |  |  | 143 | 5719.607 | 8.6517 | 5.3138
144 | 5759.604 | 8.6586 | 5.3027 |  |  | 144 | 5759.604 | 8.6586 | 5.31 |  |  | 144 | 5759.604 | 8.6586 | 5.3111
145 | 5799.601 | 8.6655 | 5.3 |  |  | 145 | 5799.601 | 8.6655 | 5.3073 |  |  | 145 | 5799.601 | 8.6655 | 5.3083
146 | 5839.599 | 8.6724 | 5.2973 |  |  | 146 | 5839.599 | 8.6724 | 5.3046 |  |  | 146 | 5839.599 | 8.6724 | 5.3056
147 | 5879.596 | 8.6792 | 5.2946 |  |  | 147 | 5879.596 | 8.6792 | 5.3019 |  |  | 147 | 5879.596 | 8.6792 | 5.3029
148 | 5919.593 | 8.686 | 5.2919 |  |  | 148 | 5919.593 | 8.686 | 5.2992 |  |  | 148 | 5919.593 | 8.686 | 5.3002
149 | 5959.59 | 8.6928 | 5.2893 |  |  | 149 | 5959.59 | 8.6928 | 5.2966 |  |  | 149 | 5959.59 | 8.6928 | 5.2976
150 | 5999.588 | 8.6994 | 5.2867 |  |  | 150 | 5999.588 | 8.6994 | 5.2939 |  |  | 150 | 5999.588 | 8.6994 | 5.295
151 | 6039.585 | 8.7061 | 5.2841 |  |  | 151 | 6039.585 | 8.7061 | 5.2913 |  |  | 151 | 6039.585 | 8.7061 | 5.2924
152 | 6079.582 | 8.7127 | 5.2816 |  |  | 152 | 6079.582 | 8.7127 | 5.2888 |  |  | 152 | 6079.582 | 8.7127 | 5.2898
153 | 6119.579 | 8.7192 | 5.2791 |  |  | 153 | 6119.579 | 8.7192 | 5.2862 |  |  | 153 | 6119.579 | 8.7192 | 5.2873
154 | 6159.577 | 8.7258 | 5.2766 |  |  | 154 | 6159.577 | 8.7258 | 5.2837 |  |  | 154 | 6159.577 | 8.7258 | 5.2848
155 | 6199.574 | 8.7322 | 5.2741 |  |  | 155 | 6199.574 | 8.7322 | 5.2812 |  |  | 155 | 6199.574 | 8.7322 | 5.2823
156 | 6239.571 | 8.7387 | 5.2717 |  |  | 156 | 6239.571 | 8.7387 | 5.2788 |  |  | 156 | 6239.571 | 8.7387 | 5.2799
157 | 6279.568 | 8.7451 | 5.2693 |  |  | 157 | 6279.568 | 8.7451 | 5.2763 |  |  | 157 | 6279.568 | 8.7451 | 5.2774
158 | 6319.566 | 8.7514 | 5.2669 |  |  | 158 | 6319.566 | 8.7514 | 5.2739 |  |  | 158 | 6319.566 | 8.7514 | 5.275
159 | 6359.563 | 8.7577 | 5.2645 |  |  | 159 | 6359.563 | 8.7577 | 5.2715 |  |  | 159 | 6359.563 | 8.7577 | 5.2726
160 | 6399.56 | 8.764 | 5.2622 |  |  | 160 | 6399.56 | 8.764 | 5.2692 |  |  | 160 | 6399.56 | 8.764 | 5.2703
161 | 6439.557 | 8.7702 | 5.2598 |  |  | 161 | 6439.557 | 8.7702 | 5.2668 |  |  | 161 | 6439.557 | 8.7702 | 5.2679
162 | 6479.555 | 8.7764 | 5.2575 |  |  | 162 | 6479.555 | 8.7764 | 5.2645 |  |  | 162 | 6479.555 | 8.7764 | 5.2656
163 | 6519.552 | 8.7826 | 5.2553 |  |  | 163 | 6519.552 | 8.7826 | 5.2622 |  |  | 163 | 6519.552 | 8.7826 | 5.2633
164 | 6559.549 | 8.7887 | 5.253 |  |  | 164 | 6559.549 | 8.7887 | 5.26 |  |  | 164 | 6559.549 | 8.7887 | 5.2611
165 | 6599.546 | 8.7948 | 5.2508 |  |  | 165 | 6599.546 | 8.7948 | 5.2577 |  |  | 165 | 6599.546 | 8.7948 | 5.2588
166 | 6639.544 | 8.8008 | 5.2486 |  |  | 166 | 6639.544 | 8.8008 | 5.2555 |  |  | 166 | 6639.544 | 8.8008 | 5.2566
167 | 6679.541 | 8.8068 | 5.2464 |  |  | 167 | 6679.541 | 8.8068 | 5.2533 |  |  | 167 | 6679.541 | 8.8068 | 5.2544
168 | 6719.538 | 8.8128 | 5.2442 |  |  | 168 | 6719.538 | 8.8128 | 5.2511 |  |  | 168 | 6719.538 | 8.8128 | 5.2522
169 | 6759.535 | 8.8187 | 5.2421 |  |  | 169 | 6759.535 | 8.8187 | 5.2489 |  |  | 169 | 6759.535 | 8.8187 | 5.2501
170 | 6799.533 | 8.8246 | 5.2399 |  |  | 170 | 6799.533 | 8.8246 | 5.2468 |  |  | 170 | 6799.533 | 8.8246 | 5.2479
171 | 6839.53 | 8.8305 | 5.2378 |  |  | 171 | 6839.53 | 8.8305 | 5.2446 |  |  | 171 | 6839.53 | 8.8305 | 5.2458
172 | 6879.527 | 8.8363 | 5.2358 |  |  | 172 | 6879.527 | 8.8363 | 5.2425 |  |  | 172 | 6879.527 | 8.8363 | 5.2437
173 | 6919.524 | 8.8421 | 5.2337 |  |  | 173 | 6919.524 | 8.8421 | 5.2405 |  |  | 173 | 6919.524 | 8.8421 | 5.2416
174 | 6959.522 | 8.8479 | 5.2317 |  |  | 174 | 6959.522 | 8.8479 | 5.2384 |  |  | 174 | 6959.522 | 8.8479 | 5.2396
175 | 6999.519 | 8.8536 | 5.2296 |  |  | 175 | 6999.519 | 8.8536 | 5.2364 |  |  | 175 | 6999.519 | 8.8536 | 5.2375
176 | 7039.516 | 8.8593 | 5.2276 |  |  | 176 | 7039.516 | 8.8593 | 5.2343 |  |  | 176 | 7039.516 | 8.8593 | 5.2355
177 | 7079.513 | 8.865 | 5.2256 |  |  | 177 | 7079.513 | 8.865 | 5.2323 |  |  | 177 | 7079.513 | 8.865 | 5.2335
178 | 7119.511 | 8.8706 | 5.2237 |  |  | 178 | 7119.511 | 8.8706 | 5.2303 |  |  | 178 | 7119.511 | 8.8706 | 5.2315
179 | 7159.508 | 8.8762 | 5.2217 |  |  | 179 | 7159.508 | 8.8762 | 5.2284 |  |  | 179 | 7159.508 | 8.8762 | 5.2296
180 | 7199.505 | 8.8818 | 5.2198 |  |  | 180 | 7199.505 | 8.8818 | 5.2264 |  |  | 180 | 7199.505 | 8.8818 | 5.2276
181 | 7239.502 | 8.8873 | 5.2179 |  |  | 181 | 7239.502 | 8.8873 | 5.2245 |  |  | 181 | 7239.502 | 8.8873 | 5.2257
182 | 7279.5 | 8.8928 | 5.216 |  |  | 182 | 7279.5 | 8.8928 | 5.2226 |  |  | 182 | 7279.5 | 8.8928 | 5.2238
183 | 7319.497 | 8.8983 | 5.2141 |  |  | 183 | 7319.497 | 8.8983 | 5.2207 |  |  | 183 | 7319.497 | 8.8983 | 5.2219
184 | 7359.494 | 8.9037 | 5.2123 |  |  | 184 | 7359.494 | 8.9037 | 5.2188 |  |  | 184 | 7359.494 | 8.9037 | 5.22
185 | 7399.491 | 8.9092 | 5.2104 |  |  | 185 | 7399.491 | 8.9092 | 5.217 |  |  | 185 | 7399.491 | 8.9092 | 5.2182
186 | 7439.489 | 8.9146 | 5.2086 |  |  | 186 | 7439.489 | 8.9146 | 5.2151 |  |  | 186 | 7439.489 | 8.9146 | 5.2163
187 | 7479.486 | 8.9199 | 5.2068 |  |  | 187 | 7479.486 | 8.9199 | 5.2133 |  |  | 187 | 7479.486 | 8.9199 | 5.2145
188 | 7519.483 | 8.9253 | 5.205 |  |  | 188 | 7519.483 | 8.9253 | 5.2115 |  |  | 188 | 7519.483 | 8.9253 | 5.2127
189 | 7559.48 | 8.9306 | 5.2032 |  |  | 189 | 7559.48 | 8.9306 | 5.2097 |  |  | 189 | 7559.48 | 8.9306 | 5.2109
190 | 7599.478 | 8.9358 | 5.2015 |  |  | 190 | 7599.478 | 8.9358 | 5.2079 |  |  | 190 | 7599.478 | 8.9358 | 5.2092
191 | 7639.475 | 8.9411 | 5.1997 |  |  | 191 | 7639.475 | 8.9411 | 5.2062 |  |  | 191 | 7639.475 | 8.9411 | 5.2074
192 | 7679.472 | 8.9463 | 5.198 |  |  | 192 | 7679.472 | 8.9463 | 5.2044 |  |  | 192 | 7679.472 | 8.9463 | 5.2057
193 | 7719.469 | 8.9515 | 5.1963 |  |  | 193 | 7719.469 | 8.9515 | 5.2027 |  |  | 193 | 7719.469 | 8.9515 | 5.204
194 | 7759.467 | 8.9567 | 5.1946 |  |  | 194 | 7759.467 | 8.9567 | 5.201 |  |  | 194 | 7759.467 | 8.9567 | 5.2023
195 | 7799.464 | 8.9618 | 5.1929 |  |  | 195 | 7799.464 | 8.9618 | 5.1993 |  |  | 195 | 7799.464 | 8.9618 | 5.2006
196 | 7839.461 | 8.9669 | 5.1913 |  |  | 196 | 7839.461 | 8.9669 | 5.1976 |  |  | 196 | 7839.461 | 8.9669 | 5.1989
197 | 7879.458 | 8.972 | 5.1896 |  |  | 197 | 7879.458 | 8.972 | 5.1959 |  |  | 197 | 7879.458 | 8.972 | 5.1972
198 | 7919.456 | 8.9771 | 5.188 |  |  | 198 | 7919.456 | 8.9771 | 5.1943 |  |  | 198 | 7919.456 | 8.9771 | 5.1956
199 | 7959.453 | 8.9821 | 5.1864 |  |  | 199 | 7959.453 | 8.9821 | 5.1927 |  |  | 199 | 7959.453 | 8.9821 | 5.194
200 | 7999.45 | 8.9871 | 5.1848 |  |  | 200 | 7999.45 | 8.9871 | 5.191 |  |  | 200 | 7999.45 | 8.9871 | 5.1923
201 | 8039.447 | 8.9921 | 5.1832 |  |  | 201 | 8039.447 | 8.9921 | 5.1894 |  |  | 201 | 8039.447 | 8.9921 | 5.1907
202 | 8079.445 | 8.9971 | 5.1816 |  |  | 202 | 8079.445 | 8.9971 | 5.1878 |  |  | 202 | 8079.445 | 8.9971 | 5.1892
203 | 8119.442 | 9.002 | 5.1801 |  |  | 203 | 8119.442 | 9.002 | 5.1863 |  |  | 203 | 8119.442 | 9.002 | 5.1876
204 | 8159.439 | 9.0069 | 5.1785 |  |  | 204 | 8159.439 | 9.0069 | 5.1847 |  |  | 204 | 8159.439 | 9.0069 | 5.186
205 | 8199.436 | 9.0118 | 5.177 |  |  | 205 | 8199.436 | 9.0118 | 5.1832 |  |  | 205 | 8199.436 | 9.0118 | 5.1845
206 | 8239.434 | 9.0167 | 5.1755 |  |  | 206 | 8239.434 | 9.0167 | 5.1816 |  |  | 206 | 8239.434 | 9.0167 | 5.183
207 | 8279.431 | 9.0215 | 5.174 |  |  | 207 | 8279.431 | 9.0215 | 5.1801 |  |  | 207 | 8279.431 | 9.0215 | 5.1814
208 | 8319.428 | 9.0263 | 5.1725 |  |  | 208 | 8319.428 | 9.0263 | 5.1786 |  |  | 208 | 8319.428 | 9.0263 | 5.1799
209 | 8359.425 | 9.0311 | 5.171 |  |  | 209 | 8359.425 | 9.0311 | 5.1771 |  |  | 209 | 8359.425 | 9.0311 | 5.1785
210 | 8399.423 | 9.0359 | 5.1695 |  |  | 210 | 8399.423 | 9.0359 | 5.1756 |  |  | 210 | 8399.423 | 9.0359 | 5.177
211 | 8439.42 | 9.0407 | 5.1681 |  |  | 211 | 8439.42 | 9.0407 | 5.1742 |  |  | 211 | 8439.42 | 9.0407 | 5.1755
212 | 8479.417 | 9.0454 | 5.1667 |  |  | 212 | 8479.417 | 9.0454 | 5.1727 |  |  | 212 | 8479.417 | 9.0454 | 5.1741
213 | 8519.414 | 9.0501 | 5.1652 |  |  | 213 | 8519.414 | 9.0501 | 5.1713 |  |  | 213 | 8519.414 | 9.0501 | 5.1726
214 | 8559.412 | 9.0548 | 5.1638 |  |  | 214 | 8559.412 | 9.0548 | 5.1698 |  |  | 214 | 8559.412 | 9.0548 | 5.1712
215 | 8599.409 | 9.0594 | 5.1624 |  |  | 215 | 8599.409 | 9.0594 | 5.1684 |  |  | 215 | 8599.409 | 9.0594 | 5.1698
216 | 8639.406 | 9.0641 | 5.161 |  |  | 216 | 8639.406 | 9.0641 | 5.167 |  |  | 216 | 8639.406 | 9.0641 | 5.1684
217 | 8679.403 | 9.0687 | 5.1597 |  |  | 217 | 8679.403 | 9.0687 | 5.1656 |  |  | 217 | 8679.403 | 9.0687 | 5.167
218 | 8719.401 | 9.0733 | 5.1583 |  |  | 218 | 8719.401 | 9.0733 | 5.1643 |  |  | 218 | 8719.401 | 9.0733 | 5.1656
219 | 8759.398 | 9.0779 | 5.1569 |  |  | 219 | 8759.398 | 9.0779 | 5.1629 |  |  | 219 | 8759.398 | 9.0779 | 5.1643
220 | 8799.395 | 9.0824 | 5.1556 |  |  | 220 | 8799.395 | 9.0824 | 5.1615 |  |  | 220 | 8799.395 | 9.0824 | 5.1629
221 | 8839.392 | 9.087 | 5.1543 |  |  | 221 | 8839.392 | 9.087 | 5.1602 |  |  | 221 | 8839.392 | 9.087 | 5.1616
222 | 8879.39 | 9.0915 | 5.153 |  |  | 222 | 8879.39 | 9.0915 | 5.1589 |  |  | 222 | 8879.39 | 9.0915 | 5.1603
223 | 8919.387 | 9.096 | 5.1517 |  |  | 223 | 8919.387 | 9.096 | 5.1575 |  |  | 223 | 8919.387 | 9.096 | 5.159
224 | 8959.384 | 9.1005 | 5.1504 |  |  | 224 | 8959.384 | 9.1005 | 5.1562 |  |  | 224 | 8959.384 | 9.1005 | 5.1577
225 | 8999.381 | 9.1049 | 5.1491 |  |  | 225 | 8999.381 | 9.1049 | 5.1549 |  |  | 225 | 8999.381 | 9.1049 | 5.1564
226 | 9039.379 | 9.1093 | 5.1478 |  |  | 226 | 9039.379 | 9.1093 | 5.1537 |  |  | 226 | 9039.379 | 9.1093 | 5.1551
227 | 9079.376 | 9.1138 | 5.1466 |  |  | 227 | 9079.376 | 9.1138 | 5.1524 |  |  | 227 | 9079.376 | 9.1138 | 5.1538
228 | 9119.373 | 9.1182 | 5.1453 |  |  | 228 | 9119.373 | 9.1182 | 5.1511 |  |  | 228 | 9119.373 | 9.1182 | 5.1526
229 | 9159.37 | 9.1225 | 5.1441 |  |  | 229 | 9159.37 | 9.1225 | 5.1499 |  |  | 229 | 9159.37 | 9.1225 | 5.1513
230 | 9199.368 | 9.1269 | 5.1428 |  |  | 230 | 9199.368 | 9.1269 | 5.1486 |  |  | 230 | 9199.368 | 9.1269 | 5.1501
231 | 9239.365 | 9.1312 | 5.1416 |  |  | 231 | 9239.365 | 9.1312 | 5.1474 |  |  | 231 | 9239.365 | 9.1312 | 5.1488
232 | 9279.362 | 9.1355 | 5.1404 |  |  | 232 | 9279.362 | 9.1355 | 5.1462 |  |  | 232 | 9279.362 | 9.1355 | 5.1476
233 | 9319.359 | 9.1398 | 5.1392 |  |  | 233 | 9319.359 | 9.1398 | 5.145 |  |  | 233 | 9319.359 | 9.1398 | 5.1464
234 | 9359.357 | 9.1441 | 5.138 |  |  | 234 | 9359.357 | 9.1441 | 5.1438 |  |  | 234 | 9359.357 | 9.1441 | 5.1452
235 | 9399.354 | 9.1484 | 5.1369 |  |  | 235 | 9399.354 | 9.1484 | 5.1426 |  |  | 235 | 9399.354 | 9.1484 | 5.144
236 | 9439.351 | 9.1526 | 5.1357 |  |  | 236 | 9439.351 | 9.1526 | 5.1414 |  |  | 236 | 9439.351 | 9.1526 | 5.1429
237 | 9479.348 | 9.1569 | 5.1346 |  |  | 237 | 9479.348 | 9.1569 | 5.1402 |  |  | 237 | 9479.348 | 9.1569 | 5.1417
238 | 9519.346 | 9.1611 | 5.1334 |  |  | 238 | 9519.346 | 9.1611 | 5.1391 |  |  | 238 | 9519.346 | 9.1611 | 5.1405
239 | 9559.343 | 9.1653 | 5.1323 |  |  | 239 | 9559.343 | 9.1653 | 5.1379 |  |  | 239 | 9559.343 | 9.1653 | 5.1394
240 | 9599.34 | 9.1694 | 5.1311 |  |  | 240 | 9599.34 | 9.1694 | 5.1368 |  |  | 240 | 9599.34 | 9.1694 | 5.1383
241 | 9639.337 | 9.1736 | 5.13 |  |  | 241 | 9639.337 | 9.1736 | 5.1356 |  |  | 241 | 9639.337 | 9.1736 | 5.1371
242 | 9679.335 | 9.1777 | 5.1289 |  |  | 242 | 9679.335 | 9.1777 | 5.1345 |  |  | 242 | 9679.335 | 9.1777 | 5.136
243 | 9719.332 | 9.1819 | 5.1278 |  |  | 243 | 9719.332 | 9.1819 | 5.1334 |  |  | 243 | 9719.332 | 9.1819 | 5.1349
244 | 9759.329 | 9.186 | 5.1267 |  |  | 244 | 9759.329 | 9.186 | 5.1323 |  |  | 244 | 9759.329 | 9.186 | 5.1338
245 | 9799.326 | 9.1901 | 5.1257 |  |  | 245 | 9799.326 | 9.1901 | 5.1312 |  |  | 245 | 9799.326 | 9.1901 | 5.1327
246 | 9839.324 | 9.1941 | 5.1246 |  |  | 246 | 9839.324 | 9.1941 | 5.1301 |  |  | 246 | 9839.324 | 9.1941 | 5.1316
247 | 9879.321 | 9.1982 | 5.1235 |  |  | 247 | 9879.321 | 9.1982 | 5.1291 |  |  | 247 | 9879.321 | 9.1982 | 5.1306
248 | 9919.318 | 9.2022 | 5.1225 |  |  | 248 | 9919.318 | 9.2022 | 5.128 |  |  | 248 | 9919.318 | 9.2022 | 5.1295
249 | 9959.315 | 9.2063 | 5.1214 |  |  | 249 | 9959.315 | 9.2063 | 5.1269 |  |  | 249 | 9959.315 | 9.2063 | 5.1285
250 | 9999.313 | 9.2103 | 5.1204 |  |  | 250 | 9999.313 | 9.2103 | 5.1259 |  |  | 250 | 9999.313 | 9.2103 | 5.1274
251 | 10039.31 | 9.2143 | 5.1194 |  |  | 251 | 10039.31 | 9.2143 | 5.1248 |  |  | 251 | 10039.31 | 9.2143 | 5.1264
252 | 10079.307 | 9.2182 | 5.1184 |  |  | 252 | 10079.307 | 9.2182 | 5.1238 |  |  | 252 | 10079.307 | 9.2182 | 5.1254
253 | 10119.304 | 9.2222 | 5.1174 |  |  | 253 | 10119.304 | 9.2222 | 5.1228 |  |  | 253 | 10119.304 | 9.2222 | 5.1243
254 | 10159.302 | 9.2261 | 5.1164 |  |  | 254 | 10159.302 | 9.2261 | 5.1218 |  |  | 254 | 10159.302 | 9.2261 | 5.1233
255 | 10199.299 | 9.2301 | 5.1154 |  |  | 255 | 10199.299 | 9.2301 | 5.1208 |  |  | 255 | 10199.299 | 9.2301 | 5.1223
256 | 10239.296 | 9.234 | 5.1144 |  |  | 256 | 10239.296 | 9.234 | 5.1198 |  |  | 256 | 10239.296 | 9.234 | 5.1213
257 | 10279.293 | 9.2379 | 5.1134 |  |  | 257 | 10279.293 | 9.2379 | 5.1188 |  |  | 257 | 10279.293 | 9.2379 | 5.1204
258 | 10319.291 | 9.2418 | 5.1124 |  |  | 258 | 10319.291 | 9.2418 | 5.1178 |  |  | 258 | 10319.291 | 9.2418 | 5.1194
259 | 10359.288 | 9.2456 | 5.1115 |  |  | 259 | 10359.288 | 9.2456 | 5.1168 |  |  | 259 | 10359.288 | 9.2456 | 5.1184
260 | 10399.285 | 9.2495 | 5.1105 |  |  | 260 | 10399.285 | 9.2495 | 5.1159 |  |  | 260 | 10399.285 | 9.2495 | 5.1174
261 | 10439.282 | 9.2533 | 5.1096 |  |  | 261 | 10439.282 | 9.2533 | 5.1149 |  |  | 261 | 10439.282 | 9.2533 | 5.1165
262 | 10479.28 | 9.2572 | 5.1086 |  |  | 262 | 10479.28 | 9.2572 | 5.114 |  |  | 262 | 10479.28 | 9.2572 | 5.1155
263 | 10519.277 | 9.261 | 5.1077 |  |  | 263 | 10519.277 | 9.261 | 5.113 |  |  | 263 | 10519.277 | 9.261 | 5.1146
264 | 10559.274 | 9.2648 | 5.1068 |  |  | 264 | 10559.274 | 9.2648 | 5.1121 |  |  | 264 | 10559.274 | 9.2648 | 5.1137
265 | 10599.271 | 9.2685 | 5.1059 |  |  | 265 | 10599.271 | 9.2685 | 5.1112 |  |  | 265 | 10599.271 | 9.2685 | 5.1128
266 | 10639.269 | 9.2723 | 5.105 |  |  | 266 | 10639.269 | 9.2723 | 5.1102 |  |  | 266 | 10639.269 | 9.2723 | 5.1118
267 | 10679.266 | 9.2761 | 5.1041 |  |  | 267 | 10679.266 | 9.2761 | 5.1093 |  |  | 267 | 10679.266 | 9.2761 | 5.1109
268 | 10719.263 | 9.2798 | 5.1032 |  |  | 268 | 10719.263 | 9.2798 | 5.1084 |  |  | 268 | 10719.263 | 9.2798 | 5.11
269 | 10759.26 | 9.2835 | 5.1023 |  |  | 269 | 10759.26 | 9.2835 | 5.1075 |  |  | 269 | 10759.26 | 9.2835 | 5.1091
270 | 10799.258 | 9.2872 | 5.1014 |  |  | 270 | 10799.258 | 9.2872 | 5.1066 |  |  | 270 | 10799.258 | 9.2872 | 5.1083
271 | 10839.255 | 9.2909 | 5.1006 |  |  | 271 | 10839.255 | 9.2909 | 5.1058 |  |  | 271 | 10839.255 | 9.2909 | 5.1074
272 | 10879.252 | 9.2946 | 5.0997 |  |  | 272 | 10879.252 | 9.2946 | 5.1049 |  |  | 272 | 10879.252 | 9.2946 | 5.1065
273 | 10919.249 | 9.2983 | 5.0988 |  |  | 273 | 10919.249 | 9.2983 | 5.104 |  |  | 273 | 10919.249 | 9.2983 | 5.1056
274 | 10959.247 | 9.3019 | 5.098 |  |  | 274 | 10959.247 | 9.3019 | 5.1032 |  |  | 274 | 10959.247 | 9.3019 | 5.1048
275 | 10999.244 | 9.3056 | 5.0972 |  |  | 275 | 10999.244 | 9.3056 | 5.1023 |  |  | 275 | 10999.244 | 9.3056 | 5.1039
276 | 11039.241 | 9.3092 | 5.0963 |  |  | 276 | 11039.241 | 9.3092 | 5.1015 |  |  | 276 | 11039.241 | 9.3092 | 5.1031
277 | 11079.238 | 9.3128 | 5.0955 |  |  | 277 | 11079.238 | 9.3128 | 5.1006 |  |  | 277 | 11079.238 | 9.3128 | 5.1022
278 | 11119.236 | 9.3164 | 5.0947 |  |  | 278 | 11119.236 | 9.3164 | 5.0998 |  |  | 278 | 11119.236 | 9.3164 | 5.1014
279 | 11159.233 | 9.32 | 5.0939 |  |  | 279 | 11159.233 | 9.32 | 5.099 |  |  | 279 | 11159.233 | 9.32 | 5.1006
280 | 11199.23 | 9.3236 | 5.093 |  |  | 280 | 11199.23 | 9.3236 | 5.0981 |  |  | 280 | 11199.23 | 9.3236 | 5.0998
281 | 11239.227 | 9.3272 | 5.0922 |  |  | 281 | 11239.227 | 9.3272 | 5.0973 |  |  | 281 | 11239.227 | 9.3272 | 5.099
282 | 11279.225 | 9.3307 | 5.0914 |  |  | 282 | 11279.225 | 9.3307 | 5.0965 |  |  | 282 | 11279.225 | 9.3307 | 5.0982
283 | 11319.222 | 9.3343 | 5.0907 |  |  | 283 | 11319.222 | 9.3343 | 5.0957 |  |  | 283 | 11319.222 | 9.3343 | 5.0974
284 | 11359.219 | 9.3378 | 5.0899 |  |  | 284 | 11359.219 | 9.3378 | 5.0949 |  |  | 284 | 11359.219 | 9.3378 | 5.0966
285 | 11399.216 | 9.3413 | 5.0891 |  |  | 285 | 11399.216 | 9.3413 | 5.0941 |  |  | 285 | 11399.216 | 9.3413 | 5.0958
286 | 11439.214 | 9.3448 | 5.0883 |  |  | 286 | 11439.214 | 9.3448 | 5.0933 |  |  | 286 | 11439.214 | 9.3448 | 5.095
287 | 11479.211 | 9.3483 | 5.0876 |  |  | 287 | 11479.211 | 9.3483 | 5.0926 |  |  | 287 | 11479.211 | 9.3483 | 5.0942
288 | 11519.208 | 9.3518 | 5.0868 |  |  | 288 | 11519.208 | 9.3518 | 5.0918 |  |  | 288 | 11519.208 | 9.3518 | 5.0935
289 | 11559.205 | 9.3552 | 5.0861 |  |  | 289 | 11559.205 | 9.3552 | 5.091 |  |  | 289 | 11559.205 | 9.3552 | 5.0927
290 | 11599.203 | 9.3587 | 5.0853 |  |  | 290 | 11599.203 | 9.3587 | 5.0903 |  |  | 290 | 11599.203 | 9.3587 | 5.092
291 | 11639.2 | 9.3621 | 5.0846 |  |  | 291 | 11639.2 | 9.3621 | 5.0895 |  |  | 291 | 11639.2 | 9.3621 | 5.0912
292 | 11679.197 | 9.3656 | 5.0838 |  |  | 292 | 11679.197 | 9.3656 | 5.0888 |  |  | 292 | 11679.197 | 9.3656 | 5.0905
293 | 11719.194 | 9.369 | 5.0831 |  |  | 293 | 11719.194 | 9.369 | 5.088 |  |  | 293 | 11719.194 | 9.369 | 5.0897
294 | 11759.192 | 9.3724 | 5.0824 |  |  | 294 | 11759.192 | 9.3724 | 5.0873 |  |  | 294 | 11759.192 | 9.3724 | 5.089
295 | 11799.189 | 9.3758 | 5.0817 |  |  | 295 | 11799.189 | 9.3758 | 5.0866 |  |  | 295 | 11799.189 | 9.3758 | 5.0883
296 | 11839.186 | 9.3792 | 5.0809 |  |  | 296 | 11839.186 | 9.3792 | 5.0858 |  |  | 296 | 11839.186 | 9.3792 | 5.0875
297 | 11879.183 | 9.3825 | 5.0802 |  |  | 297 | 11879.183 | 9.3825 | 5.0851 |  |  | 297 | 11879.183 | 9.3825 | 5.0868
298 | 11919.181 | 9.3859 | 5.0795 |  |  | 298 | 11919.181 | 9.3859 | 5.0844 |  |  | 298 | 11919.181 | 9.3859 | 5.0861
299 | 11959.178 | 9.3893 | 5.0788 |  |  | 299 | 11959.178 | 9.3893 | 5.0837 |  |  | 299 | 11959.178 | 9.3893 | 5.0854
300 | 11999.175 | 9.3926 | 5.0782 |  |  | 300 | 11999.175 | 9.3926 | 5.083 |  |  | 300 | 11999.175 | 9.3926 | 5.0847
301 | 12039.172 | 9.3959 | 5.0775 |  |  | 301 | 12039.172 | 9.3959 | 5.0823 |  |  | 301 | 12039.172 | 9.3959 | 5.084
302 | 12079.17 | 9.3992 | 5.0768 |  |  | 302 | 12079.17 | 9.3992 | 5.0816 |  |  | 302 | 12079.17 | 9.3992 | 5.0833
303 | 12119.167 | 9.4025 | 5.0761 |  |  | 303 | 12119.167 | 9.4025 | 5.0809 |  |  | 303 | 12119.167 | 9.4025 | 5.0827
304 | 12159.164 | 9.4058 | 5.0754 |  |  | 304 | 12159.164 | 9.4058 | 5.0803 |  |  | 304 | 12159.164 | 9.4058 | 5.082
305 | 12199.161 | 9.4091 | 5.0748 |  |  | 305 | 12199.161 | 9.4091 | 5.0796 |  |  | 305 | 12199.161 | 9.4091 | 5.0813
306 | 12239.159 | 9.4124 | 5.0741 |  |  | 306 | 12239.159 | 9.4124 | 5.0789 |  |  | 306 | 12239.159 | 9.4124 | 5.0806
307 | 12279.156 | 9.4157 | 5.0735 |  |  | 307 | 12279.156 | 9.4157 | 5.0783 |  |  | 307 | 12279.156 | 9.4157 | 5.08
308 | 12319.153 | 9.4189 | 5.0728 |  |  | 308 | 12319.153 | 9.4189 | 5.0776 |  |  | 308 | 12319.153 | 9.4189 | 5.0793
309 | 12359.15 | 9.4222 | 5.0722 |  |  | 309 | 12359.15 | 9.4222 | 5.0769 |  |  | 309 | 12359.15 | 9.4222 | 5.0787
310 | 12399.148 | 9.4254 | 5.0715 |  |  | 310 | 12399.148 | 9.4254 | 5.0763 |  |  | 310 | 12399.148 | 9.4254 | 5.078
311 | 12439.145 | 9.4286 | 5.0709 |  |  | 311 | 12439.145 | 9.4286 | 5.0756 |  |  | 311 | 12439.145 | 9.4286 | 5.0774
312 | 12479.142 | 9.4318 | 5.0703 |  |  | 312 | 12479.142 | 9.4318 | 5.075 |  |  | 312 | 12479.142 | 9.4318 | 5.0767
313 | 12519.139 | 9.435 | 5.0697 |  |  | 313 | 12519.139 | 9.435 | 5.0744 |  |  | 313 | 12519.139 | 9.435 | 5.0761
314 | 12559.137 | 9.4382 | 5.069 |  |  | 314 | 12559.137 | 9.4382 | 5.0737 |  |  | 314 | 12559.137 | 9.4382 | 5.0755
315 | 12599.134 | 9.4414 | 5.0684 |  |  | 315 | 12599.134 | 9.4414 | 5.0731 |  |  | 315 | 12599.134 | 9.4414 | 5.0749
316 | 12639.131 | 9.4446 | 5.0678 |  |  | 316 | 12639.131 | 9.4446 | 5.0725 |  |  | 316 | 12639.131 | 9.4446 | 5.0742
317 | 12679.128 | 9.4477 | 5.0672 |  |  | 317 | 12679.128 | 9.4477 | 5.0719 |  |  | 317 | 12679.128 | 9.4477 | 5.0736
318 | 12719.126 | 9.4509 | 5.0666 |  |  | 318 | 12719.126 | 9.4509 | 5.0713 |  |  | 318 | 12719.126 | 9.4509 | 5.073
319 | 12759.123 | 9.454 | 5.066 |  |  | 319 | 12759.123 | 9.454 | 5.0707 |  |  | 319 | 12759.123 | 9.454 | 5.0724
320 | 12799.12 | 9.4571 | 5.0654 |  |  | 320 | 12799.12 | 9.4571 | 5.0701 |  |  | 320 | 12799.12 | 9.4571 | 5.0718
321 | 12839.117 | 9.4603 | 5.0648 |  |  | 321 | 12839.117 | 9.4603 | 5.0695 |  |  | 321 | 12839.117 | 9.4603 | 5.0712
322 | 12879.115 | 9.4634 | 5.0642 |  |  | 322 | 12879.115 | 9.4634 | 5.0689 |  |  | 322 | 12879.115 | 9.4634 | 5.0706
323 | 12919.112 | 9.4665 | 5.0637 |  |  | 323 | 12919.112 | 9.4665 | 5.0683 |  |  | 323 | 12919.112 | 9.4665 | 5.07
324 | 12959.109 | 9.4696 | 5.0631 |  |  | 324 | 12959.109 | 9.4696 | 5.0677 |  |  | 324 | 12959.109 | 9.4696 | 5.0695
325 | 12999.106 | 9.4726 | 5.0625 |  |  | 325 | 12999.106 | 9.4726 | 5.0671 |  |  | 325 | 12999.106 | 9.4726 | 5.0689
326 | 13039.104 | 9.4757 | 5.062 |  |  | 326 | 13039.104 | 9.4757 | 5.0665 |  |  | 326 | 13039.104 | 9.4757 | 5.0683
327 | 13079.101 | 9.4788 | 5.0614 |  |  | 327 | 13079.101 | 9.4788 | 5.066 |  |  | 327 | 13079.101 | 9.4788 | 5.0677
328 | 13119.098 | 9.4818 | 5.0608 |  |  | 328 | 13119.098 | 9.4818 | 5.0654 |  |  | 328 | 13119.098 | 9.4818 | 5.0672
329 | 13159.095 | 9.4849 | 5.0603 |  |  | 329 | 13159.095 | 9.4849 | 5.0648 |  |  | 329 | 13159.095 | 9.4849 | 5.0666
330 | 13199.093 | 9.4879 | 5.0597 |  |  | 330 | 13199.093 | 9.4879 | 5.0643 |  |  | 330 | 13199.093 | 9.4879 | 5.066
331 | 13239.09 | 9.4909 | 5.0592 |  |  | 331 | 13239.09 | 9.4909 | 5.0637 |  |  | 331 | 13239.09 | 9.4909 | 5.0655
332 | 13279.087 | 9.4939 | 5.0586 |  |  | 332 | 13279.087 | 9.4939 | 5.0632 |  |  | 332 | 13279.087 | 9.4939 | 5.0649
333 | 13319.084 | 9.497 | 5.0581 |  |  | 333 | 13319.084 | 9.497 | 5.0626 |  |  | 333 | 13319.084 | 9.497 | 5.0644
334 | 13359.082 | 9.5 | 5.0576 |  |  | 334 | 13359.082 | 9.5 | 5.0621 |  |  | 334 | 13359.082 | 9.5 | 5.0639
335 | 13399.079 | 9.5029 | 5.057 |  |  | 335 | 13399.079 | 9.5029 | 5.0616 |  |  | 335 | 13399.079 | 9.5029 | 5.0633
336 | 13439.076 | 9.5059 | 5.0565 |  |  | 336 | 13439.076 | 9.5059 | 5.061 |  |  | 336 | 13439.076 | 9.5059 | 5.0628
337 | 13479.073 | 9.5089 | 5.056 |  |  | 337 | 13479.073 | 9.5089 | 5.0605 |  |  | 337 | 13479.073 | 9.5089 | 5.0623
338 | 13519.071 | 9.5119 | 5.0555 |  |  | 338 | 13519.071 | 9.5119 | 5.06 |  |  | 338 | 13519.071 | 9.5119 | 5.0617
339 | 13559.068 | 9.5148 | 5.055 |  |  | 339 | 13559.068 | 9.5148 | 5.0594 |  |  | 339 | 13559.068 | 9.5148 | 5.0612
340 | 13599.065 | 9.5178 | 5.0545 |  |  | 340 | 13599.065 | 9.5178 | 5.0589 |  |  | 340 | 13599.065 | 9.5178 | 5.0607
341 | 13639.062 | 9.5207 | 5.054 |  |  | 341 | 13639.062 | 9.5207 | 5.0584 |  |  | 341 | 13639.062 | 9.5207 | 5.0602
342 | 13679.06 | 9.5236 | 5.0535 |  |  | 342 | 13679.06 | 9.5236 | 5.0579 |  |  | 342 | 13679.06 | 9.5236 | 5.0597
343 | 13719.057 | 9.5265 | 5.053 |  |  | 343 | 13719.057 | 9.5265 | 5.0574 |  |  | 343 | 13719.057 | 9.5265 | 5.0592
344 | 13759.054 | 9.5295 | 5.0525 |  |  | 344 | 13759.054 | 9.5295 | 5.0569 |  |  | 344 | 13759.054 | 9.5295 | 5.0587
345 | 13799.051 | 9.5324 | 5.052 |  |  | 345 | 13799.051 | 9.5324 | 5.0564 |  |  | 345 | 13799.051 | 9.5324 | 5.0582
346 | 13839.049 | 9.5352 | 5.0515 |  |  | 346 | 13839.049 | 9.5352 | 5.0559 |  |  | 346 | 13839.049 | 9.5352 | 5.0577
347 | 13879.046 | 9.5381 | 5.051 |  |  | 347 | 13879.046 | 9.5381 | 5.0554 |  |  | 347 | 13879.046 | 9.5381 | 5.0572
348 | 13919.043 | 9.541 | 5.0505 |  |  | 348 | 13919.043 | 9.541 | 5.0549 |  |  | 348 | 13919.043 | 9.541 | 5.0567
349 | 13959.04 | 9.5439 | 5.05 |  |  | 349 | 13959.04 | 9.5439 | 5.0544 |  |  | 349 | 13959.04 | 9.5439 | 5.0562
350 | 13999.038 | 9.5467 | 5.0496 |  |  | 350 | 13999.038 | 9.5467 | 5.0539 |  |  | 350 | 13999.038 | 9.5467 | 5.0557
351 | 14039.035 | 9.5496 | 5.0491 |  |  | 351 | 14039.035 | 9.5496 | 5.0535 |  |  | 351 | 14039.035 | 9.5496 | 5.0552
352 | 14079.032 | 9.5524 | 5.0486 |  |  | 352 | 14079.032 | 9.5524 | 5.053 |  |  | 352 | 14079.032 | 9.5524 | 5.0548
353 | 14119.029 | 9.5553 | 5.0482 |  |  | 353 | 14119.029 | 9.5553 | 5.0525 |  |  | 353 | 14119.029 | 9.5553 | 5.0543
354 | 14159.027 | 9.5581 | 5.0477 |  |  | 354 | 14159.027 | 9.5581 | 5.052 |  |  | 354 | 14159.027 | 9.5581 | 5.0538
355 | 14199.024 | 9.5609 | 5.0472 |  |  | 355 | 14199.024 | 9.5609 | 5.0516 |  |  | 355 | 14199.024 | 9.5609 | 5.0534
356 | 14239.021 | 9.5637 | 5.0468 |  |  | 356 | 14239.021 | 9.5637 | 5.0511 |  |  | 356 | 14239.021 | 9.5637 | 5.0529
357 | 14279.018 | 9.5665 | 5.0463 |  |  | 357 | 14279.018 | 9.5665 | 5.0507 |  |  | 357 | 14279.018 | 9.5665 | 5.0524
358 | 14319.016 | 9.5693 | 5.0459 |  |  | 358 | 14319.016 | 9.5693 | 5.0502 |  |  | 358 | 14319.016 | 9.5693 | 5.052
359 | 14359.013 | 9.5721 | 5.0454 |  |  | 359 | 14359.013 | 9.5721 | 5.0498 |  |  | 359 | 14359.013 | 9.5721 | 5.0515
360 | 14399.01 | 9.5749 | 5.045 |  |  | 360 | 14399.01 | 9.5749 | 5.0493 |  |  | 360 | 14399.01 | 9.5749 | 5.0511
361 | 14439.007 | 9.5777 | 5.0446 |  |  | 361 | 14439.007 | 9.5777 | 5.0489 |  |  | 361 | 14439.007 | 9.5777 | 5.0506
362 | 14479.005 | 9.5805 | 5.0441 |  |  | 362 | 14479.005 | 9.5805 | 5.0484 |  |  | 362 | 14479.005 | 9.5805 | 5.0502
363 | 14519.002 | 9.5832 | 5.0437 |  |  | 363 | 14519.002 | 9.5832 | 5.048 |  |  | 363 | 14519.002 | 9.5832 | 5.0497
364 | 14558.999 | 9.586 | 5.0433 |  |  | 364 | 14558.999 | 9.586 | 5.0475 |  |  | 364 | 14558.999 | 9.586 | 5.0493
365 | 14598.996 | 9.5887 | 5.0428 |  |  | 365 | 14598.996 | 9.5887 | 5.0471 |  |  | 365 | 14598.996 | 9.5887 | 5.0489
366 | 14638.994 | 9.5914 | 5.0424 |  |  | 366 | 14638.994 | 9.5914 | 5.0467 |  |  | 366 | 14638.994 | 9.5914 | 5.0484
367 | 14678.991 | 9.5942 | 5.042 |  |  | 367 | 14678.991 | 9.5942 | 5.0462 |  |  | 367 | 14678.991 | 9.5942 | 5.048
368 | 14718.988 | 9.5969 | 5.0416 |  |  | 368 | 14718.988 | 9.5969 | 5.0458 |  |  | 368 | 14718.988 | 9.5969 | 5.0476
369 | 14758.985 | 9.5996 | 5.0412 |  |  | 369 | 14758.985 | 9.5996 | 5.0454 |  |  | 369 | 14758.985 | 9.5996 | 5.0472
370 | 14798.983 | 9.6023 | 5.0408 |  |  | 370 | 14798.983 | 9.6023 | 5.045 |  |  | 370 | 14798.983 | 9.6023 | 5.0467
371 | 14838.98 | 9.605 | 5.0403 |  |  | 371 | 14838.98 | 9.605 | 5.0446 |  |  | 371 | 14838.98 | 9.605 | 5.0463
372 | 14878.977 | 9.6077 | 5.0399 |  |  | 372 | 14878.977 | 9.6077 | 5.0442 |  |  | 372 | 14878.977 | 9.6077 | 5.0459
373 | 14918.974 | 9.6104 | 5.0395 |  |  | 373 | 14918.974 | 9.6104 | 5.0437 |  |  | 373 | 14918.974 | 9.6104 | 5.0455
374 | 14958.972 | 9.6131 | 5.0391 |  |  | 374 | 14958.972 | 9.6131 | 5.0433 |  |  | 374 | 14958.972 | 9.6131 | 5.0451
375 | 14998.969 | 9.6157 | 5.0387 |  |  | 375 | 14998.969 | 9.6157 | 5.0429 |  |  | 375 | 14998.969 | 9.6157 | 5.0447
376 | 15038.966 | 9.6184 | 5.0383 |  |  | 376 | 15038.966 | 9.6184 | 5.0425 |  |  | 376 | 15038.966 | 9.6184 | 5.0443
377 | 15078.963 | 9.6211 | 5.038 |  |  | 377 | 15078.963 | 9.6211 | 5.0421 |  |  | 377 | 15078.963 | 9.6211 | 5.0439
378 | 15118.961 | 9.6237 | 5.0376 |  |  | 378 | 15118.961 | 9.6237 | 5.0417 |  |  | 378 | 15118.961 | 9.6237 | 5.0435
379 | 15158.958 | 9.6263 | 5.0372 |  |  | 379 | 15158.958 | 9.6263 | 5.0413 |  |  | 379 | 15158.958 | 9.6263 | 5.0431
380 | 15198.955 | 9.629 | 5.0368 |  |  | 380 | 15198.955 | 9.629 | 5.0409 |  |  | 380 | 15198.955 | 9.629 | 5.0427
381 | 15238.952 | 9.6316 | 5.0364 |  |  | 381 | 15238.952 | 9.6316 | 5.0406 |  |  | 381 | 15238.952 | 9.6316 | 5.0423
382 | 15278.95 | 9.6342 | 5.036 |  |  | 382 | 15278.95 | 9.6342 | 5.0402 |  |  | 382 | 15278.95 | 9.6342 | 5.0419
383 | 15318.947 | 9.6368 | 5.0357 |  |  | 383 | 15318.947 | 9.6368 | 5.0398 |  |  | 383 | 15318.947 | 9.6368 | 5.0415
384 | 15358.944 | 9.6395 | 5.0353 |  |  | 384 | 15358.944 | 9.6395 | 5.0394 |  |  | 384 | 15358.944 | 9.6395 | 5.0412
385 | 15398.941 | 9.6421 | 5.0349 |  |  | 385 | 15398.941 | 9.6421 | 5.039 |  |  | 385 | 15398.941 | 9.6421 | 5.0408
386 | 15438.939 | 9.6446 | 5.0345 |  |  | 386 | 15438.939 | 9.6446 | 5.0387 |  |  | 386 | 15438.939 | 9.6446 | 5.0404
387 | 15478.936 | 9.6472 | 5.0342 |  |  | 387 | 15478.936 | 9.6472 | 5.0383 |  |  | 387 | 15478.936 | 9.6472 | 5.04
388 | 15518.933 | 9.6498 | 5.0338 |  |  | 388 | 15518.933 | 9.6498 | 5.0379 |  |  | 388 | 15518.933 | 9.6498 | 5.0397
389 | 15558.93 | 9.6524 | 5.0335 |  |  | 389 | 15558.93 | 9.6524 | 5.0375 |  |  | 389 | 15558.93 | 9.6524 | 5.0393
390 | 15598.928 | 9.655 | 5.0331 |  |  | 390 | 15598.928 | 9.655 | 5.0372 |  |  | 390 | 15598.928 | 9.655 | 5.0389
391 | 15638.925 | 9.6575 | 5.0327 |  |  | 391 | 15638.925 | 9.6575 | 5.0368 |  |  | 391 | 15638.925 | 9.6575 | 5.0386
392 | 15678.922 | 9.6601 | 5.0324 |  |  | 392 | 15678.922 | 9.6601 | 5.0365 |  |  | 392 | 15678.922 | 9.6601 | 5.0382
393 | 15718.919 | 9.6626 | 5.032 |  |  | 393 | 15718.919 | 9.6626 | 5.0361 |  |  | 393 | 15718.919 | 9.6626 | 5.0378
394 | 15758.917 | 9.6652 | 5.0317 |  |  | 394 | 15758.917 | 9.6652 | 5.0357 |  |  | 394 | 15758.917 | 9.6652 | 5.0375
395 | 15798.914 | 9.6677 | 5.0313 |  |  | 395 | 15798.914 | 9.6677 | 5.0354 |  |  | 395 | 15798.914 | 9.6677 | 5.0371
396 | 15838.911 | 9.6702 | 5.031 |  |  | 396 | 15838.911 | 9.6702 | 5.035 |  |  | 396 | 15838.911 | 9.6702 | 5.0368
397 | 15878.908 | 9.6727 | 5.0306 |  |  | 397 | 15878.908 | 9.6727 | 5.0347 |  |  | 397 | 15878.908 | 9.6727 | 5.0364
398 | 15918.906 | 9.6753 | 5.0303 |  |  | 398 | 15918.906 | 9.6753 | 5.0343 |  |  | 398 | 15918.906 | 9.6753 | 5.0361
399 | 15958.903 | 9.6778 | 5.03 |  |  | 399 | 15958.903 | 9.6778 | 5.034 |  |  | 399 | 15958.903 | 9.6778 | 5.0357
400 | 15998.9 | 9.6803 | 5.0296 |  |  | 400 | 15998.9 | 9.6803 | 5.0337 |  |  | 400 | 15998.9 | 9.6803 | 5.0354
401 | 16038.897 | 9.6828 | 5.0293 |  |  | 401 | 16038.897 | 9.6828 | 5.0333 |  |  | 401 | 16038.897 | 9.6828 | 5.035
402 | 16078.895 | 9.6853 | 5.029 |  |  | 402 | 16078.895 | 9.6853 | 5.033 |  |  | 402 | 16078.895 | 9.6853 | 5.0347
403 | 16118.892 | 9.6877 | 5.0286 |  |  | 403 | 16118.892 | 9.6877 | 5.0326 |  |  | 403 | 16118.892 | 9.6877 | 5.0344
404 | 16158.889 | 9.6902 | 5.0283 |  |  | 404 | 16158.889 | 9.6902 | 5.0323 |  |  | 404 | 16158.889 | 9.6902 | 5.034
405 | 16198.886 | 9.6927 | 5.028 |  |  | 405 | 16198.886 | 9.6927 | 5.032 |  |  | 405 | 16198.886 | 9.6927 | 5.0337
406 | 16238.884 | 9.6952 | 5.0277 |  |  | 406 | 16238.884 | 9.6952 | 5.0317 |  |  | 406 | 16238.884 | 9.6952 | 5.0334
407 | 16278.881 | 9.6976 | 5.0273 |  |  | 407 | 16278.881 | 9.6976 | 5.0313 |  |  | 407 | 16278.881 | 9.6976 | 5.033
408 | 16318.878 | 9.7001 | 5.027 |  |  | 408 | 16318.878 | 9.7001 | 5.031 |  |  | 408 | 16318.878 | 9.7001 | 5.0327
409 | 16358.875 | 9.7025 | 5.0267 |  |  | 409 | 16358.875 | 9.7025 | 5.0307 |  |  | 409 | 16358.875 | 9.7025 | 5.0324
410 | 16398.873 | 9.705 | 5.0264 |  |  | 410 | 16398.873 | 9.705 | 5.0304 |  |  | 410 | 16398.873 | 9.705 | 5.0321
411 | 16438.87 | 9.7074 | 5.0261 |  |  | 411 | 16438.87 | 9.7074 | 5.03 |  |  | 411 | 16438.87 | 9.7074 | 5.0317
412 | 16478.867 | 9.7098 | 5.0258 |  |  | 412 | 16478.867 | 9.7098 | 5.0297 |  |  | 412 | 16478.867 | 9.7098 | 5.0314
413 | 16518.864 | 9.7123 | 5.0255 |  |  | 413 | 16518.864 | 9.7123 | 5.0294 |  |  | 413 | 16518.864 | 9.7123 | 5.0311
414 | 16558.862 | 9.7147 | 5.0252 |  |  | 414 | 16558.862 | 9.7147 | 5.0291 |  |  | 414 | 16558.862 | 9.7147 | 5.0308
415 | 16598.859 | 9.7171 | 5.0249 |  |  | 415 | 16598.859 | 9.7171 | 5.0288 |  |  | 415 | 16598.859 | 9.7171 | 5.0305
416 | 16638.856 | 9.7195 | 5.0246 |  |  | 416 | 16638.856 | 9.7195 | 5.0285 |  |  | 416 | 16638.856 | 9.7195 | 5.0302
417 | 16678.853 | 9.7219 | 5.0242 |  |  | 417 | 16678.853 | 9.7219 | 5.0282 |  |  | 417 | 16678.853 | 9.7219 | 5.0299
418 | 16718.851 | 9.7243 | 5.024 |  |  | 418 | 16718.851 | 9.7243 | 5.0279 |  |  | 418 | 16718.851 | 9.7243 | 5.0296
419 | 16758.848 | 9.7267 | 5.0237 |  |  | 419 | 16758.848 | 9.7267 | 5.0276 |  |  | 419 | 16758.848 | 9.7267 | 5.0292
420 | 16798.845 | 9.7291 | 5.0234 |  |  | 420 | 16798.845 | 9.7291 | 5.0273 |  |  | 420 | 16798.845 | 9.7291 | 5.0289
421 | 16838.842 | 9.7314 | 5.0231 |  |  | 421 | 16838.842 | 9.7314 | 5.027 |  |  | 421 | 16838.842 | 9.7314 | 5.0286
422 | 16878.84 | 9.7338 | 5.0228 |  |  | 422 | 16878.84 | 9.7338 | 5.0267 |  |  | 422 | 16878.84 | 9.7338 | 5.0283
423 | 16918.837 | 9.7362 | 5.0225 |  |  | 423 | 16918.837 | 9.7362 | 5.0264 |  |  | 423 | 16918.837 | 9.7362 | 5.028
424 | 16958.834 | 9.7385 | 5.0222 |  |  | 424 | 16958.834 | 9.7385 | 5.0261 |  |  | 424 | 16958.834 | 9.7385 | 5.0277
425 | 16998.831 | 9.7409 | 5.0219 |  |  | 425 | 16998.831 | 9.7409 | 5.0258 |  |  | 425 | 16998.831 | 9.7409 | 5.0275
426 | 17038.829 | 9.7433 | 5.0216 |  |  | 426 | 17038.829 | 9.7433 | 5.0255 |  |  | 426 | 17038.829 | 9.7433 | 5.0272
427 | 17078.826 | 9.7456 | 5.0213 |  |  | 427 | 17078.826 | 9.7456 | 5.0252 |  |  | 427 | 17078.826 | 9.7456 | 5.0269
428 | 17118.823 | 9.7479 | 5.0211 |  |  | 428 | 17118.823 | 9.7479 | 5.0249 |  |  | 428 | 17118.823 | 9.7479 | 5.0266
429 | 17158.82 | 9.7503 | 5.0208 |  |  | 429 | 17158.82 | 9.7503 | 5.0246 |  |  | 429 | 17158.82 | 9.7503 | 5.0263
430 | 17198.818 | 9.7526 | 5.0205 |  |  | 430 | 17198.818 | 9.7526 | 5.0244 |  |  | 430 | 17198.818 | 9.7526 | 5.026
431 | 17238.815 | 9.7549 | 5.0202 |  |  | 431 | 17238.815 | 9.7549 | 5.0241 |  |  | 431 | 17238.815 | 9.7549 | 5.0257
432 | 17278.812 | 9.7572 | 5.02 |  |  | 432 | 17278.812 | 9.7572 | 5.0238 |  |  | 432 | 17278.812 | 9.7572 | 5.0254
433 | 17318.809 | 9.7595 | 5.0197 |  |  | 433 | 17318.809 | 9.7595 | 5.0235 |  |  | 433 | 17318.809 | 9.7595 | 5.0252
434 | 17358.807 | 9.7619 | 5.0194 |  |  | 434 | 17358.807 | 9.7619 | 5.0233 |  |  | 434 | 17358.807 | 9.7619 | 5.0249
435 | 17398.804 | 9.7642 | 5.0191 |  |  | 435 | 17398.804 | 9.7642 | 5.023 |  |  | 435 | 17398.804 | 9.7642 | 5.0246
436 | 17438.801 | 9.7665 | 5.0189 |  |  | 436 | 17438.801 | 9.7665 | 5.0227 |  |  | 436 | 17438.801 | 9.7665 | 5.0243
437 | 17478.798 | 9.7687 | 5.0186 |  |  | 437 | 17478.798 | 9.7687 | 5.0224 |  |  | 437 | 17478.798 | 9.7687 | 5.0241
438 | 17518.796 | 9.771 | 5.0183 |  |  | 438 | 17518.796 | 9.771 | 5.0222 |  |  | 438 | 17518.796 | 9.771 | 5.0238
439 | 17558.793 | 9.7733 | 5.0181 |  |  | 439 | 17558.793 | 9.7733 | 5.0219 |  |  | 439 | 17558.793 | 9.7733 | 5.0235
440 | 17598.79 | 9.7756 | 5.0178 |  |  | 440 | 17598.79 | 9.7756 | 5.0216 |  |  | 440 | 17598.79 | 9.7756 | 5.0232
441 | 17638.787 | 9.7779 | 5.0176 |  |  | 441 | 17638.787 | 9.7779 | 5.0214 |  |  | 441 | 17638.787 | 9.7779 | 5.023
442 | 17678.785 | 9.7801 | 5.0173 |  |  | 442 | 17678.785 | 9.7801 | 5.0211 |  |  | 442 | 17678.785 | 9.7801 | 5.0227
443 | 17718.782 | 9.7824 | 5.0171 |  |  | 443 | 17718.782 | 9.7824 | 5.0208 |  |  | 443 | 17718.782 | 9.7824 | 5.0224
444 | 17758.779 | 9.7846 | 5.0168 |  |  | 444 | 17758.779 | 9.7846 | 5.0206 |  |  | 444 | 17758.779 | 9.7846 | 5.0222
445 | 17798.776 | 9.7869 | 5.0165 |  |  | 445 | 17798.776 | 9.7869 | 5.0203 |  |  | 445 | 17798.776 | 9.7869 | 5.0219
446 | 17838.774 | 9.7891 | 5.0163 |  |  | 446 | 17838.774 | 9.7891 | 5.0201 |  |  | 446 | 17838.774 | 9.7891 | 5.0217
447 | 17878.771 | 9.7914 | 5.016 |  |  | 447 | 17878.771 | 9.7914 | 5.0198 |  |  | 447 | 17878.771 | 9.7914 | 5.0214
448 | 17918.768 | 9.7936 | 5.0158 |  |  | 448 | 17918.768 | 9.7936 | 5.0196 |  |  | 448 | 17918.768 | 9.7936 | 5.0211
449 | 17958.765 | 9.7958 | 5.0155 |  |  | 449 | 17958.765 | 9.7958 | 5.0193 |  |  | 449 | 17958.765 | 9.7958 | 5.0209
450 | 17998.763 | 9.7981 | 5.0153 |  |  | 450 | 17998.763 | 9.7981 | 5.0191 |  |  | 450 | 17998.763 | 9.7981 | 5.0206
451 | 18038.76 | 9.8003 | 5.0151 |  |  | 451 | 18038.76 | 9.8003 | 5.0188 |  |  | 451 | 18038.76 | 9.8003 | 5.0204
452 | 18078.757 | 9.8025 | 5.0148 |  |  | 452 | 18078.757 | 9.8025 | 5.0186 |  |  | 452 | 18078.757 | 9.8025 | 5.0201
453 | 18118.755 | 9.8047 | 5.0146 |  |  | 453 | 18118.755 | 9.8047 | 5.0183 |  |  | 453 | 18118.755 | 9.8047 | 5.0199
454 | 18158.752 | 9.8069 | 5.0143 |  |  | 454 | 18158.752 | 9.8069 | 5.0181 |  |  | 454 | 18158.752 | 9.8069 | 5.0196
455 | 18198.749 | 9.8091 | 5.0141 |  |  | 455 | 18198.749 | 9.8091 | 5.0178 |  |  | 455 | 18198.749 | 9.8091 | 5.0194
456 | 18238.746 | 9.8113 | 5.0139 |  |  | 456 | 18238.746 | 9.8113 | 5.0176 |  |  | 456 | 18238.746 | 9.8113 | 5.0191
457 | 18278.744 | 9.8135 | 5.0136 |  |  | 457 | 18278.744 | 9.8135 | 5.0173 |  |  | 457 | 18278.744 | 9.8135 | 5.0189
458 | 18318.741 | 9.8157 | 5.0134 |  |  | 458 | 18318.741 | 9.8157 | 5.0171 |  |  | 458 | 18318.741 | 9.8157 | 5.0186
459 | 18358.738 | 9.8179 | 5.0131 |  |  | 459 | 18358.738 | 9.8179 | 5.0169 |  |  | 459 | 18358.738 | 9.8179 | 5.0184
460 | 18398.735 | 9.82 | 5.0129 |  |  | 460 | 18398.735 | 9.82 | 5.0166 |  |  | 460 | 18398.735 | 9.82 | 5.0182
461 | 18438.733 | 9.8222 | 5.0127 |  |  | 461 | 18438.733 | 9.8222 | 5.0164 |  |  | 461 | 18438.733 | 9.8222 | 5.0179
462 | 18478.73 | 9.8244 | 5.0125 |  |  | 462 | 18478.73 | 9.8244 | 5.0162 |  |  | 462 | 18478.73 | 9.8244 | 5.0177
463 | 18518.727 | 9.8265 | 5.0122 |  |  | 463 | 18518.727 | 9.8265 | 5.0159 |  |  | 463 | 18518.727 | 9.8265 | 5.0174
464 | 18558.724 | 9.8287 | 5.012 |  |  | 464 | 18558.724 | 9.8287 | 5.0157 |  |  | 464 | 18558.724 | 9.8287 | 5.0172
465 | 18598.722 | 9.8308 | 5.0118 |  |  | 465 | 18598.722 | 9.8308 | 5.0155 |  |  | 465 | 18598.722 | 9.8308 | 5.017
466 | 18638.719 | 9.833 | 5.0115 |  |  | 466 | 18638.719 | 9.833 | 5.0152 |  |  | 466 | 18638.719 | 9.833 | 5.0167
467 | 18678.716 | 9.8351 | 5.0113 |  |  | 467 | 18678.716 | 9.8351 | 5.015 |  |  | 467 | 18678.716 | 9.8351 | 5.0165
468 | 18718.713 | 9.8373 | 5.0111 |  |  | 468 | 18718.713 | 9.8373 | 5.0148 |  |  | 468 | 18718.713 | 9.8373 | 5.0163
469 | 18758.711 | 9.8394 | 5.0109 |  |  | 469 | 18758.711 | 9.8394 | 5.0146 |  |  | 469 | 18758.711 | 9.8394 | 5.016
470 | 18798.708 | 9.8415 | 5.0107 |  |  | 470 | 18798.708 | 9.8415 | 5.0143 |  |  | 470 | 18798.708 | 9.8415 | 5.0158
471 | 18838.705 | 9.8437 | 5.0104 |  |  | 471 | 18838.705 | 9.8437 | 5.0141 |  |  | 471 | 18838.705 | 9.8437 | 5.0156
472 | 18878.702 | 9.8458 | 5.0102 |  |  | 472 | 18878.702 | 9.8458 | 5.0139 |  |  | 472 | 18878.702 | 9.8458 | 5.0154
473 | 18918.7 | 9.8479 | 5.01 |  |  | 473 | 18918.7 | 9.8479 | 5.0137 |  |  | 473 | 18918.7 | 9.8479 | 5.0151
474 | 18958.697 | 9.85 | 5.0098 |  |  | 474 | 18958.697 | 9.85 | 5.0135 |  |  | 474 | 18958.697 | 9.85 | 5.0149
475 | 18998.694 | 9.8521 | 5.0096 |  |  | 475 | 18998.694 | 9.8521 | 5.0132 |  |  | 475 | 18998.694 | 9.8521 | 5.0147
476 | 19038.691 | 9.8542 | 5.0094 |  |  | 476 | 19038.691 | 9.8542 | 5.013 |  |  | 476 | 19038.691 | 9.8542 | 5.0145
477 | 19078.689 | 9.8563 | 5.0092 |  |  | 477 | 19078.689 | 9.8563 | 5.0128 |  |  | 477 | 19078.689 | 9.8563 | 5.0142
478 | 19118.686 | 9.8584 | 5.0089 |  |  | 478 | 19118.686 | 9.8584 | 5.0126 |  |  | 478 | 19118.686 | 9.8584 | 5.014
479 | 19158.683 | 9.8605 | 5.0087 |  |  | 479 | 19158.683 | 9.8605 | 5.0124 |  |  | 479 | 19158.683 | 9.8605 | 5.0138
480 | 19198.68 | 9.8626 | 5.0085 |  |  | 480 | 19198.68 | 9.8626 | 5.0122 |  |  | 480 | 19198.68 | 9.8626 | 5.0136
481 | 19238.678 | 9.8647 | 5.0083 |  |  | 481 | 19238.678 | 9.8647 | 5.012 |  |  | 481 | 19238.678 | 9.8647 | 5.0134
482 | 19278.675 | 9.8668 | 5.0081 |  |  | 482 | 19278.675 | 9.8668 | 5.0118 |  |  | 482 | 19278.675 | 9.8668 | 5.0132
483 | 19318.672 | 9.8688 | 5.0079 |  |  | 483 | 19318.672 | 9.8688 | 5.0115 |  |  | 483 | 19318.672 | 9.8688 | 5.0129
484 | 19358.669 | 9.8709 | 5.0077 |  |  | 484 | 19358.669 | 9.8709 | 5.0113 |  |  | 484 | 19358.669 | 9.8709 | 5.0127
485 | 19398.667 | 9.873 | 5.0075 |  |  | 485 | 19398.667 | 9.873 | 5.0111 |  |  | 485 | 19398.667 | 9.873 | 5.0125
486 | 19438.664 | 9.875 | 5.0073 |  |  | 486 | 19438.664 | 9.875 | 5.0109 |  |  | 486 | 19438.664 | 9.875 | 5.0123
487 | 19478.661 | 9.8771 | 5.0071 |  |  | 487 | 19478.661 | 9.8771 | 5.0107 |  |  | 487 | 19478.661 | 9.8771 | 5.0121
488 | 19518.658 | 9.8791 | 5.0069 |  |  | 488 | 19518.658 | 9.8791 | 5.0105 |  |  | 488 | 19518.658 | 9.8791 | 5.0119
489 | 19558.656 | 9.8812 | 5.0067 |  |  | 489 | 19558.656 | 9.8812 | 5.0103 |  |  | 489 | 19558.656 | 9.8812 | 5.0117
490 | 19598.653 | 9.8832 | 5.0065 |  |  | 490 | 19598.653 | 9.8832 | 5.0101 |  |  | 490 | 19598.653 | 9.8832 | 5.0115
491 | 19638.65 | 9.8853 | 5.0063 |  |  | 491 | 19638.65 | 9.8853 | 5.0099 |  |  | 491 | 19638.65 | 9.8853 | 5.0113
492 | 19678.647 | 9.8873 | 5.0061 |  |  | 492 | 19678.647 | 9.8873 | 5.0097 |  |  | 492 | 19678.647 | 9.8873 | 5.0111
493 | 19718.645 | 9.8893 | 5.0059 |  |  | 493 | 19718.645 | 9.8893 | 5.0095 |  |  | 493 | 19718.645 | 9.8893 | 5.0109
494 | 19758.642 | 9.8913 | 5.0057 |  |  | 494 | 19758.642 | 9.8913 | 5.0093 |  |  | 494 | 19758.642 | 9.8913 | 5.0107
495 | 19798.639 | 9.8934 | 5.0055 |  |  | 495 | 19798.639 | 9.8934 | 5.0091 |  |  | 495 | 19798.639 | 9.8934 | 5.0104
496 | 19838.636 | 9.8954 | 5.0053 |  |  | 496 | 19838.636 | 9.8954 | 5.0089 |  |  | 496 | 19838.636 | 9.8954 | 5.0102
497 | 19878.634 | 9.8974 | 5.0052 |  |  | 497 | 19878.634 | 9.8974 | 5.0087 |  |  | 497 | 19878.634 | 9.8974 | 5.01
498 | 19918.631 | 9.8994 | 5.005 |  |  | 498 | 19918.631 | 9.8994 | 5.0085 |  |  | 498 | 19918.631 | 9.8994 | 5.0098
499 | 19958.628 | 9.9014 | 5.0048 |  |  | 499 | 19958.628 | 9.9014 | 5.0084 |  |  | 499 | 19958.628 | 9.9014 | 5.0096
500 | 19998.625 | 9.9034 | 5.0046 |  |  | 500 | 19998.625 | 9.9034 | 5.0082 |  |  | 500 | 19998.625 | 9.9034 | 5.0095
501 | 20038.623 | 9.9054 | 5.0044 |  |  | 501 | 20038.623 | 9.9054 | 5.008 |  |  | 501 | 20038.623 | 9.9054 | 5.0093
502 | 20078.62 | 9.9074 | 5.0042 |  |  | 502 | 20078.62 | 9.9074 | 5.0078 |  |  | 502 | 20078.62 | 9.9074 | 5.0091
503 | 20118.617 | 9.9094 | 5.004 |  |  | 503 | 20118.617 | 9.9094 | 5.0076 |  |  | 503 | 20118.617 | 9.9094 | 5.0089
504 | 20158.614 | 9.9114 | 5.0038 |  |  | 504 | 20158.614 | 9.9114 | 5.0074 |  |  | 504 | 20158.614 | 9.9114 | 5.0087
505 | 20198.612 | 9.9134 | 5.0037 |  |  | 505 | 20198.612 | 9.9134 | 5.0072 |  |  | 505 | 20198.612 | 9.9134 | 5.0085
506 | 20238.609 | 9.9153 | 5.0035 |  |  | 506 | 20238.609 | 9.9153 | 5.007 |  |  | 506 | 20238.609 | 9.9153 | 5.0083
507 | 20278.606 | 9.9173 | 5.0033 |  |  | 507 | 20278.606 | 9.9173 | 5.0069 |  |  | 507 | 20278.606 | 9.9173 | 5.0081
508 | 20318.603 | 9.9193 | 5.0031 |  |  | 508 | 20318.603 | 9.9193 | 5.0067 |  |  | 508 | 20318.603 | 9.9193 | 5.0079
509 | 20358.601 | 9.9213 | 5.0029 |  |  | 509 | 20358.601 | 9.9213 | 5.0065 |  |  | 509 | 20358.601 | 9.9213 | 5.0077
510 | 20398.598 | 9.9232 | 5.0028 |  |  | 510 | 20398.598 | 9.9232 | 5.0063 |  |  | 510 | 20398.598 | 9.9232 | 5.0075
511 | 20438.595 | 9.9252 | 5.0026 |  |  | 511 | 20438.595 | 9.9252 | 5.0061 |  |  | 511 | 20438.595 | 9.9252 | 5.0073
512 | 20478.592 | 9.9271 | 5.0024 |  |  | 512 | 20478.592 | 9.9271 | 5.0059 |  |  | 512 | 20478.592 | 9.9271 | 5.0072
513 | 20518.59 | 9.9291 | 5.0022 |  |  | 513 | 20518.59 | 9.9291 | 5.0058 |  |  | 513 | 20518.59 | 9.9291 | 5.007
514 | 20558.587 | 9.931 | 5.0021 |  |  | 514 | 20558.587 | 9.931 | 5.0056 |  |  | 514 | 20558.587 | 9.931 | 5.0068
515 | 20598.584 | 9.933 | 5.0019 |  |  | 515 | 20598.584 | 9.933 | 5.0054 |  |  | 515 | 20598.584 | 9.933 | 5.0066
516 | 20638.581 | 9.9349 | 5.0017 |  |  | 516 | 20638.581 | 9.9349 | 5.0052 |  |  | 516 | 20638.581 | 9.9349 | 5.0064
517 | 20678.579 | 9.9369 | 5.0015 |  |  | 517 | 20678.579 | 9.9369 | 5.0051 |  |  | 517 | 20678.579 | 9.9369 | 5.0062
518 | 20718.576 | 9.9388 | 5.0014 |  |  | 518 | 20718.576 | 9.9388 | 5.0049 |  |  | 518 | 20718.576 | 9.9388 | 5.006
519 | 20758.573 | 9.9407 | 5.0012 |  |  | 519 | 20758.573 | 9.9407 | 5.0047 |  |  | 519 | 20758.573 | 9.9407 | 5.0059
520 | 20798.57 | 9.9426 | 5.001 |  |  | 520 | 20798.57 | 9.9426 | 5.0045 |  |  | 520 | 20798.57 | 9.9426 | 5.0057
521 | 20838.568 | 9.9446 | 5.0009 |  |  | 521 | 20838.568 | 9.9446 | 5.0044 |  |  | 521 | 20838.568 | 9.9446 | 5.0055
522 | 20878.565 | 9.9465 | 5.0007 |  |  | 522 | 20878.565 | 9.9465 | 5.0042 |  |  | 522 | 20878.565 | 9.9465 | 5.0053
523 | 20918.562 | 9.9484 | 5.0005 |  |  | 523 | 20918.562 | 9.9484 | 5.004 |  |  | 523 | 20918.562 | 9.9484 | 5.0051
524 | 20958.559 | 9.9503 | 5.0003 |  |  | 524 | 20958.559 | 9.9503 | 5.0039 |  |  | 524 | 20958.559 | 9.9503 | 5.005
525 | 20998.557 | 9.9522 | 5.0002 |  |  | 525 | 20998.557 | 9.9522 | 5.0037 |  |  | 525 | 20998.557 | 9.9522 | 5.0048
526 | 21038.554 | 9.9541 | 5 |  |  | 526 | 21038.554 | 9.9541 | 5.0035 |  |  | 526 | 21038.554 | 9.9541 | 5.0046
527 | 21078.551 | 9.956 | 4.9999 |  |  | 527 | 21078.551 | 9.956 | 5.0034 |  |  | 527 | 21078.551 | 9.956 | 5.0044
528 | 21118.548 | 9.9579 | 4.9997 |  |  | 528 | 21118.548 | 9.9579 | 5.0032 |  |  | 528 | 21118.548 | 9.9579 | 5.0043
529 | 21158.546 | 9.9598 | 4.9995 |  |  | 529 | 21158.546 | 9.9598 | 5.003 |  |  | 529 | 21158.546 | 9.9598 | 5.0041
530 | 21198.543 | 9.9617 | 4.9994 |  |  | 530 | 21198.543 | 9.9617 | 5.0029 |  |  | 530 | 21198.543 | 9.9617 | 5.0039
531 | 21238.54 | 9.9636 | 4.9992 |  |  | 531 | 21238.54 | 9.9636 | 5.0027 |  |  | 531 | 21238.54 | 9.9636 | 5.0037
532 | 21278.537 | 9.9655 | 4.999 |  |  | 532 | 21278.537 | 9.9655 | 5.0025 |  |  | 532 | 21278.537 | 9.9655 | 5.0036
533 | 21318.535 | 9.9673 | 4.9989 |  |  | 533 | 21318.535 | 9.9673 | 5.0024 |  |  | 533 | 21318.535 | 9.9673 | 5.0034
534 | 21358.532 | 9.9692 | 4.9987 |  |  | 534 | 21358.532 | 9.9692 | 5.0022 |  |  | 534 | 21358.532 | 9.9692 | 5.0032
535 | 21398.529 | 9.9711 | 4.9986 |  |  | 535 | 21398.529 | 9.9711 | 5.002 |  |  | 535 | 21398.529 | 9.9711 | 5.0031
536 | 21438.526 | 9.9729 | 4.9984 |  |  | 536 | 21438.526 | 9.9729 | 5.0019 |  |  | 536 | 21438.526 | 9.9729 | 5.0029
537 | 21478.524 | 9.9748 | 4.9982 |  |  | 537 | 21478.524 | 9.9748 | 5.0017 |  |  | 537 | 21478.524 | 9.9748 | 5.0027
538 | 21518.521 | 9.9767 | 4.9981 |  |  | 538 | 21518.521 | 9.9767 | 5.0016 |  |  | 538 | 21518.521 | 9.9767 | 5.0026
539 | 21558.518 | 9.9785 | 4.9979 |  |  | 539 | 21558.518 | 9.9785 | 5.0014 |  |  | 539 | 21558.518 | 9.9785 | 5.0024
540 | 21598.515 | 9.9804 | 4.9978 |  |  | 540 | 21598.515 | 9.9804 | 5.0012 |  |  | 540 | 21598.515 | 9.9804 | 5.0022
541 | 21638.513 | 9.9822 | 4.9976 |  |  | 541 | 21638.513 | 9.9822 | 5.0011 |  |  | 541 | 21638.513 | 9.9822 | 5.0021
542 | 21678.51 | 9.9841 | 4.9975 |  |  | 542 | 21678.51 | 9.9841 | 5.0009 |  |  | 542 | 21678.51 | 9.9841 | 5.0019
543 | 21718.507 | 9.9859 | 4.9973 |  |  | 543 | 21718.507 | 9.9859 | 5.0008 |  |  | 543 | 21718.507 | 9.9859 | 5.0017
544 | 21758.504 | 9.9878 | 4.9972 |  |  | 544 | 21758.504 | 9.9878 | 5.0006 |  |  | 544 | 21758.504 | 9.9878 | 5.0016
545 | 21798.502 | 9.9896 | 4.997 |  |  | 545 | 21798.502 | 9.9896 | 5.0005 |  |  | 545 | 21798.502 | 9.9896 | 5.0014
546 | 21838.499 | 9.9914 | 4.9968 |  |  | 546 | 21838.499 | 9.9914 | 5.0003 |  |  | 546 | 21838.499 | 9.9914 | 5.0012
547 | 21878.496 | 9.9933 | 4.9967 |  |  | 547 | 21878.496 | 9.9933 | 5.0001 |  |  | 547 | 21878.496 | 9.9933 | 5.0011
548 | 21918.493 | 9.9951 | 4.9965 |  |  | 548 | 21918.493 | 9.9951 | 5 |  |  | 548 | 21918.493 | 9.9951 | 5.0009
549 | 21958.491 | 9.9969 | 4.9964 |  |  | 549 | 21958.491 | 9.9969 | 4.9998 |  |  | 549 | 21958.491 | 9.9969 | 5.0007
550 | 21998.488 | 9.9987 | 4.9962 |  |  | 550 | 21998.488 | 9.9987 | 4.9997 |  |  | 550 | 21998.488 | 9.9987 | 5.0006
551 | 22038.485 | 10.0005 | 4.9961 |  |  | 551 | 22038.485 | 10.0005 | 4.9995 |  |  | 551 | 22038.485 | 10.0005 | 5.0004
552 | 22078.482 | 10.0024 | 4.9959 |  |  | 552 | 22078.482 | 10.0024 | 4.9994 |  |  | 552 | 22078.482 | 10.0024 | 5.0003
553 | 22118.48 | 10.0042 | 4.9958 |  |  | 553 | 22118.48 | 10.0042 | 4.9992 |  |  | 553 | 22118.48 | 10.0042 | 5.0001
554 | 22158.477 | 10.006 | 4.9956 |  |  | 554 | 22158.477 | 10.006 | 4.9991 |  |  | 554 | 22158.477 | 10.006 | 5
555 | 22198.474 | 10.0078 | 4.9955 |  |  | 555 | 22198.474 | 10.0078 | 4.9989 |  |  | 555 | 22198.474 | 10.0078 | 4.9998
556 | 22238.471 | 10.0096 | 4.9954 |  |  | 556 | 22238.471 | 10.0096 | 4.9988 |  |  | 556 | 22238.471 | 10.0096 | 4.9996
557 | 22278.469 | 10.0114 | 4.9952 |  |  | 557 | 22278.469 | 10.0114 | 4.9986 |  |  | 557 | 22278.469 | 10.0114 | 4.9995
558 | 22318.466 | 10.0132 | 4.9951 |  |  | 558 | 22318.466 | 10.0132 | 4.9985 |  |  | 558 | 22318.466 | 10.0132 | 4.9993
559 | 22358.463 | 10.015 | 4.9949 |  |  | 559 | 22358.463 | 10.015 | 4.9983 |  |  | 559 | 22358.463 | 10.015 | 4.9992
560 | 22398.46 | 10.0167 | 4.9948 |  |  | 560 | 22398.46 | 10.0167 | 4.9982 |  |  | 560 | 22398.46 | 10.0167 | 4.999
561 | 22438.458 | 10.0185 | 4.9946 |  |  | 561 | 22438.458 | 10.0185 | 4.9981 |  |  | 561 | 22438.458 | 10.0185 | 4.9989
562 | 22478.455 | 10.0203 | 4.9945 |  |  | 562 | 22478.455 | 10.0203 | 4.9979 |  |  | 562 | 22478.455 | 10.0203 | 4.9987
563 | 22518.452 | 10.0221 | 4.9943 |  |  | 563 | 22518.452 | 10.0221 | 4.9978 |  |  | 563 | 22518.452 | 10.0221 | 4.9986
564 | 22558.449 | 10.0239 | 4.9942 |  |  | 564 | 22558.449 | 10.0239 | 4.9976 |  |  | 564 | 22558.449 | 10.0239 | 4.9984
565 | 22598.447 | 10.0256 | 4.9941 |  |  | 565 | 22598.447 | 10.0256 | 4.9975 |  |  | 565 | 22598.447 | 10.0256 | 4.9982
566 | 22638.444 | 10.0274 | 4.9939 |  |  | 566 | 22638.444 | 10.0274 | 4.9973 |  |  | 566 | 22638.444 | 10.0274 | 4.9981
567 | 22678.441 | 10.0292 | 4.9938 |  |  | 567 | 22678.441 | 10.0292 | 4.9972 |  |  | 567 | 22678.441 | 10.0292 | 4.9979
568 | 22718.438 | 10.0309 | 4.9936 |  |  | 568 | 22718.438 | 10.0309 | 4.9971 |  |  | 568 | 22718.438 | 10.0309 | 4.9978
569 | 22758.436 | 10.0327 | 4.9935 |  |  | 569 | 22758.436 | 10.0327 | 4.9969 |  |  | 569 | 22758.436 | 10.0327 | 4.9976
570 | 22798.433 | 10.0344 | 4.9934 |  |  | 570 | 22798.433 | 10.0344 | 4.9968 |  |  | 570 | 22798.433 | 10.0344 | 4.9975
571 | 22838.43 | 10.0362 | 4.9932 |  |  | 571 | 22838.43 | 10.0362 | 4.9966 |  |  | 571 | 22838.43 | 10.0362 | 4.9973
572 | 22878.427 | 10.0379 | 4.9931 |  |  | 572 | 22878.427 | 10.0379 | 4.9965 |  |  | 572 | 22878.427 | 10.0379 | 4.9972
573 | 22918.425 | 10.0397 | 4.9929 |  |  | 573 | 22918.425 | 10.0397 | 4.9964 |  |  | 573 | 22918.425 | 10.0397 | 4.997
574 | 22958.422 | 10.0414 | 4.9928 |  |  | 574 | 22958.422 | 10.0414 | 4.9962 |  |  | 574 | 22958.422 | 10.0414 | 4.9969
575 | 22998.419 | 10.0432 | 4.9927 |  |  | 575 | 22998.419 | 10.0432 | 4.9961 |  |  | 575 | 22998.419 | 10.0432 | 4.9967
576 | 23038.416 | 10.0449 | 4.9925 |  |  | 576 | 23038.416 | 10.0449 | 4.9959 |  |  | 576 | 23038.416 | 10.0449 | 4.9966
577 | 23078.414 | 10.0467 | 4.9924 |  |  | 577 | 23078.414 | 10.0467 | 4.9958 |  |  | 577 | 23078.414 | 10.0467 | 4.9965
578 | 23118.411 | 10.0484 | 4.9923 |  |  | 578 | 23118.411 | 10.0484 | 4.9957 |  |  | 578 | 23118.411 | 10.0484 | 4.9963
579 | 23158.408 | 10.0501 | 4.9921 |  |  | 579 | 23158.408 | 10.0501 | 4.9955 |  |  | 579 | 23158.408 | 10.0501 | 4.9962
580 | 23198.405 | 10.0518 | 4.992 |  |  | 580 | 23198.405 | 10.0518 | 4.9954 |  |  | 580 | 23198.405 | 10.0518 | 4.996
581 | 23238.403 | 10.0536 | 4.9919 |  |  | 581 | 23238.403 | 10.0536 | 4.9953 |  |  | 581 | 23238.403 | 10.0536 | 4.9959
582 | 23278.4 | 10.0553 | 4.9917 |  |  | 582 | 23278.4 | 10.0553 | 4.9951 |  |  | 582 | 23278.4 | 10.0553 | 4.9957
583 | 23318.397 | 10.057 | 4.9916 |  |  | 583 | 23318.397 | 10.057 | 4.995 |  |  | 583 | 23318.397 | 10.057 | 4.9956
584 | 23358.394 | 10.0587 | 4.9915 |  |  | 584 | 23358.394 | 10.0587 | 4.9949 |  |  | 584 | 23358.394 | 10.0587 | 4.9954
585 | 23398.392 | 10.0604 | 4.9913 |  |  | 585 | 23398.392 | 10.0604 | 4.9947 |  |  | 585 | 23398.392 | 10.0604 | 4.9953
586 | 23438.389 | 10.0621 | 4.9912 |  |  | 586 | 23438.389 | 10.0621 | 4.9946 |  |  | 586 | 23438.389 | 10.0621 | 4.9951
587 | 23478.386 | 10.0638 | 4.9911 |  |  | 587 | 23478.386 | 10.0638 | 4.9945 |  |  | 587 | 23478.386 | 10.0638 | 4.995
588 | 23518.383 | 10.0655 | 4.9909 |  |  | 588 | 23518.383 | 10.0655 | 4.9943 |  |  | 588 | 23518.383 | 10.0655 | 4.9949
589 | 23558.381 | 10.0672 | 4.9908 |  |  | 589 | 23558.381 | 10.0672 | 4.9942 |  |  | 589 | 23558.381 | 10.0672 | 4.9947
590 | 23598.378 | 10.0689 | 4.9907 |  |  | 590 | 23598.378 | 10.0689 | 4.9941 |  |  | 590 | 23598.378 | 10.0689 | 4.9946
591 | 23638.375 | 10.0706 | 4.9906 |  |  | 591 | 23638.375 | 10.0706 | 4.9939 |  |  | 591 | 23638.375 | 10.0706 | 4.9944
592 | 23678.372 | 10.0723 | 4.9904 |  |  | 592 | 23678.372 | 10.0723 | 4.9938 |  |  | 592 | 23678.372 | 10.0723 | 4.9943
593 | 23718.37 | 10.074 | 4.9903 |  |  | 593 | 23718.37 | 10.074 | 4.9937 |  |  | 593 | 23718.37 | 10.074 | 4.9942
594 | 23758.367 | 10.0757 | 4.9902 |  |  | 594 | 23758.367 | 10.0757 | 4.9935 |  |  | 594 | 23758.367 | 10.0757 | 4.994
595 | 23798.364 | 10.0774 | 4.99 |  |  | 595 | 23798.364 | 10.0774 | 4.9934 |  |  | 595 | 23798.364 | 10.0774 | 4.9939
596 | 23838.361 | 10.0791 | 4.9899 |  |  | 596 | 23838.361 | 10.0791 | 4.9933 |  |  | 596 | 23838.361 | 10.0791 | 4.9937
597 | 23878.359 | 10.0807 | 4.9898 |  |  | 597 | 23878.359 | 10.0807 | 4.9932 |  |  | 597 | 23878.359 | 10.0807 | 4.9936
598 | 23918.356 | 10.0824 | 4.9897 |  |  | 598 | 23918.356 | 10.0824 | 4.993 |  |  | 598 | 23918.356 | 10.0824 | 4.9935
599 | 23958.353 | 10.0841 | 4.9895 |  |  | 599 | 23958.353 | 10.0841 | 4.9929 |  |  | 599 | 23958.353 | 10.0841 | 4.9933
600 | 23998.35 | 10.0857 | 4.9894 |  |  | 600 | 23998.35 | 10.0857 | 4.9928 |  |  | 600 | 23998.35 | 10.0857 | 4.9932
601 | 24038.348 | 10.0874 | 4.9893 |  |  | 601 | 24038.348 | 10.0874 | 4.9926 |  |  | 601 | 24038.348 | 10.0874 | 4.9931
602 | 24078.345 | 10.0891 | 4.9892 |  |  | 602 | 24078.345 | 10.0891 | 4.9925 |  |  | 602 | 24078.345 | 10.0891 | 4.9929
603 | 24118.342 | 10.0907 | 4.989 |  |  | 603 | 24118.342 | 10.0907 | 4.9924 |  |  | 603 | 24118.342 | 10.0907 | 4.9928
604 | 24158.339 | 10.0924 | 4.9889 |  |  | 604 | 24158.339 | 10.0924 | 4.9923 |  |  | 604 | 24158.339 | 10.0924 | 4.9926
605 | 24198.337 | 10.094 | 4.9888 |  |  | 605 | 24198.337 | 10.094 | 4.9921 |  |  | 605 | 24198.337 | 10.094 | 4.9925
606 | 24238.334 | 10.0957 | 4.9887 |  |  | 606 | 24238.334 | 10.0957 | 4.992 |  |  | 606 | 24238.334 | 10.0957 | 4.9924
607 | 24278.331 | 10.0973 | 4.9885 |  |  | 607 | 24278.331 | 10.0973 | 4.9919 |  |  | 607 | 24278.331 | 10.0973 | 4.9922
608 | 24318.328 | 10.099 | 4.9884 |  |  | 608 | 24318.328 | 10.099 | 4.9918 |  |  | 608 | 24318.328 | 10.099 | 4.9921
609 | 24358.326 | 10.1006 | 4.9883 |  |  | 609 | 24358.326 | 10.1006 | 4.9916 |  |  | 609 | 24358.326 | 10.1006 | 4.992
610 | 24398.323 | 10.1023 | 4.9882 |  |  | 610 | 24398.323 | 10.1023 | 4.9915 |  |  | 610 | 24398.323 | 10.1023 | 4.9918
611 | 24438.32 | 10.1039 | 4.988 |  |  | 611 | 24438.32 | 10.1039 | 4.9914 |  |  | 611 | 24438.32 | 10.1039 | 4.9917
612 | 24478.317 | 10.1055 | 4.9879 |  |  | 612 | 24478.317 | 10.1055 | 4.9913 |  |  | 612 | 24478.317 | 10.1055 | 4.9916
613 | 24518.315 | 10.1072 | 4.9878 |  |  | 613 | 24518.315 | 10.1072 | 4.9912 |  |  | 613 | 24518.315 | 10.1072 | 4.9914
614 | 24558.312 | 10.1088 | 4.9877 |  |  | 614 | 24558.312 | 10.1088 | 4.991 |  |  | 614 | 24558.312 | 10.1088 | 4.9913
615 | 24598.309 | 10.1104 | 4.9876 |  |  | 615 | 24598.309 | 10.1104 | 4.9909 |  |  | 615 | 24598.309 | 10.1104 | 4.9912
616 | 24638.306 | 10.1121 | 4.9874 |  |  | 616 | 24638.306 | 10.1121 | 4.9908 |  |  | 616 | 24638.306 | 10.1121 | 4.991
617 | 24678.304 | 10.1137 | 4.9873 |  |  | 617 | 24678.304 | 10.1137 | 4.9907 |  |  | 617 | 24678.304 | 10.1137 | 4.9909
618 | 24718.301 | 10.1153 | 4.9872 |  |  | 618 | 24718.301 | 10.1153 | 4.9905 |  |  | 618 | 24718.301 | 10.1153 | 4.9908
619 | 24758.298 | 10.1169 | 4.9871 |  |  | 619 | 24758.298 | 10.1169 | 4.9904 |  |  | 619 | 24758.298 | 10.1169 | 4.9906
620 | 24798.295 | 10.1185 | 4.987 |  |  | 620 | 24798.295 | 10.1185 | 4.9903 |  |  | 620 | 24798.295 | 10.1185 | 4.9905
621 | 24838.293 | 10.1201 | 4.9868 |  |  | 621 | 24838.293 | 10.1201 | 4.9902 |  |  | 621 | 24838.293 | 10.1201 | 4.9904
622 | 24878.29 | 10.1218 | 4.9867 |  |  | 622 | 24878.29 | 10.1218 | 4.9901 |  |  | 622 | 24878.29 | 10.1218 | 4.9902
623 | 24918.287 | 10.1234 | 4.9866 |  |  | 623 | 24918.287 | 10.1234 | 4.99 |  |  | 623 | 24918.287 | 10.1234 | 4.9901
624 | 24958.284 | 10.125 | 4.9865 |  |  | 624 | 24958.284 | 10.125 | 4.9898 |  |  | 624 | 24958.284 | 10.125 | 4.99
625 | 24998.282 | 10.1266 | 4.9864 |  |  | 625 | 24998.282 | 10.1266 | 4.9897 |  |  | 625 | 24998.282 | 10.1266 | 4.9899
626 | 25038.279 | 10.1282 | 4.9863 |  |  | 626 | 25038.279 | 10.1282 | 4.9896 |  |  | 626 | 25038.279 | 10.1282 | 4.9897
627 | 25078.276 | 10.1298 | 4.9861 |  |  | 627 | 25078.276 | 10.1298 | 4.9895 |  |  | 627 | 25078.276 | 10.1298 | 4.9896
628 | 25118.273 | 10.1314 | 4.986 |  |  | 628 | 25118.273 | 10.1314 | 4.9894 |  |  | 628 | 25118.273 | 10.1314 | 4.9895
629 | 25158.271 | 10.1329 | 4.9859 |  |  | 629 | 25158.271 | 10.1329 | 4.9892 |  |  | 629 | 25158.271 | 10.1329 | 4.9893
630 | 25198.268 | 10.1345 | 4.9858 |  |  | 630 | 25198.268 | 10.1345 | 4.9891 |  |  | 630 | 25198.268 | 10.1345 | 4.9892
631 | 25238.265 | 10.1361 | 4.9857 |  |  | 631 | 25238.265 | 10.1361 | 4.989 |  |  | 631 | 25238.265 | 10.1361 | 4.9891
632 | 25278.262 | 10.1377 | 4.9856 |  |  | 632 | 25278.262 | 10.1377 | 4.9889 |  |  | 632 | 25278.262 | 10.1377 | 4.989
633 | 25318.26 | 10.1393 | 4.9854 |  |  | 633 | 25318.26 | 10.1393 | 4.9888 |  |  | 633 | 25318.26 | 10.1393 | 4.9888
634 | 25358.257 | 10.1409 | 4.9853 |  |  | 634 | 25358.257 | 10.1409 | 4.9887 |  |  | 634 | 25358.257 | 10.1409 | 4.9887
635 | 25398.254 | 10.1424 | 4.9852 |  |  | 635 | 25398.254 | 10.1424 | 4.9885 |  |  | 635 | 25398.254 | 10.1424 | 4.9886
636 | 25438.251 | 10.144 | 4.9851 |  |  | 636 | 25438.251 | 10.144 | 4.9884 |  |  | 636 | 25438.251 | 10.144 | 4.9884
637 | 25478.249 | 10.1456 | 4.985 |  |  | 637 | 25478.249 | 10.1456 | 4.9883 |  |  | 637 | 25478.249 | 10.1456 | 4.9883
638 | 25518.246 | 10.1471 | 4.9849 |  |  | 638 | 25518.246 | 10.1471 | 4.9882 |  |  | 638 | 25518.246 | 10.1471 | 4.9882
639 | 25558.243 | 10.1487 | 4.9848 |  |  | 639 | 25558.243 | 10.1487 | 4.9881 |  |  | 639 | 25558.243 | 10.1487 | 4.9881
640 | 25598.24 | 10.1503 | 4.9846 |  |  | 640 | 25598.24 | 10.1503 | 4.988 |  |  | 640 | 25598.24 | 10.1503 | 4.9879
641 | 25638.238 | 10.1518 | 4.9845 |  |  | 641 | 25638.238 | 10.1518 | 4.9879 |  |  | 641 | 25638.238 | 10.1518 | 4.9878
642 | 25678.235 | 10.1534 | 4.9844 |  |  | 642 | 25678.235 | 10.1534 | 4.9877 |  |  | 642 | 25678.235 | 10.1534 | 4.9877
643 | 25718.232 | 10.155 | 4.9843 |  |  | 643 | 25718.232 | 10.155 | 4.9876 |  |  | 643 | 25718.232 | 10.155 | 4.9876
644 | 25758.229 | 10.1565 | 4.9842 |  |  | 644 | 25758.229 | 10.1565 | 4.9875 |  |  | 644 | 25758.229 | 10.1565 | 4.9874
645 | 25798.227 | 10.1581 | 4.9841 |  |  | 645 | 25798.227 | 10.1581 | 4.9874 |  |  | 645 | 25798.227 | 10.1581 | 4.9873
646 | 25838.224 | 10.1596 | 4.984 |  |  | 646 | 25838.224 | 10.1596 | 4.9873 |  |  | 646 | 25838.224 | 10.1596 | 4.9872
647 | 25878.221 | 10.1612 | 4.9839 |  |  | 647 | 25878.221 | 10.1612 | 4.9872 |  |  | 647 | 25878.221 | 10.1612 | 4.9871
648 | 25918.218 | 10.1627 | 4.9838 |  |  | 648 | 25918.218 | 10.1627 | 4.9871 |  |  | 648 | 25918.218 | 10.1627 | 4.9869
649 | 25958.216 | 10.1642 | 4.9836 |  |  | 649 | 25958.216 | 10.1642 | 4.987 |  |  | 649 | 25958.216 | 10.1642 | 4.9868
650 | 25998.213 | 10.1658 | 4.9835 |  |  | 650 | 25998.213 | 10.1658 | 4.9868 |  |  | 650 | 25998.213 | 10.1658 | 4.9867
651 | 26038.21 | 10.1673 | 4.9834 |  |  | 651 | 26038.21 | 10.1673 | 4.9867 |  |  | 651 | 26038.21 | 10.1673 | 4.9866
652 | 26078.207 | 10.1689 | 4.9833 |  |  | 652 | 26078.207 | 10.1689 | 4.9866 |  |  | 652 | 26078.207 | 10.1689 | 4.9864
653 | 26118.205 | 10.1704 | 4.9832 |  |  | 653 | 26118.205 | 10.1704 | 4.9865 |  |  | 653 | 26118.205 | 10.1704 | 4.9863
654 | 26158.202 | 10.1719 | 4.9831 |  |  | 654 | 26158.202 | 10.1719 | 4.9864 |  |  | 654 | 26158.202 | 10.1719 | 4.9862
655 | 26198.199 | 10.1734 | 4.983 |  |  | 655 | 26198.199 | 10.1734 | 4.9863 |  |  | 655 | 26198.199 | 10.1734 | 4.9861
656 | 26238.196 | 10.175 | 4.9829 |  |  | 656 | 26238.196 | 10.175 | 4.9862 |  |  | 656 | 26238.196 | 10.175 | 4.9859
657 | 26278.194 | 10.1765 | 4.9828 |  |  | 657 | 26278.194 | 10.1765 | 4.9861 |  |  | 657 | 26278.194 | 10.1765 | 4.9858
658 | 26318.191 | 10.178 | 4.9826 |  |  | 658 | 26318.191 | 10.178 | 4.986 |  |  | 658 | 26318.191 | 10.178 | 4.9857
659 | 26358.188 | 10.1795 | 4.9825 |  |  | 659 | 26358.188 | 10.1795 | 4.9859 |  |  | 659 | 26358.188 | 10.1795 | 4.9856
660 | 26398.185 | 10.1811 | 4.9824 |  |  | 660 | 26398.185 | 10.1811 | 4.9857 |  |  | 660 | 26398.185 | 10.1811 | 4.9855
661 | 26438.183 | 10.1826 | 4.9823 |  |  | 661 | 26438.183 | 10.1826 | 4.9856 |  |  | 661 | 26438.183 | 10.1826 | 4.9853
662 | 26478.18 | 10.1841 | 4.9822 |  |  | 662 | 26478.18 | 10.1841 | 4.9855 |  |  | 662 | 26478.18 | 10.1841 | 4.9852
663 | 26518.177 | 10.1856 | 4.9821 |  |  | 663 | 26518.177 | 10.1856 | 4.9854 |  |  | 663 | 26518.177 | 10.1856 | 4.9851
664 | 26558.174 | 10.1871 | 4.982 |  |  | 664 | 26558.174 | 10.1871 | 4.9853 |  |  | 664 | 26558.174 | 10.1871 | 4.985
665 | 26598.172 | 10.1886 | 4.9819 |  |  | 665 | 26598.172 | 10.1886 | 4.9852 |  |  | 665 | 26598.172 | 10.1886 | 4.9849
666 | 26638.169 | 10.1901 | 4.9818 |  |  | 666 | 26638.169 | 10.1901 | 4.9851 |  |  | 666 | 26638.169 | 10.1901 | 4.9847
667 | 26678.166 | 10.1916 | 4.9817 |  |  | 667 | 26678.166 | 10.1916 | 4.985 |  |  | 667 | 26678.166 | 10.1916 | 4.9846
668 | 26718.163 | 10.1931 | 4.9816 |  |  | 668 | 26718.163 | 10.1931 | 4.9849 |  |  | 668 | 26718.163 | 10.1931 | 4.9845
669 | 26758.161 | 10.1946 | 4.9815 |  |  | 669 | 26758.161 | 10.1946 | 4.9848 |  |  | 669 | 26758.161 | 10.1946 | 4.9844
670 | 26798.158 | 10.1961 | 4.9814 |  |  | 670 | 26798.158 | 10.1961 | 4.9847 |  |  | 670 | 26798.158 | 10.1961 | 4.9843
671 | 26838.155 | 10.1976 | 4.9812 |  |  | 671 | 26838.155 | 10.1976 | 4.9846 |  |  | 671 | 26838.155 | 10.1976 | 4.9841
672 | 26878.152 | 10.1991 | 4.9811 |  |  | 672 | 26878.152 | 10.1991 | 4.9845 |  |  | 672 | 26878.152 | 10.1991 | 4.984
673 | 26918.15 | 10.2006 | 4.981 |  |  | 673 | 26918.15 | 10.2006 | 4.9843 |  |  | 673 | 26918.15 | 10.2006 | 4.9839
674 | 26958.147 | 10.202 | 4.9809 |  |  | 674 | 26958.147 | 10.202 | 4.9842 |  |  | 674 | 26958.147 | 10.202 | 4.9838
675 | 26998.144 | 10.2035 | 4.9808 |  |  | 675 | 26998.144 | 10.2035 | 4.9841 |  |  | 675 | 26998.144 | 10.2035 | 4.9837
676 | 27038.141 | 10.205 | 4.9807 |  |  | 676 | 27038.141 | 10.205 | 4.984 |  |  | 676 | 27038.141 | 10.205 | 4.9835
677 | 27078.139 | 10.2065 | 4.9806 |  |  | 677 | 27078.139 | 10.2065 | 4.9839 |  |  | 677 | 27078.139 | 10.2065 | 4.9834
678 | 27118.136 | 10.208 | 4.9805 |  |  | 678 | 27118.136 | 10.208 | 4.9838 |  |  | 678 | 27118.136 | 10.208 | 4.9833
679 | 27158.133 | 10.2094 | 4.9804 |  |  | 679 | 27158.133 | 10.2094 | 4.9837 |  |  | 679 | 27158.133 | 10.2094 | 4.9832
680 | 27198.13 | 10.2109 | 4.9803 |  |  | 680 | 27198.13 | 10.2109 | 4.9836 |  |  | 680 | 27198.13 | 10.2109 | 4.9831
681 | 27238.128 | 10.2124 | 4.9802 |  |  | 681 | 27238.128 | 10.2124 | 4.9835 |  |  | 681 | 27238.128 | 10.2124 | 4.9829
682 | 27278.125 | 10.2138 | 4.9801 |  |  | 682 | 27278.125 | 10.2138 | 4.9834 |  |  | 682 | 27278.125 | 10.2138 | 4.9828
683 | 27318.122 | 10.2153 | 4.98 |  |  | 683 | 27318.122 | 10.2153 | 4.9833 |  |  | 683 | 27318.122 | 10.2153 | 4.9827
684 | 27358.119 | 10.2168 | 4.9799 |  |  | 684 | 27358.119 | 10.2168 | 4.9832 |  |  | 684 | 27358.119 | 10.2168 | 4.9826
685 | 27398.117 | 10.2182 | 4.9798 |  |  | 685 | 27398.117 | 10.2182 | 4.9831 |  |  | 685 | 27398.117 | 10.2182 | 4.9825
686 | 27438.114 | 10.2197 | 4.9797 |  |  | 686 | 27438.114 | 10.2197 | 4.983 |  |  | 686 | 27438.114 | 10.2197 | 4.9823
687 | 27478.111 | 10.2211 | 4.9796 |  |  | 687 | 27478.111 | 10.2211 | 4.9829 |  |  | 687 | 27478.111 | 10.2211 | 4.9822
688 | 27518.108 | 10.2226 | 4.9795 |  |  | 688 | 27518.108 | 10.2226 | 4.9828 |  |  | 688 | 27518.108 | 10.2226 | 4.9821
689 | 27558.106 | 10.2241 | 4.9794 |  |  | 689 | 27558.106 | 10.2241 | 4.9827 |  |  | 689 | 27558.106 | 10.2241 | 4.982
690 | 27598.103 | 10.2255 | 4.9793 |  |  | 690 | 27598.103 | 10.2255 | 4.9826 |  |  | 690 | 27598.103 | 10.2255 | 4.9819
691 | 27638.1 | 10.227 | 4.9792 |  |  | 691 | 27638.1 | 10.227 | 4.9825 |  |  | 691 | 27638.1 | 10.227 | 4.9818
692 | 27678.097 | 10.2284 | 4.979 |  |  | 692 | 27678.097 | 10.2284 | 4.9824 |  |  | 692 | 27678.097 | 10.2284 | 4.9816
693 | 27718.095 | 10.2298 | 4.9789 |  |  | 693 | 27718.095 | 10.2298 | 4.9822 |  |  | 693 | 27718.095 | 10.2298 | 4.9815
694 | 27758.092 | 10.2313 | 4.9788 |  |  | 694 | 27758.092 | 10.2313 | 4.9821 |  |  | 694 | 27758.092 | 10.2313 | 4.9814
695 | 27798.089 | 10.2327 | 4.9787 |  |  | 695 | 27798.089 | 10.2327 | 4.982 |  |  | 695 | 27798.089 | 10.2327 | 4.9813
696 | 27838.086 | 10.2342 | 4.9786 |  |  | 696 | 27838.086 | 10.2342 | 4.9819 |  |  | 696 | 27838.086 | 10.2342 | 4.9812
697 | 27878.084 | 10.2356 | 4.9785 |  |  | 697 | 27878.084 | 10.2356 | 4.9818 |  |  | 697 | 27878.084 | 10.2356 | 4.9811
698 | 27918.081 | 10.237 | 4.9784 |  |  | 698 | 27918.081 | 10.237 | 4.9817 |  |  | 698 | 27918.081 | 10.237 | 4.9809
699 | 27958.078 | 10.2385 | 4.9783 |  |  | 699 | 27958.078 | 10.2385 | 4.9816 |  |  | 699 | 27958.078 | 10.2385 | 4.9808
700 | 27998.075 | 10.2399 | 4.9782 |  |  | 700 | 27998.075 | 10.2399 | 4.9815 |  |  | 700 | 27998.075 | 10.2399 | 4.9807
701 | 28038.073 | 10.2413 | 4.9781 |  |  | 701 | 28038.073 | 10.2413 | 4.9814 |  |  | 701 | 28038.073 | 10.2413 | 4.9806
702 | 28078.07 | 10.2427 | 4.978 |  |  | 702 | 28078.07 | 10.2427 | 4.9813 |  |  | 702 | 28078.07 | 10.2427 | 4.9805
703 | 28118.067 | 10.2442 | 4.9779 |  |  | 703 | 28118.067 | 10.2442 | 4.9812 |  |  | 703 | 28118.067 | 10.2442 | 4.9804
704 | 28158.064 | 10.2456 | 4.9778 |  |  | 704 | 28158.064 | 10.2456 | 4.9811 |  |  | 704 | 28158.064 | 10.2456 | 4.9802
705 | 28198.062 | 10.247 | 4.9777 |  |  | 705 | 28198.062 | 10.247 | 4.981 |  |  | 705 | 28198.062 | 10.247 | 4.9801
706 | 28238.059 | 10.2484 | 4.9776 |  |  | 706 | 28238.059 | 10.2484 | 4.9809 |  |  | 706 | 28238.059 | 10.2484 | 4.98
707 | 28278.056 | 10.2498 | 4.9775 |  |  | 707 | 28278.056 | 10.2498 | 4.9808 |  |  | 707 | 28278.056 | 10.2498 | 4.9799
708 | 28318.053 | 10.2513 | 4.9774 |  |  | 708 | 28318.053 | 10.2513 | 4.9807 |  |  | 708 | 28318.053 | 10.2513 | 4.9798
709 | 28358.051 | 10.2527 | 4.9773 |  |  | 709 | 28358.051 | 10.2527 | 4.9806 |  |  | 709 | 28358.051 | 10.2527 | 4.9797
710 | 28398.048 | 10.2541 | 4.9772 |  |  | 710 | 28398.048 | 10.2541 | 4.9805 |  |  | 710 | 28398.048 | 10.2541 | 4.9796
711 | 28438.045 | 10.2555 | 4.9771 |  |  | 711 | 28438.045 | 10.2555 | 4.9804 |  |  | 711 | 28438.045 | 10.2555 | 4.9794
712 | 28478.042 | 10.2569 | 4.977 |  |  | 712 | 28478.042 | 10.2569 | 4.9803 |  |  | 712 | 28478.042 | 10.2569 | 4.9793
713 | 28518.04 | 10.2583 | 4.9769 |  |  | 713 | 28518.04 | 10.2583 | 4.9802 |  |  | 713 | 28518.04 | 10.2583 | 4.9792
714 | 28558.037 | 10.2597 | 4.9768 |  |  | 714 | 28558.037 | 10.2597 | 4.9801 |  |  | 714 | 28558.037 | 10.2597 | 4.9791
715 | 28598.034 | 10.2611 | 4.9767 |  |  | 715 | 28598.034 | 10.2611 | 4.98 |  |  | 715 | 28598.034 | 10.2611 | 4.979
716 | 28638.031 | 10.2625 | 4.9766 |  |  | 716 | 28638.031 | 10.2625 | 4.9799 |  |  | 716 | 28638.031 | 10.2625 | 4.9789
717 | 28678.029 | 10.2639 | 4.9765 |  |  | 717 | 28678.029 | 10.2639 | 4.9798 |  |  | 717 | 28678.029 | 10.2639 | 4.9788
718 | 28718.026 | 10.2653 | 4.9764 |  |  | 718 | 28718.026 | 10.2653 | 4.9797 |  |  | 718 | 28718.026 | 10.2653 | 4.9786
719 | 28758.023 | 10.2667 | 4.9763 |  |  | 719 | 28758.023 | 10.2667 | 4.9796 |  |  | 719 | 28758.023 | 10.2667 | 4.9785
720 | 28798.02 | 10.2681 | 4.9762 |  |  | 720 | 28798.02 | 10.2681 | 4.9795 |  |  | 720 | 28798.02 | 10.2681 | 4.9784
721 | 28838.018 | 10.2694 | 4.9761 |  |  | 721 | 28838.018 | 10.2694 | 4.9794 |  |  | 721 | 28838.018 | 10.2694 | 4.9783
722 | 28878.015 | 10.2708 | 4.976 |  |  | 722 | 28878.015 | 10.2708 | 4.9793 |  |  | 722 | 28878.015 | 10.2708 | 4.9782
723 | 28918.012 | 10.2722 | 4.9759 |  |  | 723 | 28918.012 | 10.2722 | 4.9792 |  |  | 723 | 28918.012 | 10.2722 | 4.9781
724 | 28958.009 | 10.2736 | 4.9758 |  |  | 724 | 28958.009 | 10.2736 | 4.9791 |  |  | 724 | 28958.009 | 10.2736 | 4.978
725 | 28998.007 | 10.275 | 4.9757 |  |  | 725 | 28998.007 | 10.275 | 4.979 |  |  | 725 | 28998.007 | 10.275 | 4.9778
726 | 29038.004 | 10.2764 | 4.9756 |  |  | 726 | 29038.004 | 10.2764 | 4.9789 |  |  | 726 | 29038.004 | 10.2764 | 4.9777
727 | 29078.001 | 10.2777 | 4.9755 |  |  | 727 | 29078.001 | 10.2777 | 4.9788 |  |  | 727 | 29078.001 | 10.2777 | 4.9776
728 | 29117.998 | 10.2791 | 4.9754 |  |  | 728 | 29117.998 | 10.2791 | 4.9787 |  |  | 728 | 29117.998 | 10.2791 | 4.9775
729 | 29157.996 | 10.2805 | 4.9753 |  |  | 729 | 29157.996 | 10.2805 | 4.9786 |  |  | 729 | 29157.996 | 10.2805 | 4.9774
730 | 29197.993 | 10.2819 | 4.9752 |  |  | 730 | 29197.993 | 10.2819 | 4.9785 |  |  | 730 | 29197.993 | 10.2819 | 4.9773
731 | 29237.99 | 10.2832 | 4.9751 |  |  | 731 | 29237.99 | 10.2832 | 4.9784 |  |  | 731 | 29237.99 | 10.2832 | 4.9772
732 | 29277.987 | 10.2846 | 4.975 |  |  | 732 | 29277.987 | 10.2846 | 4.9783 |  |  | 732 | 29277.987 | 10.2846 | 4.977
733 | 29317.985 | 10.286 | 4.9749 |  |  | 733 | 29317.985 | 10.286 | 4.9782 |  |  | 733 | 29317.985 | 10.286 | 4.9769
734 | 29357.982 | 10.2873 | 4.9748 |  |  | 734 | 29357.982 | 10.2873 | 4.9781 |  |  | 734 | 29357.982 | 10.2873 | 4.9768
735 | 29397.979 | 10.2887 | 4.9747 |  |  | 735 | 29397.979 | 10.2887 | 4.978 |  |  | 735 | 29397.979 | 10.2887 | 4.9767
736 | 29437.976 | 10.29 | 4.9746 |  |  | 736 | 29437.976 | 10.29 | 4.9779 |  |  | 736 | 29437.976 | 10.29 | 4.9766
737 | 29477.974 | 10.2914 | 4.9745 |  |  | 737 | 29477.974 | 10.2914 | 4.9778 |  |  | 737 | 29477.974 | 10.2914 | 4.9765
738 | 29517.971 | 10.2928 | 4.9744 |  |  | 738 | 29517.971 | 10.2928 | 4.9777 |  |  | 738 | 29517.971 | 10.2928 | 4.9764
739 | 29557.968 | 10.2941 | 4.9743 |  |  | 739 | 29557.968 | 10.2941 | 4.9776 |  |  | 739 | 29557.968 | 10.2941 | 4.9763
740 | 29597.965 | 10.2955 | 4.9742 |  |  | 740 | 29597.965 | 10.2955 | 4.9775 |  |  | 740 | 29597.965 | 10.2955 | 4.9761
741 | 29637.963 | 10.2968 | 4.9741 |  |  | 741 | 29637.963 | 10.2968 | 4.9774 |  |  | 741 | 29637.963 | 10.2968 | 4.976
742 | 29677.96 | 10.2982 | 4.974 |  |  | 742 | 29677.96 | 10.2982 | 4.9773 |  |  | 742 | 29677.96 | 10.2982 | 4.9759
743 | 29717.957 | 10.2995 | 4.9739 |  |  | 743 | 29717.957 | 10.2995 | 4.9772 |  |  | 743 | 29717.957 | 10.2995 | 4.9758
744 | 29757.954 | 10.3009 | 4.9738 |  |  | 744 | 29757.954 | 10.3009 | 4.9771 |  |  | 744 | 29757.954 | 10.3009 | 4.9757
745 | 29797.952 | 10.3022 | 4.9737 |  |  | 745 | 29797.952 | 10.3022 | 4.977 |  |  | 745 | 29797.952 | 10.3022 | 4.9756
746 | 29837.949 | 10.3035 | 4.9736 |  |  | 746 | 29837.949 | 10.3035 | 4.9769 |  |  | 746 | 29837.949 | 10.3035 | 4.9755
747 | 29877.946 | 10.3049 | 4.9735 |  |  | 747 | 29877.946 | 10.3049 | 4.9768 |  |  | 747 | 29877.946 | 10.3049 | 4.9754
748 | 29917.943 | 10.3062 | 4.9734 |  |  | 748 | 29917.943 | 10.3062 | 4.9768 |  |  | 748 | 29917.943 | 10.3062 | 4.9752
749 | 29957.941 | 10.3075 | 4.9733 |  |  | 749 | 29957.941 | 10.3075 | 4.9767 |  |  | 749 | 29957.941 | 10.3075 | 4.9751
750 | 29997.938 | 10.3089 | 4.9732 |  |  | 750 | 29997.938 | 10.3089 | 4.9766 |  |  | 750 | 29997.938 | 10.3089 | 4.975
751 | 30037.935 | 10.3102 | 4.9732 |  |  | 751 | 30037.935 | 10.3102 | 4.9765 |  |  | 751 | 30037.935 | 10.3102 | 4.9749
752 | 30077.932 | 10.3115 | 4.9731 |  |  | 752 | 30077.932 | 10.3115 | 4.9764 |  |  | 752 | 30077.932 | 10.3115 | 4.9748
753 | 30117.93 | 10.3129 | 4.973 |  |  | 753 | 30117.93 | 10.3129 | 4.9763 |  |  | 753 | 30117.93 | 10.3129 | 4.9747
754 | 30157.927 | 10.3142 | 4.9729 |  |  | 754 | 30157.927 | 10.3142 | 4.9762 |  |  | 754 | 30157.927 | 10.3142 | 4.9746
755 | 30197.924 | 10.3155 | 4.9728 |  |  | 755 | 30197.924 | 10.3155 | 4.9761 |  |  | 755 | 30197.924 | 10.3155 | 4.9745
756 | 30237.921 | 10.3169 | 4.9727 |  |  | 756 | 30237.921 | 10.3169 | 4.976 |  |  | 756 | 30237.921 | 10.3169 | 4.9743
757 | 30277.919 | 10.3182 | 4.9726 |  |  | 757 | 30277.919 | 10.3182 | 4.9759 |  |  | 757 | 30277.919 | 10.3182 | 4.9742
758 | 30317.916 | 10.3195 | 4.9725 |  |  | 758 | 30317.916 | 10.3195 | 4.9758 |  |  | 758 | 30317.916 | 10.3195 | 4.9741
759 | 30357.913 | 10.3208 | 4.9724 |  |  | 759 | 30357.913 | 10.3208 | 4.9757 |  |  | 759 | 30357.913 | 10.3208 | 4.974
760 | 30397.91 | 10.3221 | 4.9723 |  |  | 760 | 30397.91 | 10.3221 | 4.9756 |  |  | 760 | 30397.91 | 10.3221 | 4.9739
761 | 30437.908 | 10.3234 | 4.9722 |  |  | 761 | 30437.908 | 10.3234 | 4.9755 |  |  | 761 | 30437.908 | 10.3234 | 4.9738
762 | 30477.905 | 10.3248 | 4.9721 |  |  | 762 | 30477.905 | 10.3248 | 4.9754 |  |  | 762 | 30477.905 | 10.3248 | 4.9737
763 | 30517.902 | 10.3261 | 4.972 |  |  | 763 | 30517.902 | 10.3261 | 4.9753 |  |  | 763 | 30517.902 | 10.3261 | 4.9736
764 | 30557.899 | 10.3274 | 4.9719 |  |  | 764 | 30557.899 | 10.3274 | 4.9752 |  |  | 764 | 30557.899 | 10.3274 | 4.9734
765 | 30597.897 | 10.3287 | 4.9718 |  |  | 765 | 30597.897 | 10.3287 | 4.9751 |  |  | 765 | 30597.897 | 10.3287 | 4.9733
766 | 30637.894 | 10.33 | 4.9717 |  |  | 766 | 30637.894 | 10.33 | 4.975 |  |  | 766 | 30637.894 | 10.33 | 4.9732
767 | 30677.891 | 10.3313 | 4.9716 |  |  | 767 | 30677.891 | 10.3313 | 4.9749 |  |  | 767 | 30677.891 | 10.3313 | 4.9731
768 | 30717.888 | 10.3326 | 4.9715 |  |  | 768 | 30717.888 | 10.3326 | 4.9748 |  |  | 768 | 30717.888 | 10.3326 | 4.973
769 | 30757.886 | 10.3339 | 4.9714 |  |  | 769 | 30757.886 | 10.3339 | 4.9747 |  |  | 769 | 30757.886 | 10.3339 | 4.9729
770 | 30797.883 | 10.3352 | 4.9713 |  |  | 770 | 30797.883 | 10.3352 | 4.9746 |  |  | 770 | 30797.883 | 10.3352 | 4.9728
771 | 30837.88 | 10.3365 | 4.9712 |  |  | 771 | 30837.88 | 10.3365 | 4.9745 |  |  | 771 | 30837.88 | 10.3365 | 4.9727
772 | 30877.877 | 10.3378 | 4.9711 |  |  | 772 | 30877.877 | 10.3378 | 4.9744 |  |  | 772 | 30877.877 | 10.3378 | 4.9725
773 | 30917.875 | 10.3391 | 4.971 |  |  | 773 | 30917.875 | 10.3391 | 4.9743 |  |  | 773 | 30917.875 | 10.3391 | 4.9724
774 | 30957.872 | 10.3404 | 4.9709 |  |  | 774 | 30957.872 | 10.3404 | 4.9742 |  |  | 774 | 30957.872 | 10.3404 | 4.9723
775 | 30997.869 | 10.3417 | 4.9708 |  |  | 775 | 30997.869 | 10.3417 | 4.9741 |  |  | 775 | 30997.869 | 10.3417 | 4.9722
776 | 31037.866 | 10.343 | 4.9707 |  |  | 776 | 31037.866 | 10.343 | 4.974 |  |  | 776 | 31037.866 | 10.343 | 4.9721
777 | 31077.864 | 10.3443 | 4.9706 |  |  | 777 | 31077.864 | 10.3443 | 4.974 |  |  | 777 | 31077.864 | 10.3443 | 4.972
778 | 31117.861 | 10.3455 | 4.9705 |  |  | 778 | 31117.861 | 10.3455 | 4.9739 |  |  | 778 | 31117.861 | 10.3455 | 4.9719
779 | 31157.858 | 10.3468 | 4.9704 |  |  | 779 | 31157.858 | 10.3468 | 4.9738 |  |  | 779 | 31157.858 | 10.3468 | 4.9718
780 | 31197.855 | 10.3481 | 4.9703 |  |  | 780 | 31197.855 | 10.3481 | 4.9737 |  |  | 780 | 31197.855 | 10.3481 | 4.9717
781 | 31237.853 | 10.3494 | 4.9702 |  |  | 781 | 31237.853 | 10.3494 | 4.9736 |  |  | 781 | 31237.853 | 10.3494 | 4.9715
782 | 31277.85 | 10.3507 | 4.9702 |  |  | 782 | 31277.85 | 10.3507 | 4.9735 |  |  | 782 | 31277.85 | 10.3507 | 4.9714
783 | 31317.847 | 10.3519 | 4.9701 |  |  | 783 | 31317.847 | 10.3519 | 4.9734 |  |  | 783 | 31317.847 | 10.3519 | 4.9713
784 | 31357.844 | 10.3532 | 4.97 |  |  | 784 | 31357.844 | 10.3532 | 4.9733 |  |  | 784 | 31357.844 | 10.3532 | 4.9712
785 | 31397.842 | 10.3545 | 4.9699 |  |  | 785 | 31397.842 | 10.3545 | 4.9732 |  |  | 785 | 31397.842 | 10.3545 | 4.9711
786 | 31437.839 | 10.3558 | 4.9698 |  |  | 786 | 31437.839 | 10.3558 | 4.9731 |  |  | 786 | 31437.839 | 10.3558 | 4.971
787 | 31477.836 | 10.357 | 4.9697 |  |  | 787 | 31477.836 | 10.357 | 4.973 |  |  | 787 | 31477.836 | 10.357 | 4.9709
788 | 31517.833 | 10.3583 | 4.9696 |  |  | 788 | 31517.833 | 10.3583 | 4.9729 |  |  | 788 | 31517.833 | 10.3583 | 4.9708
789 | 31557.831 | 10.3596 | 4.9695 |  |  | 789 | 31557.831 | 10.3596 | 4.9728 |  |  | 789 | 31557.831 | 10.3596 | 4.9707
790 | 31597.828 | 10.3608 | 4.9694 |  |  | 790 | 31597.828 | 10.3608 | 4.9727 |  |  | 790 | 31597.828 | 10.3608 | 4.9705
791 | 31637.825 | 10.3621 | 4.9693 |  |  | 791 | 31637.825 | 10.3621 | 4.9726 |  |  | 791 | 31637.825 | 10.3621 | 4.9704
792 | 31677.822 | 10.3634 | 4.9692 |  |  | 792 | 31677.822 | 10.3634 | 4.9725 |  |  | 792 | 31677.822 | 10.3634 | 4.9703
793 | 31717.82 | 10.3646 | 4.9691 |  |  | 793 | 31717.82 | 10.3646 | 4.9724 |  |  | 793 | 31717.82 | 10.3646 | 4.9702
794 | 31757.817 | 10.3659 | 4.969 |  |  | 794 | 31757.817 | 10.3659 | 4.9723 |  |  | 794 | 31757.817 | 10.3659 | 4.9701
795 | 31797.814 | 10.3672 | 4.9689 |  |  | 795 | 31797.814 | 10.3672 | 4.9722 |  |  | 795 | 31797.814 | 10.3672 | 4.97
796 | 31837.811 | 10.3684 | 4.9688 |  |  | 796 | 31837.811 | 10.3684 | 4.9721 |  |  | 796 | 31837.811 | 10.3684 | 4.9699
797 | 31877.809 | 10.3697 | 4.9687 |  |  | 797 | 31877.809 | 10.3697 | 4.972 |  |  | 797 | 31877.809 | 10.3697 | 4.9698
798 | 31917.806 | 10.3709 | 4.9686 |  |  | 798 | 31917.806 | 10.3709 | 4.9719 |  |  | 798 | 31917.806 | 10.3709 | 4.9696
799 | 31957.803 | 10.3722 | 4.9685 |  |  | 799 | 31957.803 | 10.3722 | 4.9719 |  |  | 799 | 31957.803 | 10.3722 | 4.9695
800 | 31997.8 | 10.3734 | 4.9684 |  |  | 800 | 31997.8 | 10.3734 | 4.9718 |  |  | 800 | 31997.8 | 10.3734 | 4.9694
801 | 32037.798 | 10.3747 | 4.9683 |  |  | 801 | 32037.798 | 10.3747 | 4.9717 |  |  | 801 | 32037.798 | 10.3747 | 4.9693
802 | 32077.795 | 10.3759 | 4.9682 |  |  | 802 | 32077.795 | 10.3759 | 4.9716 |  |  | 802 | 32077.795 | 10.3759 | 4.9692
803 | 32117.792 | 10.3772 | 4.9681 |  |  | 803 | 32117.792 | 10.3772 | 4.9715 |  |  | 803 | 32117.792 | 10.3772 | 4.9691
804 | 32157.789 | 10.3784 | 4.968 |  |  | 804 | 32157.789 | 10.3784 | 4.9714 |  |  | 804 | 32157.789 | 10.3784 | 4.969
805 | 32197.787 | 10.3797 | 4.9679 |  |  | 805 | 32197.787 | 10.3797 | 4.9713 |  |  | 805 | 32197.787 | 10.3797 | 4.9689
806 | 32237.784 | 10.3809 | 4.9679 |  |  | 806 | 32237.784 | 10.3809 | 4.9712 |  |  | 806 | 32237.784 | 10.3809 | 4.9688
807 | 32277.781 | 10.3821 | 4.9678 |  |  | 807 | 32277.781 | 10.3821 | 4.9711 |  |  | 807 | 32277.781 | 10.3821 | 4.9686
808 | 32317.778 | 10.3834 | 4.9677 |  |  | 808 | 32317.778 | 10.3834 | 4.971 |  |  | 808 | 32317.778 | 10.3834 | 4.9685
809 | 32357.776 | 10.3846 | 4.9676 |  |  | 809 | 32357.776 | 10.3846 | 4.9709 |  |  | 809 | 32357.776 | 10.3846 | 4.9684
810 | 32397.773 | 10.3858 | 4.9675 |  |  | 810 | 32397.773 | 10.3858 | 4.9708 |  |  | 810 | 32397.773 | 10.3858 | 4.9683
811 | 32437.77 | 10.3871 | 4.9674 |  |  | 811 | 32437.77 | 10.3871 | 4.9707 |  |  | 811 | 32437.77 | 10.3871 | 4.9682
812 | 32477.767 | 10.3883 | 4.9673 |  |  | 812 | 32477.767 | 10.3883 | 4.9706 |  |  | 812 | 32477.767 | 10.3883 | 4.9681
813 | 32517.765 | 10.3895 | 4.9672 |  |  | 813 | 32517.765 | 10.3895 | 4.9705 |  |  | 813 | 32517.765 | 10.3895 | 4.968
814 | 32557.762 | 10.3908 | 4.9671 |  |  | 814 | 32557.762 | 10.3908 | 4.9704 |  |  | 814 | 32557.762 | 10.3908 | 4.9679
815 | 32597.759 | 10.392 | 4.967 |  |  | 815 | 32597.759 | 10.392 | 4.9703 |  |  | 815 | 32597.759 | 10.392 | 4.9678
816 | 32637.756 | 10.3932 | 4.9669 |  |  | 816 | 32637.756 | 10.3932 | 4.9702 |  |  | 816 | 32637.756 | 10.3932 | 4.9676
817 | 32677.754 | 10.3944 | 4.9668 |  |  | 817 | 32677.754 | 10.3944 | 4.9701 |  |  | 817 | 32677.754 | 10.3944 | 4.9675
818 | 32717.751 | 10.3957 | 4.9667 |  |  | 818 | 32717.751 | 10.3957 | 4.9701 |  |  | 818 | 32717.751 | 10.3957 | 4.9674
819 | 32757.748 | 10.3969 | 4.9666 |  |  | 819 | 32757.748 | 10.3969 | 4.97 |  |  | 819 | 32757.748 | 10.3969 | 4.9673
820 | 32797.745 | 10.3981 | 4.9665 |  |  | 820 | 32797.745 | 10.3981 | 4.9699 |  |  | 820 | 32797.745 | 10.3981 | 4.9672
821 | 32837.743 | 10.3993 | 4.9664 |  |  | 821 | 32837.743 | 10.3993 | 4.9698 |  |  | 821 | 32837.743 | 10.3993 | 4.9671
822 | 32877.74 | 10.4006 | 4.9663 |  |  | 822 | 32877.74 | 10.4006 | 4.9697 |  |  | 822 | 32877.74 | 10.4006 | 4.967
823 | 32917.737 | 10.4018 | 4.9662 |  |  | 823 | 32917.737 | 10.4018 | 4.9696 |  |  | 823 | 32917.737 | 10.4018 | 4.9669
824 | 32957.734 | 10.403 | 4.9661 |  |  | 824 | 32957.734 | 10.403 | 4.9695 |  |  | 824 | 32957.734 | 10.403 | 4.9668
825 | 32997.732 | 10.4042 | 4.966 |  |  | 825 | 32997.732 | 10.4042 | 4.9694 |  |  | 825 | 32997.732 | 10.4042 | 4.9666
826 | 33037.729 | 10.4054 | 4.9659 |  |  | 826 | 33037.729 | 10.4054 | 4.9693 |  |  | 826 | 33037.729 | 10.4054 | 4.9665
827 | 33077.726 | 10.4066 | 4.9658 |  |  | 827 | 33077.726 | 10.4066 | 4.9692 |  |  | 827 | 33077.726 | 10.4066 | 4.9664
828 | 33117.723 | 10.4078 | 4.9658 |  |  | 828 | 33117.723 | 10.4078 | 4.9691 |  |  | 828 | 33117.723 | 10.4078 | 4.9663
829 | 33157.721 | 10.409 | 4.9657 |  |  | 829 | 33157.721 | 10.409 | 4.969 |  |  | 829 | 33157.721 | 10.409 | 4.9662
830 | 33197.718 | 10.4102 | 4.9656 |  |  | 830 | 33197.718 | 10.4102 | 4.9689 |  |  | 830 | 33197.718 | 10.4102 | 4.9661
831 | 33237.715 | 10.4114 | 4.9655 |  |  | 831 | 33237.715 | 10.4114 | 4.9688 |  |  | 831 | 33237.715 | 10.4114 | 4.966
832 | 33277.712 | 10.4126 | 4.9654 |  |  | 832 | 33277.712 | 10.4126 | 4.9687 |  |  | 832 | 33277.712 | 10.4126 | 4.9659
833 | 33317.71 | 10.4138 | 4.9653 |  |  | 833 | 33317.71 | 10.4138 | 4.9686 |  |  | 833 | 33317.71 | 10.4138 | 4.9657
834 | 33357.707 | 10.415 | 4.9652 |  |  | 834 | 33357.707 | 10.415 | 4.9685 |  |  | 834 | 33357.707 | 10.415 | 4.9656
835 | 33397.704 | 10.4162 | 4.9651 |  |  | 835 | 33397.704 | 10.4162 | 4.9684 |  |  | 835 | 33397.704 | 10.4162 | 4.9655
836 | 33437.701 | 10.4174 | 4.965 |  |  | 836 | 33437.701 | 10.4174 | 4.9683 |  |  | 836 | 33437.701 | 10.4174 | 4.9654
837 | 33477.699 | 10.4186 | 4.9649 |  |  | 837 | 33477.699 | 10.4186 | 4.9683 |  |  | 837 | 33477.699 | 10.4186 | 4.9653
838 | 33517.696 | 10.4198 | 4.9648 |  |  | 838 | 33517.696 | 10.4198 | 4.9682 |  |  | 838 | 33517.696 | 10.4198 | 4.9652
839 | 33557.693 | 10.421 | 4.9647 |  |  | 839 | 33557.693 | 10.421 | 4.9681 |  |  | 839 | 33557.693 | 10.421 | 4.9651
840 | 33597.69 | 10.4222 | 4.9646 |  |  | 840 | 33597.69 | 10.4222 | 4.968 |  |  | 840 | 33597.69 | 10.4222 | 4.965
841 | 33637.688 | 10.4234 | 4.9645 |  |  | 841 | 33637.688 | 10.4234 | 4.9679 |  |  | 841 | 33637.688 | 10.4234 | 4.9649
842 | 33677.685 | 10.4246 | 4.9644 |  |  | 842 | 33677.685 | 10.4246 | 4.9678 |  |  | 842 | 33677.685 | 10.4246 | 4.9647
843 | 33717.682 | 10.4258 | 4.9643 |  |  | 843 | 33717.682 | 10.4258 | 4.9677 |  |  | 843 | 33717.682 | 10.4258 | 4.9646
844 | 33757.679 | 10.427 | 4.9642 |  |  | 844 | 33757.679 | 10.427 | 4.9676 |  |  | 844 | 33757.679 | 10.427 | 4.9645
845 | 33797.677 | 10.4281 | 4.9641 |  |  | 845 | 33797.677 | 10.4281 | 4.9675 |  |  | 845 | 33797.677 | 10.4281 | 4.9644
846 | 33837.674 | 10.4293 | 4.964 |  |  | 846 | 33837.674 | 10.4293 | 4.9674 |  |  | 846 | 33837.674 | 10.4293 | 4.9643
847 | 33877.671 | 10.4305 | 4.9639 |  |  | 847 | 33877.671 | 10.4305 | 4.9673 |  |  | 847 | 33877.671 | 10.4305 | 4.9642
848 | 33917.668 | 10.4317 | 4.9638 |  |  | 848 | 33917.668 | 10.4317 | 4.9672 |  |  | 848 | 33917.668 | 10.4317 | 4.9641
849 | 33957.666 | 10.4329 | 4.9637 |  |  | 849 | 33957.666 | 10.4329 | 4.9671 |  |  | 849 | 33957.666 | 10.4329 | 4.964
850 | 33997.663 | 10.434 | 4.9637 |  |  | 850 | 33997.663 | 10.434 | 4.967 |  |  | 850 | 33997.663 | 10.434 | 4.9638
851 | 34037.66 | 10.4352 | 4.9636 |  |  | 851 | 34037.66 | 10.4352 | 4.9669 |  |  | 851 | 34037.66 | 10.4352 | 4.9637
852 | 34077.657 | 10.4364 | 4.9635 |  |  | 852 | 34077.657 | 10.4364 | 4.9668 |  |  | 852 | 34077.657 | 10.4364 | 4.9636
853 | 34117.655 | 10.4376 | 4.9634 |  |  | 853 | 34117.655 | 10.4376 | 4.9667 |  |  | 853 | 34117.655 | 10.4376 | 4.9635
854 | 34157.652 | 10.4387 | 4.9633 |  |  | 854 | 34157.652 | 10.4387 | 4.9666 |  |  | 854 | 34157.652 | 10.4387 | 4.9634
855 | 34197.649 | 10.4399 | 4.9632 |  |  | 855 | 34197.649 | 10.4399 | 4.9665 |  |  | 855 | 34197.649 | 10.4399 | 4.9633
856 | 34237.646 | 10.4411 | 4.9631 |  |  | 856 | 34237.646 | 10.4411 | 4.9665 |  |  | 856 | 34237.646 | 10.4411 | 4.9632
857 | 34277.644 | 10.4422 | 4.963 |  |  | 857 | 34277.644 | 10.4422 | 4.9664 |  |  | 857 | 34277.644 | 10.4422 | 4.9631
858 | 34317.641 | 10.4434 | 4.9629 |  |  | 858 | 34317.641 | 10.4434 | 4.9663 |  |  | 858 | 34317.641 | 10.4434 | 4.9629
859 | 34357.638 | 10.4446 | 4.9628 |  |  | 859 | 34357.638 | 10.4446 | 4.9662 |  |  | 859 | 34357.638 | 10.4446 | 4.9628
860 | 34397.635 | 10.4457 | 4.9627 |  |  | 860 | 34397.635 | 10.4457 | 4.9661 |  |  | 860 | 34397.635 | 10.4457 | 4.9627
861 | 34437.633 | 10.4469 | 4.9626 |  |  | 861 | 34437.633 | 10.4469 | 4.966 |  |  | 861 | 34437.633 | 10.4469 | 4.9626
862 | 34477.63 | 10.4481 | 4.9625 |  |  | 862 | 34477.63 | 10.4481 | 4.9659 |  |  | 862 | 34477.63 | 10.4481 | 4.9625
863 | 34517.627 | 10.4492 | 4.9624 |  |  | 863 | 34517.627 | 10.4492 | 4.9658 |  |  | 863 | 34517.627 | 10.4492 | 4.9624
864 | 34557.624 | 10.4504 | 4.9623 |  |  | 864 | 34557.624 | 10.4504 | 4.9657 |  |  | 864 | 34557.624 | 10.4504 | 4.9623
865 | 34597.622 | 10.4515 | 4.9622 |  |  | 865 | 34597.622 | 10.4515 | 4.9656 |  |  | 865 | 34597.622 | 10.4515 | 4.9622
866 | 34637.619 | 10.4527 | 4.9621 |  |  | 866 | 34637.619 | 10.4527 | 4.9655 |  |  | 866 | 34637.619 | 10.4527 | 4.962
867 | 34677.616 | 10.4538 | 4.962 |  |  | 867 | 34677.616 | 10.4538 | 4.9654 |  |  | 867 | 34677.616 | 10.4538 | 4.9619
868 | 34717.613 | 10.455 | 4.9619 |  |  | 868 | 34717.613 | 10.455 | 4.9653 |  |  | 868 | 34717.613 | 10.455 | 4.9618
869 | 34757.611 | 10.4562 | 4.9618 |  |  | 869 | 34757.611 | 10.4562 | 4.9652 |  |  | 869 | 34757.611 | 10.4562 | 4.9617
870 | 34797.608 | 10.4573 | 4.9617 |  |  | 870 | 34797.608 | 10.4573 | 4.9651 |  |  | 870 | 34797.608 | 10.4573 | 4.9616
871 | 34837.605 | 10.4585 | 4.9616 |  |  | 871 | 34837.605 | 10.4585 | 4.965 |  |  | 871 | 34837.605 | 10.4585 | 4.9615
872 | 34877.602 | 10.4596 | 4.9615 |  |  | 872 | 34877.602 | 10.4596 | 4.9649 |  |  | 872 | 34877.602 | 10.4596 | 4.9614
873 | 34917.6 | 10.4607 | 4.9615 |  |  | 873 | 34917.6 | 10.4607 | 4.9648 |  |  | 873 | 34917.6 | 10.4607 | 4.9613
874 | 34957.597 | 10.4619 | 4.9614 |  |  | 874 | 34957.597 | 10.4619 | 4.9647 |  |  | 874 | 34957.597 | 10.4619 | 4.9611
875 | 34997.594 | 10.463 | 4.9613 |  |  | 875 | 34997.594 | 10.463 | 4.9647 |  |  | 875 | 34997.594 | 10.463 | 4.961
876 | 35037.591 | 10.4642 | 4.9612 |  |  | 876 | 35037.591 | 10.4642 | 4.9646 |  |  | 876 | 35037.591 | 10.4642 | 4.9609
877 | 35077.589 | 10.4653 | 4.9611 |  |  | 877 | 35077.589 | 10.4653 | 4.9645 |  |  | 877 | 35077.589 | 10.4653 | 4.9608
878 | 35117.586 | 10.4665 | 4.961 |  |  | 878 | 35117.586 | 10.4665 | 4.9644 |  |  | 878 | 35117.586 | 10.4665 | 4.9607
879 | 35157.583 | 10.4676 | 4.9609 |  |  | 879 | 35157.583 | 10.4676 | 4.9643 |  |  | 879 | 35157.583 | 10.4676 | 4.9606
880 | 35197.58 | 10.4687 | 4.9608 |  |  | 880 | 35197.58 | 10.4687 | 4.9642 |  |  | 880 | 35197.58 | 10.4687 | 4.9605
881 | 35237.578 | 10.4699 | 4.9607 |  |  | 881 | 35237.578 | 10.4699 | 4.9641 |  |  | 881 | 35237.578 | 10.4699 | 4.9604
882 | 35277.575 | 10.471 | 4.9606 |  |  | 882 | 35277.575 | 10.471 | 4.964 |  |  | 882 | 35277.575 | 10.471 | 4.9602
883 | 35317.572 | 10.4721 | 4.9605 |  |  | 883 | 35317.572 | 10.4721 | 4.9639 |  |  | 883 | 35317.572 | 10.4721 | 4.9601
884 | 35357.569 | 10.4733 | 4.9604 |  |  | 884 | 35357.569 | 10.4733 | 4.9638 |  |  | 884 | 35357.569 | 10.4733 | 4.96
885 | 35397.567 | 10.4744 | 4.9603 |  |  | 885 | 35397.567 | 10.4744 | 4.9637 |  |  | 885 | 35397.567 | 10.4744 | 4.9599
886 | 35437.564 | 10.4755 | 4.9602 |  |  | 886 | 35437.564 | 10.4755 | 4.9636 |  |  | 886 | 35437.564 | 10.4755 | 4.9598
887 | 35477.561 | 10.4767 | 4.9601 |  |  | 887 | 35477.561 | 10.4767 | 4.9635 |  |  | 887 | 35477.561 | 10.4767 | 4.9597
888 | 35517.558 | 10.4778 | 4.96 |  |  | 888 | 35517.558 | 10.4778 | 4.9634 |  |  | 888 | 35517.558 | 10.4778 | 4.9596
889 | 35557.556 | 10.4789 | 4.9599 |  |  | 889 | 35557.556 | 10.4789 | 4.9633 |  |  | 889 | 35557.556 | 10.4789 | 4.9594
890 | 35597.553 | 10.48 | 4.9598 |  |  | 890 | 35597.553 | 10.48 | 4.9632 |  |  | 890 | 35597.553 | 10.48 | 4.9593
891 | 35637.55 | 10.4812 | 4.9597 |  |  | 891 | 35637.55 | 10.4812 | 4.9631 |  |  | 891 | 35637.55 | 10.4812 | 4.9592
892 | 35677.547 | 10.4823 | 4.9596 |  |  | 892 | 35677.547 | 10.4823 | 4.963 |  |  | 892 | 35677.547 | 10.4823 | 4.9591
893 | 35717.545 | 10.4834 | 4.9595 |  |  | 893 | 35717.545 | 10.4834 | 4.9629 |  |  | 893 | 35717.545 | 10.4834 | 4.959
894 | 35757.542 | 10.4845 | 4.9594 |  |  | 894 | 35757.542 | 10.4845 | 4.9628 |  |  | 894 | 35757.542 | 10.4845 | 4.9589
895 | 35797.539 | 10.4856 | 4.9593 |  |  | 895 | 35797.539 | 10.4856 | 4.9628 |  |  | 895 | 35797.539 | 10.4856 | 4.9588
896 | 35837.536 | 10.4868 | 4.9592 |  |  | 896 | 35837.536 | 10.4868 | 4.9627 |  |  | 896 | 35837.536 | 10.4868 | 4.9587
897 | 35877.534 | 10.4879 | 4.9591 |  |  | 897 | 35877.534 | 10.4879 | 4.9626 |  |  | 897 | 35877.534 | 10.4879 | 4.9585
898 | 35917.531 | 10.489 | 4.9591 |  |  | 898 | 35917.531 | 10.489 | 4.9625 |  |  | 898 | 35917.531 | 10.489 | 4.9584
899 | 35957.528 | 10.4901 | 4.959 |  |  | 899 | 35957.528 | 10.4901 | 4.9624 |  |  | 899 | 35957.528 | 10.4901 | 4.9583
900 | 35997.525 | 10.4912 | 4.9589 |  |  | 900 | 35997.525 | 10.4912 | 4.9623 |  |  | 900 | 35997.525 | 10.4912 | 4.9582
901 | 36037.523 | 10.4923 | 4.9588 |  |  | 901 | 36037.523 | 10.4923 | 4.9622 |  |  | 901 | 36037.523 | 10.4923 | 4.9581
902 | 36077.52 | 10.4934 | 4.9587 |  |  | 902 | 36077.52 | 10.4934 | 4.9621 |  |  | 902 | 36077.52 | 10.4934 | 4.958
903 | 36117.517 | 10.4945 | 4.9586 |  |  | 903 | 36117.517 | 10.4945 | 4.962 |  |  | 903 | 36117.517 | 10.4945 | 4.9579
904 | 36157.515 | 10.4956 | 4.9585 |  |  | 904 | 36157.515 | 10.4956 | 4.9619 |  |  | 904 | 36157.515 | 10.4956 | 4.9577
905 | 36197.512 | 10.4967 | 4.9584 |  |  | 905 | 36197.512 | 10.4967 | 4.9618 |  |  | 905 | 36197.512 | 10.4967 | 4.9576
906 | 36237.509 | 10.4978 | 4.9583 |  |  | 906 | 36237.509 | 10.4978 | 4.9617 |  |  | 906 | 36237.509 | 10.4978 | 4.9575
907 | 36277.506 | 10.499 | 4.9582 |  |  | 907 | 36277.506 | 10.499 | 4.9616 |  |  | 907 | 36277.506 | 10.499 | 4.9574
908 | 36317.504 | 10.5001 | 4.9581 |  |  | 908 | 36317.504 | 10.5001 | 4.9615 |  |  | 908 | 36317.504 | 10.5001 | 4.9573
909 | 36357.501 | 10.5012 | 4.958 |  |  | 909 | 36357.501 | 10.5012 | 4.9614 |  |  | 909 | 36357.501 | 10.5012 | 4.9572
910 | 36397.498 | 10.5023 | 4.9579 |  |  | 910 | 36397.498 | 10.5023 | 4.9613 |  |  | 910 | 36397.498 | 10.5023 | 4.9571
911 | 36437.495 | 10.5034 | 4.9578 |  |  | 911 | 36437.495 | 10.5034 | 4.9612 |  |  | 911 | 36437.495 | 10.5034 | 4.9569
912 | 36477.493 | 10.5045 | 4.9577 |  |  | 912 | 36477.493 | 10.5045 | 4.9611 |  |  | 912 | 36477.493 | 10.5045 | 4.9568
913 | 36517.49 | 10.5055 | 4.9576 |  |  | 913 | 36517.49 | 10.5055 | 4.961 |  |  | 913 | 36517.49 | 10.5055 | 4.9567
914 | 36557.487 | 10.5066 | 4.9575 |  |  | 914 | 36557.487 | 10.5066 | 4.9609 |  |  | 914 | 36557.487 | 10.5066 | 4.9566
915 | 36597.484 | 10.5077 | 4.9574 |  |  | 915 | 36597.484 | 10.5077 | 4.9608 |  |  | 915 | 36597.484 | 10.5077 | 4.9565
916 | 36637.482 | 10.5088 | 4.9573 |  |  | 916 | 36637.482 | 10.5088 | 4.9607 |  |  | 916 | 36637.482 | 10.5088 | 4.9564
917 | 36677.479 | 10.5099 | 4.9572 |  |  | 917 | 36677.479 | 10.5099 | 4.9606 |  |  | 917 | 36677.479 | 10.5099 | 4.9563
918 | 36717.476 | 10.511 | 4.9571 |  |  | 918 | 36717.476 | 10.511 | 4.9606 |  |  | 918 | 36717.476 | 10.511 | 4.9561
919 | 36757.473 | 10.5121 | 4.957 |  |  | 919 | 36757.473 | 10.5121 | 4.9605 |  |  | 919 | 36757.473 | 10.5121 | 4.956
920 | 36797.471 | 10.5132 | 4.9569 |  |  | 920 | 36797.471 | 10.5132 | 4.9604 |  |  | 920 | 36797.471 | 10.5132 | 4.9559
921 | 36837.468 | 10.5143 | 4.9568 |  |  | 921 | 36837.468 | 10.5143 | 4.9603 |  |  | 921 | 36837.468 | 10.5143 | 4.9558
922 | 36877.465 | 10.5154 | 4.9567 |  |  | 922 | 36877.465 | 10.5154 | 4.9602 |  |  | 922 | 36877.465 | 10.5154 | 4.9557
923 | 36917.462 | 10.5164 | 4.9566 |  |  | 923 | 36917.462 | 10.5164 | 4.9601 |  |  | 923 | 36917.462 | 10.5164 | 4.9556
924 | 36957.46 | 10.5175 | 4.9565 |  |  | 924 | 36957.46 | 10.5175 | 4.96 |  |  | 924 | 36957.46 | 10.5175 | 4.9555
925 | 36997.457 | 10.5186 | 4.9564 |  |  | 925 | 36997.457 | 10.5186 | 4.9599 |  |  | 925 | 36997.457 | 10.5186 | 4.9553
926 | 37037.454 | 10.5197 | 4.9563 |  |  | 926 | 37037.454 | 10.5197 | 4.9598 |  |  | 926 | 37037.454 | 10.5197 | 4.9552
927 | 37077.451 | 10.5208 | 4.9562 |  |  | 927 | 37077.451 | 10.5208 | 4.9597 |  |  | 927 | 37077.451 | 10.5208 | 4.9551
928 | 37117.449 | 10.5218 | 4.9561 |  |  | 928 | 37117.449 | 10.5218 | 4.9596 |  |  | 928 | 37117.449 | 10.5218 | 4.955
929 | 37157.446 | 10.5229 | 4.956 |  |  | 929 | 37157.446 | 10.5229 | 4.9595 |  |  | 929 | 37157.446 | 10.5229 | 4.9549
930 | 37197.443 | 10.524 | 4.956 |  |  | 930 | 37197.443 | 10.524 | 4.9594 |  |  | 930 | 37197.443 | 10.524 | 4.9548
931 | 37237.44 | 10.5251 | 4.9559 |  |  | 931 | 37237.44 | 10.5251 | 4.9593 |  |  | 931 | 37237.44 | 10.5251 | 4.9547
932 | 37277.438 | 10.5261 | 4.9558 |  |  | 932 | 37277.438 | 10.5261 | 4.9592 |  |  | 932 | 37277.438 | 10.5261 | 4.9545
933 | 37317.435 | 10.5272 | 4.9557 |  |  | 933 | 37317.435 | 10.5272 | 4.9591 |  |  | 933 | 37317.435 | 10.5272 | 4.9544
934 | 37357.432 | 10.5283 | 4.9556 |  |  | 934 | 37357.432 | 10.5283 | 4.959 |  |  | 934 | 37357.432 | 10.5283 | 4.9543
935 | 37397.429 | 10.5294 | 4.9555 |  |  | 935 | 37397.429 | 10.5294 | 4.9589 |  |  | 935 | 37397.429 | 10.5294 | 4.9542
936 | 37437.427 | 10.5304 | 4.9554 |  |  | 936 | 37437.427 | 10.5304 | 4.9588 |  |  | 936 | 37437.427 | 10.5304 | 4.9541
937 | 37477.424 | 10.5315 | 4.9553 |  |  | 937 | 37477.424 | 10.5315 | 4.9587 |  |  | 937 | 37477.424 | 10.5315 | 4.954
938 | 37517.421 | 10.5326 | 4.9552 |  |  | 938 | 37517.421 | 10.5326 | 4.9586 |  |  | 938 | 37517.421 | 10.5326 | 4.9538
939 | 37557.418 | 10.5336 | 4.9551 |  |  | 939 | 37557.418 | 10.5336 | 4.9585 |  |  | 939 | 37557.418 | 10.5336 | 4.9537
940 | 37597.416 | 10.5347 | 4.955 |  |  | 940 | 37597.416 | 10.5347 | 4.9584 |  |  | 940 | 37597.416 | 10.5347 | 4.9536
941 | 37637.413 | 10.5358 | 4.9549 |  |  | 941 | 37637.413 | 10.5358 | 4.9583 |  |  | 941 | 37637.413 | 10.5358 | 4.9535
942 | 37677.41 | 10.5368 | 4.9548 |  |  | 942 | 37677.41 | 10.5368 | 4.9582 |  |  | 942 | 37677.41 | 10.5368 | 4.9534
943 | 37717.407 | 10.5379 | 4.9547 |  |  | 943 | 37717.407 | 10.5379 | 4.9581 |  |  | 943 | 37717.407 | 10.5379 | 4.9533
944 | 37757.405 | 10.5389 | 4.9546 |  |  | 944 | 37757.405 | 10.5389 | 4.9581 |  |  | 944 | 37757.405 | 10.5389 | 4.9532
945 | 37797.402 | 10.54 | 4.9545 |  |  | 945 | 37797.402 | 10.54 | 4.958 |  |  | 945 | 37797.402 | 10.54 | 4.953
946 | 37837.399 | 10.5411 | 4.9544 |  |  | 946 | 37837.399 | 10.5411 | 4.9579 |  |  | 946 | 37837.399 | 10.5411 | 4.9529
947 | 37877.396 | 10.5421 | 4.9543 |  |  | 947 | 37877.396 | 10.5421 | 4.9578 |  |  | 947 | 37877.396 | 10.5421 | 4.9528
948 | 37917.394 | 10.5432 | 4.9542 |  |  | 948 | 37917.394 | 10.5432 | 4.9577 |  |  | 948 | 37917.394 | 10.5432 | 4.9527
949 | 37957.391 | 10.5442 | 4.9541 |  |  | 949 | 37957.391 | 10.5442 | 4.9576 |  |  | 949 | 37957.391 | 10.5442 | 4.9526
950 | 37997.388 | 10.5453 | 4.954 |  |  | 950 | 37997.388 | 10.5453 | 4.9575 |  |  | 950 | 37997.388 | 10.5453 | 4.9525
951 | 38037.385 | 10.5463 | 4.9539 |  |  | 951 | 38037.385 | 10.5463 | 4.9574 |  |  | 951 | 38037.385 | 10.5463 | 4.9523
952 | 38077.383 | 10.5474 | 4.9538 |  |  | 952 | 38077.383 | 10.5474 | 4.9573 |  |  | 952 | 38077.383 | 10.5474 | 4.9522
953 | 38117.38 | 10.5484 | 4.9537 |  |  | 953 | 38117.38 | 10.5484 | 4.9572 |  |  | 953 | 38117.38 | 10.5484 | 4.9521
954 | 38157.377 | 10.5495 | 4.9536 |  |  | 954 | 38157.377 | 10.5495 | 4.9571 |  |  | 954 | 38157.377 | 10.5495 | 4.952
955 | 38197.374 | 10.5505 | 4.9535 |  |  | 955 | 38197.374 | 10.5505 | 4.957 |  |  | 955 | 38197.374 | 10.5505 | 4.9519
956 | 38237.372 | 10.5516 | 4.9534 |  |  | 956 | 38237.372 | 10.5516 | 4.9569 |  |  | 956 | 38237.372 | 10.5516 | 4.9518
957 | 38277.369 | 10.5526 | 4.9533 |  |  | 957 | 38277.369 | 10.5526 | 4.9568 |  |  | 957 | 38277.369 | 10.5526 | 4.9516
958 | 38317.366 | 10.5537 | 4.9532 |  |  | 958 | 38317.366 | 10.5537 | 4.9567 |  |  | 958 | 38317.366 | 10.5537 | 4.9515
959 | 38357.363 | 10.5547 | 4.9531 |  |  | 959 | 38357.363 | 10.5547 | 4.9566 |  |  | 959 | 38357.363 | 10.5547 | 4.9514
960 | 38397.361 | 10.5557 | 4.953 |  |  | 960 | 38397.361 | 10.5557 | 4.9565 |  |  | 960 | 38397.361 | 10.5557 | 4.9513
961 | 38437.358 | 10.5568 | 4.9529 |  |  | 961 | 38437.358 | 10.5568 | 4.9564 |  |  | 961 | 38437.358 | 10.5568 | 4.9512
962 | 38477.355 | 10.5578 | 4.9528 |  |  | 962 | 38477.355 | 10.5578 | 4.9563 |  |  | 962 | 38477.355 | 10.5578 | 4.9511
963 | 38517.352 | 10.5589 | 4.9527 |  |  | 963 | 38517.352 | 10.5589 | 4.9562 |  |  | 963 | 38517.352 | 10.5589 | 4.9509
964 | 38557.35 | 10.5599 | 4.9526 |  |  | 964 | 38557.35 | 10.5599 | 4.9561 |  |  | 964 | 38557.35 | 10.5599 | 4.9508
965 | 38597.347 | 10.5609 | 4.9525 |  |  | 965 | 38597.347 | 10.5609 | 4.956 |  |  | 965 | 38597.347 | 10.5609 | 4.9507
966 | 38637.344 | 10.562 | 4.9524 |  |  | 966 | 38637.344 | 10.562 | 4.9559 |  |  | 966 | 38637.344 | 10.562 | 4.9506
967 | 38677.341 | 10.563 | 4.9523 |  |  | 967 | 38677.341 | 10.563 | 4.9558 |  |  | 967 | 38677.341 | 10.563 | 4.9505
968 | 38717.339 | 10.564 | 4.9522 |  |  | 968 | 38717.339 | 10.564 | 4.9557 |  |  | 968 | 38717.339 | 10.564 | 4.9504
969 | 38757.336 | 10.5651 | 4.9521 |  |  | 969 | 38757.336 | 10.5651 | 4.9556 |  |  | 969 | 38757.336 | 10.5651 | 4.9502
970 | 38797.333 | 10.5661 | 4.952 |  |  | 970 | 38797.333 | 10.5661 | 4.9555 |  |  | 970 | 38797.333 | 10.5661 | 4.9501
971 | 38837.33 | 10.5671 | 4.9519 |  |  | 971 | 38837.33 | 10.5671 | 4.9554 |  |  | 971 | 38837.33 | 10.5671 | 4.95
972 | 38877.328 | 10.5682 | 4.9518 |  |  | 972 | 38877.328 | 10.5682 | 4.9553 |  |  | 972 | 38877.328 | 10.5682 | 4.9499
973 | 38917.325 | 10.5692 | 4.9517 |  |  | 973 | 38917.325 | 10.5692 | 4.9552 |  |  | 973 | 38917.325 | 10.5692 | 4.9498
974 | 38957.322 | 10.5702 | 4.9516 |  |  | 974 | 38957.322 | 10.5702 | 4.9551 |  |  | 974 | 38957.322 | 10.5702 | 4.9497
975 | 38997.319 | 10.5712 | 4.9515 |  |  | 975 | 38997.319 | 10.5712 | 4.955 |  |  | 975 | 38997.319 | 10.5712 | 4.9495
976 | 39037.317 | 10.5723 | 4.9514 |  |  | 976 | 39037.317 | 10.5723 | 4.9549 |  |  | 976 | 39037.317 | 10.5723 | 4.9494
977 | 39077.314 | 10.5733 | 4.9513 |  |  | 977 | 39077.314 | 10.5733 | 4.9548 |  |  | 977 | 39077.314 | 10.5733 | 4.9493
978 | 39117.311 | 10.5743 | 4.9512 |  |  | 978 | 39117.311 | 10.5743 | 4.9547 |  |  | 978 | 39117.311 | 10.5743 | 4.9492
979 | 39157.308 | 10.5753 | 4.9511 |  |  | 979 | 39157.308 | 10.5753 | 4.9547 |  |  | 979 | 39157.308 | 10.5753 | 4.9491
980 | 39197.306 | 10.5764 | 4.951 |  |  | 980 | 39197.306 | 10.5764 | 4.9546 |  |  | 980 | 39197.306 | 10.5764 | 4.949
981 | 39237.303 | 10.5774 | 4.9509 |  |  | 981 | 39237.303 | 10.5774 | 4.9545 |  |  | 981 | 39237.303 | 10.5774 | 4.9488
982 | 39277.3 | 10.5784 | 4.9508 |  |  | 982 | 39277.3 | 10.5784 | 4.9544 |  |  | 982 | 39277.3 | 10.5784 | 4.9487
983 | 39317.297 | 10.5794 | 4.9507 |  |  | 983 | 39317.297 | 10.5794 | 4.9543 |  |  | 983 | 39317.297 | 10.5794 | 4.9486
984 | 39357.295 | 10.5804 | 4.9506 |  |  | 984 | 39357.295 | 10.5804 | 4.9542 |  |  | 984 | 39357.295 | 10.5804 | 4.9485
985 | 39397.292 | 10.5815 | 4.9505 |  |  | 985 | 39397.292 | 10.5815 | 4.9541 |  |  | 985 | 39397.292 | 10.5815 | 4.9484
986 | 39437.289 | 10.5825 | 4.9504 |  |  | 986 | 39437.289 | 10.5825 | 4.954 |  |  | 986 | 39437.289 | 10.5825 | 4.9483
987 | 39477.286 | 10.5835 | 4.9504 |  |  | 987 | 39477.286 | 10.5835 | 4.9539 |  |  | 987 | 39477.286 | 10.5835 | 4.9481
988 | 39517.284 | 10.5845 | 4.9503 |  |  | 988 | 39517.284 | 10.5845 | 4.9538 |  |  | 988 | 39517.284 | 10.5845 | 4.948
989 | 39557.281 | 10.5855 | 4.9502 |  |  | 989 | 39557.281 | 10.5855 | 4.9537 |  |  | 989 | 39557.281 | 10.5855 | 4.9479
990 | 39597.278 | 10.5865 | 4.9501 |  |  | 990 | 39597.278 | 10.5865 | 4.9536 |  |  | 990 | 39597.278 | 10.5865 | 4.9478

## Run10
MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR | [5 empty] | MOSAIQ 4.0 TSR
File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run10\1 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run10\2 01 001.RAW |  |  | File Name: |  |  | D:\cranfield\Alex Grenyer BAE_UQSS\Thermography2\Run10\3 01 001.RAW
Cursor Location: |  |  | (178, 193) |  |  | Cursor Location: |  |  | (179, 193) |  |  | Cursor Location: |  |  | (179, 193)
Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15) |  |  | Cursor Size: |  |  | (15 x 15)
Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev | Frame | Time ms | ln(Time) | Cursor1 | Mean | StdDev
1 | 39.997 | 3.6888 | 7.751 | 5.1348 | 0.32712 | 1 | 39.997 | 3.6888 | 7.7559 | 5.1337 | 0.32859 | 1 | 39.997 | 3.6888 | 7.759 | 5.1393 | 0.32719
2 | 79.995 | 4.382 | 7.4193 |  |  | 2 | 79.995 | 4.382 | 7.4219 |  |  | 2 | 79.995 | 4.382 | 7.4258
3 | 119.992 | 4.7874 | 7.2258 |  |  | 3 | 119.992 | 4.7874 | 7.2293 |  |  | 3 | 119.992 | 4.7874 | 7.2314
4 | 159.989 | 5.0751 | 7.0863 |  |  | 4 | 159.989 | 5.0751 | 7.0905 |  |  | 4 | 159.989 | 5.0751 | 7.0915
5 | 199.986 | 5.2982 | 6.9765 |  |  | 5 | 199.986 | 5.2982 | 6.981 |  |  | 5 | 199.986 | 5.2982 | 6.9816
6 | 239.984 | 5.4806 | 6.8858 |  |  | 6 | 239.984 | 5.4806 | 6.8905 |  |  | 6 | 239.984 | 5.4806 | 6.8909
7 | 279.981 | 5.6347 | 6.8085 |  |  | 7 | 279.981 | 5.6347 | 6.8132 |  |  | 7 | 279.981 | 5.6347 | 6.8136
8 | 319.978 | 5.7683 | 6.7411 |  |  | 8 | 319.978 | 5.7683 | 6.7458 |  |  | 8 | 319.978 | 5.7683 | 6.7463
9 | 359.975 | 5.886 | 6.6814 |  |  | 9 | 359.975 | 5.886 | 6.686 |  |  | 9 | 359.975 | 5.886 | 6.6867
10 | 399.973 | 5.9914 | 6.6278 |  |  | 10 | 399.973 | 5.9914 | 6.6323 |  |  | 10 | 399.973 | 5.9914 | 6.6331
11 | 439.97 | 6.0867 | 6.5792 |  |  | 11 | 439.97 | 6.0867 | 6.5835 |  |  | 11 | 439.97 | 6.0867 | 6.5846
12 | 479.967 | 6.1737 | 6.5347 |  |  | 12 | 479.967 | 6.1737 | 6.5389 |  |  | 12 | 479.967 | 6.1737 | 6.5402
13 | 519.964 | 6.2538 | 6.4938 |  |  | 13 | 519.964 | 6.2538 | 6.4978 |  |  | 13 | 519.964 | 6.2538 | 6.4992
14 | 559.962 | 6.3279 | 6.4558 |  |  | 14 | 559.962 | 6.3279 | 6.4597 |  |  | 14 | 559.962 | 6.3279 | 6.4613
15 | 599.959 | 6.3969 | 6.4204 |  |  | 15 | 599.959 | 6.3969 | 6.4242 |  |  | 15 | 599.959 | 6.3969 | 6.4259
16 | 639.956 | 6.4614 | 6.3873 |  |  | 16 | 639.956 | 6.4614 | 6.391 |  |  | 16 | 639.956 | 6.4614 | 6.3928
17 | 679.953 | 6.522 | 6.3562 |  |  | 17 | 679.953 | 6.522 | 6.3598 |  |  | 17 | 679.953 | 6.522 | 6.3617
18 | 719.951 | 6.5792 | 6.3268 |  |  | 18 | 719.951 | 6.5792 | 6.3304 |  |  | 18 | 719.951 | 6.5792 | 6.3324
19 | 759.948 | 6.6332 | 6.2991 |  |  | 19 | 759.948 | 6.6332 | 6.3025 |  |  | 19 | 759.948 | 6.6332 | 6.3046
20 | 799.945 | 6.6845 | 6.2727 |  |  | 20 | 799.945 | 6.6845 | 6.2761 |  |  | 20 | 799.945 | 6.6845 | 6.2783
21 | 839.942 | 6.7333 | 6.2477 |  |  | 21 | 839.942 | 6.7333 | 6.251 |  |  | 21 | 839.942 | 6.7333 | 6.2532
22 | 879.94 | 6.7799 | 6.2238 |  |  | 22 | 879.94 | 6.7799 | 6.227 |  |  | 22 | 879.94 | 6.7799 | 6.2293
23 | 919.937 | 6.8243 | 6.201 |  |  | 23 | 919.937 | 6.8243 | 6.2041 |  |  | 23 | 919.937 | 6.8243 | 6.2065
24 | 959.934 | 6.8669 | 6.1791 |  |  | 24 | 959.934 | 6.8669 | 6.1822 |  |  | 24 | 959.934 | 6.8669 | 6.1846
25 | 999.931 | 6.9077 | 6.1582 |  |  | 25 | 999.931 | 6.9077 | 6.1612 |  |  | 25 | 999.931 | 6.9077 | 6.1636
26 | 1039.929 | 6.9469 | 6.138 |  |  | 26 | 1039.929 | 6.9469 | 6.1411 |  |  | 26 | 1039.929 | 6.9469 | 6.1435
27 | 1079.926 | 6.9846 | 6.1187 |  |  | 27 | 1079.926 | 6.9846 | 6.1217 |  |  | 27 | 1079.926 | 6.9846 | 6.1241
28 | 1119.923 | 7.021 | 6.1 |  |  | 28 | 1119.923 | 7.021 | 6.103 |  |  | 28 | 1119.923 | 7.021 | 6.1054
29 | 1159.92 | 7.0561 | 6.0821 |  |  | 29 | 1159.92 | 7.0561 | 6.085 |  |  | 29 | 1159.92 | 7.0561 | 6.0874
30 | 1199.918 | 7.09 | 6.0647 |  |  | 30 | 1199.918 | 7.09 | 6.0676 |  |  | 30 | 1199.918 | 7.09 | 6.07
31 | 1239.915 | 7.1228 | 6.0479 |  |  | 31 | 1239.915 | 7.1228 | 6.0508 |  |  | 31 | 1239.915 | 7.1228 | 6.0532
32 | 1279.912 | 7.1545 | 6.0317 |  |  | 32 | 1279.912 | 7.1545 | 6.0345 |  |  | 32 | 1279.912 | 7.1545 | 6.0369
33 | 1319.909 | 7.1853 | 6.0159 |  |  | 33 | 1319.909 | 7.1853 | 6.0188 |  |  | 33 | 1319.909 | 7.1853 | 6.0212
34 | 1359.907 | 7.2152 | 6.0007 |  |  | 34 | 1359.907 | 7.2152 | 6.0035 |  |  | 34 | 1359.907 | 7.2152 | 6.0059
35 | 1399.904 | 7.2442 | 5.9859 |  |  | 35 | 1399.904 | 7.2442 | 5.9887 |  |  | 35 | 1399.904 | 7.2442 | 5.991
36 | 1439.901 | 7.2723 | 5.9715 |  |  | 36 | 1439.901 | 7.2723 | 5.9743 |  |  | 36 | 1439.901 | 7.2723 | 5.9766
37 | 1479.898 | 7.2997 | 5.9575 |  |  | 37 | 1479.898 | 7.2997 | 5.9603 |  |  | 37 | 1479.898 | 7.2997 | 5.9626
38 | 1519.896 | 7.3264 | 5.9439 |  |  | 38 | 1519.896 | 7.3264 | 5.9467 |  |  | 38 | 1519.896 | 7.3264 | 5.949
39 | 1559.893 | 7.3524 | 5.9307 |  |  | 39 | 1559.893 | 7.3524 | 5.9334 |  |  | 39 | 1559.893 | 7.3524 | 5.9358
40 | 1599.89 | 7.3777 | 5.9178 |  |  | 40 | 1599.89 | 7.3777 | 5.9205 |  |  | 40 | 1599.89 | 7.3777 | 5.9228
41 | 1639.887 | 7.4024 | 5.9052 |  |  | 41 | 1639.887 | 7.4024 | 5.908 |  |  | 41 | 1639.887 | 7.4024 | 5.9103
42 | 1679.885 | 7.4265 | 5.893 |  |  | 42 | 1679.885 | 7.4265 | 5.8957 |  |  | 42 | 1679.885 | 7.4265 | 5.898
43 | 1719.882 | 7.45 | 5.881 |  |  | 43 | 1719.882 | 7.45 | 5.8838 |  |  | 43 | 1719.882 | 7.45 | 5.886
44 | 1759.879 | 7.473 | 5.8694 |  |  | 44 | 1759.879 | 7.473 | 5.8721 |  |  | 44 | 1759.879 | 7.473 | 5.8743
45 | 1799.876 | 7.4955 | 5.858 |  |  | 45 | 1799.876 | 7.4955 | 5.8607 |  |  | 45 | 1799.876 | 7.4955 | 5.8629
46 | 1839.874 | 7.5175 | 5.8469 |  |  | 46 | 1839.874 | 7.5175 | 5.8496 |  |  | 46 | 1839.874 | 7.5175 | 5.8518
47 | 1879.871 | 7.539 | 5.836 |  |  | 47 | 1879.871 | 7.539 | 5.8387 |  |  | 47 | 1879.871 | 7.539 | 5.8409
48 | 1919.868 | 7.56 | 5.8254 |  |  | 48 | 1919.868 | 7.56 | 5.828 |  |  | 48 | 1919.868 | 7.56 | 5.8302
49 | 1959.865 | 7.5806 | 5.8149 |  |  | 49 | 1959.865 | 7.5806 | 5.8176 |  |  | 49 | 1959.865 | 7.5806 | 5.8198
50 | 1999.863 | 7.6008 | 5.8048 |  |  | 50 | 1999.863 | 7.6008 | 5.8074 |  |  | 50 | 1999.863 | 7.6008 | 5.8096
51 | 2039.86 | 7.6206 | 5.7948 |  |  | 51 | 2039.86 | 7.6206 | 5.7974 |  |  | 51 | 2039.86 | 7.6206 | 5.7996
52 | 2079.857 | 7.6401 | 5.785 |  |  | 52 | 2079.857 | 7.6401 | 5.7877 |  |  | 52 | 2079.857 | 7.6401 | 5.7898
53 | 2119.854 | 7.6591 | 5.7755 |  |  | 53 | 2119.854 | 7.6591 | 5.7781 |  |  | 53 | 2119.854 | 7.6591 | 5.7802
54 | 2159.852 | 7.6778 | 5.7661 |  |  | 54 | 2159.852 | 7.6778 | 5.7687 |  |  | 54 | 2159.852 | 7.6778 | 5.7708
55 | 2199.849 | 7.6961 | 5.7569 |  |  | 55 | 2199.849 | 7.6961 | 5.7595 |  |  | 55 | 2199.849 | 7.6961 | 5.7616
56 | 2239.846 | 7.7142 | 5.7479 |  |  | 56 | 2239.846 | 7.7142 | 5.7505 |  |  | 56 | 2239.846 | 7.7142 | 5.7526
57 | 2279.843 | 7.7319 | 5.739 |  |  | 57 | 2279.843 | 7.7319 | 5.7416 |  |  | 57 | 2279.843 | 7.7319 | 5.7437
58 | 2319.841 | 7.7493 | 5.7303 |  |  | 58 | 2319.841 | 7.7493 | 5.733 |  |  | 58 | 2319.841 | 7.7493 | 5.735
59 | 2359.838 | 7.7663 | 5.7218 |  |  | 59 | 2359.838 | 7.7663 | 5.7244 |  |  | 59 | 2359.838 | 7.7663 | 5.7265
60 | 2399.835 | 7.7832 | 5.7135 |  |  | 60 | 2399.835 | 7.7832 | 5.7161 |  |  | 60 | 2399.835 | 7.7832 | 5.7181
61 | 2439.832 | 7.7997 | 5.7053 |  |  | 61 | 2439.832 | 7.7997 | 5.7078 |  |  | 61 | 2439.832 | 7.7997 | 5.7099
62 | 2479.83 | 7.8159 | 5.6972 |  |  | 62 | 2479.83 | 7.8159 | 5.6998 |  |  | 62 | 2479.83 | 7.8159 | 5.7018
63 | 2519.827 | 7.8319 | 5.6893 |  |  | 63 | 2519.827 | 7.8319 | 5.6918 |  |  | 63 | 2519.827 | 7.8319 | 5.6939
64 | 2559.824 | 7.8477 | 5.6815 |  |  | 64 | 2559.824 | 7.8477 | 5.684 |  |  | 64 | 2559.824 | 7.8477 | 5.6861
65 | 2599.821 | 7.8632 | 5.6738 |  |  | 65 | 2599.821 | 7.8632 | 5.6764 |  |  | 65 | 2599.821 | 7.8632 | 5.6784
66 | 2639.819 | 7.8785 | 5.6663 |  |  | 66 | 2639.819 | 7.8785 | 5.6688 |  |  | 66 | 2639.819 | 7.8785 | 5.6708
67 | 2679.816 | 7.8935 | 5.6589 |  |  | 67 | 2679.816 | 7.8935 | 5.6614 |  |  | 67 | 2679.816 | 7.8935 | 5.6634
68 | 2719.813 | 7.9083 | 5.6516 |  |  | 68 | 2719.813 | 7.9083 | 5.6541 |  |  | 68 | 2719.813 | 7.9083 | 5.6562
69 | 2759.81 | 7.9229 | 5.6445 |  |  | 69 | 2759.81 | 7.9229 | 5.647 |  |  | 69 | 2759.81 | 7.9229 | 5.649
70 | 2799.808 | 7.9373 | 5.6374 |  |  | 70 | 2799.808 | 7.9373 | 5.6399 |  |  | 70 | 2799.808 | 7.9373 | 5.6419
71 | 2839.805 | 7.9515 | 5.6305 |  |  | 71 | 2839.805 | 7.9515 | 5.633 |  |  | 71 | 2839.805 | 7.9515 | 5.635
72 | 2879.802 | 7.9655 | 5.6237 |  |  | 72 | 2879.802 | 7.9655 | 5.6261 |  |  | 72 | 2879.802 | 7.9655 | 5.6282
73 | 2919.799 | 7.9793 | 5.617 |  |  | 73 | 2919.799 | 7.9793 | 5.6194 |  |  | 73 | 2919.799 | 7.9793 | 5.6214
74 | 2959.797 | 7.9929 | 5.6103 |  |  | 74 | 2959.797 | 7.9929 | 5.6128 |  |  | 74 | 2959.797 | 7.9929 | 5.6148
75 | 2999.794 | 8.0063 | 5.6038 |  |  | 75 | 2999.794 | 8.0063 | 5.6063 |  |  | 75 | 2999.794 | 8.0063 | 5.6083
76 | 3039.791 | 8.0195 | 5.5974 |  |  | 76 | 3039.791 | 8.0195 | 5.5998 |  |  | 76 | 3039.791 | 8.0195 | 5.6019
77 | 3079.788 | 8.0326 | 5.5911 |  |  | 77 | 3079.788 | 8.0326 | 5.5935 |  |  | 77 | 3079.788 | 8.0326 | 5.5956
78 | 3119.786 | 8.0455 | 5.5849 |  |  | 78 | 3119.786 | 8.0455 | 5.5873 |  |  | 78 | 3119.786 | 8.0455 | 5.5893
79 | 3159.783 | 8.0583 | 5.5788 |  |  | 79 | 3159.783 | 8.0583 | 5.5811 |  |  | 79 | 3159.783 | 8.0583 | 5.5832
80 | 3199.78 | 8.0708 | 5.5727 |  |  | 80 | 3199.78 | 8.0708 | 5.5751 |  |  | 80 | 3199.78 | 8.0708 | 5.5772
81 | 3239.777 | 8.0833 | 5.5668 |  |  | 81 | 3239.777 | 8.0833 | 5.5691 |  |  | 81 | 3239.777 | 8.0833 | 5.5712
82 | 3279.775 | 8.0955 | 5.5609 |  |  | 82 | 3279.775 | 8.0955 | 5.5633 |  |  | 82 | 3279.775 | 8.0955 | 5.5653
83 | 3319.772 | 8.1077 | 5.5551 |  |  | 83 | 3319.772 | 8.1077 | 5.5575 |  |  | 83 | 3319.772 | 8.1077 | 5.5595
84 | 3359.769 | 8.1196 | 5.5494 |  |  | 84 | 3359.769 | 8.1196 | 5.5517 |  |  | 84 | 3359.769 | 8.1196 | 5.5538
85 | 3399.766 | 8.1315 | 5.5438 |  |  | 85 | 3399.766 | 8.1315 | 5.5461 |  |  | 85 | 3399.766 | 8.1315 | 5.5482
86 | 3439.764 | 8.1432 | 5.5383 |  |  | 86 | 3439.764 | 8.1432 | 5.5406 |  |  | 86 | 3439.764 | 8.1432 | 5.5427
87 | 3479.761 | 8.1547 | 5.5328 |  |  | 87 | 3479.761 | 8.1547 | 5.5351 |  |  | 87 | 3479.761 | 8.1547 | 5.5372
88 | 3519.758 | 8.1661 | 5.5274 |  |  | 88 | 3519.758 | 8.1661 | 5.5297 |  |  | 88 | 3519.758 | 8.1661 | 5.5318
89 | 3559.755 | 8.1774 | 5.5221 |  |  | 89 | 3559.755 | 8.1774 | 5.5243 |  |  | 89 | 3559.755 | 8.1774 | 5.5265
90 | 3599.753 | 8.1886 | 5.5169 |  |  | 90 | 3599.753 | 8.1886 | 5.5191 |  |  | 90 | 3599.753 | 8.1886 | 5.5212
91 | 3639.75 | 8.1997 | 5.5117 |  |  | 91 | 3639.75 | 8.1997 | 5.5139 |  |  | 91 | 3639.75 | 8.1997 | 5.5161
92 | 3679.747 | 8.2106 | 5.5066 |  |  | 92 | 3679.747 | 8.2106 | 5.5088 |  |  | 92 | 3679.747 | 8.2106 | 5.511
93 | 3719.744 | 8.2214 | 5.5016 |  |  | 93 | 3719.744 | 8.2214 | 5.5037 |  |  | 93 | 3719.744 | 8.2214 | 5.5059
94 | 3759.742 | 8.2321 | 5.4966 |  |  | 94 | 3759.742 | 8.2321 | 5.4987 |  |  | 94 | 3759.742 | 8.2321 | 5.501
95 | 3799.739 | 8.2427 | 5.4917 |  |  | 95 | 3799.739 | 8.2427 | 5.4938 |  |  | 95 | 3799.739 | 8.2427 | 5.496
96 | 3839.736 | 8.2532 | 5.4868 |  |  | 96 | 3839.736 | 8.2532 | 5.4889 |  |  | 96 | 3839.736 | 8.2532 | 5.4912
97 | 3879.733 | 8.2635 | 5.4821 |  |  | 97 | 3879.733 | 8.2635 | 5.4841 |  |  | 97 | 3879.733 | 8.2635 | 5.4864
98 | 3919.731 | 8.2738 | 5.4773 |  |  | 98 | 3919.731 | 8.2738 | 5.4794 |  |  | 98 | 3919.731 | 8.2738 | 5.4817
99 | 3959.728 | 8.2839 | 5.4727 |  |  | 99 | 3959.728 | 8.2839 | 5.4747 |  |  | 99 | 3959.728 | 8.2839 | 5.477
100 | 3999.725 | 8.294 | 5.4681 |  |  | 100 | 3999.725 | 8.294 | 5.4701 |  |  | 100 | 3999.725 | 8.294 | 5.4724
101 | 4039.722 | 8.3039 | 5.4636 |  |  | 101 | 4039.722 | 8.3039 | 5.4655 |  |  | 101 | 4039.722 | 8.3039 | 5.4679
102 | 4079.72 | 8.3138 | 5.4591 |  |  | 102 | 4079.72 | 8.3138 | 5.461 |  |  | 102 | 4079.72 | 8.3138 | 5.4634
103 | 4119.717 | 8.3235 | 5.4546 |  |  | 103 | 4119.717 | 8.3235 | 5.4566 |  |  | 103 | 4119.717 | 8.3235 | 5.459
104 | 4159.714 | 8.3332 | 5.4503 |  |  | 104 | 4159.714 | 8.3332 | 5.4522 |  |  | 104 | 4159.714 | 8.3332 | 5.4546
105 | 4199.711 | 8.3428 | 5.4459 |  |  | 105 | 4199.711 | 8.3428 | 5.4478 |  |  | 105 | 4199.711 | 8.3428 | 5.4503
106 | 4239.709 | 8.3522 | 5.4417 |  |  | 106 | 4239.709 | 8.3522 | 5.4436 |  |  | 106 | 4239.709 | 8.3522 | 5.446
107 | 4279.706 | 8.3616 | 5.4375 |  |  | 107 | 4279.706 | 8.3616 | 5.4393 |  |  | 107 | 4279.706 | 8.3616 | 5.4418
108 | 4319.703 | 8.3709 | 5.4333 |  |  | 108 | 4319.703 | 8.3709 | 5.4351 |  |  | 108 | 4319.703 | 8.3709 | 5.4376
109 | 4359.7 | 8.3802 | 5.4292 |  |  | 109 | 4359.7 | 8.3802 | 5.431 |  |  | 109 | 4359.7 | 8.3802 | 5.4335
110 | 4399.698 | 8.3893 | 5.4251 |  |  | 110 | 4399.698 | 8.3893 | 5.4269 |  |  | 110 | 4399.698 | 8.3893 | 5.4294
111 | 4439.695 | 8.3983 | 5.4211 |  |  | 111 | 4439.695 | 8.3983 | 5.4229 |  |  | 111 | 4439.695 | 8.3983 | 5.4254
112 | 4479.692 | 8.4073 | 5.4171 |  |  | 112 | 4479.692 | 8.4073 | 5.4189 |  |  | 112 | 4479.692 | 8.4073 | 5.4215
113 | 4519.689 | 8.4162 | 5.4132 |  |  | 113 | 4519.689 | 8.4162 | 5.4149 |  |  | 113 | 4519.689 | 8.4162 | 5.4175
114 | 4559.687 | 8.425 | 5.4093 |  |  | 114 | 4559.687 | 8.425 | 5.411 |  |  | 114 | 4559.687 | 8.425 | 5.4137
115 | 4599.684 | 8.4337 | 5.4055 |  |  | 115 | 4599.684 | 8.4337 | 5.4072 |  |  | 115 | 4599.684 | 8.4337 | 5.4098
116 | 4639.681 | 8.4424 | 5.4017 |  |  | 116 | 4639.681 | 8.4424 | 5.4034 |  |  | 116 | 4639.681 | 8.4424 | 5.406
117 | 4679.678 | 8.451 | 5.398 |  |  | 117 | 4679.678 | 8.451 | 5.3996 |  |  | 117 | 4679.678 | 8.451 | 5.4023
118 | 4719.676 | 8.4595 | 5.3943 |  |  | 118 | 4719.676 | 8.4595 | 5.3959 |  |  | 118 | 4719.676 | 8.4595 | 5.3986
119 | 4759.673 | 8.4679 | 5.3906 |  |  | 119 | 4759.673 | 8.4679 | 5.3922 |  |  | 119 | 4759.673 | 8.4679 | 5.3949
120 | 4799.67 | 8.4763 | 5.387 |  |  | 120 | 4799.67 | 8.4763 | 5.3886 |  |  | 120 | 4799.67 | 8.4763 | 5.3913
121 | 4839.667 | 8.4846 | 5.3834 |  |  | 121 | 4839.667 | 8.4846 | 5.385 |  |  | 121 | 4839.667 | 8.4846 | 5.3877
122 | 4879.665 | 8.4928 | 5.3799 |  |  | 122 | 4879.665 | 8.4928 | 5.3814 |  |  | 122 | 4879.665 | 8.4928 | 5.3842
123 | 4919.662 | 8.501 | 5.3764 |  |  | 123 | 4919.662 | 8.501 | 5.3779 |  |  | 123 | 4919.662 | 8.501 | 5.3807
124 | 4959.659 | 8.5091 | 5.3729 |  |  | 124 | 4959.659 | 8.5091 | 5.3744 |  |  | 124 | 4959.659 | 8.5091 | 5.3773
125 | 4999.656 | 8.5171 | 5.3695 |  |  | 125 | 4999.656 | 8.5171 | 5.371 |  |  | 125 | 4999.656 | 8.5171 | 5.3738
126 | 5039.654 | 8.5251 | 5.3661 |  |  | 126 | 5039.654 | 8.5251 | 5.3676 |  |  | 126 | 5039.654 | 8.5251 | 5.3705
127 | 5079.651 | 8.533 | 5.3628 |  |  | 127 | 5079.651 | 8.533 | 5.3642 |  |  | 127 | 5079.651 | 8.533 | 5.3671
128 | 5119.648 | 8.5408 | 5.3595 |  |  | 128 | 5119.648 | 8.5408 | 5.3609 |  |  | 128 | 5119.648 | 8.5408 | 5.3638
129 | 5159.645 | 8.5486 | 5.3562 |  |  | 129 | 5159.645 | 8.5486 | 5.3576 |  |  | 129 | 5159.645 | 8.5486 | 5.3605
130 | 5199.643 | 8.5563 | 5.353 |  |  | 130 | 5199.643 | 8.5563 | 5.3543 |  |  | 130 | 5199.643 | 8.5563 | 5.3573
131 | 5239.64 | 8.564 | 5.3498 |  |  | 131 | 5239.64 | 8.564 | 5.3511 |  |  | 131 | 5239.64 | 8.564 | 5.3541
132 | 5279.637 | 8.5716 | 5.3466 |  |  | 132 | 5279.637 | 8.5716 | 5.3479 |  |  | 132 | 5279.637 | 8.5716 | 5.351
133 | 5319.634 | 8.5792 | 5.3435 |  |  | 133 | 5319.634 | 8.5792 | 5.3448 |  |  | 133 | 5319.634 | 8.5792 | 5.3478
134 | 5359.632 | 8.5867 | 5.3404 |  |  | 134 | 5359.632 | 8.5867 | 5.3417 |  |  | 134 | 5359.632 | 8.5867 | 5.3447
135 | 5399.629 | 8.5941 | 5.3374 |  |  | 135 | 5399.629 | 8.5941 | 5.3386 |  |  | 135 | 5399.629 | 8.5941 | 5.3417
136 | 5439.626 | 8.6015 | 5.3343 |  |  | 136 | 5439.626 | 8.6015 | 5.3355 |  |  | 136 | 5439.626 | 8.6015 | 5.3386
137 | 5479.623 | 8.6088 | 5.3313 |  |  | 137 | 5479.623 | 8.6088 | 5.3325 |  |  | 137 | 5479.623 | 8.6088 | 5.3356
138 | 5519.621 | 8.6161 | 5.3284 |  |  | 138 | 5519.621 | 8.6161 | 5.3295 |  |  | 138 | 5519.621 | 8.6161 | 5.3327
139 | 5559.618 | 8.6233 | 5.3254 |  |  | 139 | 5559.618 | 8.6233 | 5.3265 |  |  | 139 | 5559.618 | 8.6233 | 5.3297
140 | 5599.615 | 8.6305 | 5.3225 |  |  | 140 | 5599.615 | 8.6305 | 5.3236 |  |  | 140 | 5599.615 | 8.6305 | 5.3268
141 | 5639.612 | 8.6376 | 5.3197 |  |  | 141 | 5639.612 | 8.6376 | 5.3207 |  |  | 141 | 5639.612 | 8.6376 | 5.324
142 | 5679.61 | 8.6446 | 5.3168 |  |  | 142 | 5679.61 | 8.6446 | 5.3179 |  |  | 142 | 5679.61 | 8.6446 | 5.3211
143 | 5719.607 | 8.6517 | 5.314 |  |  | 143 | 5719.607 | 8.6517 | 5.315 |  |  | 143 | 5719.607 | 8.6517 | 5.3183
144 | 5759.604 | 8.6586 | 5.3112 |  |  | 144 | 5759.604 | 8.6586 | 5.3122 |  |  | 144 | 5759.604 | 8.6586 | 5.3155
145 | 5799.601 | 8.6655 | 5.3085 |  |  | 145 | 5799.601 | 8.6655 | 5.3094 |  |  | 145 | 5799.601 | 8.6655 | 5.3128
146 | 5839.599 | 8.6724 | 5.3057 |  |  | 146 | 5839.599 | 8.6724 | 5.3067 |  |  | 146 | 5839.599 | 8.6724 | 5.31
147 | 5879.596 | 8.6792 | 5.303 |  |  | 147 | 5879.596 | 8.6792 | 5.304 |  |  | 147 | 5879.596 | 8.6792 | 5.3073
148 | 5919.593 | 8.686 | 5.3004 |  |  | 148 | 5919.593 | 8.686 | 5.3013 |  |  | 148 | 5919.593 | 8.686 | 5.3047
149 | 5959.59 | 8.6928 | 5.2977 |  |  | 149 | 5959.59 | 8.6928 | 5.2986 |  |  | 149 | 5959.59 | 8.6928 | 5.302
150 | 5999.588 | 8.6994 | 5.2951 |  |  | 150 | 5999.588 | 8.6994 | 5.296 |  |  | 150 | 5999.588 | 8.6994 | 5.2994
151 | 6039.585 | 8.7061 | 5.2925 |  |  | 151 | 6039.585 | 8.7061 | 5.2934 |  |  | 151 | 6039.585 | 8.7061 | 5.2968
152 | 6079.582 | 8.7127 | 5.29 |  |  | 152 | 6079.582 | 8.7127 | 5.2908 |  |  | 152 | 6079.582 | 8.7127 | 5.2943
153 | 6119.579 | 8.7192 | 5.2874 |  |  | 153 | 6119.579 | 8.7192 | 5.2882 |  |  | 153 | 6119.579 | 8.7192 | 5.2917
154 | 6159.577 | 8.7258 | 5.2849 |  |  | 154 | 6159.577 | 8.7258 | 5.2857 |  |  | 154 | 6159.577 | 8.7258 | 5.2892
155 | 6199.574 | 8.7322 | 5.2824 |  |  | 155 | 6199.574 | 8.7322 | 5.2832 |  |  | 155 | 6199.574 | 8.7322 | 5.2867
156 | 6239.571 | 8.7387 | 5.28 |  |  | 156 | 6239.571 | 8.7387 | 5.2807 |  |  | 156 | 6239.571 | 8.7387 | 5.2843
157 | 6279.568 | 8.7451 | 5.2775 |  |  | 157 | 6279.568 | 8.7451 | 5.2782 |  |  | 157 | 6279.568 | 8.7451 | 5.2818
158 | 6319.566 | 8.7514 | 5.2751 |  |  | 158 | 6319.566 | 8.7514 | 5.2758 |  |  | 158 | 6319.566 | 8.7514 | 5.2794
159 | 6359.563 | 8.7577 | 5.2727 |  |  | 159 | 6359.563 | 8.7577 | 5.2734 |  |  | 159 | 6359.563 | 8.7577 | 5.277
160 | 6399.56 | 8.764 | 5.2704 |  |  | 160 | 6399.56 | 8.764 | 5.271 |  |  | 160 | 6399.56 | 8.764 | 5.2747
161 | 6439.557 | 8.7702 | 5.268 |  |  | 161 | 6439.557 | 8.7702 | 5.2686 |  |  | 161 | 6439.557 | 8.7702 | 5.2723
162 | 6479.555 | 8.7764 | 5.2657 |  |  | 162 | 6479.555 | 8.7764 | 5.2663 |  |  | 162 | 6479.555 | 8.7764 | 5.27
163 | 6519.552 | 8.7826 | 5.2634 |  |  | 163 | 6519.552 | 8.7826 | 5.264 |  |  | 163 | 6519.552 | 8.7826 | 5.2677
164 | 6559.549 | 8.7887 | 5.2612 |  |  | 164 | 6559.549 | 8.7887 | 5.2617 |  |  | 164 | 6559.549 | 8.7887 | 5.2654
165 | 6599.546 | 8.7948 | 5.2589 |  |  | 165 | 6599.546 | 8.7948 | 5.2594 |  |  | 165 | 6599.546 | 8.7948 | 5.2632
166 | 6639.544 | 8.8008 | 5.2567 |  |  | 166 | 6639.544 | 8.8008 | 5.2572 |  |  | 166 | 6639.544 | 8.8008 | 5.261
167 | 6679.541 | 8.8068 | 5.2545 |  |  | 167 | 6679.541 | 8.8068 | 5.2549 |  |  | 167 | 6679.541 | 8.8068 | 5.2588
168 | 6719.538 | 8.8128 | 5.2523 |  |  | 168 | 6719.538 | 8.8128 | 5.2527 |  |  | 168 | 6719.538 | 8.8128 | 5.2566
169 | 6759.535 | 8.8187 | 5.2501 |  |  | 169 | 6759.535 | 8.8187 | 5.2505 |  |  | 169 | 6759.535 | 8.8187 | 5.2544
170 | 6799.533 | 8.8246 | 5.248 |  |  | 170 | 6799.533 | 8.8246 | 5.2484 |  |  | 170 | 6799.533 | 8.8246 | 5.2523
171 | 6839.53 | 8.8305 | 5.2459 |  |  | 171 | 6839.53 | 8.8305 | 5.2462 |  |  | 171 | 6839.53 | 8.8305 | 5.2501
172 | 6879.527 | 8.8363 | 5.2438 |  |  | 172 | 6879.527 | 8.8363 | 5.2441 |  |  | 172 | 6879.527 | 8.8363 | 5.248
173 | 6919.524 | 8.8421 | 5.2417 |  |  | 173 | 6919.524 | 8.8421 | 5.242 |  |  | 173 | 6919.524 | 8.8421 | 5.246
174 | 6959.522 | 8.8479 | 5.2396 |  |  | 174 | 6959.522 | 8.8479 | 5.2399 |  |  | 174 | 6959.522 | 8.8479 | 5.2439
175 | 6999.519 | 8.8536 | 5.2376 |  |  | 175 | 6999.519 | 8.8536 | 5.2379 |  |  | 175 | 6999.519 | 8.8536 | 5.2418
176 | 7039.516 | 8.8593 | 5.2356 |  |  | 176 | 7039.516 | 8.8593 | 5.2358 |  |  | 176 | 7039.516 | 8.8593 | 5.2398
177 | 7079.513 | 8.865 | 5.2336 |  |  | 177 | 7079.513 | 8.865 | 5.2338 |  |  | 177 | 7079.513 | 8.865 | 5.2378
178 | 7119.511 | 8.8706 | 5.2316 |  |  | 178 | 7119.511 | 8.8706 | 5.2318 |  |  | 178 | 7119.511 | 8.8706 | 5.2358
179 | 7159.508 | 8.8762 | 5.2296 |  |  | 179 | 7159.508 | 8.8762 | 5.2298 |  |  | 179 | 7159.508 | 8.8762 | 5.2339
180 | 7199.505 | 8.8818 | 5.2277 |  |  | 180 | 7199.505 | 8.8818 | 5.2278 |  |  | 180 | 7199.505 | 8.8818 | 5.2319
181 | 7239.502 | 8.8873 | 5.2257 |  |  | 181 | 7239.502 | 8.8873 | 5.2259 |  |  | 181 | 7239.502 | 8.8873 | 5.23
182 | 7279.5 | 8.8928 | 5.2238 |  |  | 182 | 7279.5 | 8.8928 | 5.2239 |  |  | 182 | 7279.5 | 8.8928 | 5.2281
183 | 7319.497 | 8.8983 | 5.2219 |  |  | 183 | 7319.497 | 8.8983 | 5.222 |  |  | 183 | 7319.497 | 8.8983 | 5.2262
184 | 7359.494 | 8.9037 | 5.2201 |  |  | 184 | 7359.494 | 8.9037 | 5.2201 |  |  | 184 | 7359.494 | 8.9037 | 5.2243
185 | 7399.491 | 8.9092 | 5.2182 |  |  | 185 | 7399.491 | 8.9092 | 5.2183 |  |  | 185 | 7399.491 | 8.9092 | 5.2224
186 | 7439.489 | 8.9146 | 5.2164 |  |  | 186 | 7439.489 | 8.9146 | 5.2164 |  |  | 186 | 7439.489 | 8.9146 | 5.2206
187 | 7479.486 | 8.9199 | 5.2145 |  |  | 187 | 7479.486 | 8.9199 | 5.2146 |  |  | 187 | 7479.486 | 8.9199 | 5.2188
188 | 7519.483 | 8.9253 | 5.2127 |  |  | 188 | 7519.483 | 8.9253 | 5.2127 |  |  | 188 | 7519.483 | 8.9253 | 5.217
189 | 7559.48 | 8.9306 | 5.2109 |  |  | 189 | 7559.48 | 8.9306 | 5.2109 |  |  | 189 | 7559.48 | 8.9306 | 5.2152
190 | 7599.478 | 8.9358 | 5.2092 |  |  | 190 | 7599.478 | 8.9358 | 5.2091 |  |  | 190 | 7599.478 | 8.9358 | 5.2134
191 | 7639.475 | 8.9411 | 5.2074 |  |  | 191 | 7639.475 | 8.9411 | 5.2073 |  |  | 191 | 7639.475 | 8.9411 | 5.2117
192 | 7679.472 | 8.9463 | 5.2057 |  |  | 192 | 7679.472 | 8.9463 | 5.2056 |  |  | 192 | 7679.472 | 8.9463 | 5.2099
193 | 7719.469 | 8.9515 | 5.2039 |  |  | 193 | 7719.469 | 8.9515 | 5.2038 |  |  | 193 | 7719.469 | 8.9515 | 5.2082
194 | 7759.467 | 8.9567 | 5.2022 |  |  | 194 | 7759.467 | 8.9567 | 5.2021 |  |  | 194 | 7759.467 | 8.9567 | 5.2065
195 | 7799.464 | 8.9618 | 5.2005 |  |  | 195 | 7799.464 | 8.9618 | 5.2004 |  |  | 195 | 7799.464 | 8.9618 | 5.2048
196 | 7839.461 | 8.9669 | 5.1989 |  |  | 196 | 7839.461 | 8.9669 | 5.1987 |  |  | 196 | 7839.461 | 8.9669 | 5.2031
197 | 7879.458 | 8.972 | 5.1972 |  |  | 197 | 7879.458 | 8.972 | 5.197 |  |  | 197 | 7879.458 | 8.972 | 5.2014
198 | 7919.456 | 8.9771 | 5.1956 |  |  | 198 | 7919.456 | 8.9771 | 5.1954 |  |  | 198 | 7919.456 | 8.9771 | 5.1998
199 | 7959.453 | 8.9821 | 5.1939 |  |  | 199 | 7959.453 | 8.9821 | 5.1937 |  |  | 199 | 7959.453 | 8.9821 | 5.1982
200 | 7999.45 | 8.9871 | 5.1923 |  |  | 200 | 7999.45 | 8.9871 | 5.1921 |  |  | 200 | 7999.45 | 8.9871 | 5.1965
201 | 8039.447 | 8.9921 | 5.1907 |  |  | 201 | 8039.447 | 8.9921 | 5.1905 |  |  | 201 | 8039.447 | 8.9921 | 5.1949
202 | 8079.445 | 8.9971 | 5.1891 |  |  | 202 | 8079.445 | 8.9971 | 5.1888 |  |  | 202 | 8079.445 | 8.9971 | 5.1933
203 | 8119.442 | 9.002 | 5.1876 |  |  | 203 | 8119.442 | 9.002 | 5.1873 |  |  | 203 | 8119.442 | 9.002 | 5.1918
204 | 8159.439 | 9.0069 | 5.186 |  |  | 204 | 8159.439 | 9.0069 | 5.1857 |  |  | 204 | 8159.439 | 9.0069 | 5.1902
205 | 8199.436 | 9.0118 | 5.1844 |  |  | 205 | 8199.436 | 9.0118 | 5.1841 |  |  | 205 | 8199.436 | 9.0118 | 5.1886
206 | 8239.434 | 9.0167 | 5.1829 |  |  | 206 | 8239.434 | 9.0167 | 5.1826 |  |  | 206 | 8239.434 | 9.0167 | 5.1871
207 | 8279.431 | 9.0215 | 5.1814 |  |  | 207 | 8279.431 | 9.0215 | 5.181 |  |  | 207 | 8279.431 | 9.0215 | 5.1856
208 | 8319.428 | 9.0263 | 5.1799 |  |  | 208 | 8319.428 | 9.0263 | 5.1795 |  |  | 208 | 8319.428 | 9.0263 | 5.1841
209 | 8359.425 | 9.0311 | 5.1784 |  |  | 209 | 8359.425 | 9.0311 | 5.178 |  |  | 209 | 8359.425 | 9.0311 | 5.1826
210 | 8399.423 | 9.0359 | 5.1769 |  |  | 210 | 8399.423 | 9.0359 | 5.1765 |  |  | 210 | 8399.423 | 9.0359 | 5.1811
211 | 8439.42 | 9.0407 | 5.1755 |  |  | 211 | 8439.42 | 9.0407 | 5.175 |  |  | 211 | 8439.42 | 9.0407 | 5.1796
212 | 8479.417 | 9.0454 | 5.174 |  |  | 212 | 8479.417 | 9.0454 | 5.1735 |  |  | 212 | 8479.417 | 9.0454 | 5.1782
213 | 8519.414 | 9.0501 | 5.1726 |  |  | 213 | 8519.414 | 9.0501 | 5.1721 |  |  | 213 | 8519.414 | 9.0501 | 5.1767
214 | 8559.412 | 9.0548 | 5.1711 |  |  | 214 | 8559.412 | 9.0548 | 5.1706 |  |  | 214 | 8559.412 | 9.0548 | 5.1753
215 | 8599.409 | 9.0594 | 5.1697 |  |  | 215 | 8599.409 | 9.0594 | 5.1692 |  |  | 215 | 8599.409 | 9.0594 | 5.1739
216 | 8639.406 | 9.0641 | 5.1683 |  |  | 216 | 8639.406 | 9.0641 | 5.1678 |  |  | 216 | 8639.406 | 9.0641 | 5.1725
217 | 8679.403 | 9.0687 | 5.1669 |  |  | 217 | 8679.403 | 9.0687 | 5.1664 |  |  | 217 | 8679.403 | 9.0687 | 5.1711
218 | 8719.401 | 9.0733 | 5.1656 |  |  | 218 | 8719.401 | 9.0733 | 5.165 |  |  | 218 | 8719.401 | 9.0733 | 5.1697
219 | 8759.398 | 9.0779 | 5.1642 |  |  | 219 | 8759.398 | 9.0779 | 5.1636 |  |  | 219 | 8759.398 | 9.0779 | 5.1684
220 | 8799.395 | 9.0824 | 5.1628 |  |  | 220 | 8799.395 | 9.0824 | 5.1622 |  |  | 220 | 8799.395 | 9.0824 | 5.167
221 | 8839.392 | 9.087 | 5.1615 |  |  | 221 | 8839.392 | 9.087 | 5.1609 |  |  | 221 | 8839.392 | 9.087 | 5.1657
222 | 8879.39 | 9.0915 | 5.1602 |  |  | 222 | 8879.39 | 9.0915 | 5.1595 |  |  | 222 | 8879.39 | 9.0915 | 5.1643
223 | 8919.387 | 9.096 | 5.1589 |  |  | 223 | 8919.387 | 9.096 | 5.1582 |  |  | 223 | 8919.387 | 9.096 | 5.163
224 | 8959.384 | 9.1005 | 5.1576 |  |  | 224 | 8959.384 | 9.1005 | 5.1569 |  |  | 224 | 8959.384 | 9.1005 | 5.1617
225 | 8999.381 | 9.1049 | 5.1563 |  |  | 225 | 8999.381 | 9.1049 | 5.1556 |  |  | 225 | 8999.381 | 9.1049 | 5.1604
226 | 9039.379 | 9.1093 | 5.155 |  |  | 226 | 9039.379 | 9.1093 | 5.1543 |  |  | 226 | 9039.379 | 9.1093 | 5.1591
227 | 9079.376 | 9.1138 | 5.1537 |  |  | 227 | 9079.376 | 9.1138 | 5.153 |  |  | 227 | 9079.376 | 9.1138 | 5.1578
228 | 9119.373 | 9.1182 | 5.1524 |  |  | 228 | 9119.373 | 9.1182 | 5.1517 |  |  | 228 | 9119.373 | 9.1182 | 5.1566
229 | 9159.37 | 9.1225 | 5.1512 |  |  | 229 | 9159.37 | 9.1225 | 5.1504 |  |  | 229 | 9159.37 | 9.1225 | 5.1553
230 | 9199.368 | 9.1269 | 5.15 |  |  | 230 | 9199.368 | 9.1269 | 5.1492 |  |  | 230 | 9199.368 | 9.1269 | 5.1541
231 | 9239.365 | 9.1312 | 5.1487 |  |  | 231 | 9239.365 | 9.1312 | 5.1479 |  |  | 231 | 9239.365 | 9.1312 | 5.1528
232 | 9279.362 | 9.1355 | 5.1475 |  |  | 232 | 9279.362 | 9.1355 | 5.1467 |  |  | 232 | 9279.362 | 9.1355 | 5.1516
233 | 9319.359 | 9.1398 | 5.1463 |  |  | 233 | 9319.359 | 9.1398 | 5.1455 |  |  | 233 | 9319.359 | 9.1398 | 5.1504
234 | 9359.357 | 9.1441 | 5.1451 |  |  | 234 | 9359.357 | 9.1441 | 5.1443 |  |  | 234 | 9359.357 | 9.1441 | 5.1492
235 | 9399.354 | 9.1484 | 5.1439 |  |  | 235 | 9399.354 | 9.1484 | 5.1431 |  |  | 235 | 9399.354 | 9.1484 | 5.148
236 | 9439.351 | 9.1526 | 5.1427 |  |  | 236 | 9439.351 | 9.1526 | 5.1419 |  |  | 236 | 9439.351 | 9.1526 | 5.1468
237 | 9479.348 | 9.1569 | 5.1416 |  |  | 237 | 9479.348 | 9.1569 | 5.1407 |  |  | 237 | 9479.348 | 9.1569 | 5.1456
238 | 9519.346 | 9.1611 | 5.1404 |  |  | 238 | 9519.346 | 9.1611 | 5.1395 |  |  | 238 | 9519.346 | 9.1611 | 5.1445
239 | 9559.343 | 9.1653 | 5.1393 |  |  | 239 | 9559.343 | 9.1653 | 5.1383 |  |  | 239 | 9559.343 | 9.1653 | 5.1433
240 | 9599.34 | 9.1694 | 5.1381 |  |  | 240 | 9599.34 | 9.1694 | 5.1372 |  |  | 240 | 9599.34 | 9.1694 | 5.1422
241 | 9639.337 | 9.1736 | 5.137 |  |  | 241 | 9639.337 | 9.1736 | 5.136 |  |  | 241 | 9639.337 | 9.1736 | 5.1411
242 | 9679.335 | 9.1777 | 5.1359 |  |  | 242 | 9679.335 | 9.1777 | 5.1349 |  |  | 242 | 9679.335 | 9.1777 | 5.1399
243 | 9719.332 | 9.1819 | 5.1348 |  |  | 243 | 9719.332 | 9.1819 | 5.1338 |  |  | 243 | 9719.332 | 9.1819 | 5.1388
244 | 9759.329 | 9.186 | 5.1337 |  |  | 244 | 9759.329 | 9.186 | 5.1327 |  |  | 244 | 9759.329 | 9.186 | 5.1377
245 | 9799.326 | 9.1901 | 5.1326 |  |  | 245 | 9799.326 | 9.1901 | 5.1316 |  |  | 245 | 9799.326 | 9.1901 | 5.1366
246 | 9839.324 | 9.1941 | 5.1315 |  |  | 246 | 9839.324 | 9.1941 | 5.1305 |  |  | 246 | 9839.324 | 9.1941 | 5.1355
247 | 9879.321 | 9.1982 | 5.1304 |  |  | 247 | 9879.321 | 9.1982 | 5.1294 |  |  | 247 | 9879.321 | 9.1982 | 5.1344
248 | 9919.318 | 9.2022 | 5.1293 |  |  | 248 | 9919.318 | 9.2022 | 5.1283 |  |  | 248 | 9919.318 | 9.2022 | 5.1334
249 | 9959.315 | 9.2063 | 5.1283 |  |  | 249 | 9959.315 | 9.2063 | 5.1272 |  |  | 249 | 9959.315 | 9.2063 | 5.1323
250 | 9999.313 | 9.2103 | 5.1272 |  |  | 250 | 9999.313 | 9.2103 | 5.1262 |  |  | 250 | 9999.313 | 9.2103 | 5.1313
251 | 10039.31 | 9.2143 | 5.1262 |  |  | 251 | 10039.31 | 9.2143 | 5.1251 |  |  | 251 | 10039.31 | 9.2143 | 5.1302
252 | 10079.307 | 9.2182 | 5.1252 |  |  | 252 | 10079.307 | 9.2182 | 5.1241 |  |  | 252 | 10079.307 | 9.2182 | 5.1292
253 | 10119.304 | 9.2222 | 5.1242 |  |  | 253 | 10119.304 | 9.2222 | 5.123 |  |  | 253 | 10119.304 | 9.2222 | 5.1282
254 | 10159.302 | 9.2261 | 5.1231 |  |  | 254 | 10159.302 | 9.2261 | 5.122 |  |  | 254 | 10159.302 | 9.2261 | 5.1271
255 | 10199.299 | 9.2301 | 5.1221 |  |  | 255 | 10199.299 | 9.2301 | 5.121 |  |  | 255 | 10199.299 | 9.2301 | 5.1261
256 | 10239.296 | 9.234 | 5.1211 |  |  | 256 | 10239.296 | 9.234 | 5.12 |  |  | 256 | 10239.296 | 9.234 | 5.1251
257 | 10279.293 | 9.2379 | 5.1201 |  |  | 257 | 10279.293 | 9.2379 | 5.119 |  |  | 257 | 10279.293 | 9.2379 | 5.1241
258 | 10319.291 | 9.2418 | 5.1192 |  |  | 258 | 10319.291 | 9.2418 | 5.118 |  |  | 258 | 10319.291 | 9.2418 | 5.1232
259 | 10359.288 | 9.2456 | 5.1182 |  |  | 259 | 10359.288 | 9.2456 | 5.117 |  |  | 259 | 10359.288 | 9.2456 | 5.1222
260 | 10399.285 | 9.2495 | 5.1172 |  |  | 260 | 10399.285 | 9.2495 | 5.116 |  |  | 260 | 10399.285 | 9.2495 | 5.1212
261 | 10439.282 | 9.2533 | 5.1163 |  |  | 261 | 10439.282 | 9.2533 | 5.1151 |  |  | 261 | 10439.282 | 9.2533 | 5.1203
262 | 10479.28 | 9.2572 | 5.1153 |  |  | 262 | 10479.28 | 9.2572 | 5.1141 |  |  | 262 | 10479.28 | 9.2572 | 5.1193
263 | 10519.277 | 9.261 | 5.1144 |  |  | 263 | 10519.277 | 9.261 | 5.1131 |  |  | 263 | 10519.277 | 9.261 | 5.1184
264 | 10559.274 | 9.2648 | 5.1135 |  |  | 264 | 10559.274 | 9.2648 | 5.1122 |  |  | 264 | 10559.274 | 9.2648 | 5.1174
265 | 10599.271 | 9.2685 | 5.1125 |  |  | 265 | 10599.271 | 9.2685 | 5.1113 |  |  | 265 | 10599.271 | 9.2685 | 5.1165
266 | 10639.269 | 9.2723 | 5.1116 |  |  | 266 | 10639.269 | 9.2723 | 5.1103 |  |  | 266 | 10639.269 | 9.2723 | 5.1156
267 | 10679.266 | 9.2761 | 5.1107 |  |  | 267 | 10679.266 | 9.2761 | 5.1094 |  |  | 267 | 10679.266 | 9.2761 | 5.1146
268 | 10719.263 | 9.2798 | 5.1098 |  |  | 268 | 10719.263 | 9.2798 | 5.1085 |  |  | 268 | 10719.263 | 9.2798 | 5.1137
269 | 10759.26 | 9.2835 | 5.1089 |  |  | 269 | 10759.26 | 9.2835 | 5.1076 |  |  | 269 | 10759.26 | 9.2835 | 5.1128
270 | 10799.258 | 9.2872 | 5.108 |  |  | 270 | 10799.258 | 9.2872 | 5.1067 |  |  | 270 | 10799.258 | 9.2872 | 5.1119
271 | 10839.255 | 9.2909 | 5.1071 |  |  | 271 | 10839.255 | 9.2909 | 5.1058 |  |  | 271 | 10839.255 | 9.2909 | 5.1111
272 | 10879.252 | 9.2946 | 5.1063 |  |  | 272 | 10879.252 | 9.2946 | 5.1049 |  |  | 272 | 10879.252 | 9.2946 | 5.1102
273 | 10919.249 | 9.2983 | 5.1054 |  |  | 273 | 10919.249 | 9.2983 | 5.104 |  |  | 273 | 10919.249 | 9.2983 | 5.1093
274 | 10959.247 | 9.3019 | 5.1045 |  |  | 274 | 10959.247 | 9.3019 | 5.1032 |  |  | 274 | 10959.247 | 9.3019 | 5.1084
275 | 10999.244 | 9.3056 | 5.1037 |  |  | 275 | 10999.244 | 9.3056 | 5.1023 |  |  | 275 | 10999.244 | 9.3056 | 5.1076
276 | 11039.241 | 9.3092 | 5.1028 |  |  | 276 | 11039.241 | 9.3092 | 5.1014 |  |  | 276 | 11039.241 | 9.3092 | 5.1067
277 | 11079.238 | 9.3128 | 5.102 |  |  | 277 | 11079.238 | 9.3128 | 5.1006 |  |  | 277 | 11079.238 | 9.3128 | 5.1059
278 | 11119.236 | 9.3164 | 5.1011 |  |  | 278 | 11119.236 | 9.3164 | 5.0997 |  |  | 278 | 11119.236 | 9.3164 | 5.105
279 | 11159.233 | 9.32 | 5.1003 |  |  | 279 | 11159.233 | 9.32 | 5.0989 |  |  | 279 | 11159.233 | 9.32 | 5.1042
280 | 11199.23 | 9.3236 | 5.0995 |  |  | 280 | 11199.23 | 9.3236 | 5.0981 |  |  | 280 | 11199.23 | 9.3236 | 5.1034
281 | 11239.227 | 9.3272 | 5.0987 |  |  | 281 | 11239.227 | 9.3272 | 5.0972 |  |  | 281 | 11239.227 | 9.3272 | 5.1026
282 | 11279.225 | 9.3307 | 5.0979 |  |  | 282 | 11279.225 | 9.3307 | 5.0964 |  |  | 282 | 11279.225 | 9.3307 | 5.1018
283 | 11319.222 | 9.3343 | 5.0971 |  |  | 283 | 11319.222 | 9.3343 | 5.0956 |  |  | 283 | 11319.222 | 9.3343 | 5.101
284 | 11359.219 | 9.3378 | 5.0963 |  |  | 284 | 11359.219 | 9.3378 | 5.0948 |  |  | 284 | 11359.219 | 9.3378 | 5.1002
285 | 11399.216 | 9.3413 | 5.0955 |  |  | 285 | 11399.216 | 9.3413 | 5.094 |  |  | 285 | 11399.216 | 9.3413 | 5.0994
286 | 11439.214 | 9.3448 | 5.0947 |  |  | 286 | 11439.214 | 9.3448 | 5.0932 |  |  | 286 | 11439.214 | 9.3448 | 5.0986
287 | 11479.211 | 9.3483 | 5.0939 |  |  | 287 | 11479.211 | 9.3483 | 5.0924 |  |  | 287 | 11479.211 | 9.3483 | 5.0978
288 | 11519.208 | 9.3518 | 5.0932 |  |  | 288 | 11519.208 | 9.3518 | 5.0917 |  |  | 288 | 11519.208 | 9.3518 | 5.097
289 | 11559.205 | 9.3552 | 5.0924 |  |  | 289 | 11559.205 | 9.3552 | 5.0909 |  |  | 289 | 11559.205 | 9.3552 | 5.0962
290 | 11599.203 | 9.3587 | 5.0916 |  |  | 290 | 11599.203 | 9.3587 | 5.0901 |  |  | 290 | 11599.203 | 9.3587 | 5.0955
291 | 11639.2 | 9.3621 | 5.0909 |  |  | 291 | 11639.2 | 9.3621 | 5.0894 |  |  | 291 | 11639.2 | 9.3621 | 5.0947
292 | 11679.197 | 9.3656 | 5.0901 |  |  | 292 | 11679.197 | 9.3656 | 5.0886 |  |  | 292 | 11679.197 | 9.3656 | 5.094
293 | 11719.194 | 9.369 | 5.0894 |  |  | 293 | 11719.194 | 9.369 | 5.0878 |  |  | 293 | 11719.194 | 9.369 | 5.0932
294 | 11759.192 | 9.3724 | 5.0887 |  |  | 294 | 11759.192 | 9.3724 | 5.0871 |  |  | 294 | 11759.192 | 9.3724 | 5.0925
295 | 11799.189 | 9.3758 | 5.0879 |  |  | 295 | 11799.189 | 9.3758 | 5.0864 |  |  | 295 | 11799.189 | 9.3758 | 5.0918
296 | 11839.186 | 9.3792 | 5.0872 |  |  | 296 | 11839.186 | 9.3792 | 5.0856 |  |  | 296 | 11839.186 | 9.3792 | 5.091
297 | 11879.183 | 9.3825 | 5.0865 |  |  | 297 | 11879.183 | 9.3825 | 5.0849 |  |  | 297 | 11879.183 | 9.3825 | 5.0903
298 | 11919.181 | 9.3859 | 5.0858 |  |  | 298 | 11919.181 | 9.3859 | 5.0842 |  |  | 298 | 11919.181 | 9.3859 | 5.0896
299 | 11959.178 | 9.3893 | 5.0851 |  |  | 299 | 11959.178 | 9.3893 | 5.0835 |  |  | 299 | 11959.178 | 9.3893 | 5.0889
300 | 11999.175 | 9.3926 | 5.0844 |  |  | 300 | 11999.175 | 9.3926 | 5.0828 |  |  | 300 | 11999.175 | 9.3926 | 5.0882
301 | 12039.172 | 9.3959 | 5.0837 |  |  | 301 | 12039.172 | 9.3959 | 5.082 |  |  | 301 | 12039.172 | 9.3959 | 5.0875
302 | 12079.17 | 9.3992 | 5.083 |  |  | 302 | 12079.17 | 9.3992 | 5.0813 |  |  | 302 | 12079.17 | 9.3992 | 5.0868
303 | 12119.167 | 9.4025 | 5.0823 |  |  | 303 | 12119.167 | 9.4025 | 5.0807 |  |  | 303 | 12119.167 | 9.4025 | 5.0861
304 | 12159.164 | 9.4058 | 5.0816 |  |  | 304 | 12159.164 | 9.4058 | 5.08 |  |  | 304 | 12159.164 | 9.4058 | 5.0854
305 | 12199.161 | 9.4091 | 5.0809 |  |  | 305 | 12199.161 | 9.4091 | 5.0793 |  |  | 305 | 12199.161 | 9.4091 | 5.0847
306 | 12239.159 | 9.4124 | 5.0803 |  |  | 306 | 12239.159 | 9.4124 | 5.0786 |  |  | 306 | 12239.159 | 9.4124 | 5.084
307 | 12279.156 | 9.4157 | 5.0796 |  |  | 307 | 12279.156 | 9.4157 | 5.0779 |  |  | 307 | 12279.156 | 9.4157 | 5.0834
308 | 12319.153 | 9.4189 | 5.0789 |  |  | 308 | 12319.153 | 9.4189 | 5.0773 |  |  | 308 | 12319.153 | 9.4189 | 5.0827
309 | 12359.15 | 9.4222 | 5.0783 |  |  | 309 | 12359.15 | 9.4222 | 5.0766 |  |  | 309 | 12359.15 | 9.4222 | 5.082
310 | 12399.148 | 9.4254 | 5.0776 |  |  | 310 | 12399.148 | 9.4254 | 5.0759 |  |  | 310 | 12399.148 | 9.4254 | 5.0814
311 | 12439.145 | 9.4286 | 5.077 |  |  | 311 | 12439.145 | 9.4286 | 5.0753 |  |  | 311 | 12439.145 | 9.4286 | 5.0807
312 | 12479.142 | 9.4318 | 5.0763 |  |  | 312 | 12479.142 | 9.4318 | 5.0746 |  |  | 312 | 12479.142 | 9.4318 | 5.0801
313 | 12519.139 | 9.435 | 5.0757 |  |  | 313 | 12519.139 | 9.435 | 5.074 |  |  | 313 | 12519.139 | 9.435 | 5.0794
314 | 12559.137 | 9.4382 | 5.0751 |  |  | 314 | 12559.137 | 9.4382 | 5.0734 |  |  | 314 | 12559.137 | 9.4382 | 5.0788
315 | 12599.134 | 9.4414 | 5.0745 |  |  | 315 | 12599.134 | 9.4414 | 5.0727 |  |  | 315 | 12599.134 | 9.4414 | 5.0782
316 | 12639.131 | 9.4446 | 5.0738 |  |  | 316 | 12639.131 | 9.4446 | 5.0721 |  |  | 316 | 12639.131 | 9.4446 | 5.0776
317 | 12679.128 | 9.4477 | 5.0732 |  |  | 317 | 12679.128 | 9.4477 | 5.0715 |  |  | 317 | 12679.128 | 9.4477 | 5.0769
318 | 12719.126 | 9.4509 | 5.0726 |  |  | 318 | 12719.126 | 9.4509 | 5.0709 |  |  | 318 | 12719.126 | 9.4509 | 5.0763
319 | 12759.123 | 9.454 | 5.072 |  |  | 319 | 12759.123 | 9.454 | 5.0702 |  |  | 319 | 12759.123 | 9.454 | 5.0757
320 | 12799.12 | 9.4571 | 5.0714 |  |  | 320 | 12799.12 | 9.4571 | 5.0696 |  |  | 320 | 12799.12 | 9.4571 | 5.0751
321 | 12839.117 | 9.4603 | 5.0708 |  |  | 321 | 12839.117 | 9.4603 | 5.069 |  |  | 321 | 12839.117 | 9.4603 | 5.0745
322 | 12879.115 | 9.4634 | 5.0702 |  |  | 322 | 12879.115 | 9.4634 | 5.0684 |  |  | 322 | 12879.115 | 9.4634 | 5.0739
323 | 12919.112 | 9.4665 | 5.0696 |  |  | 323 | 12919.112 | 9.4665 | 5.0678 |  |  | 323 | 12919.112 | 9.4665 | 5.0733
324 | 12959.109 | 9.4696 | 5.069 |  |  | 324 | 12959.109 | 9.4696 | 5.0672 |  |  | 324 | 12959.109 | 9.4696 | 5.0727
325 | 12999.106 | 9.4726 | 5.0684 |  |  | 325 | 12999.106 | 9.4726 | 5.0666 |  |  | 325 | 12999.106 | 9.4726 | 5.0721
326 | 13039.104 | 9.4757 | 5.0679 |  |  | 326 | 13039.104 | 9.4757 | 5.0661 |  |  | 326 | 13039.104 | 9.4757 | 5.0715
327 | 13079.101 | 9.4788 | 5.0673 |  |  | 327 | 13079.101 | 9.4788 | 5.0655 |  |  | 327 | 13079.101 | 9.4788 | 5.071
328 | 13119.098 | 9.4818 | 5.0667 |  |  | 328 | 13119.098 | 9.4818 | 5.0649 |  |  | 328 | 13119.098 | 9.4818 | 5.0704
329 | 13159.095 | 9.4849 | 5.0662 |  |  | 329 | 13159.095 | 9.4849 | 5.0643 |  |  | 329 | 13159.095 | 9.4849 | 5.0698
330 | 13199.093 | 9.4879 | 5.0656 |  |  | 330 | 13199.093 | 9.4879 | 5.0638 |  |  | 330 | 13199.093 | 9.4879 | 5.0693
331 | 13239.09 | 9.4909 | 5.065 |  |  | 331 | 13239.09 | 9.4909 | 5.0632 |  |  | 331 | 13239.09 | 9.4909 | 5.0687
332 | 13279.087 | 9.4939 | 5.0645 |  |  | 332 | 13279.087 | 9.4939 | 5.0626 |  |  | 332 | 13279.087 | 9.4939 | 5.0681
333 | 13319.084 | 9.497 | 5.0639 |  |  | 333 | 13319.084 | 9.497 | 5.0621 |  |  | 333 | 13319.084 | 9.497 | 5.0676
334 | 13359.082 | 9.5 | 5.0634 |  |  | 334 | 13359.082 | 9.5 | 5.0615 |  |  | 334 | 13359.082 | 9.5 | 5.067
335 | 13399.079 | 9.5029 | 5.0628 |  |  | 335 | 13399.079 | 9.5029 | 5.061 |  |  | 335 | 13399.079 | 9.5029 | 5.0665
336 | 13439.076 | 9.5059 | 5.0623 |  |  | 336 | 13439.076 | 9.5059 | 5.0604 |  |  | 336 | 13439.076 | 9.5059 | 5.0659
337 | 13479.073 | 9.5089 | 5.0618 |  |  | 337 | 13479.073 | 9.5089 | 5.0599 |  |  | 337 | 13479.073 | 9.5089 | 5.0654
338 | 13519.071 | 9.5119 | 5.0612 |  |  | 338 | 13519.071 | 9.5119 | 5.0594 |  |  | 338 | 13519.071 | 9.5119 | 5.0649
339 | 13559.068 | 9.5148 | 5.0607 |  |  | 339 | 13559.068 | 9.5148 | 5.0588 |  |  | 339 | 13559.068 | 9.5148 | 5.0643
340 | 13599.065 | 9.5178 | 5.0602 |  |  | 340 | 13599.065 | 9.5178 | 5.0583 |  |  | 340 | 13599.065 | 9.5178 | 5.0638
341 | 13639.062 | 9.5207 | 5.0597 |  |  | 341 | 13639.062 | 9.5207 | 5.0578 |  |  | 341 | 13639.062 | 9.5207 | 5.0633
342 | 13679.06 | 9.5236 | 5.0592 |  |  | 342 | 13679.06 | 9.5236 | 5.0573 |  |  | 342 | 13679.06 | 9.5236 | 5.0628
343 | 13719.057 | 9.5265 | 5.0587 |  |  | 343 | 13719.057 | 9.5265 | 5.0568 |  |  | 343 | 13719.057 | 9.5265 | 5.0623
344 | 13759.054 | 9.5295 | 5.0582 |  |  | 344 | 13759.054 | 9.5295 | 5.0562 |  |  | 344 | 13759.054 | 9.5295 | 5.0618
345 | 13799.051 | 9.5324 | 5.0576 |  |  | 345 | 13799.051 | 9.5324 | 5.0557 |  |  | 345 | 13799.051 | 9.5324 | 5.0612
346 | 13839.049 | 9.5352 | 5.0571 |  |  | 346 | 13839.049 | 9.5352 | 5.0552 |  |  | 346 | 13839.049 | 9.5352 | 5.0607
347 | 13879.046 | 9.5381 | 5.0567 |  |  | 347 | 13879.046 | 9.5381 | 5.0547 |  |  | 347 | 13879.046 | 9.5381 | 5.0602
348 | 13919.043 | 9.541 | 5.0562 |  |  | 348 | 13919.043 | 9.541 | 5.0542 |  |  | 348 | 13919.043 | 9.541 | 5.0597
349 | 13959.04 | 9.5439 | 5.0557 |  |  | 349 | 13959.04 | 9.5439 | 5.0537 |  |  | 349 | 13959.04 | 9.5439 | 5.0592
350 | 13999.038 | 9.5467 | 5.0552 |  |  | 350 | 13999.038 | 9.5467 | 5.0533 |  |  | 350 | 13999.038 | 9.5467 | 5.0588
351 | 14039.035 | 9.5496 | 5.0547 |  |  | 351 | 14039.035 | 9.5496 | 5.0528 |  |  | 351 | 14039.035 | 9.5496 | 5.0583
352 | 14079.032 | 9.5524 | 5.0542 |  |  | 352 | 14079.032 | 9.5524 | 5.0523 |  |  | 352 | 14079.032 | 9.5524 | 5.0578
353 | 14119.029 | 9.5553 | 5.0537 |  |  | 353 | 14119.029 | 9.5553 | 5.0518 |  |  | 353 | 14119.029 | 9.5553 | 5.0573
354 | 14159.027 | 9.5581 | 5.0533 |  |  | 354 | 14159.027 | 9.5581 | 5.0513 |  |  | 354 | 14159.027 | 9.5581 | 5.0568
355 | 14199.024 | 9.5609 | 5.0528 |  |  | 355 | 14199.024 | 9.5609 | 5.0509 |  |  | 355 | 14199.024 | 9.5609 | 5.0564
356 | 14239.021 | 9.5637 | 5.0523 |  |  | 356 | 14239.021 | 9.5637 | 5.0504 |  |  | 356 | 14239.021 | 9.5637 | 5.0559
357 | 14279.018 | 9.5665 | 5.0519 |  |  | 357 | 14279.018 | 9.5665 | 5.0499 |  |  | 357 | 14279.018 | 9.5665 | 5.0554
358 | 14319.016 | 9.5693 | 5.0514 |  |  | 358 | 14319.016 | 9.5693 | 5.0495 |  |  | 358 | 14319.016 | 9.5693 | 5.055
359 | 14359.013 | 9.5721 | 5.051 |  |  | 359 | 14359.013 | 9.5721 | 5.049 |  |  | 359 | 14359.013 | 9.5721 | 5.0545
360 | 14399.01 | 9.5749 | 5.0505 |  |  | 360 | 14399.01 | 9.5749 | 5.0485 |  |  | 360 | 14399.01 | 9.5749 | 5.054
361 | 14439.007 | 9.5777 | 5.0501 |  |  | 361 | 14439.007 | 9.5777 | 5.0481 |  |  | 361 | 14439.007 | 9.5777 | 5.0536
362 | 14479.005 | 9.5805 | 5.0496 |  |  | 362 | 14479.005 | 9.5805 | 5.0476 |  |  | 362 | 14479.005 | 9.5805 | 5.0531
363 | 14519.002 | 9.5832 | 5.0492 |  |  | 363 | 14519.002 | 9.5832 | 5.0472 |  |  | 363 | 14519.002 | 9.5832 | 5.0527
364 | 14558.999 | 9.586 | 5.0487 |  |  | 364 | 14558.999 | 9.586 | 5.0467 |  |  | 364 | 14558.999 | 9.586 | 5.0522
365 | 14598.996 | 9.5887 | 5.0483 |  |  | 365 | 14598.996 | 9.5887 | 5.0463 |  |  | 365 | 14598.996 | 9.5887 | 5.0518
366 | 14638.994 | 9.5914 | 5.0479 |  |  | 366 | 14638.994 | 9.5914 | 5.0459 |  |  | 366 | 14638.994 | 9.5914 | 5.0514
367 | 14678.991 | 9.5942 | 5.0474 |  |  | 367 | 14678.991 | 9.5942 | 5.0454 |  |  | 367 | 14678.991 | 9.5942 | 5.0509
368 | 14718.988 | 9.5969 | 5.047 |  |  | 368 | 14718.988 | 9.5969 | 5.045 |  |  | 368 | 14718.988 | 9.5969 | 5.0505
369 | 14758.985 | 9.5996 | 5.0466 |  |  | 369 | 14758.985 | 9.5996 | 5.0446 |  |  | 369 | 14758.985 | 9.5996 | 5.0501
370 | 14798.983 | 9.6023 | 5.0461 |  |  | 370 | 14798.983 | 9.6023 | 5.0441 |  |  | 370 | 14798.983 | 9.6023 | 5.0496
371 | 14838.98 | 9.605 | 5.0457 |  |  | 371 | 14838.98 | 9.605 | 5.0437 |  |  | 371 | 14838.98 | 9.605 | 5.0492
372 | 14878.977 | 9.6077 | 5.0453 |  |  | 372 | 14878.977 | 9.6077 | 5.0433 |  |  | 372 | 14878.977 | 9.6077 | 5.0488
373 | 14918.974 | 9.6104 | 5.0449 |  |  | 373 | 14918.974 | 9.6104 | 5.0429 |  |  | 373 | 14918.974 | 9.6104 | 5.0484
374 | 14958.972 | 9.6131 | 5.0445 |  |  | 374 | 14958.972 | 9.6131 | 5.0425 |  |  | 374 | 14958.972 | 9.6131 | 5.048
375 | 14998.969 | 9.6157 | 5.0441 |  |  | 375 | 14998.969 | 9.6157 | 5.042 |  |  | 375 | 14998.969 | 9.6157 | 5.0476
376 | 15038.966 | 9.6184 | 5.0437 |  |  | 376 | 15038.966 | 9.6184 | 5.0416 |  |  | 376 | 15038.966 | 9.6184 | 5.0471
377 | 15078.963 | 9.6211 | 5.0433 |  |  | 377 | 15078.963 | 9.6211 | 5.0412 |  |  | 377 | 15078.963 | 9.6211 | 5.0467
378 | 15118.961 | 9.6237 | 5.0429 |  |  | 378 | 15118.961 | 9.6237 | 5.0408 |  |  | 378 | 15118.961 | 9.6237 | 5.0463
379 | 15158.958 | 9.6263 | 5.0425 |  |  | 379 | 15158.958 | 9.6263 | 5.0404 |  |  | 379 | 15158.958 | 9.6263 | 5.0459
380 | 15198.955 | 9.629 | 5.0421 |  |  | 380 | 15198.955 | 9.629 | 5.04 |  |  | 380 | 15198.955 | 9.629 | 5.0455
381 | 15238.952 | 9.6316 | 5.0417 |  |  | 381 | 15238.952 | 9.6316 | 5.0396 |  |  | 381 | 15238.952 | 9.6316 | 5.0451
382 | 15278.95 | 9.6342 | 5.0413 |  |  | 382 | 15278.95 | 9.6342 | 5.0392 |  |  | 382 | 15278.95 | 9.6342 | 5.0447
383 | 15318.947 | 9.6368 | 5.0409 |  |  | 383 | 15318.947 | 9.6368 | 5.0388 |  |  | 383 | 15318.947 | 9.6368 | 5.0443
384 | 15358.944 | 9.6395 | 5.0405 |  |  | 384 | 15358.944 | 9.6395 | 5.0385 |  |  | 384 | 15358.944 | 9.6395 | 5.044
385 | 15398.941 | 9.6421 | 5.0401 |  |  | 385 | 15398.941 | 9.6421 | 5.0381 |  |  | 385 | 15398.941 | 9.6421 | 5.0436
386 | 15438.939 | 9.6446 | 5.0397 |  |  | 386 | 15438.939 | 9.6446 | 5.0377 |  |  | 386 | 15438.939 | 9.6446 | 5.0432
387 | 15478.936 | 9.6472 | 5.0394 |  |  | 387 | 15478.936 | 9.6472 | 5.0373 |  |  | 387 | 15478.936 | 9.6472 | 5.0428
388 | 15518.933 | 9.6498 | 5.039 |  |  | 388 | 15518.933 | 9.6498 | 5.0369 |  |  | 388 | 15518.933 | 9.6498 | 5.0424
389 | 15558.93 | 9.6524 | 5.0386 |  |  | 389 | 15558.93 | 9.6524 | 5.0366 |  |  | 389 | 15558.93 | 9.6524 | 5.0421
390 | 15598.928 | 9.655 | 5.0382 |  |  | 390 | 15598.928 | 9.655 | 5.0362 |  |  | 390 | 15598.928 | 9.655 | 5.0417
391 | 15638.925 | 9.6575 | 5.0379 |  |  | 391 | 15638.925 | 9.6575 | 5.0358 |  |  | 391 | 15638.925 | 9.6575 | 5.0413
392 | 15678.922 | 9.6601 | 5.0375 |  |  | 392 | 15678.922 | 9.6601 | 5.0354 |  |  | 392 | 15678.922 | 9.6601 | 5.0409
393 | 15718.919 | 9.6626 | 5.0371 |  |  | 393 | 15718.919 | 9.6626 | 5.0351 |  |  | 393 | 15718.919 | 9.6626 | 5.0406
394 | 15758.917 | 9.6652 | 5.0368 |  |  | 394 | 15758.917 | 9.6652 | 5.0347 |  |  | 394 | 15758.917 | 9.6652 | 5.0402
395 | 15798.914 | 9.6677 | 5.0364 |  |  | 395 | 15798.914 | 9.6677 | 5.0344 |  |  | 395 | 15798.914 | 9.6677 | 5.0398
396 | 15838.911 | 9.6702 | 5.0361 |  |  | 396 | 15838.911 | 9.6702 | 5.034 |  |  | 396 | 15838.911 | 9.6702 | 5.0395
397 | 15878.908 | 9.6727 | 5.0357 |  |  | 397 | 15878.908 | 9.6727 | 5.0336 |  |  | 397 | 15878.908 | 9.6727 | 5.0391
398 | 15918.906 | 9.6753 | 5.0354 |  |  | 398 | 15918.906 | 9.6753 | 5.0333 |  |  | 398 | 15918.906 | 9.6753 | 5.0388
399 | 15958.903 | 9.6778 | 5.035 |  |  | 399 | 15958.903 | 9.6778 | 5.0329 |  |  | 399 | 15958.903 | 9.6778 | 5.0384
400 | 15998.9 | 9.6803 | 5.0347 |  |  | 400 | 15998.9 | 9.6803 | 5.0326 |  |  | 400 | 15998.9 | 9.6803 | 5.0381
401 | 16038.897 | 9.6828 | 5.0343 |  |  | 401 | 16038.897 | 9.6828 | 5.0322 |  |  | 401 | 16038.897 | 9.6828 | 5.0377
402 | 16078.895 | 9.6853 | 5.034 |  |  | 402 | 16078.895 | 9.6853 | 5.0319 |  |  | 402 | 16078.895 | 9.6853 | 5.0374
403 | 16118.892 | 9.6877 | 5.0336 |  |  | 403 | 16118.892 | 9.6877 | 5.0315 |  |  | 403 | 16118.892 | 9.6877 | 5.037
404 | 16158.889 | 9.6902 | 5.0333 |  |  | 404 | 16158.889 | 9.6902 | 5.0312 |  |  | 404 | 16158.889 | 9.6902 | 5.0367
405 | 16198.886 | 9.6927 | 5.033 |  |  | 405 | 16198.886 | 9.6927 | 5.0309 |  |  | 405 | 16198.886 | 9.6927 | 5.0364
406 | 16238.884 | 9.6952 | 5.0326 |  |  | 406 | 16238.884 | 9.6952 | 5.0305 |  |  | 406 | 16238.884 | 9.6952 | 5.036
407 | 16278.881 | 9.6976 | 5.0323 |  |  | 407 | 16278.881 | 9.6976 | 5.0302 |  |  | 407 | 16278.881 | 9.6976 | 5.0357
408 | 16318.878 | 9.7001 | 5.032 |  |  | 408 | 16318.878 | 9.7001 | 5.0299 |  |  | 408 | 16318.878 | 9.7001 | 5.0353
409 | 16358.875 | 9.7025 | 5.0316 |  |  | 409 | 16358.875 | 9.7025 | 5.0295 |  |  | 409 | 16358.875 | 9.7025 | 5.035
410 | 16398.873 | 9.705 | 5.0313 |  |  | 410 | 16398.873 | 9.705 | 5.0292 |  |  | 410 | 16398.873 | 9.705 | 5.0347
411 | 16438.87 | 9.7074 | 5.031 |  |  | 411 | 16438.87 | 9.7074 | 5.0289 |  |  | 411 | 16438.87 | 9.7074 | 5.0344
412 | 16478.867 | 9.7098 | 5.0307 |  |  | 412 | 16478.867 | 9.7098 | 5.0286 |  |  | 412 | 16478.867 | 9.7098 | 5.034
413 | 16518.864 | 9.7123 | 5.0303 |  |  | 413 | 16518.864 | 9.7123 | 5.0282 |  |  | 413 | 16518.864 | 9.7123 | 5.0337
414 | 16558.862 | 9.7147 | 5.03 |  |  | 414 | 16558.862 | 9.7147 | 5.0279 |  |  | 414 | 16558.862 | 9.7147 | 5.0334
415 | 16598.859 | 9.7171 | 5.0297 |  |  | 415 | 16598.859 | 9.7171 | 5.0276 |  |  | 415 | 16598.859 | 9.7171 | 5.0331
416 | 16638.856 | 9.7195 | 5.0294 |  |  | 416 | 16638.856 | 9.7195 | 5.0273 |  |  | 416 | 16638.856 | 9.7195 | 5.0328
417 | 16678.853 | 9.7219 | 5.0291 |  |  | 417 | 16678.853 | 9.7219 | 5.027 |  |  | 417 | 16678.853 | 9.7219 | 5.0324
418 | 16718.851 | 9.7243 | 5.0288 |  |  | 418 | 16718.851 | 9.7243 | 5.0267 |  |  | 418 | 16718.851 | 9.7243 | 5.0321
419 | 16758.848 | 9.7267 | 5.0285 |  |  | 419 | 16758.848 | 9.7267 | 5.0263 |  |  | 419 | 16758.848 | 9.7267 | 5.0318
420 | 16798.845 | 9.7291 | 5.0281 |  |  | 420 | 16798.845 | 9.7291 | 5.026 |  |  | 420 | 16798.845 | 9.7291 | 5.0315
421 | 16838.842 | 9.7314 | 5.0278 |  |  | 421 | 16838.842 | 9.7314 | 5.0257 |  |  | 421 | 16838.842 | 9.7314 | 5.0312
422 | 16878.84 | 9.7338 | 5.0275 |  |  | 422 | 16878.84 | 9.7338 | 5.0254 |  |  | 422 | 16878.84 | 9.7338 | 5.0309
423 | 16918.837 | 9.7362 | 5.0272 |  |  | 423 | 16918.837 | 9.7362 | 5.0251 |  |  | 423 | 16918.837 | 9.7362 | 5.0306
424 | 16958.834 | 9.7385 | 5.0269 |  |  | 424 | 16958.834 | 9.7385 | 5.0248 |  |  | 424 | 16958.834 | 9.7385 | 5.0303
425 | 16998.831 | 9.7409 | 5.0266 |  |  | 425 | 16998.831 | 9.7409 | 5.0245 |  |  | 425 | 16998.831 | 9.7409 | 5.03
426 | 17038.829 | 9.7433 | 5.0263 |  |  | 426 | 17038.829 | 9.7433 | 5.0242 |  |  | 426 | 17038.829 | 9.7433 | 5.0297
427 | 17078.826 | 9.7456 | 5.026 |  |  | 427 | 17078.826 | 9.7456 | 5.0239 |  |  | 427 | 17078.826 | 9.7456 | 5.0294
428 | 17118.823 | 9.7479 | 5.0258 |  |  | 428 | 17118.823 | 9.7479 | 5.0236 |  |  | 428 | 17118.823 | 9.7479 | 5.0291
429 | 17158.82 | 9.7503 | 5.0255 |  |  | 429 | 17158.82 | 9.7503 | 5.0233 |  |  | 429 | 17158.82 | 9.7503 | 5.0288
430 | 17198.818 | 9.7526 | 5.0252 |  |  | 430 | 17198.818 | 9.7526 | 5.023 |  |  | 430 | 17198.818 | 9.7526 | 5.0285
431 | 17238.815 | 9.7549 | 5.0249 |  |  | 431 | 17238.815 | 9.7549 | 5.0228 |  |  | 431 | 17238.815 | 9.7549 | 5.0282
432 | 17278.812 | 9.7572 | 5.0246 |  |  | 432 | 17278.812 | 9.7572 | 5.0225 |  |  | 432 | 17278.812 | 9.7572 | 5.0279
433 | 17318.809 | 9.7595 | 5.0243 |  |  | 433 | 17318.809 | 9.7595 | 5.0222 |  |  | 433 | 17318.809 | 9.7595 | 5.0277
434 | 17358.807 | 9.7619 | 5.024 |  |  | 434 | 17358.807 | 9.7619 | 5.0219 |  |  | 434 | 17358.807 | 9.7619 | 5.0274
435 | 17398.804 | 9.7642 | 5.0237 |  |  | 435 | 17398.804 | 9.7642 | 5.0216 |  |  | 435 | 17398.804 | 9.7642 | 5.0271
436 | 17438.801 | 9.7665 | 5.0235 |  |  | 436 | 17438.801 | 9.7665 | 5.0213 |  |  | 436 | 17438.801 | 9.7665 | 5.0268
437 | 17478.798 | 9.7687 | 5.0232 |  |  | 437 | 17478.798 | 9.7687 | 5.0211 |  |  | 437 | 17478.798 | 9.7687 | 5.0265
438 | 17518.796 | 9.771 | 5.0229 |  |  | 438 | 17518.796 | 9.771 | 5.0208 |  |  | 438 | 17518.796 | 9.771 | 5.0262
439 | 17558.793 | 9.7733 | 5.0226 |  |  | 439 | 17558.793 | 9.7733 | 5.0205 |  |  | 439 | 17558.793 | 9.7733 | 5.026
440 | 17598.79 | 9.7756 | 5.0224 |  |  | 440 | 17598.79 | 9.7756 | 5.0202 |  |  | 440 | 17598.79 | 9.7756 | 5.0257
441 | 17638.787 | 9.7779 | 5.0221 |  |  | 441 | 17638.787 | 9.7779 | 5.02 |  |  | 441 | 17638.787 | 9.7779 | 5.0254
442 | 17678.785 | 9.7801 | 5.0218 |  |  | 442 | 17678.785 | 9.7801 | 5.0197 |  |  | 442 | 17678.785 | 9.7801 | 5.0251
443 | 17718.782 | 9.7824 | 5.0216 |  |  | 443 | 17718.782 | 9.7824 | 5.0194 |  |  | 443 | 17718.782 | 9.7824 | 5.0249
444 | 17758.779 | 9.7846 | 5.0213 |  |  | 444 | 17758.779 | 9.7846 | 5.0192 |  |  | 444 | 17758.779 | 9.7846 | 5.0246
445 | 17798.776 | 9.7869 | 5.021 |  |  | 445 | 17798.776 | 9.7869 | 5.0189 |  |  | 445 | 17798.776 | 9.7869 | 5.0243
446 | 17838.774 | 9.7891 | 5.0208 |  |  | 446 | 17838.774 | 9.7891 | 5.0186 |  |  | 446 | 17838.774 | 9.7891 | 5.0241
447 | 17878.771 | 9.7914 | 5.0205 |  |  | 447 | 17878.771 | 9.7914 | 5.0184 |  |  | 447 | 17878.771 | 9.7914 | 5.0238
448 | 17918.768 | 9.7936 | 5.0202 |  |  | 448 | 17918.768 | 9.7936 | 5.0181 |  |  | 448 | 17918.768 | 9.7936 | 5.0236
449 | 17958.765 | 9.7958 | 5.02 |  |  | 449 | 17958.765 | 9.7958 | 5.0178 |  |  | 449 | 17958.765 | 9.7958 | 5.0233
450 | 17998.763 | 9.7981 | 5.0197 |  |  | 450 | 17998.763 | 9.7981 | 5.0176 |  |  | 450 | 17998.763 | 9.7981 | 5.023
451 | 18038.76 | 9.8003 | 5.0195 |  |  | 451 | 18038.76 | 9.8003 | 5.0173 |  |  | 451 | 18038.76 | 9.8003 | 5.0228
452 | 18078.757 | 9.8025 | 5.0192 |  |  | 452 | 18078.757 | 9.8025 | 5.0171 |  |  | 452 | 18078.757 | 9.8025 | 5.0225
453 | 18118.755 | 9.8047 | 5.0189 |  |  | 453 | 18118.755 | 9.8047 | 5.0168 |  |  | 453 | 18118.755 | 9.8047 | 5.0223
454 | 18158.752 | 9.8069 | 5.0187 |  |  | 454 | 18158.752 | 9.8069 | 5.0166 |  |  | 454 | 18158.752 | 9.8069 | 5.022
455 | 18198.749 | 9.8091 | 5.0184 |  |  | 455 | 18198.749 | 9.8091 | 5.0163 |  |  | 455 | 18198.749 | 9.8091 | 5.0218
456 | 18238.746 | 9.8113 | 5.0182 |  |  | 456 | 18238.746 | 9.8113 | 5.0161 |  |  | 456 | 18238.746 | 9.8113 | 5.0215
457 | 18278.744 | 9.8135 | 5.0179 |  |  | 457 | 18278.744 | 9.8135 | 5.0158 |  |  | 457 | 18278.744 | 9.8135 | 5.0213
458 | 18318.741 | 9.8157 | 5.0177 |  |  | 458 | 18318.741 | 9.8157 | 5.0156 |  |  | 458 | 18318.741 | 9.8157 | 5.021
459 | 18358.738 | 9.8179 | 5.0174 |  |  | 459 | 18358.738 | 9.8179 | 5.0153 |  |  | 459 | 18358.738 | 9.8179 | 5.0208
460 | 18398.735 | 9.82 | 5.0172 |  |  | 460 | 18398.735 | 9.82 | 5.0151 |  |  | 460 | 18398.735 | 9.82 | 5.0205
461 | 18438.733 | 9.8222 | 5.017 |  |  | 461 | 18438.733 | 9.8222 | 5.0148 |  |  | 461 | 18438.733 | 9.8222 | 5.0203
462 | 18478.73 | 9.8244 | 5.0167 |  |  | 462 | 18478.73 | 9.8244 | 5.0146 |  |  | 462 | 18478.73 | 9.8244 | 5.02
463 | 18518.727 | 9.8265 | 5.0165 |  |  | 463 | 18518.727 | 9.8265 | 5.0143 |  |  | 463 | 18518.727 | 9.8265 | 5.0198
464 | 18558.724 | 9.8287 | 5.0162 |  |  | 464 | 18558.724 | 9.8287 | 5.0141 |  |  | 464 | 18558.724 | 9.8287 | 5.0195
465 | 18598.722 | 9.8308 | 5.016 |  |  | 465 | 18598.722 | 9.8308 | 5.0139 |  |  | 465 | 18598.722 | 9.8308 | 5.0193
466 | 18638.719 | 9.833 | 5.0158 |  |  | 466 | 18638.719 | 9.833 | 5.0136 |  |  | 466 | 18638.719 | 9.833 | 5.0191
467 | 18678.716 | 9.8351 | 5.0155 |  |  | 467 | 18678.716 | 9.8351 | 5.0134 |  |  | 467 | 18678.716 | 9.8351 | 5.0188
468 | 18718.713 | 9.8373 | 5.0153 |  |  | 468 | 18718.713 | 9.8373 | 5.0132 |  |  | 468 | 18718.713 | 9.8373 | 5.0186
469 | 18758.711 | 9.8394 | 5.0151 |  |  | 469 | 18758.711 | 9.8394 | 5.0129 |  |  | 469 | 18758.711 | 9.8394 | 5.0184
470 | 18798.708 | 9.8415 | 5.0148 |  |  | 470 | 18798.708 | 9.8415 | 5.0127 |  |  | 470 | 18798.708 | 9.8415 | 5.0181
471 | 18838.705 | 9.8437 | 5.0146 |  |  | 471 | 18838.705 | 9.8437 | 5.0125 |  |  | 471 | 18838.705 | 9.8437 | 5.0179
472 | 18878.702 | 9.8458 | 5.0144 |  |  | 472 | 18878.702 | 9.8458 | 5.0122 |  |  | 472 | 18878.702 | 9.8458 | 5.0177
473 | 18918.7 | 9.8479 | 5.0141 |  |  | 473 | 18918.7 | 9.8479 | 5.012 |  |  | 473 | 18918.7 | 9.8479 | 5.0174
474 | 18958.697 | 9.85 | 5.0139 |  |  | 474 | 18958.697 | 9.85 | 5.0118 |  |  | 474 | 18958.697 | 9.85 | 5.0172
475 | 18998.694 | 9.8521 | 5.0137 |  |  | 475 | 18998.694 | 9.8521 | 5.0116 |  |  | 475 | 18998.694 | 9.8521 | 5.017
476 | 19038.691 | 9.8542 | 5.0135 |  |  | 476 | 19038.691 | 9.8542 | 5.0113 |  |  | 476 | 19038.691 | 9.8542 | 5.0168
477 | 19078.689 | 9.8563 | 5.0132 |  |  | 477 | 19078.689 | 9.8563 | 5.0111 |  |  | 477 | 19078.689 | 9.8563 | 5.0165
478 | 19118.686 | 9.8584 | 5.013 |  |  | 478 | 19118.686 | 9.8584 | 5.0109 |  |  | 478 | 19118.686 | 9.8584 | 5.0163
479 | 19158.683 | 9.8605 | 5.0128 |  |  | 479 | 19158.683 | 9.8605 | 5.0107 |  |  | 479 | 19158.683 | 9.8605 | 5.0161
480 | 19198.68 | 9.8626 | 5.0126 |  |  | 480 | 19198.68 | 9.8626 | 5.0104 |  |  | 480 | 19198.68 | 9.8626 | 5.0159
481 | 19238.678 | 9.8647 | 5.0123 |  |  | 481 | 19238.678 | 9.8647 | 5.0102 |  |  | 481 | 19238.678 | 9.8647 | 5.0157
482 | 19278.675 | 9.8668 | 5.0121 |  |  | 482 | 19278.675 | 9.8668 | 5.01 |  |  | 482 | 19278.675 | 9.8668 | 5.0154
483 | 19318.672 | 9.8688 | 5.0119 |  |  | 483 | 19318.672 | 9.8688 | 5.0098 |  |  | 483 | 19318.672 | 9.8688 | 5.0152
484 | 19358.669 | 9.8709 | 5.0117 |  |  | 484 | 19358.669 | 9.8709 | 5.0096 |  |  | 484 | 19358.669 | 9.8709 | 5.015
485 | 19398.667 | 9.873 | 5.0115 |  |  | 485 | 19398.667 | 9.873 | 5.0094 |  |  | 485 | 19398.667 | 9.873 | 5.0148
486 | 19438.664 | 9.875 | 5.0113 |  |  | 486 | 19438.664 | 9.875 | 5.0091 |  |  | 486 | 19438.664 | 9.875 | 5.0146
487 | 19478.661 | 9.8771 | 5.011 |  |  | 487 | 19478.661 | 9.8771 | 5.0089 |  |  | 487 | 19478.661 | 9.8771 | 5.0144
488 | 19518.658 | 9.8791 | 5.0108 |  |  | 488 | 19518.658 | 9.8791 | 5.0087 |  |  | 488 | 19518.658 | 9.8791 | 5.0141
489 | 19558.656 | 9.8812 | 5.0106 |  |  | 489 | 19558.656 | 9.8812 | 5.0085 |  |  | 489 | 19558.656 | 9.8812 | 5.0139
490 | 19598.653 | 9.8832 | 5.0104 |  |  | 490 | 19598.653 | 9.8832 | 5.0083 |  |  | 490 | 19598.653 | 9.8832 | 5.0137
491 | 19638.65 | 9.8853 | 5.0102 |  |  | 491 | 19638.65 | 9.8853 | 5.0081 |  |  | 491 | 19638.65 | 9.8853 | 5.0135
492 | 19678.647 | 9.8873 | 5.01 |  |  | 492 | 19678.647 | 9.8873 | 5.0079 |  |  | 492 | 19678.647 | 9.8873 | 5.0133
493 | 19718.645 | 9.8893 | 5.0098 |  |  | 493 | 19718.645 | 9.8893 | 5.0077 |  |  | 493 | 19718.645 | 9.8893 | 5.0131
494 | 19758.642 | 9.8913 | 5.0096 |  |  | 494 | 19758.642 | 9.8913 | 5.0075 |  |  | 494 | 19758.642 | 9.8913 | 5.0129
495 | 19798.639 | 9.8934 | 5.0094 |  |  | 495 | 19798.639 | 9.8934 | 5.0073 |  |  | 495 | 19798.639 | 9.8934 | 5.0127
496 | 19838.636 | 9.8954 | 5.0092 |  |  | 496 | 19838.636 | 9.8954 | 5.0071 |  |  | 496 | 19838.636 | 9.8954 | 5.0125
497 | 19878.634 | 9.8974 | 5.009 |  |  | 497 | 19878.634 | 9.8974 | 5.0069 |  |  | 497 | 19878.634 | 9.8974 | 5.0123
498 | 19918.631 | 9.8994 | 5.0088 |  |  | 498 | 19918.631 | 9.8994 | 5.0066 |  |  | 498 | 19918.631 | 9.8994 | 5.0121
499 | 19958.628 | 9.9014 | 5.0086 |  |  | 499 | 19958.628 | 9.9014 | 5.0064 |  |  | 499 | 19958.628 | 9.9014 | 5.0119
500 | 19998.625 | 9.9034 | 5.0084 |  |  | 500 | 19998.625 | 9.9034 | 5.0062 |  |  | 500 | 19998.625 | 9.9034 | 5.0117
501 | 20038.623 | 9.9054 | 5.0082 |  |  | 501 | 20038.623 | 9.9054 | 5.006 |  |  | 501 | 20038.623 | 9.9054 | 5.0115
502 | 20078.62 | 9.9074 | 5.008 |  |  | 502 | 20078.62 | 9.9074 | 5.0059 |  |  | 502 | 20078.62 | 9.9074 | 5.0113
503 | 20118.617 | 9.9094 | 5.0078 |  |  | 503 | 20118.617 | 9.9094 | 5.0057 |  |  | 503 | 20118.617 | 9.9094 | 5.0111
504 | 20158.614 | 9.9114 | 5.0076 |  |  | 504 | 20158.614 | 9.9114 | 5.0055 |  |  | 504 | 20158.614 | 9.9114 | 5.0109
505 | 20198.612 | 9.9134 | 5.0074 |  |  | 505 | 20198.612 | 9.9134 | 5.0053 |  |  | 505 | 20198.612 | 9.9134 | 5.0107
506 | 20238.609 | 9.9153 | 5.0072 |  |  | 506 | 20238.609 | 9.9153 | 5.0051 |  |  | 506 | 20238.609 | 9.9153 | 5.0105
507 | 20278.606 | 9.9173 | 5.007 |  |  | 507 | 20278.606 | 9.9173 | 5.0049 |  |  | 507 | 20278.606 | 9.9173 | 5.0103
508 | 20318.603 | 9.9193 | 5.0068 |  |  | 508 | 20318.603 | 9.9193 | 5.0047 |  |  | 508 | 20318.603 | 9.9193 | 5.0101
509 | 20358.601 | 9.9213 | 5.0066 |  |  | 509 | 20358.601 | 9.9213 | 5.0045 |  |  | 509 | 20358.601 | 9.9213 | 5.0099
510 | 20398.598 | 9.9232 | 5.0064 |  |  | 510 | 20398.598 | 9.9232 | 5.0043 |  |  | 510 | 20398.598 | 9.9232 | 5.0097
511 | 20438.595 | 9.9252 | 5.0062 |  |  | 511 | 20438.595 | 9.9252 | 5.0041 |  |  | 511 | 20438.595 | 9.9252 | 5.0095
512 | 20478.592 | 9.9271 | 5.006 |  |  | 512 | 20478.592 | 9.9271 | 5.0039 |  |  | 512 | 20478.592 | 9.9271 | 5.0094
513 | 20518.59 | 9.9291 | 5.0058 |  |  | 513 | 20518.59 | 9.9291 | 5.0037 |  |  | 513 | 20518.59 | 9.9291 | 5.0092
514 | 20558.587 | 9.931 | 5.0056 |  |  | 514 | 20558.587 | 9.931 | 5.0035 |  |  | 514 | 20558.587 | 9.931 | 5.009
515 | 20598.584 | 9.933 | 5.0054 |  |  | 515 | 20598.584 | 9.933 | 5.0034 |  |  | 515 | 20598.584 | 9.933 | 5.0088
516 | 20638.581 | 9.9349 | 5.0053 |  |  | 516 | 20638.581 | 9.9349 | 5.0032 |  |  | 516 | 20638.581 | 9.9349 | 5.0086
517 | 20678.579 | 9.9369 | 5.0051 |  |  | 517 | 20678.579 | 9.9369 | 5.003 |  |  | 517 | 20678.579 | 9.9369 | 5.0084
518 | 20718.576 | 9.9388 | 5.0049 |  |  | 518 | 20718.576 | 9.9388 | 5.0028 |  |  | 518 | 20718.576 | 9.9388 | 5.0082
519 | 20758.573 | 9.9407 | 5.0047 |  |  | 519 | 20758.573 | 9.9407 | 5.0026 |  |  | 519 | 20758.573 | 9.9407 | 5.0081
520 | 20798.57 | 9.9426 | 5.0045 |  |  | 520 | 20798.57 | 9.9426 | 5.0024 |  |  | 520 | 20798.57 | 9.9426 | 5.0079
521 | 20838.568 | 9.9446 | 5.0043 |  |  | 521 | 20838.568 | 9.9446 | 5.0022 |  |  | 521 | 20838.568 | 9.9446 | 5.0077
522 | 20878.565 | 9.9465 | 5.0042 |  |  | 522 | 20878.565 | 9.9465 | 5.0021 |  |  | 522 | 20878.565 | 9.9465 | 5.0075
523 | 20918.562 | 9.9484 | 5.004 |  |  | 523 | 20918.562 | 9.9484 | 5.0019 |  |  | 523 | 20918.562 | 9.9484 | 5.0073
524 | 20958.559 | 9.9503 | 5.0038 |  |  | 524 | 20958.559 | 9.9503 | 5.0017 |  |  | 524 | 20958.559 | 9.9503 | 5.0071
525 | 20998.557 | 9.9522 | 5.0036 |  |  | 525 | 20998.557 | 9.9522 | 5.0015 |  |  | 525 | 20998.557 | 9.9522 | 5.007
526 | 21038.554 | 9.9541 | 5.0034 |  |  | 526 | 21038.554 | 9.9541 | 5.0013 |  |  | 526 | 21038.554 | 9.9541 | 5.0068
527 | 21078.551 | 9.956 | 5.0032 |  |  | 527 | 21078.551 | 9.956 | 5.0012 |  |  | 527 | 21078.551 | 9.956 | 5.0066
528 | 21118.548 | 9.9579 | 5.0031 |  |  | 528 | 21118.548 | 9.9579 | 5.001 |  |  | 528 | 21118.548 | 9.9579 | 5.0064
529 | 21158.546 | 9.9598 | 5.0029 |  |  | 529 | 21158.546 | 9.9598 | 5.0008 |  |  | 529 | 21158.546 | 9.9598 | 5.0063
530 | 21198.543 | 9.9617 | 5.0027 |  |  | 530 | 21198.543 | 9.9617 | 5.0006 |  |  | 530 | 21198.543 | 9.9617 | 5.0061
531 | 21238.54 | 9.9636 | 5.0025 |  |  | 531 | 21238.54 | 9.9636 | 5.0005 |  |  | 531 | 21238.54 | 9.9636 | 5.0059
532 | 21278.537 | 9.9655 | 5.0024 |  |  | 532 | 21278.537 | 9.9655 | 5.0003 |  |  | 532 | 21278.537 | 9.9655 | 5.0057
533 | 21318.535 | 9.9673 | 5.0022 |  |  | 533 | 21318.535 | 9.9673 | 5.0001 |  |  | 533 | 21318.535 | 9.9673 | 5.0056
534 | 21358.532 | 9.9692 | 5.002 |  |  | 534 | 21358.532 | 9.9692 | 4.9999 |  |  | 534 | 21358.532 | 9.9692 | 5.0054
535 | 21398.529 | 9.9711 | 5.0018 |  |  | 535 | 21398.529 | 9.9711 | 4.9998 |  |  | 535 | 21398.529 | 9.9711 | 5.0052
536 | 21438.526 | 9.9729 | 5.0017 |  |  | 536 | 21438.526 | 9.9729 | 4.9996 |  |  | 536 | 21438.526 | 9.9729 | 5.0051
537 | 21478.524 | 9.9748 | 5.0015 |  |  | 537 | 21478.524 | 9.9748 | 4.9994 |  |  | 537 | 21478.524 | 9.9748 | 5.0049
538 | 21518.521 | 9.9767 | 5.0013 |  |  | 538 | 21518.521 | 9.9767 | 4.9993 |  |  | 538 | 21518.521 | 9.9767 | 5.0047
539 | 21558.518 | 9.9785 | 5.0012 |  |  | 539 | 21558.518 | 9.9785 | 4.9991 |  |  | 539 | 21558.518 | 9.9785 | 5.0045
540 | 21598.515 | 9.9804 | 5.001 |  |  | 540 | 21598.515 | 9.9804 | 4.9989 |  |  | 540 | 21598.515 | 9.9804 | 5.0044
541 | 21638.513 | 9.9822 | 5.0008 |  |  | 541 | 21638.513 | 9.9822 | 4.9987 |  |  | 541 | 21638.513 | 9.9822 | 5.0042
542 | 21678.51 | 9.9841 | 5.0006 |  |  | 542 | 21678.51 | 9.9841 | 4.9986 |  |  | 542 | 21678.51 | 9.9841 | 5.004
543 | 21718.507 | 9.9859 | 5.0005 |  |  | 543 | 21718.507 | 9.9859 | 4.9984 |  |  | 543 | 21718.507 | 9.9859 | 5.0039
544 | 21758.504 | 9.9878 | 5.0003 |  |  | 544 | 21758.504 | 9.9878 | 4.9982 |  |  | 544 | 21758.504 | 9.9878 | 5.0037
545 | 21798.502 | 9.9896 | 5.0001 |  |  | 545 | 21798.502 | 9.9896 | 4.9981 |  |  | 545 | 21798.502 | 9.9896 | 5.0036
546 | 21838.499 | 9.9914 | 5 |  |  | 546 | 21838.499 | 9.9914 | 4.9979 |  |  | 546 | 21838.499 | 9.9914 | 5.0034
547 | 21878.496 | 9.9933 | 4.9998 |  |  | 547 | 21878.496 | 9.9933 | 4.9978 |  |  | 547 | 21878.496 | 9.9933 | 5.0032
548 | 21918.493 | 9.9951 | 4.9996 |  |  | 548 | 21918.493 | 9.9951 | 4.9976 |  |  | 548 | 21918.493 | 9.9951 | 5.0031
549 | 21958.491 | 9.9969 | 4.9995 |  |  | 549 | 21958.491 | 9.9969 | 4.9974 |  |  | 549 | 21958.491 | 9.9969 | 5.0029
550 | 21998.488 | 9.9987 | 4.9993 |  |  | 550 | 21998.488 | 9.9987 | 4.9973 |  |  | 550 | 21998.488 | 9.9987 | 5.0027
551 | 22038.485 | 10.0005 | 4.9992 |  |  | 551 | 22038.485 | 10.0005 | 4.9971 |  |  | 551 | 22038.485 | 10.0005 | 5.0026
552 | 22078.482 | 10.0024 | 4.999 |  |  | 552 | 22078.482 | 10.0024 | 4.9969 |  |  | 552 | 22078.482 | 10.0024 | 5.0024
553 | 22118.48 | 10.0042 | 4.9988 |  |  | 553 | 22118.48 | 10.0042 | 4.9968 |  |  | 553 | 22118.48 | 10.0042 | 5.0023
554 | 22158.477 | 10.006 | 4.9987 |  |  | 554 | 22158.477 | 10.006 | 4.9966 |  |  | 554 | 22158.477 | 10.006 | 5.0021
555 | 22198.474 | 10.0078 | 4.9985 |  |  | 555 | 22198.474 | 10.0078 | 4.9965 |  |  | 555 | 22198.474 | 10.0078 | 5.0019
556 | 22238.471 | 10.0096 | 4.9983 |  |  | 556 | 22238.471 | 10.0096 | 4.9963 |  |  | 556 | 22238.471 | 10.0096 | 5.0018
557 | 22278.469 | 10.0114 | 4.9982 |  |  | 557 | 22278.469 | 10.0114 | 4.9961 |  |  | 557 | 22278.469 | 10.0114 | 5.0016
558 | 22318.466 | 10.0132 | 4.998 |  |  | 558 | 22318.466 | 10.0132 | 4.996 |  |  | 558 | 22318.466 | 10.0132 | 5.0015
559 | 22358.463 | 10.015 | 4.9979 |  |  | 559 | 22358.463 | 10.015 | 4.9958 |  |  | 559 | 22358.463 | 10.015 | 5.0013
560 | 22398.46 | 10.0167 | 4.9977 |  |  | 560 | 22398.46 | 10.0167 | 4.9957 |  |  | 560 | 22398.46 | 10.0167 | 5.0012
561 | 22438.458 | 10.0185 | 4.9975 |  |  | 561 | 22438.458 | 10.0185 | 4.9955 |  |  | 561 | 22438.458 | 10.0185 | 5.001
562 | 22478.455 | 10.0203 | 4.9974 |  |  | 562 | 22478.455 | 10.0203 | 4.9954 |  |  | 562 | 22478.455 | 10.0203 | 5.0009
563 | 22518.452 | 10.0221 | 4.9972 |  |  | 563 | 22518.452 | 10.0221 | 4.9952 |  |  | 563 | 22518.452 | 10.0221 | 5.0007
564 | 22558.449 | 10.0239 | 4.9971 |  |  | 564 | 22558.449 | 10.0239 | 4.995 |  |  | 564 | 22558.449 | 10.0239 | 5.0005
565 | 22598.447 | 10.0256 | 4.9969 |  |  | 565 | 22598.447 | 10.0256 | 4.9949 |  |  | 565 | 22598.447 | 10.0256 | 5.0004
566 | 22638.444 | 10.0274 | 4.9968 |  |  | 566 | 22638.444 | 10.0274 | 4.9947 |  |  | 566 | 22638.444 | 10.0274 | 5.0002
567 | 22678.441 | 10.0292 | 4.9966 |  |  | 567 | 22678.441 | 10.0292 | 4.9946 |  |  | 567 | 22678.441 | 10.0292 | 5.0001
568 | 22718.438 | 10.0309 | 4.9965 |  |  | 568 | 22718.438 | 10.0309 | 4.9944 |  |  | 568 | 22718.438 | 10.0309 | 4.9999
569 | 22758.436 | 10.0327 | 4.9963 |  |  | 569 | 22758.436 | 10.0327 | 4.9943 |  |  | 569 | 22758.436 | 10.0327 | 4.9998
570 | 22798.433 | 10.0344 | 4.9961 |  |  | 570 | 22798.433 | 10.0344 | 4.9941 |  |  | 570 | 22798.433 | 10.0344 | 4.9996
571 | 22838.43 | 10.0362 | 4.996 |  |  | 571 | 22838.43 | 10.0362 | 4.994 |  |  | 571 | 22838.43 | 10.0362 | 4.9995
572 | 22878.427 | 10.0379 | 4.9958 |  |  | 572 | 22878.427 | 10.0379 | 4.9938 |  |  | 572 | 22878.427 | 10.0379 | 4.9993
573 | 22918.425 | 10.0397 | 4.9957 |  |  | 573 | 22918.425 | 10.0397 | 4.9937 |  |  | 573 | 22918.425 | 10.0397 | 4.9992
574 | 22958.422 | 10.0414 | 4.9955 |  |  | 574 | 22958.422 | 10.0414 | 4.9935 |  |  | 574 | 22958.422 | 10.0414 | 4.9991
575 | 22998.419 | 10.0432 | 4.9954 |  |  | 575 | 22998.419 | 10.0432 | 4.9934 |  |  | 575 | 22998.419 | 10.0432 | 4.9989
576 | 23038.416 | 10.0449 | 4.9952 |  |  | 576 | 23038.416 | 10.0449 | 4.9932 |  |  | 576 | 23038.416 | 10.0449 | 4.9988
577 | 23078.414 | 10.0467 | 4.9951 |  |  | 577 | 23078.414 | 10.0467 | 4.9931 |  |  | 577 | 23078.414 | 10.0467 | 4.9986
578 | 23118.411 | 10.0484 | 4.9949 |  |  | 578 | 23118.411 | 10.0484 | 4.9929 |  |  | 578 | 23118.411 | 10.0484 | 4.9985
579 | 23158.408 | 10.0501 | 4.9948 |  |  | 579 | 23158.408 | 10.0501 | 4.9928 |  |  | 579 | 23158.408 | 10.0501 | 4.9983
580 | 23198.405 | 10.0518 | 4.9946 |  |  | 580 | 23198.405 | 10.0518 | 4.9926 |  |  | 580 | 23198.405 | 10.0518 | 4.9982
581 | 23238.403 | 10.0536 | 4.9945 |  |  | 581 | 23238.403 | 10.0536 | 4.9925 |  |  | 581 | 23238.403 | 10.0536 | 4.998
582 | 23278.4 | 10.0553 | 4.9943 |  |  | 582 | 23278.4 | 10.0553 | 4.9923 |  |  | 582 | 23278.4 | 10.0553 | 4.9979
583 | 23318.397 | 10.057 | 4.9942 |  |  | 583 | 23318.397 | 10.057 | 4.9922 |  |  | 583 | 23318.397 | 10.057 | 4.9977
584 | 23358.394 | 10.0587 | 4.994 |  |  | 584 | 23358.394 | 10.0587 | 4.992 |  |  | 584 | 23358.394 | 10.0587 | 4.9976
585 | 23398.392 | 10.0604 | 4.9939 |  |  | 585 | 23398.392 | 10.0604 | 4.9919 |  |  | 585 | 23398.392 | 10.0604 | 4.9975
586 | 23438.389 | 10.0621 | 4.9938 |  |  | 586 | 23438.389 | 10.0621 | 4.9917 |  |  | 586 | 23438.389 | 10.0621 | 4.9973
587 | 23478.386 | 10.0638 | 4.9936 |  |  | 587 | 23478.386 | 10.0638 | 4.9916 |  |  | 587 | 23478.386 | 10.0638 | 4.9972
588 | 23518.383 | 10.0655 | 4.9935 |  |  | 588 | 23518.383 | 10.0655 | 4.9915 |  |  | 588 | 23518.383 | 10.0655 | 4.997
589 | 23558.381 | 10.0672 | 4.9933 |  |  | 589 | 23558.381 | 10.0672 | 4.9913 |  |  | 589 | 23558.381 | 10.0672 | 4.9969
590 | 23598.378 | 10.0689 | 4.9932 |  |  | 590 | 23598.378 | 10.0689 | 4.9912 |  |  | 590 | 23598.378 | 10.0689 | 4.9968
591 | 23638.375 | 10.0706 | 4.993 |  |  | 591 | 23638.375 | 10.0706 | 4.991 |  |  | 591 | 23638.375 | 10.0706 | 4.9966
592 | 23678.372 | 10.0723 | 4.9929 |  |  | 592 | 23678.372 | 10.0723 | 4.9909 |  |  | 592 | 23678.372 | 10.0723 | 4.9965
593 | 23718.37 | 10.074 | 4.9927 |  |  | 593 | 23718.37 | 10.074 | 4.9907 |  |  | 593 | 23718.37 | 10.074 | 4.9963
594 | 23758.367 | 10.0757 | 4.9926 |  |  | 594 | 23758.367 | 10.0757 | 4.9906 |  |  | 594 | 23758.367 | 10.0757 | 4.9962
595 | 23798.364 | 10.0774 | 4.9925 |  |  | 595 | 23798.364 | 10.0774 | 4.9905 |  |  | 595 | 23798.364 | 10.0774 | 4.9961
596 | 23838.361 | 10.0791 | 4.9923 |  |  | 596 | 23838.361 | 10.0791 | 4.9903 |  |  | 596 | 23838.361 | 10.0791 | 4.9959
597 | 23878.359 | 10.0807 | 4.9922 |  |  | 597 | 23878.359 | 10.0807 | 4.9902 |  |  | 597 | 23878.359 | 10.0807 | 4.9958
598 | 23918.356 | 10.0824 | 4.992 |  |  | 598 | 23918.356 | 10.0824 | 4.99 |  |  | 598 | 23918.356 | 10.0824 | 4.9956
599 | 23958.353 | 10.0841 | 4.9919 |  |  | 599 | 23958.353 | 10.0841 | 4.9899 |  |  | 599 | 23958.353 | 10.0841 | 4.9955
600 | 23998.35 | 10.0857 | 4.9918 |  |  | 600 | 23998.35 | 10.0857 | 4.9898 |  |  | 600 | 23998.35 | 10.0857 | 4.9954
601 | 24038.348 | 10.0874 | 4.9916 |  |  | 601 | 24038.348 | 10.0874 | 4.9896 |  |  | 601 | 24038.348 | 10.0874 | 4.9952
602 | 24078.345 | 10.0891 | 4.9915 |  |  | 602 | 24078.345 | 10.0891 | 4.9895 |  |  | 602 | 24078.345 | 10.0891 | 4.9951
603 | 24118.342 | 10.0907 | 4.9913 |  |  | 603 | 24118.342 | 10.0907 | 4.9893 |  |  | 603 | 24118.342 | 10.0907 | 4.995
604 | 24158.339 | 10.0924 | 4.9912 |  |  | 604 | 24158.339 | 10.0924 | 4.9892 |  |  | 604 | 24158.339 | 10.0924 | 4.9948
605 | 24198.337 | 10.094 | 4.9911 |  |  | 605 | 24198.337 | 10.094 | 4.9891 |  |  | 605 | 24198.337 | 10.094 | 4.9947
606 | 24238.334 | 10.0957 | 4.9909 |  |  | 606 | 24238.334 | 10.0957 | 4.9889 |  |  | 606 | 24238.334 | 10.0957 | 4.9946
607 | 24278.331 | 10.0973 | 4.9908 |  |  | 607 | 24278.331 | 10.0973 | 4.9888 |  |  | 607 | 24278.331 | 10.0973 | 4.9944
608 | 24318.328 | 10.099 | 4.9906 |  |  | 608 | 24318.328 | 10.099 | 4.9887 |  |  | 608 | 24318.328 | 10.099 | 4.9943
609 | 24358.326 | 10.1006 | 4.9905 |  |  | 609 | 24358.326 | 10.1006 | 4.9885 |  |  | 609 | 24358.326 | 10.1006 | 4.9942
610 | 24398.323 | 10.1023 | 4.9904 |  |  | 610 | 24398.323 | 10.1023 | 4.9884 |  |  | 610 | 24398.323 | 10.1023 | 4.994
611 | 24438.32 | 10.1039 | 4.9902 |  |  | 611 | 24438.32 | 10.1039 | 4.9882 |  |  | 611 | 24438.32 | 10.1039 | 4.9939
612 | 24478.317 | 10.1055 | 4.9901 |  |  | 612 | 24478.317 | 10.1055 | 4.9881 |  |  | 612 | 24478.317 | 10.1055 | 4.9938
613 | 24518.315 | 10.1072 | 4.99 |  |  | 613 | 24518.315 | 10.1072 | 4.988 |  |  | 613 | 24518.315 | 10.1072 | 4.9936
614 | 24558.312 | 10.1088 | 4.9898 |  |  | 614 | 24558.312 | 10.1088 | 4.9878 |  |  | 614 | 24558.312 | 10.1088 | 4.9935
615 | 24598.309 | 10.1104 | 4.9897 |  |  | 615 | 24598.309 | 10.1104 | 4.9877 |  |  | 615 | 24598.309 | 10.1104 | 4.9934
616 | 24638.306 | 10.1121 | 4.9895 |  |  | 616 | 24638.306 | 10.1121 | 4.9876 |  |  | 616 | 24638.306 | 10.1121 | 4.9933
617 | 24678.304 | 10.1137 | 4.9894 |  |  | 617 | 24678.304 | 10.1137 | 4.9874 |  |  | 617 | 24678.304 | 10.1137 | 4.9931
618 | 24718.301 | 10.1153 | 4.9893 |  |  | 618 | 24718.301 | 10.1153 | 4.9873 |  |  | 618 | 24718.301 | 10.1153 | 4.993
619 | 24758.298 | 10.1169 | 4.9891 |  |  | 619 | 24758.298 | 10.1169 | 4.9872 |  |  | 619 | 24758.298 | 10.1169 | 4.9929
620 | 24798.295 | 10.1185 | 4.989 |  |  | 620 | 24798.295 | 10.1185 | 4.987 |  |  | 620 | 24798.295 | 10.1185 | 4.9927
621 | 24838.293 | 10.1201 | 4.9889 |  |  | 621 | 24838.293 | 10.1201 | 4.9869 |  |  | 621 | 24838.293 | 10.1201 | 4.9926
622 | 24878.29 | 10.1218 | 4.9887 |  |  | 622 | 24878.29 | 10.1218 | 4.9868 |  |  | 622 | 24878.29 | 10.1218 | 4.9925
623 | 24918.287 | 10.1234 | 4.9886 |  |  | 623 | 24918.287 | 10.1234 | 4.9866 |  |  | 623 | 24918.287 | 10.1234 | 4.9924
624 | 24958.284 | 10.125 | 4.9885 |  |  | 624 | 24958.284 | 10.125 | 4.9865 |  |  | 624 | 24958.284 | 10.125 | 4.9922
625 | 24998.282 | 10.1266 | 4.9883 |  |  | 625 | 24998.282 | 10.1266 | 4.9864 |  |  | 625 | 24998.282 | 10.1266 | 4.9921
626 | 25038.279 | 10.1282 | 4.9882 |  |  | 626 | 25038.279 | 10.1282 | 4.9862 |  |  | 626 | 25038.279 | 10.1282 | 4.992
627 | 25078.276 | 10.1298 | 4.9881 |  |  | 627 | 25078.276 | 10.1298 | 4.9861 |  |  | 627 | 25078.276 | 10.1298 | 4.9918
628 | 25118.273 | 10.1314 | 4.9879 |  |  | 628 | 25118.273 | 10.1314 | 4.986 |  |  | 628 | 25118.273 | 10.1314 | 4.9917
629 | 25158.271 | 10.1329 | 4.9878 |  |  | 629 | 25158.271 | 10.1329 | 4.9859 |  |  | 629 | 25158.271 | 10.1329 | 4.9916
630 | 25198.268 | 10.1345 | 4.9877 |  |  | 630 | 25198.268 | 10.1345 | 4.9857 |  |  | 630 | 25198.268 | 10.1345 | 4.9915
631 | 25238.265 | 10.1361 | 4.9875 |  |  | 631 | 25238.265 | 10.1361 | 4.9856 |  |  | 631 | 25238.265 | 10.1361 | 4.9913
632 | 25278.262 | 10.1377 | 4.9874 |  |  | 632 | 25278.262 | 10.1377 | 4.9855 |  |  | 632 | 25278.262 | 10.1377 | 4.9912
633 | 25318.26 | 10.1393 | 4.9873 |  |  | 633 | 25318.26 | 10.1393 | 4.9853 |  |  | 633 | 25318.26 | 10.1393 | 4.9911
634 | 25358.257 | 10.1409 | 4.9872 |  |  | 634 | 25358.257 | 10.1409 | 4.9852 |  |  | 634 | 25358.257 | 10.1409 | 4.991
635 | 25398.254 | 10.1424 | 4.987 |  |  | 635 | 25398.254 | 10.1424 | 4.9851 |  |  | 635 | 25398.254 | 10.1424 | 4.9908
636 | 25438.251 | 10.144 | 4.9869 |  |  | 636 | 25438.251 | 10.144 | 4.9849 |  |  | 636 | 25438.251 | 10.144 | 4.9907
637 | 25478.249 | 10.1456 | 4.9868 |  |  | 637 | 25478.249 | 10.1456 | 4.9848 |  |  | 637 | 25478.249 | 10.1456 | 4.9906
638 | 25518.246 | 10.1471 | 4.9866 |  |  | 638 | 25518.246 | 10.1471 | 4.9847 |  |  | 638 | 25518.246 | 10.1471 | 4.9905
639 | 25558.243 | 10.1487 | 4.9865 |  |  | 639 | 25558.243 | 10.1487 | 4.9846 |  |  | 639 | 25558.243 | 10.1487 | 4.9903
640 | 25598.24 | 10.1503 | 4.9864 |  |  | 640 | 25598.24 | 10.1503 | 4.9844 |  |  | 640 | 25598.24 | 10.1503 | 4.9902
641 | 25638.238 | 10.1518 | 4.9862 |  |  | 641 | 25638.238 | 10.1518 | 4.9843 |  |  | 641 | 25638.238 | 10.1518 | 4.9901
642 | 25678.235 | 10.1534 | 4.9861 |  |  | 642 | 25678.235 | 10.1534 | 4.9842 |  |  | 642 | 25678.235 | 10.1534 | 4.99
643 | 25718.232 | 10.155 | 4.986 |  |  | 643 | 25718.232 | 10.155 | 4.9841 |  |  | 643 | 25718.232 | 10.155 | 4.9899
644 | 25758.229 | 10.1565 | 4.9859 |  |  | 644 | 25758.229 | 10.1565 | 4.9839 |  |  | 644 | 25758.229 | 10.1565 | 4.9897
645 | 25798.227 | 10.1581 | 4.9857 |  |  | 645 | 25798.227 | 10.1581 | 4.9838 |  |  | 645 | 25798.227 | 10.1581 | 4.9896
646 | 25838.224 | 10.1596 | 4.9856 |  |  | 646 | 25838.224 | 10.1596 | 4.9837 |  |  | 646 | 25838.224 | 10.1596 | 4.9895
647 | 25878.221 | 10.1612 | 4.9855 |  |  | 647 | 25878.221 | 10.1612 | 4.9835 |  |  | 647 | 25878.221 | 10.1612 | 4.9894
648 | 25918.218 | 10.1627 | 4.9854 |  |  | 648 | 25918.218 | 10.1627 | 4.9834 |  |  | 648 | 25918.218 | 10.1627 | 4.9893
649 | 25958.216 | 10.1642 | 4.9852 |  |  | 649 | 25958.216 | 10.1642 | 4.9833 |  |  | 649 | 25958.216 | 10.1642 | 4.9891
650 | 25998.213 | 10.1658 | 4.9851 |  |  | 650 | 25998.213 | 10.1658 | 4.9832 |  |  | 650 | 25998.213 | 10.1658 | 4.989
651 | 26038.21 | 10.1673 | 4.985 |  |  | 651 | 26038.21 | 10.1673 | 4.983 |  |  | 651 | 26038.21 | 10.1673 | 4.9889
652 | 26078.207 | 10.1689 | 4.9848 |  |  | 652 | 26078.207 | 10.1689 | 4.9829 |  |  | 652 | 26078.207 | 10.1689 | 4.9888
653 | 26118.205 | 10.1704 | 4.9847 |  |  | 653 | 26118.205 | 10.1704 | 4.9828 |  |  | 653 | 26118.205 | 10.1704 | 4.9887
654 | 26158.202 | 10.1719 | 4.9846 |  |  | 654 | 26158.202 | 10.1719 | 4.9827 |  |  | 654 | 26158.202 | 10.1719 | 4.9885
655 | 26198.199 | 10.1734 | 4.9845 |  |  | 655 | 26198.199 | 10.1734 | 4.9825 |  |  | 655 | 26198.199 | 10.1734 | 4.9884
656 | 26238.196 | 10.175 | 4.9843 |  |  | 656 | 26238.196 | 10.175 | 4.9824 |  |  | 656 | 26238.196 | 10.175 | 4.9883
657 | 26278.194 | 10.1765 | 4.9842 |  |  | 657 | 26278.194 | 10.1765 | 4.9823 |  |  | 657 | 26278.194 | 10.1765 | 4.9882
658 | 26318.191 | 10.178 | 4.9841 |  |  | 658 | 26318.191 | 10.178 | 4.9822 |  |  | 658 | 26318.191 | 10.178 | 4.9881
659 | 26358.188 | 10.1795 | 4.984 |  |  | 659 | 26358.188 | 10.1795 | 4.982 |  |  | 659 | 26358.188 | 10.1795 | 4.9879
660 | 26398.185 | 10.1811 | 4.9838 |  |  | 660 | 26398.185 | 10.1811 | 4.9819 |  |  | 660 | 26398.185 | 10.1811 | 4.9878
661 | 26438.183 | 10.1826 | 4.9837 |  |  | 661 | 26438.183 | 10.1826 | 4.9818 |  |  | 661 | 26438.183 | 10.1826 | 4.9877
662 | 26478.18 | 10.1841 | 4.9836 |  |  | 662 | 26478.18 | 10.1841 | 4.9817 |  |  | 662 | 26478.18 | 10.1841 | 4.9876
663 | 26518.177 | 10.1856 | 4.9835 |  |  | 663 | 26518.177 | 10.1856 | 4.9816 |  |  | 663 | 26518.177 | 10.1856 | 4.9875
664 | 26558.174 | 10.1871 | 4.9833 |  |  | 664 | 26558.174 | 10.1871 | 4.9814 |  |  | 664 | 26558.174 | 10.1871 | 4.9874
665 | 26598.172 | 10.1886 | 4.9832 |  |  | 665 | 26598.172 | 10.1886 | 4.9813 |  |  | 665 | 26598.172 | 10.1886 | 4.9872
666 | 26638.169 | 10.1901 | 4.9831 |  |  | 666 | 26638.169 | 10.1901 | 4.9812 |  |  | 666 | 26638.169 | 10.1901 | 4.9871
667 | 26678.166 | 10.1916 | 4.983 |  |  | 667 | 26678.166 | 10.1916 | 4.9811 |  |  | 667 | 26678.166 | 10.1916 | 4.987
668 | 26718.163 | 10.1931 | 4.9829 |  |  | 668 | 26718.163 | 10.1931 | 4.9809 |  |  | 668 | 26718.163 | 10.1931 | 4.9869
669 | 26758.161 | 10.1946 | 4.9827 |  |  | 669 | 26758.161 | 10.1946 | 4.9808 |  |  | 669 | 26758.161 | 10.1946 | 4.9868
670 | 26798.158 | 10.1961 | 4.9826 |  |  | 670 | 26798.158 | 10.1961 | 4.9807 |  |  | 670 | 26798.158 | 10.1961 | 4.9867
671 | 26838.155 | 10.1976 | 4.9825 |  |  | 671 | 26838.155 | 10.1976 | 4.9806 |  |  | 671 | 26838.155 | 10.1976 | 4.9865
672 | 26878.152 | 10.1991 | 4.9824 |  |  | 672 | 26878.152 | 10.1991 | 4.9804 |  |  | 672 | 26878.152 | 10.1991 | 4.9864
673 | 26918.15 | 10.2006 | 4.9822 |  |  | 673 | 26918.15 | 10.2006 | 4.9803 |  |  | 673 | 26918.15 | 10.2006 | 4.9863
674 | 26958.147 | 10.202 | 4.9821 |  |  | 674 | 26958.147 | 10.202 | 4.9802 |  |  | 674 | 26958.147 | 10.202 | 4.9862
675 | 26998.144 | 10.2035 | 4.982 |  |  | 675 | 26998.144 | 10.2035 | 4.9801 |  |  | 675 | 26998.144 | 10.2035 | 4.9861
676 | 27038.141 | 10.205 | 4.9819 |  |  | 676 | 27038.141 | 10.205 | 4.98 |  |  | 676 | 27038.141 | 10.205 | 4.986
677 | 27078.139 | 10.2065 | 4.9817 |  |  | 677 | 27078.139 | 10.2065 | 4.9798 |  |  | 677 | 27078.139 | 10.2065 | 4.9859
678 | 27118.136 | 10.208 | 4.9816 |  |  | 678 | 27118.136 | 10.208 | 4.9797 |  |  | 678 | 27118.136 | 10.208 | 4.9857
679 | 27158.133 | 10.2094 | 4.9815 |  |  | 679 | 27158.133 | 10.2094 | 4.9796 |  |  | 679 | 27158.133 | 10.2094 | 4.9856
680 | 27198.13 | 10.2109 | 4.9814 |  |  | 680 | 27198.13 | 10.2109 | 4.9795 |  |  | 680 | 27198.13 | 10.2109 | 4.9855
681 | 27238.128 | 10.2124 | 4.9813 |  |  | 681 | 27238.128 | 10.2124 | 4.9794 |  |  | 681 | 27238.128 | 10.2124 | 4.9854
682 | 27278.125 | 10.2138 | 4.9811 |  |  | 682 | 27278.125 | 10.2138 | 4.9792 |  |  | 682 | 27278.125 | 10.2138 | 4.9853
683 | 27318.122 | 10.2153 | 4.981 |  |  | 683 | 27318.122 | 10.2153 | 4.9791 |  |  | 683 | 27318.122 | 10.2153 | 4.9852
684 | 27358.119 | 10.2168 | 4.9809 |  |  | 684 | 27358.119 | 10.2168 | 4.979 |  |  | 684 | 27358.119 | 10.2168 | 4.9851
685 | 27398.117 | 10.2182 | 4.9808 |  |  | 685 | 27398.117 | 10.2182 | 4.9789 |  |  | 685 | 27398.117 | 10.2182 | 4.985
686 | 27438.114 | 10.2197 | 4.9807 |  |  | 686 | 27438.114 | 10.2197 | 4.9788 |  |  | 686 | 27438.114 | 10.2197 | 4.9848
687 | 27478.111 | 10.2211 | 4.9805 |  |  | 687 | 27478.111 | 10.2211 | 4.9786 |  |  | 687 | 27478.111 | 10.2211 | 4.9847
688 | 27518.108 | 10.2226 | 4.9804 |  |  | 688 | 27518.108 | 10.2226 | 4.9785 |  |  | 688 | 27518.108 | 10.2226 | 4.9846
689 | 27558.106 | 10.2241 | 4.9803 |  |  | 689 | 27558.106 | 10.2241 | 4.9784 |  |  | 689 | 27558.106 | 10.2241 | 4.9845
690 | 27598.103 | 10.2255 | 4.9802 |  |  | 690 | 27598.103 | 10.2255 | 4.9783 |  |  | 690 | 27598.103 | 10.2255 | 4.9844
691 | 27638.1 | 10.227 | 4.9801 |  |  | 691 | 27638.1 | 10.227 | 4.9782 |  |  | 691 | 27638.1 | 10.227 | 4.9843
692 | 27678.097 | 10.2284 | 4.9799 |  |  | 692 | 27678.097 | 10.2284 | 4.978 |  |  | 692 | 27678.097 | 10.2284 | 4.9842
693 | 27718.095 | 10.2298 | 4.9798 |  |  | 693 | 27718.095 | 10.2298 | 4.9779 |  |  | 693 | 27718.095 | 10.2298 | 4.9841
694 | 27758.092 | 10.2313 | 4.9797 |  |  | 694 | 27758.092 | 10.2313 | 4.9778 |  |  | 694 | 27758.092 | 10.2313 | 4.9839
695 | 27798.089 | 10.2327 | 4.9796 |  |  | 695 | 27798.089 | 10.2327 | 4.9777 |  |  | 695 | 27798.089 | 10.2327 | 4.9838
696 | 27838.086 | 10.2342 | 4.9795 |  |  | 696 | 27838.086 | 10.2342 | 4.9776 |  |  | 696 | 27838.086 | 10.2342 | 4.9837
697 | 27878.084 | 10.2356 | 4.9793 |  |  | 697 | 27878.084 | 10.2356 | 4.9775 |  |  | 697 | 27878.084 | 10.2356 | 4.9836
698 | 27918.081 | 10.237 | 4.9792 |  |  | 698 | 27918.081 | 10.237 | 4.9773 |  |  | 698 | 27918.081 | 10.237 | 4.9835
699 | 27958.078 | 10.2385 | 4.9791 |  |  | 699 | 27958.078 | 10.2385 | 4.9772 |  |  | 699 | 27958.078 | 10.2385 | 4.9834
700 | 27998.075 | 10.2399 | 4.979 |  |  | 700 | 27998.075 | 10.2399 | 4.9771 |  |  | 700 | 27998.075 | 10.2399 | 4.9833
701 | 28038.073 | 10.2413 | 4.9789 |  |  | 701 | 28038.073 | 10.2413 | 4.977 |  |  | 701 | 28038.073 | 10.2413 | 4.9832
702 | 28078.07 | 10.2427 | 4.9787 |  |  | 702 | 28078.07 | 10.2427 | 4.9769 |  |  | 702 | 28078.07 | 10.2427 | 4.9831
703 | 28118.067 | 10.2442 | 4.9786 |  |  | 703 | 28118.067 | 10.2442 | 4.9767 |  |  | 703 | 28118.067 | 10.2442 | 4.9829
704 | 28158.064 | 10.2456 | 4.9785 |  |  | 704 | 28158.064 | 10.2456 | 4.9766 |  |  | 704 | 28158.064 | 10.2456 | 4.9828
705 | 28198.062 | 10.247 | 4.9784 |  |  | 705 | 28198.062 | 10.247 | 4.9765 |  |  | 705 | 28198.062 | 10.247 | 4.9827
706 | 28238.059 | 10.2484 | 4.9783 |  |  | 706 | 28238.059 | 10.2484 | 4.9764 |  |  | 706 | 28238.059 | 10.2484 | 4.9826
707 | 28278.056 | 10.2498 | 4.9782 |  |  | 707 | 28278.056 | 10.2498 | 4.9763 |  |  | 707 | 28278.056 | 10.2498 | 4.9825
708 | 28318.053 | 10.2513 | 4.978 |  |  | 708 | 28318.053 | 10.2513 | 4.9762 |  |  | 708 | 28318.053 | 10.2513 | 4.9824
709 | 28358.051 | 10.2527 | 4.9779 |  |  | 709 | 28358.051 | 10.2527 | 4.976 |  |  | 709 | 28358.051 | 10.2527 | 4.9823
710 | 28398.048 | 10.2541 | 4.9778 |  |  | 710 | 28398.048 | 10.2541 | 4.9759 |  |  | 710 | 28398.048 | 10.2541 | 4.9822
711 | 28438.045 | 10.2555 | 4.9777 |  |  | 711 | 28438.045 | 10.2555 | 4.9758 |  |  | 711 | 28438.045 | 10.2555 | 4.9821
712 | 28478.042 | 10.2569 | 4.9776 |  |  | 712 | 28478.042 | 10.2569 | 4.9757 |  |  | 712 | 28478.042 | 10.2569 | 4.982
713 | 28518.04 | 10.2583 | 4.9774 |  |  | 713 | 28518.04 | 10.2583 | 4.9756 |  |  | 713 | 28518.04 | 10.2583 | 4.9819
714 | 28558.037 | 10.2597 | 4.9773 |  |  | 714 | 28558.037 | 10.2597 | 4.9755 |  |  | 714 | 28558.037 | 10.2597 | 4.9817
715 | 28598.034 | 10.2611 | 4.9772 |  |  | 715 | 28598.034 | 10.2611 | 4.9753 |  |  | 715 | 28598.034 | 10.2611 | 4.9816
716 | 28638.031 | 10.2625 | 4.9771 |  |  | 716 | 28638.031 | 10.2625 | 4.9752 |  |  | 716 | 28638.031 | 10.2625 | 4.9815
717 | 28678.029 | 10.2639 | 4.977 |  |  | 717 | 28678.029 | 10.2639 | 4.9751 |  |  | 717 | 28678.029 | 10.2639 | 4.9814
718 | 28718.026 | 10.2653 | 4.9769 |  |  | 718 | 28718.026 | 10.2653 | 4.975 |  |  | 718 | 28718.026 | 10.2653 | 4.9813
719 | 28758.023 | 10.2667 | 4.9767 |  |  | 719 | 28758.023 | 10.2667 | 4.9749 |  |  | 719 | 28758.023 | 10.2667 | 4.9812
720 | 28798.02 | 10.2681 | 4.9766 |  |  | 720 | 28798.02 | 10.2681 | 4.9748 |  |  | 720 | 28798.02 | 10.2681 | 4.9811
721 | 28838.018 | 10.2694 | 4.9765 |  |  | 721 | 28838.018 | 10.2694 | 4.9746 |  |  | 721 | 28838.018 | 10.2694 | 4.981
722 | 28878.015 | 10.2708 | 4.9764 |  |  | 722 | 28878.015 | 10.2708 | 4.9745 |  |  | 722 | 28878.015 | 10.2708 | 4.9809
723 | 28918.012 | 10.2722 | 4.9763 |  |  | 723 | 28918.012 | 10.2722 | 4.9744 |  |  | 723 | 28918.012 | 10.2722 | 4.9808
724 | 28958.009 | 10.2736 | 4.9762 |  |  | 724 | 28958.009 | 10.2736 | 4.9743 |  |  | 724 | 28958.009 | 10.2736 | 4.9807
725 | 28998.007 | 10.275 | 4.976 |  |  | 725 | 28998.007 | 10.275 | 4.9742 |  |  | 725 | 28998.007 | 10.275 | 4.9806
726 | 29038.004 | 10.2764 | 4.9759 |  |  | 726 | 29038.004 | 10.2764 | 4.9741 |  |  | 726 | 29038.004 | 10.2764 | 4.9805
727 | 29078.001 | 10.2777 | 4.9758 |  |  | 727 | 29078.001 | 10.2777 | 4.9739 |  |  | 727 | 29078.001 | 10.2777 | 4.9803
728 | 29117.998 | 10.2791 | 4.9757 |  |  | 728 | 29117.998 | 10.2791 | 4.9738 |  |  | 728 | 29117.998 | 10.2791 | 4.9802
729 | 29157.996 | 10.2805 | 4.9756 |  |  | 729 | 29157.996 | 10.2805 | 4.9737 |  |  | 729 | 29157.996 | 10.2805 | 4.9801
730 | 29197.993 | 10.2819 | 4.9755 |  |  | 730 | 29197.993 | 10.2819 | 4.9736 |  |  | 730 | 29197.993 | 10.2819 | 4.98
731 | 29237.99 | 10.2832 | 4.9753 |  |  | 731 | 29237.99 | 10.2832 | 4.9735 |  |  | 731 | 29237.99 | 10.2832 | 4.9799
732 | 29277.987 | 10.2846 | 4.9752 |  |  | 732 | 29277.987 | 10.2846 | 4.9734 |  |  | 732 | 29277.987 | 10.2846 | 4.9798
733 | 29317.985 | 10.286 | 4.9751 |  |  | 733 | 29317.985 | 10.286 | 4.9732 |  |  | 733 | 29317.985 | 10.286 | 4.9797
734 | 29357.982 | 10.2873 | 4.975 |  |  | 734 | 29357.982 | 10.2873 | 4.9731 |  |  | 734 | 29357.982 | 10.2873 | 4.9796
735 | 29397.979 | 10.2887 | 4.9749 |  |  | 735 | 29397.979 | 10.2887 | 4.973 |  |  | 735 | 29397.979 | 10.2887 | 4.9795
736 | 29437.976 | 10.29 | 4.9748 |  |  | 736 | 29437.976 | 10.29 | 4.9729 |  |  | 736 | 29437.976 | 10.29 | 4.9794
737 | 29477.974 | 10.2914 | 4.9747 |  |  | 737 | 29477.974 | 10.2914 | 4.9728 |  |  | 737 | 29477.974 | 10.2914 | 4.9793
738 | 29517.971 | 10.2928 | 4.9745 |  |  | 738 | 29517.971 | 10.2928 | 4.9727 |  |  | 738 | 29517.971 | 10.2928 | 4.9792
739 | 29557.968 | 10.2941 | 4.9744 |  |  | 739 | 29557.968 | 10.2941 | 4.9726 |  |  | 739 | 29557.968 | 10.2941 | 4.9791
740 | 29597.965 | 10.2955 | 4.9743 |  |  | 740 | 29597.965 | 10.2955 | 4.9724 |  |  | 740 | 29597.965 | 10.2955 | 4.979
741 | 29637.963 | 10.2968 | 4.9742 |  |  | 741 | 29637.963 | 10.2968 | 4.9723 |  |  | 741 | 29637.963 | 10.2968 | 4.9789
742 | 29677.96 | 10.2982 | 4.9741 |  |  | 742 | 29677.96 | 10.2982 | 4.9722 |  |  | 742 | 29677.96 | 10.2982 | 4.9788
743 | 29717.957 | 10.2995 | 4.974 |  |  | 743 | 29717.957 | 10.2995 | 4.9721 |  |  | 743 | 29717.957 | 10.2995 | 4.9786
744 | 29757.954 | 10.3009 | 4.9738 |  |  | 744 | 29757.954 | 10.3009 | 4.972 |  |  | 744 | 29757.954 | 10.3009 | 4.9785
745 | 29797.952 | 10.3022 | 4.9737 |  |  | 745 | 29797.952 | 10.3022 | 4.9719 |  |  | 745 | 29797.952 | 10.3022 | 4.9784
746 | 29837.949 | 10.3035 | 4.9736 |  |  | 746 | 29837.949 | 10.3035 | 4.9717 |  |  | 746 | 29837.949 | 10.3035 | 4.9783
747 | 29877.946 | 10.3049 | 4.9735 |  |  | 747 | 29877.946 | 10.3049 | 4.9716 |  |  | 747 | 29877.946 | 10.3049 | 4.9782
748 | 29917.943 | 10.3062 | 4.9734 |  |  | 748 | 29917.943 | 10.3062 | 4.9715 |  |  | 748 | 29917.943 | 10.3062 | 4.9781
749 | 29957.941 | 10.3075 | 4.9733 |  |  | 749 | 29957.941 | 10.3075 | 4.9714 |  |  | 749 | 29957.941 | 10.3075 | 4.978
750 | 29997.938 | 10.3089 | 4.9732 |  |  | 750 | 29997.938 | 10.3089 | 4.9713 |  |  | 750 | 29997.938 | 10.3089 | 4.9779
751 | 30037.935 | 10.3102 | 4.973 |  |  | 751 | 30037.935 | 10.3102 | 4.9712 |  |  | 751 | 30037.935 | 10.3102 | 4.9778
752 | 30077.932 | 10.3115 | 4.9729 |  |  | 752 | 30077.932 | 10.3115 | 4.9711 |  |  | 752 | 30077.932 | 10.3115 | 4.9777
753 | 30117.93 | 10.3129 | 4.9728 |  |  | 753 | 30117.93 | 10.3129 | 4.9709 |  |  | 753 | 30117.93 | 10.3129 | 4.9776
754 | 30157.927 | 10.3142 | 4.9727 |  |  | 754 | 30157.927 | 10.3142 | 4.9708 |  |  | 754 | 30157.927 | 10.3142 | 4.9775
755 | 30197.924 | 10.3155 | 4.9726 |  |  | 755 | 30197.924 | 10.3155 | 4.9707 |  |  | 755 | 30197.924 | 10.3155 | 4.9774
756 | 30237.921 | 10.3169 | 4.9725 |  |  | 756 | 30237.921 | 10.3169 | 4.9706 |  |  | 756 | 30237.921 | 10.3169 | 4.9773
757 | 30277.919 | 10.3182 | 4.9724 |  |  | 757 | 30277.919 | 10.3182 | 4.9705 |  |  | 757 | 30277.919 | 10.3182 | 4.9772
758 | 30317.916 | 10.3195 | 4.9722 |  |  | 758 | 30317.916 | 10.3195 | 4.9704 |  |  | 758 | 30317.916 | 10.3195 | 4.9771
759 | 30357.913 | 10.3208 | 4.9721 |  |  | 759 | 30357.913 | 10.3208 | 4.9703 |  |  | 759 | 30357.913 | 10.3208 | 4.977
760 | 30397.91 | 10.3221 | 4.972 |  |  | 760 | 30397.91 | 10.3221 | 4.9701 |  |  | 760 | 30397.91 | 10.3221 | 4.9769
761 | 30437.908 | 10.3234 | 4.9719 |  |  | 761 | 30437.908 | 10.3234 | 4.97 |  |  | 761 | 30437.908 | 10.3234 | 4.9768
762 | 30477.905 | 10.3248 | 4.9718 |  |  | 762 | 30477.905 | 10.3248 | 4.9699 |  |  | 762 | 30477.905 | 10.3248 | 4.9767
763 | 30517.902 | 10.3261 | 4.9717 |  |  | 763 | 30517.902 | 10.3261 | 4.9698 |  |  | 763 | 30517.902 | 10.3261 | 4.9766
764 | 30557.899 | 10.3274 | 4.9716 |  |  | 764 | 30557.899 | 10.3274 | 4.9697 |  |  | 764 | 30557.899 | 10.3274 | 4.9765
765 | 30597.897 | 10.3287 | 4.9714 |  |  | 765 | 30597.897 | 10.3287 | 4.9696 |  |  | 765 | 30597.897 | 10.3287 | 4.9763
766 | 30637.894 | 10.33 | 4.9713 |  |  | 766 | 30637.894 | 10.33 | 4.9695 |  |  | 766 | 30637.894 | 10.33 | 4.9762
767 | 30677.891 | 10.3313 | 4.9712 |  |  | 767 | 30677.891 | 10.3313 | 4.9693 |  |  | 767 | 30677.891 | 10.3313 | 4.9761
768 | 30717.888 | 10.3326 | 4.9711 |  |  | 768 | 30717.888 | 10.3326 | 4.9692 |  |  | 768 | 30717.888 | 10.3326 | 4.976
769 | 30757.886 | 10.3339 | 4.971 |  |  | 769 | 30757.886 | 10.3339 | 4.9691 |  |  | 769 | 30757.886 | 10.3339 | 4.9759
770 | 30797.883 | 10.3352 | 4.9709 |  |  | 770 | 30797.883 | 10.3352 | 4.969 |  |  | 770 | 30797.883 | 10.3352 | 4.9758
771 | 30837.88 | 10.3365 | 4.9708 |  |  | 771 | 30837.88 | 10.3365 | 4.9689 |  |  | 771 | 30837.88 | 10.3365 | 4.9757
772 | 30877.877 | 10.3378 | 4.9706 |  |  | 772 | 30877.877 | 10.3378 | 4.9688 |  |  | 772 | 30877.877 | 10.3378 | 4.9756
773 | 30917.875 | 10.3391 | 4.9705 |  |  | 773 | 30917.875 | 10.3391 | 4.9687 |  |  | 773 | 30917.875 | 10.3391 | 4.9755
774 | 30957.872 | 10.3404 | 4.9704 |  |  | 774 | 30957.872 | 10.3404 | 4.9685 |  |  | 774 | 30957.872 | 10.3404 | 4.9754
775 | 30997.869 | 10.3417 | 4.9703 |  |  | 775 | 30997.869 | 10.3417 | 4.9684 |  |  | 775 | 30997.869 | 10.3417 | 4.9753
776 | 31037.866 | 10.343 | 4.9702 |  |  | 776 | 31037.866 | 10.343 | 4.9683 |  |  | 776 | 31037.866 | 10.343 | 4.9752
777 | 31077.864 | 10.3443 | 4.9701 |  |  | 777 | 31077.864 | 10.3443 | 4.9682 |  |  | 777 | 31077.864 | 10.3443 | 4.9751
778 | 31117.861 | 10.3455 | 4.97 |  |  | 778 | 31117.861 | 10.3455 | 4.9681 |  |  | 778 | 31117.861 | 10.3455 | 4.975
779 | 31157.858 | 10.3468 | 4.9698 |  |  | 779 | 31157.858 | 10.3468 | 4.968 |  |  | 779 | 31157.858 | 10.3468 | 4.9749
780 | 31197.855 | 10.3481 | 4.9697 |  |  | 780 | 31197.855 | 10.3481 | 4.9679 |  |  | 780 | 31197.855 | 10.3481 | 4.9748
781 | 31237.853 | 10.3494 | 4.9696 |  |  | 781 | 31237.853 | 10.3494 | 4.9677 |  |  | 781 | 31237.853 | 10.3494 | 4.9747
782 | 31277.85 | 10.3507 | 4.9695 |  |  | 782 | 31277.85 | 10.3507 | 4.9676 |  |  | 782 | 31277.85 | 10.3507 | 4.9746
783 | 31317.847 | 10.3519 | 4.9694 |  |  | 783 | 31317.847 | 10.3519 | 4.9675 |  |  | 783 | 31317.847 | 10.3519 | 4.9745
784 | 31357.844 | 10.3532 | 4.9693 |  |  | 784 | 31357.844 | 10.3532 | 4.9674 |  |  | 784 | 31357.844 | 10.3532 | 4.9744
785 | 31397.842 | 10.3545 | 4.9692 |  |  | 785 | 31397.842 | 10.3545 | 4.9673 |  |  | 785 | 31397.842 | 10.3545 | 4.9743
786 | 31437.839 | 10.3558 | 4.969 |  |  | 786 | 31437.839 | 10.3558 | 4.9672 |  |  | 786 | 31437.839 | 10.3558 | 4.9742
787 | 31477.836 | 10.357 | 4.9689 |  |  | 787 | 31477.836 | 10.357 | 4.9671 |  |  | 787 | 31477.836 | 10.357 | 4.9741
788 | 31517.833 | 10.3583 | 4.9688 |  |  | 788 | 31517.833 | 10.3583 | 4.9669 |  |  | 788 | 31517.833 | 10.3583 | 4.974
789 | 31557.831 | 10.3596 | 4.9687 |  |  | 789 | 31557.831 | 10.3596 | 4.9668 |  |  | 789 | 31557.831 | 10.3596 | 4.9739
790 | 31597.828 | 10.3608 | 4.9686 |  |  | 790 | 31597.828 | 10.3608 | 4.9667 |  |  | 790 | 31597.828 | 10.3608 | 4.9738
791 | 31637.825 | 10.3621 | 4.9685 |  |  | 791 | 31637.825 | 10.3621 | 4.9666 |  |  | 791 | 31637.825 | 10.3621 | 4.9737
792 | 31677.822 | 10.3634 | 4.9684 |  |  | 792 | 31677.822 | 10.3634 | 4.9665 |  |  | 792 | 31677.822 | 10.3634 | 4.9736
793 | 31717.82 | 10.3646 | 4.9682 |  |  | 793 | 31717.82 | 10.3646 | 4.9664 |  |  | 793 | 31717.82 | 10.3646 | 4.9735
794 | 31757.817 | 10.3659 | 4.9681 |  |  | 794 | 31757.817 | 10.3659 | 4.9663 |  |  | 794 | 31757.817 | 10.3659 | 4.9734
795 | 31797.814 | 10.3672 | 4.968 |  |  | 795 | 31797.814 | 10.3672 | 4.9661 |  |  | 795 | 31797.814 | 10.3672 | 4.9733
796 | 31837.811 | 10.3684 | 4.9679 |  |  | 796 | 31837.811 | 10.3684 | 4.966 |  |  | 796 | 31837.811 | 10.3684 | 4.9732
797 | 31877.809 | 10.3697 | 4.9678 |  |  | 797 | 31877.809 | 10.3697 | 4.9659 |  |  | 797 | 31877.809 | 10.3697 | 4.9731
798 | 31917.806 | 10.3709 | 4.9677 |  |  | 798 | 31917.806 | 10.3709 | 4.9658 |  |  | 798 | 31917.806 | 10.3709 | 4.9729
799 | 31957.803 | 10.3722 | 4.9676 |  |  | 799 | 31957.803 | 10.3722 | 4.9657 |  |  | 799 | 31957.803 | 10.3722 | 4.9728
800 | 31997.8 | 10.3734 | 4.9675 |  |  | 800 | 31997.8 | 10.3734 | 4.9656 |  |  | 800 | 31997.8 | 10.3734 | 4.9727
801 | 32037.798 | 10.3747 | 4.9673 |  |  | 801 | 32037.798 | 10.3747 | 4.9655 |  |  | 801 | 32037.798 | 10.3747 | 4.9726
802 | 32077.795 | 10.3759 | 4.9672 |  |  | 802 | 32077.795 | 10.3759 | 4.9653 |  |  | 802 | 32077.795 | 10.3759 | 4.9725
803 | 32117.792 | 10.3772 | 4.9671 |  |  | 803 | 32117.792 | 10.3772 | 4.9652 |  |  | 803 | 32117.792 | 10.3772 | 4.9724
804 | 32157.789 | 10.3784 | 4.967 |  |  | 804 | 32157.789 | 10.3784 | 4.9651 |  |  | 804 | 32157.789 | 10.3784 | 4.9723
805 | 32197.787 | 10.3797 | 4.9669 |  |  | 805 | 32197.787 | 10.3797 | 4.965 |  |  | 805 | 32197.787 | 10.3797 | 4.9722
806 | 32237.784 | 10.3809 | 4.9668 |  |  | 806 | 32237.784 | 10.3809 | 4.9649 |  |  | 806 | 32237.784 | 10.3809 | 4.9721
807 | 32277.781 | 10.3821 | 4.9667 |  |  | 807 | 32277.781 | 10.3821 | 4.9648 |  |  | 807 | 32277.781 | 10.3821 | 4.972
808 | 32317.778 | 10.3834 | 4.9665 |  |  | 808 | 32317.778 | 10.3834 | 4.9647 |  |  | 808 | 32317.778 | 10.3834 | 4.9719
809 | 32357.776 | 10.3846 | 4.9664 |  |  | 809 | 32357.776 | 10.3846 | 4.9646 |  |  | 809 | 32357.776 | 10.3846 | 4.9718
810 | 32397.773 | 10.3858 | 4.9663 |  |  | 810 | 32397.773 | 10.3858 | 4.9644 |  |  | 810 | 32397.773 | 10.3858 | 4.9717
811 | 32437.77 | 10.3871 | 4.9662 |  |  | 811 | 32437.77 | 10.3871 | 4.9643 |  |  | 811 | 32437.77 | 10.3871 | 4.9716
812 | 32477.767 | 10.3883 | 4.9661 |  |  | 812 | 32477.767 | 10.3883 | 4.9642 |  |  | 812 | 32477.767 | 10.3883 | 4.9715
813 | 32517.765 | 10.3895 | 4.966 |  |  | 813 | 32517.765 | 10.3895 | 4.9641 |  |  | 813 | 32517.765 | 10.3895 | 4.9714
814 | 32557.762 | 10.3908 | 4.9659 |  |  | 814 | 32557.762 | 10.3908 | 4.964 |  |  | 814 | 32557.762 | 10.3908 | 4.9713
815 | 32597.759 | 10.392 | 4.9657 |  |  | 815 | 32597.759 | 10.392 | 4.9639 |  |  | 815 | 32597.759 | 10.392 | 4.9712
816 | 32637.756 | 10.3932 | 4.9656 |  |  | 816 | 32637.756 | 10.3932 | 4.9638 |  |  | 816 | 32637.756 | 10.3932 | 4.9711
817 | 32677.754 | 10.3944 | 4.9655 |  |  | 817 | 32677.754 | 10.3944 | 4.9636 |  |  | 817 | 32677.754 | 10.3944 | 4.971
818 | 32717.751 | 10.3957 | 4.9654 |  |  | 818 | 32717.751 | 10.3957 | 4.9635 |  |  | 818 | 32717.751 | 10.3957 | 4.9709
819 | 32757.748 | 10.3969 | 4.9653 |  |  | 819 | 32757.748 | 10.3969 | 4.9634 |  |  | 819 | 32757.748 | 10.3969 | 4.9708
820 | 32797.745 | 10.3981 | 4.9652 |  |  | 820 | 32797.745 | 10.3981 | 4.9633 |  |  | 820 | 32797.745 | 10.3981 | 4.9707
821 | 32837.743 | 10.3993 | 4.9651 |  |  | 821 | 32837.743 | 10.3993 | 4.9632 |  |  | 821 | 32837.743 | 10.3993 | 4.9706
822 | 32877.74 | 10.4006 | 4.965 |  |  | 822 | 32877.74 | 10.4006 | 4.9631 |  |  | 822 | 32877.74 | 10.4006 | 4.9705
823 | 32917.737 | 10.4018 | 4.9648 |  |  | 823 | 32917.737 | 10.4018 | 4.963 |  |  | 823 | 32917.737 | 10.4018 | 4.9704
824 | 32957.734 | 10.403 | 4.9647 |  |  | 824 | 32957.734 | 10.403 | 4.9628 |  |  | 824 | 32957.734 | 10.403 | 4.9703
825 | 32997.732 | 10.4042 | 4.9646 |  |  | 825 | 32997.732 | 10.4042 | 4.9627 |  |  | 825 | 32997.732 | 10.4042 | 4.9702
826 | 33037.729 | 10.4054 | 4.9645 |  |  | 826 | 33037.729 | 10.4054 | 4.9626 |  |  | 826 | 33037.729 | 10.4054 | 4.9701
827 | 33077.726 | 10.4066 | 4.9644 |  |  | 827 | 33077.726 | 10.4066 | 4.9625 |  |  | 827 | 33077.726 | 10.4066 | 4.97
828 | 33117.723 | 10.4078 | 4.9643 |  |  | 828 | 33117.723 | 10.4078 | 4.9624 |  |  | 828 | 33117.723 | 10.4078 | 4.9699
829 | 33157.721 | 10.409 | 4.9642 |  |  | 829 | 33157.721 | 10.409 | 4.9623 |  |  | 829 | 33157.721 | 10.409 | 4.9698
830 | 33197.718 | 10.4102 | 4.964 |  |  | 830 | 33197.718 | 10.4102 | 4.9622 |  |  | 830 | 33197.718 | 10.4102 | 4.9697
831 | 33237.715 | 10.4114 | 4.9639 |  |  | 831 | 33237.715 | 10.4114 | 4.962 |  |  | 831 | 33237.715 | 10.4114 | 4.9696
832 | 33277.712 | 10.4126 | 4.9638 |  |  | 832 | 33277.712 | 10.4126 | 4.9619 |  |  | 832 | 33277.712 | 10.4126 | 4.9695
833 | 33317.71 | 10.4138 | 4.9637 |  |  | 833 | 33317.71 | 10.4138 | 4.9618 |  |  | 833 | 33317.71 | 10.4138 | 4.9694
834 | 33357.707 | 10.415 | 4.9636 |  |  | 834 | 33357.707 | 10.415 | 4.9617 |  |  | 834 | 33357.707 | 10.415 | 4.9693
835 | 33397.704 | 10.4162 | 4.9635 |  |  | 835 | 33397.704 | 10.4162 | 4.9616 |  |  | 835 | 33397.704 | 10.4162 | 4.9692
836 | 33437.701 | 10.4174 | 4.9634 |  |  | 836 | 33437.701 | 10.4174 | 4.9615 |  |  | 836 | 33437.701 | 10.4174 | 4.9691
837 | 33477.699 | 10.4186 | 4.9632 |  |  | 837 | 33477.699 | 10.4186 | 4.9613 |  |  | 837 | 33477.699 | 10.4186 | 4.969
838 | 33517.696 | 10.4198 | 4.9631 |  |  | 838 | 33517.696 | 10.4198 | 4.9612 |  |  | 838 | 33517.696 | 10.4198 | 4.9689
839 | 33557.693 | 10.421 | 4.963 |  |  | 839 | 33557.693 | 10.421 | 4.9611 |  |  | 839 | 33557.693 | 10.421 | 4.9688
840 | 33597.69 | 10.4222 | 4.9629 |  |  | 840 | 33597.69 | 10.4222 | 4.961 |  |  | 840 | 33597.69 | 10.4222 | 4.9687
841 | 33637.688 | 10.4234 | 4.9628 |  |  | 841 | 33637.688 | 10.4234 | 4.9609 |  |  | 841 | 33637.688 | 10.4234 | 4.9686
842 | 33677.685 | 10.4246 | 4.9627 |  |  | 842 | 33677.685 | 10.4246 | 4.9608 |  |  | 842 | 33677.685 | 10.4246 | 4.9685
843 | 33717.682 | 10.4258 | 4.9626 |  |  | 843 | 33717.682 | 10.4258 | 4.9607 |  |  | 843 | 33717.682 | 10.4258 | 4.9684
844 | 33757.679 | 10.427 | 4.9624 |  |  | 844 | 33757.679 | 10.427 | 4.9605 |  |  | 844 | 33757.679 | 10.427 | 4.9683
845 | 33797.677 | 10.4281 | 4.9623 |  |  | 845 | 33797.677 | 10.4281 | 4.9604 |  |  | 845 | 33797.677 | 10.4281 | 4.9682
846 | 33837.674 | 10.4293 | 4.9622 |  |  | 846 | 33837.674 | 10.4293 | 4.9603 |  |  | 846 | 33837.674 | 10.4293 | 4.9681
847 | 33877.671 | 10.4305 | 4.9621 |  |  | 847 | 33877.671 | 10.4305 | 4.9602 |  |  | 847 | 33877.671 | 10.4305 | 4.968
848 | 33917.668 | 10.4317 | 4.962 |  |  | 848 | 33917.668 | 10.4317 | 4.9601 |  |  | 848 | 33917.668 | 10.4317 | 4.9679
849 | 33957.666 | 10.4329 | 4.9619 |  |  | 849 | 33957.666 | 10.4329 | 4.96 |  |  | 849 | 33957.666 | 10.4329 | 4.9678
850 | 33997.663 | 10.434 | 4.9618 |  |  | 850 | 33997.663 | 10.434 | 4.9599 |  |  | 850 | 33997.663 | 10.434 | 4.9677
851 | 34037.66 | 10.4352 | 4.9616 |  |  | 851 | 34037.66 | 10.4352 | 4.9597 |  |  | 851 | 34037.66 | 10.4352 | 4.9676
852 | 34077.657 | 10.4364 | 4.9615 |  |  | 852 | 34077.657 | 10.4364 | 4.9596 |  |  | 852 | 34077.657 | 10.4364 | 4.9675
853 | 34117.655 | 10.4376 | 4.9614 |  |  | 853 | 34117.655 | 10.4376 | 4.9595 |  |  | 853 | 34117.655 | 10.4376 | 4.9674
854 | 34157.652 | 10.4387 | 4.9613 |  |  | 854 | 34157.652 | 10.4387 | 4.9594 |  |  | 854 | 34157.652 | 10.4387 | 4.9673
855 | 34197.649 | 10.4399 | 4.9612 |  |  | 855 | 34197.649 | 10.4399 | 4.9593 |  |  | 855 | 34197.649 | 10.4399 | 4.9672
856 | 34237.646 | 10.4411 | 4.9611 |  |  | 856 | 34237.646 | 10.4411 | 4.9592 |  |  | 856 | 34237.646 | 10.4411 | 4.967
857 | 34277.644 | 10.4422 | 4.961 |  |  | 857 | 34277.644 | 10.4422 | 4.959 |  |  | 857 | 34277.644 | 10.4422 | 4.9669
858 | 34317.641 | 10.4434 | 4.9609 |  |  | 858 | 34317.641 | 10.4434 | 4.9589 |  |  | 858 | 34317.641 | 10.4434 | 4.9668
859 | 34357.638 | 10.4446 | 4.9607 |  |  | 859 | 34357.638 | 10.4446 | 4.9588 |  |  | 859 | 34357.638 | 10.4446 | 4.9667
860 | 34397.635 | 10.4457 | 4.9606 |  |  | 860 | 34397.635 | 10.4457 | 4.9587 |  |  | 860 | 34397.635 | 10.4457 | 4.9666
861 | 34437.633 | 10.4469 | 4.9605 |  |  | 861 | 34437.633 | 10.4469 | 4.9586 |  |  | 861 | 34437.633 | 10.4469 | 4.9665
862 | 34477.63 | 10.4481 | 4.9604 |  |  | 862 | 34477.63 | 10.4481 | 4.9585 |  |  | 862 | 34477.63 | 10.4481 | 4.9664
863 | 34517.627 | 10.4492 | 4.9603 |  |  | 863 | 34517.627 | 10.4492 | 4.9584 |  |  | 863 | 34517.627 | 10.4492 | 4.9663
864 | 34557.624 | 10.4504 | 4.9602 |  |  | 864 | 34557.624 | 10.4504 | 4.9582 |  |  | 864 | 34557.624 | 10.4504 | 4.9662
865 | 34597.622 | 10.4515 | 4.96 |  |  | 865 | 34597.622 | 10.4515 | 4.9581 |  |  | 865 | 34597.622 | 10.4515 | 4.9661
866 | 34637.619 | 10.4527 | 4.9599 |  |  | 866 | 34637.619 | 10.4527 | 4.958 |  |  | 866 | 34637.619 | 10.4527 | 4.966
867 | 34677.616 | 10.4538 | 4.9598 |  |  | 867 | 34677.616 | 10.4538 | 4.9579 |  |  | 867 | 34677.616 | 10.4538 | 4.9659
868 | 34717.613 | 10.455 | 4.9597 |  |  | 868 | 34717.613 | 10.455 | 4.9578 |  |  | 868 | 34717.613 | 10.455 | 4.9658
869 | 34757.611 | 10.4562 | 4.9596 |  |  | 869 | 34757.611 | 10.4562 | 4.9577 |  |  | 869 | 34757.611 | 10.4562 | 4.9657
870 | 34797.608 | 10.4573 | 4.9595 |  |  | 870 | 34797.608 | 10.4573 | 4.9575 |  |  | 870 | 34797.608 | 10.4573 | 4.9656
871 | 34837.605 | 10.4585 | 4.9594 |  |  | 871 | 34837.605 | 10.4585 | 4.9574 |  |  | 871 | 34837.605 | 10.4585 | 4.9655
872 | 34877.602 | 10.4596 | 4.9592 |  |  | 872 | 34877.602 | 10.4596 | 4.9573 |  |  | 872 | 34877.602 | 10.4596 | 4.9654
873 | 34917.6 | 10.4607 | 4.9591 |  |  | 873 | 34917.6 | 10.4607 | 4.9572 |  |  | 873 | 34917.6 | 10.4607 | 4.9653
874 | 34957.597 | 10.4619 | 4.959 |  |  | 874 | 34957.597 | 10.4619 | 4.9571 |  |  | 874 | 34957.597 | 10.4619 | 4.9652
875 | 34997.594 | 10.463 | 4.9589 |  |  | 875 | 34997.594 | 10.463 | 4.957 |  |  | 875 | 34997.594 | 10.463 | 4.9651
876 | 35037.591 | 10.4642 | 4.9588 |  |  | 876 | 35037.591 | 10.4642 | 4.9568 |  |  | 876 | 35037.591 | 10.4642 | 4.965
877 | 35077.589 | 10.4653 | 4.9587 |  |  | 877 | 35077.589 | 10.4653 | 4.9567 |  |  | 877 | 35077.589 | 10.4653 | 4.9649
878 | 35117.586 | 10.4665 | 4.9586 |  |  | 878 | 35117.586 | 10.4665 | 4.9566 |  |  | 878 | 35117.586 | 10.4665 | 4.9648
879 | 35157.583 | 10.4676 | 4.9584 |  |  | 879 | 35157.583 | 10.4676 | 4.9565 |  |  | 879 | 35157.583 | 10.4676 | 4.9647
880 | 35197.58 | 10.4687 | 4.9583 |  |  | 880 | 35197.58 | 10.4687 | 4.9564 |  |  | 880 | 35197.58 | 10.4687 | 4.9646
881 | 35237.578 | 10.4699 | 4.9582 |  |  | 881 | 35237.578 | 10.4699 | 4.9563 |  |  | 881 | 35237.578 | 10.4699 | 4.9645
882 | 35277.575 | 10.471 | 4.9581 |  |  | 882 | 35277.575 | 10.471 | 4.9562 |  |  | 882 | 35277.575 | 10.471 | 4.9644
883 | 35317.572 | 10.4721 | 4.958 |  |  | 883 | 35317.572 | 10.4721 | 4.956 |  |  | 883 | 35317.572 | 10.4721 | 4.9643
884 | 35357.569 | 10.4733 | 4.9579 |  |  | 884 | 35357.569 | 10.4733 | 4.9559 |  |  | 884 | 35357.569 | 10.4733 | 4.9642
885 | 35397.567 | 10.4744 | 4.9578 |  |  | 885 | 35397.567 | 10.4744 | 4.9558 |  |  | 885 | 35397.567 | 10.4744 | 4.9641
886 | 35437.564 | 10.4755 | 4.9576 |  |  | 886 | 35437.564 | 10.4755 | 4.9557 |  |  | 886 | 35437.564 | 10.4755 | 4.964
887 | 35477.561 | 10.4767 | 4.9575 |  |  | 887 | 35477.561 | 10.4767 | 4.9556 |  |  | 887 | 35477.561 | 10.4767 | 4.9639
888 | 35517.558 | 10.4778 | 4.9574 |  |  | 888 | 35517.558 | 10.4778 | 4.9555 |  |  | 888 | 35517.558 | 10.4778 | 4.9638
889 | 35557.556 | 10.4789 | 4.9573 |  |  | 889 | 35557.556 | 10.4789 | 4.9553 |  |  | 889 | 35557.556 | 10.4789 | 4.9637
890 | 35597.553 | 10.48 | 4.9572 |  |  | 890 | 35597.553 | 10.48 | 4.9552 |  |  | 890 | 35597.553 | 10.48 | 4.9636
891 | 35637.55 | 10.4812 | 4.9571 |  |  | 891 | 35637.55 | 10.4812 | 4.9551 |  |  | 891 | 35637.55 | 10.4812 | 4.9635
892 | 35677.547 | 10.4823 | 4.9569 |  |  | 892 | 35677.547 | 10.4823 | 4.955 |  |  | 892 | 35677.547 | 10.4823 | 4.9634
893 | 35717.545 | 10.4834 | 4.9568 |  |  | 893 | 35717.545 | 10.4834 | 4.9549 |  |  | 893 | 35717.545 | 10.4834 | 4.9633
894 | 35757.542 | 10.4845 | 4.9567 |  |  | 894 | 35757.542 | 10.4845 | 4.9548 |  |  | 894 | 35757.542 | 10.4845 | 4.9632
895 | 35797.539 | 10.4856 | 4.9566 |  |  | 895 | 35797.539 | 10.4856 | 4.9546 |  |  | 895 | 35797.539 | 10.4856 | 4.9631
896 | 35837.536 | 10.4868 | 4.9565 |  |  | 896 | 35837.536 | 10.4868 | 4.9545 |  |  | 896 | 35837.536 | 10.4868 | 4.963
897 | 35877.534 | 10.4879 | 4.9564 |  |  | 897 | 35877.534 | 10.4879 | 4.9544 |  |  | 897 | 35877.534 | 10.4879 | 4.9629
898 | 35917.531 | 10.489 | 4.9563 |  |  | 898 | 35917.531 | 10.489 | 4.9543 |  |  | 898 | 35917.531 | 10.489 | 4.9628
899 | 35957.528 | 10.4901 | 4.9561 |  |  | 899 | 35957.528 | 10.4901 | 4.9542 |  |  | 899 | 35957.528 | 10.4901 | 4.9627
900 | 35997.525 | 10.4912 | 4.956 |  |  | 900 | 35997.525 | 10.4912 | 4.9541 |  |  | 900 | 35997.525 | 10.4912 | 4.9626
901 | 36037.523 | 10.4923 | 4.9559 |  |  | 901 | 36037.523 | 10.4923 | 4.9539 |  |  | 901 | 36037.523 | 10.4923 | 4.9625
902 | 36077.52 | 10.4934 | 4.9558 |  |  | 902 | 36077.52 | 10.4934 | 4.9538 |  |  | 902 | 36077.52 | 10.4934 | 4.9624
903 | 36117.517 | 10.4945 | 4.9557 |  |  | 903 | 36117.517 | 10.4945 | 4.9537 |  |  | 903 | 36117.517 | 10.4945 | 4.9623
904 | 36157.515 | 10.4956 | 4.9556 |  |  | 904 | 36157.515 | 10.4956 | 4.9536 |  |  | 904 | 36157.515 | 10.4956 | 4.9622
905 | 36197.512 | 10.4967 | 4.9554 |  |  | 905 | 36197.512 | 10.4967 | 4.9535 |  |  | 905 | 36197.512 | 10.4967 | 4.9621
906 | 36237.509 | 10.4978 | 4.9553 |  |  | 906 | 36237.509 | 10.4978 | 4.9533 |  |  | 906 | 36237.509 | 10.4978 | 4.962
907 | 36277.506 | 10.499 | 4.9552 |  |  | 907 | 36277.506 | 10.499 | 4.9532 |  |  | 907 | 36277.506 | 10.499 | 4.9619
908 | 36317.504 | 10.5001 | 4.9551 |  |  | 908 | 36317.504 | 10.5001 | 4.9531 |  |  | 908 | 36317.504 | 10.5001 | 4.9618
909 | 36357.501 | 10.5012 | 4.955 |  |  | 909 | 36357.501 | 10.5012 | 4.953 |  |  | 909 | 36357.501 | 10.5012 | 4.9617
910 | 36397.498 | 10.5023 | 4.9549 |  |  | 910 | 36397.498 | 10.5023 | 4.9529 |  |  | 910 | 36397.498 | 10.5023 | 4.9616
911 | 36437.495 | 10.5034 | 4.9548 |  |  | 911 | 36437.495 | 10.5034 | 4.9528 |  |  | 911 | 36437.495 | 10.5034 | 4.9615
912 | 36477.493 | 10.5045 | 4.9546 |  |  | 912 | 36477.493 | 10.5045 | 4.9526 |  |  | 912 | 36477.493 | 10.5045 | 4.9614
913 | 36517.49 | 10.5055 | 4.9545 |  |  | 913 | 36517.49 | 10.5055 | 4.9525 |  |  | 913 | 36517.49 | 10.5055 | 4.9613
914 | 36557.487 | 10.5066 | 4.9544 |  |  | 914 | 36557.487 | 10.5066 | 4.9524 |  |  | 914 | 36557.487 | 10.5066 | 4.9612
915 | 36597.484 | 10.5077 | 4.9543 |  |  | 915 | 36597.484 | 10.5077 | 4.9523 |  |  | 915 | 36597.484 | 10.5077 | 4.9611
916 | 36637.482 | 10.5088 | 4.9542 |  |  | 916 | 36637.482 | 10.5088 | 4.9522 |  |  | 916 | 36637.482 | 10.5088 | 4.961
917 | 36677.479 | 10.5099 | 4.9541 |  |  | 917 | 36677.479 | 10.5099 | 4.9521 |  |  | 917 | 36677.479 | 10.5099 | 4.9609
918 | 36717.476 | 10.511 | 4.9539 |  |  | 918 | 36717.476 | 10.511 | 4.9519 |  |  | 918 | 36717.476 | 10.511 | 4.9607
919 | 36757.473 | 10.5121 | 4.9538 |  |  | 919 | 36757.473 | 10.5121 | 4.9518 |  |  | 919 | 36757.473 | 10.5121 | 4.9606
920 | 36797.471 | 10.5132 | 4.9537 |  |  | 920 | 36797.471 | 10.5132 | 4.9517 |  |  | 920 | 36797.471 | 10.5132 | 4.9605
921 | 36837.468 | 10.5143 | 4.9536 |  |  | 921 | 36837.468 | 10.5143 | 4.9516 |  |  | 921 | 36837.468 | 10.5143 | 4.9604
922 | 36877.465 | 10.5154 | 4.9535 |  |  | 922 | 36877.465 | 10.5154 | 4.9515 |  |  | 922 | 36877.465 | 10.5154 | 4.9603
923 | 36917.462 | 10.5164 | 4.9534 |  |  | 923 | 36917.462 | 10.5164 | 4.9513 |  |  | 923 | 36917.462 | 10.5164 | 4.9602
924 | 36957.46 | 10.5175 | 4.9532 |  |  | 924 | 36957.46 | 10.5175 | 4.9512 |  |  | 924 | 36957.46 | 10.5175 | 4.9601
925 | 36997.457 | 10.5186 | 4.9531 |  |  | 925 | 36997.457 | 10.5186 | 4.9511 |  |  | 925 | 36997.457 | 10.5186 | 4.96
926 | 37037.454 | 10.5197 | 4.953 |  |  | 926 | 37037.454 | 10.5197 | 4.951 |  |  | 926 | 37037.454 | 10.5197 | 4.9599
927 | 37077.451 | 10.5208 | 4.9529 |  |  | 927 | 37077.451 | 10.5208 | 4.9509 |  |  | 927 | 37077.451 | 10.5208 | 4.9598
928 | 37117.449 | 10.5218 | 4.9528 |  |  | 928 | 37117.449 | 10.5218 | 4.9508 |  |  | 928 | 37117.449 | 10.5218 | 4.9597
929 | 37157.446 | 10.5229 | 4.9527 |  |  | 929 | 37157.446 | 10.5229 | 4.9506 |  |  | 929 | 37157.446 | 10.5229 | 4.9596
930 | 37197.443 | 10.524 | 4.9525 |  |  | 930 | 37197.443 | 10.524 | 4.9505 |  |  | 930 | 37197.443 | 10.524 | 4.9595
931 | 37237.44 | 10.5251 | 4.9524 |  |  | 931 | 37237.44 | 10.5251 | 4.9504 |  |  | 931 | 37237.44 | 10.5251 | 4.9594
932 | 37277.438 | 10.5261 | 4.9523 |  |  | 932 | 37277.438 | 10.5261 | 4.9503 |  |  | 932 | 37277.438 | 10.5261 | 4.9593
933 | 37317.435 | 10.5272 | 4.9522 |  |  | 933 | 37317.435 | 10.5272 | 4.9502 |  |  | 933 | 37317.435 | 10.5272 | 4.9592
934 | 37357.432 | 10.5283 | 4.9521 |  |  | 934 | 37357.432 | 10.5283 | 4.95 |  |  | 934 | 37357.432 | 10.5283 | 4.9591
935 | 37397.429 | 10.5294 | 4.952 |  |  | 935 | 37397.429 | 10.5294 | 4.9499 |  |  | 935 | 37397.429 | 10.5294 | 4.959
936 | 37437.427 | 10.5304 | 4.9518 |  |  | 936 | 37437.427 | 10.5304 | 4.9498 |  |  | 936 | 37437.427 | 10.5304 | 4.9589
937 | 37477.424 | 10.5315 | 4.9517 |  |  | 937 | 37477.424 | 10.5315 | 4.9497 |  |  | 937 | 37477.424 | 10.5315 | 4.9588
938 | 37517.421 | 10.5326 | 4.9516 |  |  | 938 | 37517.421 | 10.5326 | 4.9496 |  |  | 938 | 37517.421 | 10.5326 | 4.9587
939 | 37557.418 | 10.5336 | 4.9515 |  |  | 939 | 37557.418 | 10.5336 | 4.9494 |  |  | 939 | 37557.418 | 10.5336 | 4.9586
940 | 37597.416 | 10.5347 | 4.9514 |  |  | 940 | 37597.416 | 10.5347 | 4.9493 |  |  | 940 | 37597.416 | 10.5347 | 4.9585
941 | 37637.413 | 10.5358 | 4.9513 |  |  | 941 | 37637.413 | 10.5358 | 4.9492 |  |  | 941 | 37637.413 | 10.5358 | 4.9584
942 | 37677.41 | 10.5368 | 4.9511 |  |  | 942 | 37677.41 | 10.5368 | 4.9491 |  |  | 942 | 37677.41 | 10.5368 | 4.9583
943 | 37717.407 | 10.5379 | 4.951 |  |  | 943 | 37717.407 | 10.5379 | 4.949 |  |  | 943 | 37717.407 | 10.5379 | 4.9582
944 | 37757.405 | 10.5389 | 4.9509 |  |  | 944 | 37757.405 | 10.5389 | 4.9489 |  |  | 944 | 37757.405 | 10.5389 | 4.9581
945 | 37797.402 | 10.54 | 4.9508 |  |  | 945 | 37797.402 | 10.54 | 4.9487 |  |  | 945 | 37797.402 | 10.54 | 4.958
946 | 37837.399 | 10.5411 | 4.9507 |  |  | 946 | 37837.399 | 10.5411 | 4.9486 |  |  | 946 | 37837.399 | 10.5411 | 4.9579
947 | 37877.396 | 10.5421 | 4.9506 |  |  | 947 | 37877.396 | 10.5421 | 4.9485 |  |  | 947 | 37877.396 | 10.5421 | 4.9578
948 | 37917.394 | 10.5432 | 4.9504 |  |  | 948 | 37917.394 | 10.5432 | 4.9484 |  |  | 948 | 37917.394 | 10.5432 | 4.9577
949 | 37957.391 | 10.5442 | 4.9503 |  |  | 949 | 37957.391 | 10.5442 | 4.9483 |  |  | 949 | 37957.391 | 10.5442 | 4.9576
950 | 37997.388 | 10.5453 | 4.9502 |  |  | 950 | 37997.388 | 10.5453 | 4.9481 |  |  | 950 | 37997.388 | 10.5453 | 4.9575
951 | 38037.385 | 10.5463 | 4.9501 |  |  | 951 | 38037.385 | 10.5463 | 4.948 |  |  | 951 | 38037.385 | 10.5463 | 4.9574
952 | 38077.383 | 10.5474 | 4.95 |  |  | 952 | 38077.383 | 10.5474 | 4.9479 |  |  | 952 | 38077.383 | 10.5474 | 4.9573
953 | 38117.38 | 10.5484 | 4.9498 |  |  | 953 | 38117.38 | 10.5484 | 4.9478 |  |  | 953 | 38117.38 | 10.5484 | 4.9572
954 | 38157.377 | 10.5495 | 4.9497 |  |  | 954 | 38157.377 | 10.5495 | 4.9477 |  |  | 954 | 38157.377 | 10.5495 | 4.9571
955 | 38197.374 | 10.5505 | 4.9496 |  |  | 955 | 38197.374 | 10.5505 | 4.9475 |  |  | 955 | 38197.374 | 10.5505 | 4.957
956 | 38237.372 | 10.5516 | 4.9495 |  |  | 956 | 38237.372 | 10.5516 | 4.9474 |  |  | 956 | 38237.372 | 10.5516 | 4.9569
957 | 38277.369 | 10.5526 | 4.9494 |  |  | 957 | 38277.369 | 10.5526 | 4.9473 |  |  | 957 | 38277.369 | 10.5526 | 4.9568
958 | 38317.366 | 10.5537 | 4.9493 |  |  | 958 | 38317.366 | 10.5537 | 4.9472 |  |  | 958 | 38317.366 | 10.5537 | 4.9567
959 | 38357.363 | 10.5547 | 4.9491 |  |  | 959 | 38357.363 | 10.5547 | 4.9471 |  |  | 959 | 38357.363 | 10.5547 | 4.9565
960 | 38397.361 | 10.5557 | 4.949 |  |  | 960 | 38397.361 | 10.5557 | 4.9469 |  |  | 960 | 38397.361 | 10.5557 | 4.9564
961 | 38437.358 | 10.5568 | 4.9489 |  |  | 961 | 38437.358 | 10.5568 | 4.9468 |  |  | 961 | 38437.358 | 10.5568 | 4.9563
962 | 38477.355 | 10.5578 | 4.9488 |  |  | 962 | 38477.355 | 10.5578 | 4.9467 |  |  | 962 | 38477.355 | 10.5578 | 4.9562
963 | 38517.352 | 10.5589 | 4.9487 |  |  | 963 | 38517.352 | 10.5589 | 4.9466 |  |  | 963 | 38517.352 | 10.5589 | 4.9561
964 | 38557.35 | 10.5599 | 4.9486 |  |  | 964 | 38557.35 | 10.5599 | 4.9465 |  |  | 964 | 38557.35 | 10.5599 | 4.956
965 | 38597.347 | 10.5609 | 4.9484 |  |  | 965 | 38597.347 | 10.5609 | 4.9463 |  |  | 965 | 38597.347 | 10.5609 | 4.9559
966 | 38637.344 | 10.562 | 4.9483 |  |  | 966 | 38637.344 | 10.562 | 4.9462 |  |  | 966 | 38637.344 | 10.562 | 4.9558
967 | 38677.341 | 10.563 | 4.9482 |  |  | 967 | 38677.341 | 10.563 | 4.9461 |  |  | 967 | 38677.341 | 10.563 | 4.9557
968 | 38717.339 | 10.564 | 4.9481 |  |  | 968 | 38717.339 | 10.564 | 4.946 |  |  | 968 | 38717.339 | 10.564 | 4.9556
969 | 38757.336 | 10.5651 | 4.948 |  |  | 969 | 38757.336 | 10.5651 | 4.9458 |  |  | 969 | 38757.336 | 10.5651 | 4.9555
970 | 38797.333 | 10.5661 | 4.9478 |  |  | 970 | 38797.333 | 10.5661 | 4.9457 |  |  | 970 | 38797.333 | 10.5661 | 4.9554
971 | 38837.33 | 10.5671 | 4.9477 |  |  | 971 | 38837.33 | 10.5671 | 4.9456 |  |  | 971 | 38837.33 | 10.5671 | 4.9553
972 | 38877.328 | 10.5682 | 4.9476 |  |  | 972 | 38877.328 | 10.5682 | 4.9455 |  |  | 972 | 38877.328 | 10.5682 | 4.9552
973 | 38917.325 | 10.5692 | 4.9475 |  |  | 973 | 38917.325 | 10.5692 | 4.9454 |  |  | 973 | 38917.325 | 10.5692 | 4.9551
974 | 38957.322 | 10.5702 | 4.9474 |  |  | 974 | 38957.322 | 10.5702 | 4.9452 |  |  | 974 | 38957.322 | 10.5702 | 4.955
975 | 38997.319 | 10.5712 | 4.9472 |  |  | 975 | 38997.319 | 10.5712 | 4.9451 |  |  | 975 | 38997.319 | 10.5712 | 4.9549
976 | 39037.317 | 10.5723 | 4.9471 |  |  | 976 | 39037.317 | 10.5723 | 4.945 |  |  | 976 | 39037.317 | 10.5723 | 4.9548
977 | 39077.314 | 10.5733 | 4.947 |  |  | 977 | 39077.314 | 10.5733 | 4.9449 |  |  | 977 | 39077.314 | 10.5733 | 4.9547
978 | 39117.311 | 10.5743 | 4.9469 |  |  | 978 | 39117.311 | 10.5743 | 4.9448 |  |  | 978 | 39117.311 | 10.5743 | 4.9546
979 | 39157.308 | 10.5753 | 4.9468 |  |  | 979 | 39157.308 | 10.5753 | 4.9446 |  |  | 979 | 39157.308 | 10.5753 | 4.9545
980 | 39197.306 | 10.5764 | 4.9467 |  |  | 980 | 39197.306 | 10.5764 | 4.9445 |  |  | 980 | 39197.306 | 10.5764 | 4.9544
981 | 39237.303 | 10.5774 | 4.9465 |  |  | 981 | 39237.303 | 10.5774 | 4.9444 |  |  | 981 | 39237.303 | 10.5774 | 4.9543
982 | 39277.3 | 10.5784 | 4.9464 |  |  | 982 | 39277.3 | 10.5784 | 4.9443 |  |  | 982 | 39277.3 | 10.5784 | 4.9542
983 | 39317.297 | 10.5794 | 4.9463 |  |  | 983 | 39317.297 | 10.5794 | 4.9442 |  |  | 983 | 39317.297 | 10.5794 | 4.9541
984 | 39357.295 | 10.5804 | 4.9462 |  |  | 984 | 39357.295 | 10.5804 | 4.944 |  |  | 984 | 39357.295 | 10.5804 | 4.954
985 | 39397.292 | 10.5815 | 4.9461 |  |  | 985 | 39397.292 | 10.5815 | 4.9439 |  |  | 985 | 39397.292 | 10.5815 | 4.9539
986 | 39437.289 | 10.5825 | 4.9459 |  |  | 986 | 39437.289 | 10.5825 | 4.9438 |  |  | 986 | 39437.289 | 10.5825 | 4.9538
987 | 39477.286 | 10.5835 | 4.9458 |  |  | 987 | 39477.286 | 10.5835 | 4.9437 |  |  | 987 | 39477.286 | 10.5835 | 4.9537
988 | 39517.284 | 10.5845 | 4.9457 |  |  | 988 | 39517.284 | 10.5845 | 4.9435 |  |  | 988 | 39517.284 | 10.5845 | 4.9536
989 | 39557.281 | 10.5855 | 4.9456 |  |  | 989 | 39557.281 | 10.5855 | 4.9434 |  |  | 989 | 39557.281 | 10.5855 | 4.9534
990 | 39597.278 | 10.5865 | 4.9455 |  |  | 990 | 39597.278 | 10.5865 | 4.9433 |  |  | 990 | 39597.278 | 10.5865 | 4.9533